## Change Log
Document Change Log
Initals | Date | Edited Existing Data? | Added New Data? | Description
PR | 2017-10-31 | No | Yes | Change Log Started
PR | 2017-10-31 | Yes | No | expanded PMs in FS2_N0108 rack
PR | 2017-11-13 | Yes | No | Added Patch panel details to FS2_n0108 profile monitors
PR | 2017-11-14 | Yes | No | Cleaned up File; added Lattice File location, standardized formatting, removed HV Bias Planning page
PR | 2018-01-03 | Yes |  | Changed Configuration of the Faraday Cups in FE_N0106
PR | 2018-01-04 | Yes |  | Audited Device D line number using the working lattice file. ONLY devices up to the RFQ were looked at. Changes are highlited in green
PR | 2018-08-08 | Yes | No | Some typos in the Pico inputs were found in Rack N0604 and now fixed
PR | 2018-08-15 |  | Yes | Add all additional to complete beamline
PR | 2018-09-05 | Yes |  | Adjusted some incorrect iseg asignments in newer racks (nothing 'in use' was changed)
PR | 2018-09-11 | Yes |  | Fixed errors with FE FC assigment
PR | 2018-09-13 | Yes |  | Full Aduit of all D-line numbers
PR | 2018-09-19 | Yes |  | Added Column for Device Type for easier parsing
PR | 2018-10-30 |  | Yes | Added Wa03 Beam Stop
PR | 11/16/218 | Yes |  | Changed mislabeled FS1 BPM from FS1_BBS:BPM_D2436 to FS1_BTS:BPM_D2436
PR | 2018-11-21 | Yes |  | Adjusted BPM Ref in LS1_N0206 to match what is connected
PR | 2019-01-22 |  | Yes | Added Temp Pico card in MTCA 2
PR | 2019-03-01 | Yes |  | Changed Flapper PMs bias; Fixed a cable number in FS1_N0302; Fixed print area
PR | 2019-03-13 | Yes |  | Changed FE_N0106 for the change in biasing of the FFC
PR | 2019-06-04 | Yes | Yes | Added newly 're-installed' HMRs ; Added Master Sheet
PR | 2019-06-12 | Yes |  | Changed devices to be consistant across mutiple sheets
PR | 2019-06-24 | Yes |  | Changed all biasing signal's function to 'HV Bias'; Changed the Slot number of an iseg card in LS2_N4202
PR | 2019-07-01 | Yes |  | Allocated spot for new FC in FS1; rack LS2_N0808
PR | 2019-07-25 |  | Yes | Added BSM current wire feedback to Pico9 in FS2 N0302
PR | 2019-08-26 | Yes | Yes | Updated wiring for new FC in LS2_N0808
PR | 2019-09-13 | Yes |  | Removed FPGDB fields for PMs that should have not been there
PR | 2019-09-16 | Yes |  | Changed Signal Assinment for PM D998
PR | 2019-10-16 | Yes | Yes | Add BCM cables to document; Added device type field for PM bias and GND plugs.
SC | 2019-10-18 | Yes |  | Changed sheet tab names -- shorter. Changed BD D1224 bias cable description.
PR | 2019-10-25 |  | Yes | Add new line items for bias patch cables and ground plugs
PR | 2019-10-25 | Yes |  | Changed Channel Assignment of WA03 Beam dump
PR | 2019-10-29 | Yes |  | Changed Termination Device column to label items as panels or boxes
PR | 2019-10-29 | Yes |  | Swapped Channel assignment for FC_D0814 and FC_D0796
PR | 2019-11-18 | Yes |  | Added some missing CPC connection information on Profile Monitors in mutilple racks
PR | 2019-12-03 | Yes |  | Swapped Channel assignment for FC_D0814 and FC_D0796 back to original assignment to ready
PR | 2020-01-03 | Yes |  | Changed IC_D4180->D4081; Changed BD to FC; modifed FS1 FC name; changed IC in MTCA15
PR | 2020-01-06 | Yes |  | Added FS2_FC
PR | 2020-01-13 | Yes |  | Fix FS1 FC device type from HMR -> FC
PR | 2020-01-20 | Yes |  | Fixed name of FS2 BPMs
PR | 2020-01-22 | Yes | Yes | Added new Pico in mTCA3 for FS2 Faraday Cup; MOVED WA03 FC from Channel 1 to 0 on its Pico.
PR | 2020-02-04 | Yes |  | Edited some LS2 BPM REF cabling to match the actual reference we are using. Also Adjusted some incorrect CableDB #'s ofr LS2 HMRs
PR | 2020-02-05 | Yes |  | Fixed HMR Heading in LS2_N1708
PR | 2020-02-12 | Yes |  | Added PM D2552 using the cables from D2504
PR | 2020-02-20 | Yes | Yes | Made MANY small changes to MTCA1-MTCA9 & MTCA20 based on what cables are actually in the Racks Example - most changes were reflect which BPM Refs were actually used.
PR | 2020-03-04 |  | Yes | Added two BDS Ion Chambers to BDS_N0105

## Sept '18 LatticeLayout
Last updated on (see details on the "ChangeInfo" sheet): | [5 empty] | 2018-01-17
BRANCH 1 Artemis_B Source Line | [20 empty] | BRANCH 2 High Power Source Line
System | Sub- system | Device | Beam line position (dm) | Name | Device Type | Element Name | Minimum Beam Pipe Inner Diameter (mm) | Element Flange to Flange Length (m) | PLACE HOLDER | Element Effective Length (m) | Global Coordinate element center Xc (m) | Global Coordinate element center Yc (m) | Global Coordinate element center Zc (m) | Global Coordinate Z (m) | Global Coordinate Y (m) | Global Coordinate X (m) | Accumulated center-to-center Length (m) | Accumulated end-to-end Length (m) | Comment | Distance from Artemis source (m) | System | Sub- system | Device | Beam line position (dm) | Name | Device Type | Element Name | Minimum Beam Pipe Inner Diameter (mm) | Element Flange to Flange Length (m) | PLACE HOLDER | Element Effective Length (m) | Global Coordinate element center Xc (m) | Global Coordinate element center Yc (m) | Global Coordinate element center Zc (m) | Global Coordinate Z (m) | Global Coordinate Y (m) | Global Coordinate X (m) | Accumulated center-to-center Length (m) | Accumulated end-to-end Length (m) | Comment
[20 empty] | 68.4786 | FE | ISRC2 | ELC1 | 664 | FE_ISRC2:ELC1_D0664 |  | VENUS electrode |  |  |  | 0 | 368.5533 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 368.5533 | 66.4319 | 66.4319
[31 empty] | 0.025397 | 368.5406 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 368.5279 | 66.4446 | 66.4573
[21 empty] | FE | ISRC2 | ELC2 | 665 | FE_ISRC2:ELC2_D0665 |  | VENUS extraction/puller | 150 |  |  | 0 | 368.5279 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 368.5279 | 66.4573 | 66.4573
[31 empty] | 0.33074 | 368.3626 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 368.1972 | 66.6227 | 66.788
[31 empty] | 0 | 368.1972 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 368.1972 | 66.788 | 66.788
[31 empty] | 0 | 368.1972 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 368.1972 | 66.788 | 66.788
[21 empty] | FE | ISRC2 | SOLR | 670 | FE_ISRC2:SOLR_D0670 | S4s | solenoid | 150 |  |  | 0.33815 | 368.0281 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 367.8591 | 66.9571 | 67.1262
[31 empty] | 0.17229 | 367.7729 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 367.6868 | 67.2123 | 67.2985
[21 empty] | FE | ISRC2 | TMP | 673 | FE_ISRC2:TMP_D0673 | TMP_L | turbo pump, 2050 l/s |  |  |  | 0 | 367.6868 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 367.6868 | 67.2985 | 67.2985
[31 empty] | 0.20348 | 367.585 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 367.4833 | 67.4002 | 67.502
[21 empty] | FE | ISRC2 | GV | 675 | FE_ISRC2:GV_D0675 | GV_L | gate valve, 10" x 8" | 150 |  |  | 0.080264 | 367.4432 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 367.403 | 67.5421 | 67.5822
[31 empty] | 0 | 367.403 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 367.403 | 67.5822 | 67.5822
[21 empty] | FE | ISRC2 | ACC | 677 | FE_ISRC2:ACC_D0677 | ACC | acceleration column | 150 |  |  | 0.26772 | 367.2692 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 367.1353 | 67.7161 | 67.8499
FE | ISRC1 | ELC1 | 685 | FE_ISRC1:ELC1_D0685 |  | extraction electrode inner wall |  |  |  | 0 | 370.7066 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 370.7066 | 68.4786 | 68.4786 |  | 0 | [6 empty] | suppression electrode | 150 |  |  | 0.024638 | 367.123 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 367.1107 | 67.8623 | 67.8746
[6 empty] | 8 mm extraction hole |  |  |  | 0.0023657 | 370.7054 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 370.7042 | 68.4798 | 68.481 |  | 0.0011829 | [6 empty] | bellows | 150 |  |  | 0.016422 | 367.1025 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 367.0942 | 67.8828 | 67.891
[6 empty] | artemis_b extraction conical wall |  |  |  | 0.011632 | 370.6984 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 370.6926 | 68.4868 | 68.4926 |  | 0.0081818 | [10 empty] | 0 | 367.0942 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 367.0942 | 67.891 | 67.891
[6 empty] | extraction mounting plate |  |  |  | 0 | 370.6926 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 370.6926 | 68.4926 | 68.4926 |  | 0.013998 | FE | SCS2 | DCHV | 680 | FE_SCS2:DCHV_D0680 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 366.999 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 366.9037 | 67.9862 | 68.0815
[6 empty] | gap (puller & extraction hole) |  |  |  | 0.038348 | 370.6734 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 370.6543 | 68.5118 | 68.531 |  | 0.033172 | [7 empty] | 150 |  |  | 0.0254 | 366.891 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 366.8783 | 68.0942 | 68.1069
FE | ISRC1 | ELC2 | 685 | FE_ISRC1:ELC2_D0685 | PULL_B | puller electrode (bias) |  |  |  | 0.0044341 | 370.652 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 370.6498 | 68.5332 | 68.5354 |  | 0.054563 | FE | SCS2 | SOLR | 683 | FE_SCS2:SOLR_D0683 | S4b | solenoid | 150 |  |  | 0.4 | 366.6783 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 366.4783 | 68.3069 | 68.5069
[6 empty] | gap (puller main & bias) |  |  |  | 0.0047569 | 370.6475 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 370.6451 | 68.5378 | 68.5402 |  | 0.059159 | [7 empty] | 150 |  |  | 0.125 | 366.4158 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 366.3533 | 68.5694 | 68.6319
FE | ISRC1 | ELC3 | 685 | FE_ISRC1:ELC3_D0685 | PULL_M | puller electrode (main) |  |  |  | 0.0044341 | 370.6429 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 370.6406 | 68.5424 | 68.5446 |  | 0.063754 | [10 empty] | 0 | 366.3533 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 366.3533 | 68.6319 | 68.6319
[6 empty] | puller tube |  |  |  | 0.13195 | 370.5747 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 370.5087 | 68.6106 | 68.6766 |  | 0.13195 | FE | SCS2 | TMP | 686 | FE_SCS2:TMP_D0686 | TMP_M | turbo pump, 700 l/s | 150 |  |  | 0 | 366.3533 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 366.3533 | 68.6319 | 68.6319
[6 empty] | extraction box |  |  |  | 0.036269 | 370.4905 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 370.4724 | 68.6947 | 68.7128 |  | 0.21606 | [10 empty] | 0 | 366.3533 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 366.3533 | 68.6319 | 68.6319
FE | ISRC1 | TMP | 687 | FE_ISRC1:TMP_D0687 | TMP_M | turbo pump, 700 l/s |  |  |  | 0 | 370.4724 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 370.4724 | 68.7128 | 68.7128 |  | 0.23419 | [7 empty] | 150 |  |  | 0.11975 | 366.2935 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 366.2336 | 68.6918 | 68.7516
[6 empty] | extraction box |  |  |  | 0.1768 | 370.384 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 370.2956 | 68.8012 | 68.8896 |  | 0.32259 | [7 empty] | 150 |  |  | 0 | 366.2336 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 366.2336 | 68.7516 | 68.7516
FE | ISRC1 | SOLR | 690 | FE_ISRC1:SOLR_D0690 | S0 | solenoid | 150 |  |  | 0.26032 | 370.1655 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 370.0353 | 69.0198 | 69.1499 | 40cm effective length | 0.54116 | FE | SCS2 | DCHV | 688 | FE_SCS2:DCHV_D0688 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 366.1383 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 366.0431 | 68.8469 | 68.9421
[6 empty] | bellow | 150 |  |  | 0.23421 | 369.9182 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 369.8011 | 69.2671 | 69.3842 |  | 0.78843 | [10 empty] | 0 | 366.0431 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 366.0431 | 68.9421 | 68.9421
FE | ISRC1 | DCHV | 695 | FE_ISRC1:DCHV_D0695 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 369.7058 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 369.6106 | 69.4794 | 69.5747 |  | 1.0008 | [10 empty] | 0 | 366.0431 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 366.0431 | 68.9421 | 68.9421
[6 empty] | tube | 150 |  |  | 0.02686 | 369.5971 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 369.5837 | 69.5881 | 69.6015 |  | 1.1095 | [7 empty] | 150 |  |  | 0.13975 | 365.9732 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 365.9033 | 69.012 | 69.0819
FE | ISRC1 | GV | 696 | FE_ISRC1:GV_D0696 | GV_L | gate valve | 150 |  |  | 0.080264 | 369.5436 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 369.5035 | 69.6417 | 69.6818 |  | 1.163 | FE | SCS2 | DH | 696 | FE_SCS2:DH_D0696 | D5 | dipole-90deg, D5-1 | 120 X 150 |  |  | 1 | 365.585 | 418.9447 | 51.6647 | 51.6647 | 419.263 | 365.2667 | 69.5819 | 70.0819
[7 empty] | 150 |  |  | 0 | 369.5035 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 369.5035 | 69.6818 | 69.6818 | extra length due to a 10" CFF on GV_D0696 | 1.2032 | FE | SRC2 | TMP | 701 | FE_SRC2:TMP_D0701 | TMP_M | turbo pump, 700 l/s |  |  |  | 0 | 365.2667 | 419.263 | 51.6647 | 51.6647 | 419.263 | 365.2667 | 70.0819 | 70.0819
FE | ISRC1 | ACC | 698 | FE_ISRC1:ACC_D0698 |  | acceleration column | 150 |  |  | 0.26772 | 369.3696 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 369.2357 | 69.8156 | 69.9495 |  | 1.337 | [10 empty] | 0.06705 | 365.2667 | 419.2965 | 51.6647 | 51.6647 | 419.33 | 365.2667 | 70.1154 | 70.1489
FE | ISRC1 | ACC | 700 | FE_ISRC1:ACC_D0700 |  | suppression electrode | 150 |  |  | 0.024638 | 369.2234 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 369.2111 | 69.9618 | 69.9741 |  | 1.4832 | FE | SCS2 | DCHV | 702 | FE_SCS2:DCHV_D0702 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 365.2667 | 419.4253 | 51.6647 | 51.6647 | 419.5205 | 365.2667 | 70.2442 | 70.3394
[10 empty] | 0.016854 | 369.2027 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 369.1942 | 69.9826 | 69.991 |  | 1.5039 | [7 empty] | 150 |  |  | 0.09245 | 365.2667 | 419.5668 | 51.6647 | 51.6647 | 419.613 | 365.2667 | 70.3857 | 70.4319
[10 empty] | 0.1905 | 369.099 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 369.0037 | 70.0862 | 70.1815 | Moved to D0695 location | 1.6076 | FE | SCS2 | QHE | 705 | FE_SCS2:QHE_D0705 | Q7 | ES_quad_X | 150 |  |  | 0.2 | 365.2667 | 419.713 | 51.6647 | 51.6647 | 419.813 | 365.2667 | 70.5319 | 70.6319 | TRIPLET HOUSING
[7 empty] | 150 |  |  | 0.0254 | 368.991 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 368.9783 | 70.1942 | 70.2069 |  | 1.7156 | [7 empty] | 150 |  |  | 0.135 | 365.2667 | 419.8805 | 51.6647 | 51.6647 | 419.948 | 365.2667 | 70.6994 | 70.7669 | TRIPLET HOUSING
FE | SCS1 | SOLR | 704 | FE_SCS1:SOLR_D0704 | S4b | solenoid | 150 |  |  | 0.4 | 368.7783 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 368.5783 | 70.4069 | 70.6069 | 40cm effective length | 1.9283 | FE | SCS2 | QVE | 709 | FE_SCS2:QVE_D0709 | Q7 | ES_quad_Y | 150 |  |  | 0.2 | 365.2667 | 420.048 | 51.6647 | 51.6647 | 420.148 | 365.2667 | 70.8669 | 70.9669 | TRIPLET HOUSING
[7 empty] | 150 |  |  | 0.125 | 368.5158 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 368.4533 | 70.6694 | 70.7319 |  | 2.1908 | [7 empty] | 150 |  |  | 0.135 | 365.2667 | 420.2155 | 51.6647 | 51.6647 | 420.283 | 365.2667 | 71.0344 | 71.1019 | TRIPLET HOUSING
FE | SCS1 | TMP | 707 | FE_SCS1:TMP_D0707 | TMP_M | turbo pump, 700 l/s |  |  |  | 0 | 368.4533 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 368.4533 | 70.7319 | 70.7319 |  | 2.2533 | FE | SCS2 | QHE | 712 | FE_SCS2:QHE_D0712 | Q7 | ES_quad_X | 150 |  |  | 0.2 | 365.2667 | 420.383 | 51.6647 | 51.6647 | 420.483 | 365.2667 | 71.2019 | 71.3019 | TRIPLET HOUSING
[7 empty] | 150 |  |  | 0.11975 | 368.3935 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 368.3336 | 70.7918 | 70.8516 |  | 2.3131 | [10 empty] | 0 | 365.2667 | 420.483 | 51.6647 | 51.6647 | 420.483 | 365.2667 | 71.3019 | 71.3019
[10 empty] | 0 | 368.3336 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 368.3336 | 70.8516 | 70.8516 |  | 2.373 | [6 empty] | bellow+box | 150 |  | 0.28 | 0.22821 | 365.2667 | 420.5971 | 51.6647 | 51.6647 | 420.7112 | 365.2667 | 71.416 | 71.5301
FE | SCS1 | DCHV | 709 | FE_SCS1:DCHV_D0709 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 368.2383 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 368.1431 | 70.9469 | 71.0421 | dimension to yoke not flange | 2.4683 | FE | SCS2 | SLT | 715 | FE_SCS2:SLT_D0715 |  | four jaw collimator | 150 |  |  | 0 | 365.2667 | 420.7112 | 51.6647 | 51.6647 | 420.7112 | 365.2667 | 71.5301 | 71.5301
[7 empty] | 150 |  |  | 0.13975 | 368.0732 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 368.0033 | 71.112 | 71.1819 |  | 2.6334 | [7 empty] | 150 |  | 0.2159 | 0.021 | 365.2667 | 420.7217 | 51.6647 | 51.6647 | 420.7322 | 365.2667 | 71.5406 | 71.5511
FE | SCS1 | DH | 717 | FE_SCS1:DH_D0717 | D5 | dipole-90deg, D5-1 | 120 x 150 |  |  | 1 | 367.685 | 427.2545 | 51.6647 | 51.6647 | 426.9362 | 367.3667 | 71.6819 | 72.1819 |  | 3.2033 | FE | SCS2 | TMP | 716 | FE_SCS2:TMP_D0716 | TMP_M | turbo pump, 700 l/s |  |  |  | 0 | 365.2667 | 420.7322 | 51.6647 | 51.6647 | 420.7322 | 365.2667 | 71.5511 | 71.5511
FE | SCS1 | TMP | 717 | FE_SCS1:TMP_D0717 | TMP_M | turbo pump, 700 l/s |  |  |  | 0 | 367.3667 | 426.9362 | 51.6647 | 51.6647 | 426.9362 | 367.3667 | 72.1819 | 72.1819 | special: pump port on the side of magnet | 3.7033 | [10 empty] | 0.11811 | 365.2667 | 420.7913 | 51.6647 | 51.6647 | 420.8503 | 365.2667 | 71.6102 | 71.6692
[7 empty] | 150 |  |  | 0 | 367.3667 | 426.9362 | 51.6647 | 51.6647 | 426.9362 | 367.3667 | 72.1819 | 72.1819 | straight port | 3.7033 | FE | SCS2 | FC | 717 | FE_SCS2:FC_D0717 | FC2 | Faraday cup | 150 |  |  | 0 | 365.2667 | 420.8503 | 51.6647 | 51.6647 | 420.8503 | 365.2667 | 71.6692 | 71.6692
[7 empty] | 150 |  |  | 0 | 367.3667 | 426.9362 | 51.6647 | 51.6647 | 426.9362 | 367.3667 | 72.1819 | 72.1819 |  | 3.7033 | [7 empty] | 150 |  |  | 0.18034 | 365.2667 | 420.9405 | 51.6647 | 51.6647 | 421.0306 | 365.2667 | 71.7594 | 71.8496
[7 empty] | 150 |  |  | 0.06705 | 367.3667 | 426.9027 | 51.6647 | 51.6647 | 426.8692 | 367.3667 | 72.2154 | 72.2489 | dimension to yoke not flange | 3.7368 | FE | SCS2 | VD | 718 | FE_SCS2:VD_D0718 | V1 | viewer | 150 |  |  | 0 | 365.2667 | 421.0306 | 51.6647 | 51.6647 | 421.0306 | 365.2667 | 71.8496 | 71.8496
FE | SCS1 | DCHV | 723 | FE_SCS1:DCHV_D0723 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 367.3667 | 426.7739 | 51.6647 | 51.6647 | 426.6787 | 367.3667 | 72.3442 | 72.4394 |  | 3.8656 | FE | SCS2 | EMS | 718 | FE_SCS2:EMS_D0718 | EM1 | emittance scanner (x & y) | 150 |  |  | 0 | 365.2667 | 421.0306 | 51.6647 | 51.6647 | 421.0306 | 365.2667 | 71.8496 | 71.8496
[7 empty] | 150 |  |  | 0.09245 | 367.3667 | 426.6324 | 51.6647 | 51.6647 | 426.5862 | 367.3667 | 72.4857 | 72.5319 |  | 4.007 | [7 empty] | 150 |  | 0.2641 | 0.21234 | 365.2667 | 421.1368 | 51.6647 | 51.6647 | 421.243 | 365.2667 | 71.9557 | 72.0619
FE | SCS1 | QHE | 726 | FE_SCS1:QHE_D0726 | Q7 | ES_quad_X | 150 |  |  | 0.2 | 367.3667 | 426.4862 | 51.6647 | 51.6647 | 426.3862 | 367.3667 | 72.6319 | 72.7319 | TRIPLET HOUSING | 4.1533 | FE | SCS2 | QHE | 722 | FE_SCS2:QHE_D0722 | Q7 | ES_quad_X | 150 |  |  | 0.2 | 365.2667 | 421.343 | 51.6647 | 51.6647 | 421.443 | 365.2667 | 72.1619 | 72.2619 | TRIPLET HOUSING
[7 empty] | 150 |  |  | 0.135 | 367.3667 | 426.3187 | 51.6647 | 51.6647 | 426.2512 | 367.3667 | 72.7994 | 72.8669 | TRIPLET HOUSING | 4.3208 | [7 empty] | 150 |  |  | 0.135 | 365.2667 | 421.5105 | 51.6647 | 51.6647 | 421.578 | 365.2667 | 72.3294 | 72.3969 | TRIPLET HOUSING
FE | SCS1 | QVE | 730 | FE_SCS1:QVE_D0730 | Q7 | ES_quad_Y | 150 |  |  | 0.2 | 367.3667 | 426.1512 | 51.6647 | 51.6647 | 426.0512 | 367.3667 | 72.9669 | 73.0669 | TRIPLET HOUSING | 4.4883 | FE | SCS2 | QVE | 725 | FE_SCS2:QVE_D0725 | Q7 | ES_quad_Y | 150 |  |  | 0.2 | 365.2667 | 421.678 | 51.6647 | 51.6647 | 421.778 | 365.2667 | 72.4969 | 72.5969 | TRIPLET HOUSING
[7 empty] | 150 |  |  | 0.135 | 367.3667 | 425.9837 | 51.6647 | 51.6647 | 425.9162 | 367.3667 | 73.1344 | 73.2019 | TRIPLET HOUSING | 4.6558 | [7 empty] | 150 |  |  | 0.135 | 365.2667 | 421.8455 | 51.6647 | 51.6647 | 421.913 | 365.2667 | 72.6644 | 72.7319 | TRIPLET HOUSING
FE | SCS1 | QHE | 733 | FE_SCS1:QHE_D0733 | Q7 | ES_quad_X | 150 |  |  | 0.2 | 367.3667 | 425.8162 | 51.6647 | 51.6647 | 425.7162 | 367.3667 | 73.3019 | 73.4019 | TRIPLET HOUSING | 4.8233 | FE | SCS2 | QHE | 728 | FE_SCS2:QHE_D0728 | Q7 | ES_quad_X | 150 |  |  | 0.2 | 365.2667 | 422.013 | 51.6647 | 51.6647 | 422.113 | 365.2667 | 72.8319 | 72.9319 | TRIPLET HOUSING
[6 empty] | bellow+box | 150 |  |  | 0.22821 | 367.3667 | 425.6021 | 51.6647 | 51.6647 | 425.488 | 367.3667 | 73.516 | 73.6301 |  | 5.0374 | [7 empty] | 150 |  | 0.06975 | 0.09245 | 365.2667 | 422.1592 | 51.6647 | 51.6647 | 422.2054 | 365.2667 | 72.9781 | 73.0243
FE | SCS1 | SLT | 736 | FE_SCS1:SLT_D0736 |  | four jaw collimator | 150 |  |  | 0 | 367.3667 | 425.488 | 51.6647 | 51.6647 | 425.488 | 367.3667 | 73.6301 | 73.6301 |  | 5.1515 | [10 empty] | 0 | 365.2667 | 422.2054 | 51.6647 | 51.6647 | 422.2054 | 365.2667 | 73.0243 | 73.0243
[7 empty] | 150 |  |  | 0.021 | 367.3667 | 425.4775 | 51.6647 | 51.6647 | 425.467 | 367.3667 | 73.6406 | 73.6511 |  | 5.162 | [10 empty] | 0 | 365.2667 | 422.2054 | 51.6647 | 51.6647 | 422.2054 | 365.2667 | 73.0243 | 73.0243
FE | SCS1 | TMP | 737 | FE_SCS1:TMP_D0737 | TMP_M | turbo pump, 700 l/s |  |  |  | 0 | 367.3667 | 425.467 | 51.6647 | 51.6647 | 425.467 | 367.3667 | 73.6511 | 73.6511 |  | 5.1725 | [10 empty] | 0 | 365.2667 | 422.2054 | 51.6647 | 51.6647 | 422.2054 | 365.2667 | 73.0243 | 73.0243
[6 empty] | box | 150 |  |  | 0.11811 | 367.3667 | 425.408 | 51.6647 | 51.6647 | 425.3489 | 367.3667 | 73.7102 | 73.7692 |  | 5.2315 | FE | SCS2 | DCHV | 731 | FE_SCS2:DCHV_D0731 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 365.2667 | 422.3007 | 51.6647 | 51.6647 | 422.3959 | 365.2667 | 73.1196 | 73.2148
FE | SCS1 | FC | 738 | FE_SCS1:FC_D0738 | FC2 | Faraday cup | 150 |  |  | 0 | 367.3667 | 425.3489 | 51.6647 | 51.6647 | 425.3489 | 367.3667 | 73.7692 | 73.7692 |  | 5.2906 | [9 empty] | 0.08975 | 0.06705 | 365.2667 | 422.4295 | 51.6647 | 51.6647 | 422.463 | 365.2667 | 73.2484 | 73.2819
[7 empty] | 150 |  |  | 0.18034 | 367.3667 | 425.2587 | 51.6647 | 51.6647 | 425.1686 | 367.3667 | 73.8594 | 73.9496 |  | 5.3808 | FE | SCS2 | DH | 738 | FE_SCS2:DH_D0738 | D5 | dipole-90deg, D5-2 | 120 X 150 |  |  | 1 | 365.585 | 422.7813 | 51.6647 | 51.6647 | 423.0996 | 365.9033 | 73.7819 | 74.2819
FE | SCS1 | VD | 739 | FE_SCS1:VD_D0739 | V1 | viewer | 150 |  |  | 0 | 367.3667 | 425.1686 | 51.6647 | 51.6647 | 425.1686 | 367.3667 | 73.9496 | 73.9496 |  | 5.4709 | FE | SCS2 | TMP | 743 | FE_SCS2:TMP_D0743 | TMP_M | turbo pump, 700 l/s | 150 |  |  | 0 | 365.9033 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 365.9033 | 74.2819 | 74.2819
FE | SCS1 | EMS | 739 | FE_SCS1:EMS_D0739 | EM1 | emittance scanner (x & y) | 150 |  |  | 0 | 367.3667 | 425.1686 | 51.6647 | 51.6647 | 425.1686 | 367.3667 | 73.9496 | 73.9496 |  | 5.4709 | [9 empty] | 0.225 | 0.225 | 366.0158 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 366.1283 | 74.3944 | 74.5069
[6 empty] | bellow+box | 150 |  |  | 0.21234 | 367.3667 | 425.0624 | 51.6647 | 51.6647 | 424.9562 | 367.3667 | 74.0557 | 74.1619 |  | 5.5771 | FE | SCS2 | QVE | 746 | FE_SCS2:QVE_D0746 | Q7 | ES_quad_Y | 150 |  |  | 0.2 | 366.2283 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 366.3283 | 74.6069 | 74.7069 | DOUBLET HOUSING
FE | SCS1 | QHE | 743 | FE_SCS1:QHE_D0743 | Q7 | ES_quad_X | 150 |  |  | 0.2 | 367.3667 | 424.8562 | 51.6647 | 51.6647 | 424.7562 | 367.3667 | 74.2619 | 74.3619 | TRIPLET HOUSING | 5.7833 | [7 empty] | 150 |  |  | 0.135 | 366.3958 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 366.4633 | 74.7744 | 74.8419 | DOUBLET HOUSING
[7 empty] | 150 |  |  | 0.135 | 367.3667 | 424.6887 | 51.6647 | 51.6647 | 424.6212 | 367.3667 | 74.4294 | 74.4969 | TRIPLET HOUSING | 5.9508 | FE | SCS2 | QHE | 749 | FE_SCS2:QHE_D0749 | Q7 | ES_quad_X | 150 |  |  | 0.2 | 366.5633 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 366.6633 | 74.9419 | 75.0419 | DOUBLET HOUSING
FE | SCS1 | QVE | 746 | FE_SCS1:QVE_D0746 | Q7 | ES_quad_Y | 150 |  |  | 0.2 | 367.3667 | 424.5212 | 51.6647 | 51.6647 | 424.4212 | 367.3667 | 74.5969 | 74.6969 | TRIPLET HOUSING | 6.1183 | [9 empty] | 0.20331 | 0.19159 | 366.7591 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 366.8549 | 75.1377 | 75.2335
[7 empty] | 150 |  |  | 0.135 | 367.3667 | 424.3537 | 51.6647 | 51.6647 | 424.2862 | 367.3667 | 74.7644 | 74.8319 | TRIPLET HOUSING | 6.2858 | FE | SCS2 | PM | 752 | FE_SCS2:PM_D0752 | PM1 | profile monitor (large fork) | 150 |  |  | 0 | 366.8549 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 366.8549 | 75.2335 | 75.2335 | and 700 l/s turbo pump
FE | SCS1 | QHE | 749 | FE_SCS1:QHE_D0749 | Q7 | ES_quad_X | 150 |  |  | 0.2 | 367.3667 | 424.1862 | 51.6647 | 51.6647 | 424.0862 | 367.3667 | 74.9319 | 75.0319 | TRIPLET HOUSING | 6.4533 | FE | SCS2 | TMP | 752 | FE_SCS2:TMP_D0752 | TMP_M | turbo pump, 700 l/s | 150 |  |  | 0 | 366.8549 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 366.8549 | 75.2335 | 75.2335
[6 empty] | bellow | 150 |  |  | 0.14959 | 367.3667 | 424.0114 | 51.6647 | 51.6647 | 423.9366 | 367.3667 | 75.1067 | 75.1815 |  | 6.6281 | [7 empty] | 150 |  | 0.10475 | 0.12327 | 366.9166 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 366.9782 | 75.2951 | 75.3567
FE | SCS1 | GV | 752 | FE_SCS1:GV_D0752 | GV_L | gate valve | 150 |  |  | 0.080264 | 367.3667 | 423.8965 | 51.6647 | 51.6647 | 423.8564 | 367.3667 | 75.2216 | 75.2618 |  | 6.743 | [7 empty] | 150 |  |  | 0 | 366.9782 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 366.9782 | 75.3567 | 75.3567
[7 empty] | 150 |  |  | 0.12014 | 367.3667 | 423.7963 | 51.6647 | 51.6647 | 423.7362 | 367.3667 | 75.3218 | 75.3819 | dimension to yoke not flange | 6.8432 | FE | SCS2 | DCHV | 755 | FE_SCS2:DCHV_D0755 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 367.0734 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 367.1687 | 75.452 | 75.5472
FE | LEBT | DH | 759 | FE_LEBT:DH_D0759 | D5 | dipole-90deg, D5-2 | 120 x 150 |  |  | 1 | 367.685 | 423.4179 | 51.6647 | 51.6647 | 423.0996 | 368.0033 | 75.8819 | 76.3819 |  | 7.4033 | [9 empty] | 0.0067937 | 0 | 367.1687 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 367.1687 | 75.5472 | 75.5472
[7 empty] | 150 |  |  | 0.225 | 368.1158 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 368.2283 | 76.4944 | 76.6069 | dimension to yoke not flange | 8.0158 | FE | SCS2 | GV | 756 | FE_SCS2:GV_D0756 | GV_L | gate valve | 150 |  |  | 0.080264 | 367.2088 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 367.249 | 75.5874 | 75.6275
FE | LEBT | QVE | 767 | FE_LEBT:QVE_D0767 | Q7 | ES_quad_Y | 150 |  |  | 0.2 | 368.3283 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 368.4283 | 76.7069 | 76.8069 | DOUBLET HOUSING | 8.2283 | [6 empty] | reducer / tie-in to LEBT | 150 |  |  | 0.75438 | 367.6262 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 368.0033 | 76.0047 | 76.3819
[7 empty] | 150 |  |  | 0.135 | 368.4958 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 368.5633 | 76.8744 | 76.9419 | DOUBLET HOUSING | 8.3958 | [14 empty] | 51.6647 | 423.0996 | 368.0033
FE | LEBT | QHE | 770 | FE_LEBT:QHE_D0770 | Q7 | ES_quad_X | 150 |  |  | 0.2 | 368.6633 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 368.7633 | 77.0419 | 77.1419 | DOUBLET HOUSING | 8.5633
[7 empty] | 150 |  |  | 0.10225 | 368.8145 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 368.8656 | 77.193 | 77.2441 |  | 8.7144
[7 empty] | 150 |  |  | 0 | 368.8656 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 368.8656 | 77.2441 | 77.2441 |  | 8.7655
FE | LEBT | DCHV | 773 | FE_LEBT:DCHV_D0773 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 368.9608 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 369.0561 | 77.3394 | 77.4346 |  | 8.8608
[7 empty] | 150 |  |  | 0.10225 | 369.1072 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 369.1583 | 77.4858 | 77.5369 |  | 9.0071
FE | LEBT | QHE | 776 | FE_LEBT:QHE_D0776 | Q7 | ES_quad_X | 150 |  |  | 0.2 | 369.2583 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 369.3583 | 77.6369 | 77.7369 | DOUBLET HOUSING | 9.1583
[7 empty] | 150 |  |  | 0.135 | 369.4258 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 369.4933 | 77.8044 | 77.8719 | DOUBLET HOUSING | 9.3258
FE | LEBT | QVE | 780 | FE_LEBT:QVE_D0780 | Q7 | ES_quad_Y | 150 |  |  | 0.2 | 369.5933 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 369.6933 | 77.9719 | 78.0719 | DOUBLET HOUSING | 9.4933
[7 empty] | 150 |  |  | 0.19159 | 369.7891 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 369.8849 | 78.1677 | 78.2635 |  | 9.6891
FE | LEBT | PM | 783 | FE_LEBT:PM_D0783 | PM1 | profile monitor (large fork) | 150 |  |  | 0 | 369.8849 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 369.8849 | 78.2635 | 78.2635 |  | 9.7849
FE | LEBT | TMP | 783 | FE_LEBT:TMP_D0783 | TMP_M | pump, 700 l/s |  |  |  | 0 | 369.8849 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 369.8849 | 78.2635 | 78.2635 |  | 9.7849
[7 empty] | 150 |  |  | 0.21841 | 369.9941 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 370.1033 | 78.3727 | 78.4819 |  | 9.8941
FE | LEBT | SOLR | 787 | FE_LEBT:SOLR_D0787 | S4 | solenoid | 150 |  |  | 0.4 | 370.3033 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 370.5033 | 78.6819 | 78.8819 | 40cm effective length | 10.2033
[7 empty] | 150 |  |  | 0.0254 | 370.516 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 370.5287 | 78.8946 | 78.9073 |  | 10.416
FE | LEBT | DCHV | 790 | FE_LEBT:DCHV_D0790 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 370.624 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 370.7192 | 79.0025 | 79.0978 |  | 10.5239
[7 empty] | 150 |  |  | 0.060413 | 370.7495 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 370.7797 | 79.128 | 79.1582 |  | 10.6494
FE | LEBT | CHP | 792 | FE_LEBT:CHP_D0792 |  | chopper | 70 x 150 |  |  | 0.1 | 370.8297 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 370.8797 | 79.2082 | 79.2582 | 10cm deflector plates | 10.7296
[6 empty] | box+tube | 150 |  |  | 0.20349 | 370.9814 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 371.0831 | 79.36 | 79.4617 |  | 10.8813
[7 empty] | 150 |  |  | 0 | 371.0831 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 371.0831 | 79.4617 | 79.4617 |  | 10.9831
FE | LEBT | NEGP | 795 | FE_LEBT:NEGP_D0795 | NEG-0 | getter pump, 3500 l/s |  |  |  | 0 | 371.0831 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 371.0831 | 79.4617 | 79.4617 |  | 10.9831
[7 empty] | 150 |  |  | 0.11717 | 371.1417 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 371.2003 | 79.5203 | 79.5789 |  | 11.0417
FE | LEBT | AP | 796 | FE_LEBT:AP_D0796 |  | apertures | 150 |  |  | 0 | 371.2003 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 371.2003 | 79.5789 | 79.5789 |  | 11.1002
[6 empty] | bellow+tube/box | 150 |  |  | 0.025375 | 371.213 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 371.2257 | 79.5916 | 79.6042 |  | 11.1129
FE | LEBT | FC | 796 | FE_LEBT:FC_D0796 | FC1 | Faraday cup | 150 |  |  | 0 | 371.2257 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 371.2257 | 79.6042 | 79.6042 |  | 11.1256
[7 empty] | 150 |  |  | 0.3528 | 371.4021 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 371.5785 | 79.7806 | 79.957 |  | 11.302
FE | LEBT | SOLR | 802 | FE_LEBT:SOLR_D0802 | S4 | solenoid | 150 |  |  | 0.4 | 371.7785 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 371.9785 | 80.157 | 80.357 | 40cm effective length | 11.6784
[7 empty] | 150 |  |  | 0.0254 | 371.9912 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 372.0039 | 80.3697 | 80.3824 |  | 11.8911
FE | LEBT | DCHV | 805 | FE_LEBT:DCHV_D0805 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 372.0991 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 372.1944 | 80.4777 | 80.5729 |  | 11.9991
[7 empty] | 150 |  |  | 0.16071 | 372.2748 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 372.3551 | 80.6533 | 80.7337 |  | 12.1747
FE | LEBT | AP | 807 | FE_LEBT:AP_D0807 |  | apertures | 150 |  |  | 0 | 372.3551 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 372.3551 | 80.7337 | 80.7337 |  | 12.255
[7 empty] | 150 |  |  | 0.042926 | 372.3766 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 372.398 | 80.7551 | 80.7766 |  | 12.2765
FE | LEBT | PM | 808 | FE_LEBT:PM_D0808 | PM1 | profile monitor (large fork) | 150 |  |  | 0 | 372.398 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 372.398 | 80.7766 | 80.7766 |  | 12.298
[7 empty] | 150 |  |  | 0.3954 | 372.5957 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 372.7934 | 80.9743 | 81.172 |  | 12.4957
FE | LEBT | EMS | 812 | FE_LEBT:EMS_D0812 | EM2 | Pepper Pot | 150 |  |  | 0 | 372.7934 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 372.7934 | 81.172 | 81.172 |  | 12.6934
FE | LEBT | VD | 812 | FE_LEBT:VD_D0812 | V1 | viewer | 150 |  |  | 0 | 372.7934 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 372.7934 | 81.172 | 81.172 |  | 12.6934
[7 empty] | 150 |  |  | 0.20828 | 372.8976 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 373.0017 | 81.2761 | 81.3803 |  | 12.7975
FE | LEBT | NEGP | 814 | FE_LEBT:NEGP_D0814 | NEG-0 | getter pump, 3500 l/s |  |  |  | 0 | 373.0017 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 373.0017 | 81.3803 | 81.3803 |  | 12.9016
FE | LEBT | FC | 814 | FE_LEBT:FC_D0814 | FC1 | Faraday cup | 150 |  |  | 0 | 373.0017 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 373.0017 | 81.3803 | 81.3803 |  | 12.9016
[7 empty] | 150 |  |  | 0.21693 | 373.1102 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 373.2186 | 81.4887 | 81.5972 | 40cm effective length | 13.0101
FE | LEBT | SOLR | 818 | FE_LEBT:SOLR_D0818 | S4 | solenoid | 150 |  |  | 0.4 | 373.4186 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 373.6186 | 81.7972 | 81.9972 |  | 13.3186
[7 empty] | 150 |  |  | 0.0254 | 373.6313 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 373.644 | 82.0099 | 82.0226 |  | 13.5313
FE | LEBT | DCHV | 821 | FE_LEBT:DCHV_D0821 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 373.7393 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 373.8345 | 82.1178 | 82.2131 |  | 13.6392
[7 empty] | 150 |  |  | 0.13973 | 373.9044 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 373.9743 | 82.283 | 82.3528 |  | 13.8043
FE | LEBT | BD | 824 | FE_LEBT:BD_D0824 | dump | beam plug (low power) |  |  |  | 0 | 373.9743 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 373.9743 | 82.3528 | 82.3528 |  | 13.8742
[7 empty] | 150 |  |  | 0.038798 | 373.9937 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 374.0131 | 82.3722 | 82.3916 |  | 13.8936
FE | LEBT | PM | 824 | FE_LEBT:PM_D0824 | PM1 | profile monitor (large fork) | 150 |  |  | 0 | 374.0131 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 374.0131 | 82.3916 | 82.3916 |  | 13.913
[7 empty] | 150 |  |  | 0.2018 | 374.114 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 374.2149 | 82.4925 | 82.5934 |  | 14.0139
FE | LEBT | GV | 826 | FE_LEBT:GV_D0826 | GV_L | gate valve | 150 |  |  | 0.080264 | 374.255 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 374.2951 | 82.6336 | 82.6737 |  | 14.1549
[7 empty] | 150 |  |  | 0.12351 | 374.3569 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 374.4186 | 82.7354 | 82.7972 | dimension to yoke not flange | 14.2568
FE | LEBT | DVE | 833 | FE_LEBT:DVE_D0833 | D7 | ES_bender 1 | 100 x 150 |  |  | 1 | 374.737 | 423.0996 | 51.3464 | 51.0281 | 423.0996 | 375.0553 | 83.2972 | 83.7972 |  | 14.8186
FE | LEBT | TMP | 833 | FE_LEBT:TMP_D0833 | TMP_M | pump, 700 l/s |  |  |  | 0 | 375.0553 | 423.0996 | 51.0281 | 51.0281 | 423.0996 | 375.0553 | 83.7972 | 83.7972 | special: pump port on the side of e-bender | 15.3186
[7 empty] | 150 |  |  | 0 | 375.0553 | 423.0996 | 50.9682 | 50.9084 | 423.0996 | 375.0553 | 83.7972 | 83.7972 |  | 15.3186
[7 empty] | 150 |  |  | 0.11975 | 375.0553 | 423.0996 | 50.9084 | 50.9084 | 423.0996 | 375.0553 | 83.8571 | 83.9169 | dimension to yoke not flange | 15.3784
FE | LEBT | DCHV | 840 | FE_LEBT:DCHV_D0840 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 375.0553 | 423.0996 | 50.8131 | 50.7179 | 423.0996 | 375.0553 | 84.0122 | 84.1074 |  | 15.5336
[7 empty] | 150 |  |  | 0.23975 | 375.0553 | 423.0996 | 50.598 | 50.4781 | 423.0996 | 375.0553 | 84.2273 | 84.3472 |  | 15.7487
FE | LEBT | QHE | 844 | FE_LEBT:QHE_D0844 | Q7 | ES_quad_X | 150 |  |  | 0.2 | 375.0553 | 423.0996 | 50.3781 | 50.2781 | 423.0996 | 375.0553 | 84.4472 | 84.5472 | TRIPLET HOUSING | 15.9686
[7 empty] | 150 |  |  | 0.135 | 375.0553 | 423.0996 | 50.2106 | 50.1431 | 423.0996 | 375.0553 | 84.6147 | 84.6822 | TRIPLET HOUSING | 16.1361
FE | LEBT | QVE | 848 | FE_LEBT:QVE_D0848 | Q7 | ES_quad_Y | 150 |  |  | 0.2 | 375.0553 | 423.0996 | 50.0431 | 49.9431 | 423.0996 | 375.0553 | 84.7822 | 84.8822 | TRIPLET HOUSING | 16.3036
[7 empty] | 150 |  |  | 0.135 | 375.0553 | 423.0996 | 49.8756 | 49.8081 | 423.0996 | 375.0553 | 84.9497 | 85.0172 | TRIPLET HOUSING | 16.4711
FE | LEBT | QHE | 851 | FE_LEBT:QHE_D0851 | Q7 | ES_quad_X | 150 |  |  | 0.2 | 375.0553 | 423.0996 | 49.7081 | 49.6081 | 423.0996 | 375.0553 | 85.1172 | 85.2172 | TRIPLET HOUSING | 16.6386
[6 empty] | bellow ? | 150 |  |  | 0.34302 | 375.0553 | 423.0996 | 49.4366 | 49.2651 | 423.0996 | 375.0553 | 85.3887 | 85.5602 |  | 16.9101
FE | LEBT | PM | 856 | FE_LEBT:PM_D0856 | PM1 | profile monitor (large fork) | 150 |  |  | 0 | 375.0553 | 423.0996 | 49.2651 | 49.2651 | 423.0996 | 375.0553 | 85.5602 | 85.5602 |  | 17.0816
[6 empty] | bellow+tube | 150 |  |  | 0.73973 | 375.0553 | 423.0996 | 48.8952 | 48.5254 | 423.0996 | 375.0553 | 85.9301 | 86.2999 |  | 17.4514
[6 empty] | 4 way cross | 150 |  |  | 0 | 375.0553 | 423.0996 | 48.5254 | 48.5254 | 423.0996 | 375.0553 | 86.2999 | 86.2999 |  | 17.8213
FE | LEBT | NEGP | 863 | FE_LEBT:NEGP_D0863 | NEG-0 | getter pump, 3500 l/s |  |  |  | 0 | 375.0553 | 423.0996 | 48.5254 | 48.5254 | 423.0996 | 375.0553 | 86.2999 | 86.2999 |  | 17.8213
[7 empty] | 150 |  |  | 0.42075 | 375.0553 | 423.0996 | 48.315 | 48.1046 | 423.0996 | 375.0553 | 86.5103 | 86.7207 |  | 18.0317
FE | LEBT | DCHV | 868 | FE_LEBT:DCHV_D0868 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 375.0553 | 423.0996 | 48.0094 | 47.9141 | 423.0996 | 375.0553 | 86.8159 | 86.9112 |  | 18.3373
[7 empty] | 150 |  |  | 0.08475 | 375.0553 | 423.0996 | 47.8717 | 47.8294 | 423.0996 | 375.0553 | 86.9536 | 86.9959 |  | 18.4749
FE | LEBT | QVE | 871 | FE_LEBT:QVE_D0871 | Q7 | ES_quad_Y | 150 |  |  | 0.2 | 375.0553 | 423.0996 | 47.7294 | 47.6294 | 423.0996 | 375.0553 | 87.0959 | 87.1959 | TRIPLET HOUSING | 18.6173
[7 empty] | 150 |  |  | 0.135 | 375.0553 | 423.0996 | 47.5619 | 47.4944 | 423.0996 | 375.0553 | 87.2634 | 87.3309 | TRIPLET HOUSING | 18.7848
FE | LEBT | QHE | 874 | FE_LEBT:QHE_D0874 | Q7 | ES_quad_X | 150 |  |  | 0.2 | 375.0553 | 423.0996 | 47.3944 | 47.2944 | 423.0996 | 375.0553 | 87.4309 | 87.5309 | TRIPLET HOUSING | 18.9523
[7 empty] | 150 |  |  | 0.135 | 375.0553 | 423.0996 | 47.2269 | 47.1594 | 423.0996 | 375.0553 | 87.5984 | 87.6659 | TRIPLET HOUSING | 19.1198
FE | LEBT | QVE | 878 | FE_LEBT:QVE_D0878 | Q7 | ES_quad_Y | 150 |  |  | 0.2 | 375.0553 | 423.0996 | 47.0594 | 46.9594 | 423.0996 | 375.0553 | 87.7659 | 87.8659 | TRIPLET HOUSING | 19.2873
[7 empty] | 150 |  |  | 0.08475 | 375.0553 | 423.0996 | 46.917 | 46.8746 | 423.0996 | 375.0553 | 87.9083 | 87.9507 |  | 19.4297
FE | LEBT | DCHV | 880 | FE_LEBT:DCHV_D0880 | C4 | corrector+bellow+tube | 150 |  |  | 0.1905 | 375.0553 | 423.0996 | 46.7794 | 46.6841 | 423.0996 | 375.0553 | 88.0459 | 88.1412 |  | 19.5673
[7 empty] | 150 |  |  | 0.37948 | 375.0553 | 423.0996 | 46.4944 | 46.3046 | 423.0996 | 375.0553 | 88.3309 | 88.5207 |  | 19.8523
FE | LEBT | PM | 885 | FE_LEBT:PM_D0885 | PM1 | profile monitor (large fork) | 150 |  |  | 0 | 375.0553 | 423.0996 | 46.3046 | 46.3046 | 423.0996 | 375.0553 | 88.5207 | 88.5207 |  | 20.042
[7 empty] | 150 |  |  | 0.058903 | 375.0553 | 423.0996 | 46.2752 | 46.2457 | 423.0996 | 375.0553 | 88.5501 | 88.5796 |  | 20.0715
[6 empty] | 4 way cross ?? | 150 |  |  | 0 | 375.0553 | 423.0996 | 46.2457 | 46.2457 | 423.0996 | 375.0553 | 88.5796 | 88.5796 |  | 20.1009
FE | LEBT | NEGP | 886 | FE_LEBT:NEGP_D0886 | NEG-0 | getter pump, 3500 l/s |  |  |  | 0 | 375.0553 | 423.0996 | 46.2457 | 46.2457 | 423.0996 | 375.0553 | 88.5796 | 88.5796 |  | 20.1009
[6 empty] | tube | 150 |  |  | 0.44637 | 375.0553 | 423.0996 | 46.0226 | 45.7994 | 423.0996 | 375.0553 | 88.8028 | 89.0259 |  | 20.3241
FE | LEBT | QVE | 891 | FE_LEBT:QVE_D0891 | Q7 | ES_quad_Y | 150 |  |  | 0.2 | 375.0553 | 423.0996 | 45.6994 | 45.5994 | 423.0996 | 375.0553 | 89.1259 | 89.2259 | TRIPLET HOUSING | 20.6473
[7 empty] | 150 |  |  | 0.135 | 375.0553 | 423.0996 | 45.5319 | 45.4644 | 423.0996 | 375.0553 | 89.2934 | 89.3609 | TRIPLET HOUSING | 20.8148
FE | LEBT | QHE | 895 | FE_LEBT:QHE_D0895 | Q7 | ES_quad_X | 150 |  |  | 0.2 | 375.0553 | 423.0996 | 45.3644 | 45.2644 | 423.0996 | 375.0553 | 89.4609 | 89.5609 | TRIPLET HOUSING | 20.9823
[7 empty] | 150 |  |  | 0.135 | 375.0553 | 423.0996 | 45.1969 | 45.1294 | 423.0996 | 375.0553 | 89.6284 | 89.6959 | TRIPLET HOUSING | 21.1498
FE | LEBT | QVE | 898 | FE_LEBT:QVE_D0898 | Q7 | ES_quad_Y | 150 |  |  | 0.2 | 375.0553 | 423.0996 | 45.0294 | 44.9294 | 423.0996 | 375.0553 | 89.7959 | 89.8959 | TRIPLET HOUSING | 21.3173
[7 empty] | 150 |  |  | 0.08475 | 375.0553 | 423.0996 | 44.887 | 44.8446 | 423.0996 | 375.0553 | 89.9383 | 89.9807 |  | 21.4597
FE | LEBT | DCHV | 901 | FE_LEBT:DCHV_D0901 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 375.0553 | 423.0996 | 44.7494 | 44.6541 | 423.0996 | 375.0553 | 90.0759 | 90.1712 |  | 21.5973
[7 empty] | 150 |  |  | 0.42077 | 375.0553 | 423.0996 | 44.4437 | 44.2334 | 423.0996 | 375.0553 | 90.3816 | 90.592 |  | 21.903
[6 empty] | 4 way cross ?? | 150 |  |  | 0 | 375.0553 | 423.0996 | 44.2334 | 44.2334 | 423.0996 | 375.0553 | 90.592 | 90.592 |  | 22.1133
FE | LEBT | NEGP | 906 | FE_LEBT:NEGP_D0906 | NEG-0 | getter pump, 3500 l/s |  |  |  | 0 | 375.0553 | 423.0996 | 44.2334 | 44.2334 | 423.0996 | 375.0553 | 90.592 | 90.592 |  | 22.1133
[6 empty] | tube | 150 |  |  | 0.57746 | 375.0553 | 423.0996 | 43.9446 | 43.6559 | 423.0996 | 375.0553 | 90.8807 | 91.1694 |  | 22.4021
[6 empty] | 4 way cross | 150 |  |  | 0 | 375.0553 | 423.0996 | 43.6559 | 43.6559 | 423.0996 | 375.0553 | 91.1694 | 91.1694 |  | 22.6908
[7 empty] | 150 |  |  | 0 | 375.0553 | 423.0996 | 43.6559 | 43.6559 | 423.0996 | 375.0553 | 91.1694 | 91.1694 |  | 22.6908
FE | LEBT | PM | 912 | FE_LEBT:PM_D0912 | PM1 | profile monitor (large fork) | 150 |  |  | 0 | 375.0553 | 423.0996 | 43.6559 | 43.6559 | 423.0996 | 375.0553 | 91.1694 | 91.1694 |  | 22.6908
[6 empty] | bellow | 150 |  |  | 0.50527 | 375.0553 | 423.0996 | 43.4033 | 43.1506 | 423.0996 | 375.0553 | 91.4221 | 91.6747 |  | 22.9434
FE | LEBT | QHE | 918 | FE_LEBT:QHE_D0918 | Q7 | ES_quad_X | 150 |  |  | 0.2 | 375.0553 | 423.0996 | 43.0506 | 42.9506 | 423.0996 | 375.0553 | 91.7747 | 91.8747 | TRIPLET HOUSING | 23.2961
[7 empty] | 150 |  |  | 0.135 | 375.0553 | 423.0996 | 42.8831 | 42.8156 | 423.0996 | 375.0553 | 91.9422 | 92.0097 | TRIPLET HOUSING | 23.4636
FE | LEBT | QVE | 921 | FE_LEBT:QVE_D0921 | Q7 | ES_quad_Y | 150 |  |  | 0.2 | 375.0553 | 423.0996 | 42.7156 | 42.6156 | 423.0996 | 375.0553 | 92.1097 | 92.2097 | TRIPLET HOUSING | 23.6311
[7 empty] | 150 |  |  | 0.135 | 375.0553 | 423.0996 | 42.5481 | 42.4806 | 423.0996 | 375.0553 | 92.2772 | 92.3447 | TRIPLET HOUSING | 23.7986
FE | LEBT | QHE | 924 | FE_LEBT:QHE_D0924 | Q7 | ES_quad_X | 150 |  |  | 0.2 | 375.0553 | 423.0996 | 42.3806 | 42.2806 | 423.0996 | 375.0553 | 92.4447 | 92.5447 | TRIPLET HOUSING | 23.9661
[7 empty] | 150 |  |  | 0.23975 | 375.0553 | 423.0996 | 42.1607 | 42.0409 | 423.0996 | 375.0553 | 92.6646 | 92.7844 |  | 24.1859
FE | LEBT | DCHV | 929 | FE_LEBT:DCHV_D0929 | C4 | corrector+bellow+tube | 150 |  |  | 0.1905 | 375.0553 | 423.0996 | 41.9456 | 41.8504 | 423.0996 | 375.0553 | 92.8797 | 92.9749 |  | 24.4011
[7 empty] | 150 |  |  | 0.11975 | 375.0553 | 423.0996 | 41.7905 | 41.7306 | 423.0996 | 375.0553 | 93.0348 | 93.0947 | dimension to yoke not flange | 24.5562
FE | LEBT | DVE | 936 | FE_LEBT:DVE_D0936 | D7 | ES_bender 2 | 100 x 150 |  |  | 1 | 374.737 | 423.0996 | 41.4123 | 41.094 | 423.0996 | 374.4186 | 93.5947 | 94.0947 |  | 25.1161
FE | LEBT | NEGP | 936 | FE_LEBT:NEGP_D0936 | NEG-0 | getter pump, 3500 l/s |  |  |  | 0 | 374.737 | 423.0996 | 41.094 | 41.094 | 423.0996 | 375.0553 | 94.0947 | 94.0947 | special: pump port on the side of e-bender | 25.6161
[7 empty] | 150 |  |  | 3e-06 | 374.737 | 423.0996 | 41.094 | 41.094 | 423.0996 | 374.4186 | 94.0947 | 94.0947 | adjustment for residuals | 25.6161
[7 empty] | 150 |  |  | 0.12351 | 374.3569 | 423.0996 | 41.094 | 41.094 | 423.0996 | 374.2951 | 94.1565 | 94.2182 | dimension to yoke not flange | 25.6778
FE | LEBT | GV | 943 | FE_LEBT:GV_D0943 | GV_L | gate valve | 150 |  |  | 0.080264 | 374.255 | 423.0996 | 41.094 | 41.094 | 423.0996 | 374.2149 | 94.2583 | 94.2985 |  | 25.7797
[7 empty] | 150 |  |  | 0.2018 | 374.114 | 423.0996 | 41.094 | 41.094 | 423.0996 | 374.0131 | 94.3994 | 94.5003 |  | 25.9207
FE | LEBT | PM | 945 | FE_LEBT:PM_D0945 | PM1 | profile monitor (large fork) | 150 |  |  | 0 | 374.0131 | 423.0996 | 41.094 | 41.094 | 423.0996 | 374.0131 | 94.5003 | 94.5003 |  | 26.0216
[7 empty] | 150 |  |  | 0.17853 | 373.9238 | 423.0996 | 41.094 | 41.094 | 423.0996 | 373.8345 | 94.5895 | 94.6788 |  | 26.1109
FE | LEBT | DCHV | 948 | FE_LEBT:DCHV_D0948 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 373.7393 | 423.0996 | 41.094 | 41.094 | 423.0996 | 373.644 | 94.774 | 94.8693 |  | 26.2954
[7 empty] | 150 |  |  | 0.0254 | 373.6313 | 423.0996 | 41.094 | 41.094 | 423.0996 | 373.6186 | 94.882 | 94.8947 |  | 26.4034
FE | LEBT | SOLR | 951 | FE_LEBT:SOLR_D0951 | S4 | solenoid | 150 |  |  | 0.4 | 373.4186 | 423.0996 | 41.094 | 41.094 | 423.0996 | 373.2186 | 95.0947 | 95.2947 | 40cm effective length | 26.6161
[6 empty] | bellow+tube | 150 |  |  | 0.21693 | 373.1102 | 423.0996 | 41.094 | 41.094 | 423.0996 | 373.0017 | 95.4032 | 95.5116 |  | 26.9245
FE | LEBT | NEGP | 955 | FE_LEBT:NEGP_D0955 | NEG-0 | getter pump, 3500 l/s |  |  |  | 0 | 373.0017 | 423.0996 | 41.094 | 41.094 | 423.0996 | 373.0017 | 95.5116 | 95.5116 |  | 27.033
[7 empty] | 150 |  |  | 0.18546 | 372.909 | 423.0996 | 41.094 | 41.094 | 423.0996 | 372.8163 | 95.6044 | 95.6971 |  | 27.1257
FE | LEBT | ATT | 957 | FE_LEBT:ATT_D0957 | ATT1 | beam attenuator | 150 |  |  | 0 | 372.8163 | 423.0996 | 41.094 | 41.094 | 423.0996 | 372.8163 | 95.6971 | 95.6971 |  | 27.2185
[7 empty] | 150 |  |  | 0.073025 | 372.7797 | 423.0996 | 41.094 | 41.094 | 423.0996 | 372.7432 | 95.7336 | 95.7701 |  | 27.255
FE | LEBT | ATT | 958 | FE_LEBT:ATT_D0958 | ATT1 | beam attenuator | 150 |  |  | 0 | 372.7432 | 423.0996 | 41.094 | 41.094 | 423.0996 | 372.7432 | 95.7701 | 95.7701 |  | 27.2915
[7 empty] | 150 |  |  | 0.34527 | 372.5706 | 423.0996 | 41.094 | 41.094 | 423.0996 | 372.398 | 95.9427 | 96.1154 |  | 27.4641
FE | LEBT | PM | 961 | FE_LEBT:PM_D0961 | PM1 | profile monitor (large fork) | 150 |  |  | 0 | 372.398 | 423.0996 | 41.094 | 41.094 | 423.0996 | 372.398 | 96.1154 | 96.1154 |  | 27.6368
[7 empty] | 150 |  |  | 0.20357 | 372.2962 | 423.0996 | 41.094 | 41.094 | 423.0996 | 372.1944 | 96.2172 | 96.3189 |  | 27.7385
FE | LEBT | DCHV | 964 | FE_LEBT:DCHV_D0964 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 372.0991 | 423.0996 | 41.094 | 41.094 | 423.0996 | 372.0039 | 96.4142 | 96.5094 |  | 27.9356
[7 empty] | 150 |  |  | 0.0254 | 371.9912 | 423.0996 | 41.094 | 41.094 | 423.0996 | 371.9785 | 96.5221 | 96.5348 |  | 28.0435
[7 empty] | 150 |  |  | 0 | 371.9785 | 423.0996 | 41.094 | 41.094 | 423.0996 | 371.9785 | 96.5348 | 96.5348 |  | 28.0562
FE | LEBT | SOLR | 967 | FE_LEBT:SOLR_D0967 | S4 | solenoid | 150 |  |  | 0.4 | 371.7785 | 423.0996 | 41.094 | 41.094 | 423.0996 | 371.5785 | 96.7348 | 96.9348 | 40cm effective length | 28.2562
[6 empty] | bellow | 150 |  |  | 0.21855 | 371.4692 | 423.0996 | 41.094 | 41.094 | 423.0996 | 371.3599 | 97.0441 | 97.1534 |  | 28.5655
FE | LEBT | PM | 972 | FE_LEBT:PM_D0972 | PM1 | profile monitor (large fork) | 150 |  |  | 0 | 371.3599 | 423.0996 | 41.094 | 41.094 | 423.0996 | 371.3599 | 97.1534 | 97.1534 | no longer has aperture | 28.6748
[7 empty] | 150 |  |  | 0.28661 | 371.2166 | 423.0996 | 41.094 | 41.094 | 423.0996 | 371.0733 | 97.2967 | 97.44 |  | 28.8181
FE | LEBT | ATT | 974 | FE_LEBT:ATT_D0974 | ATT1 | beam attenuator | 150 |  |  | 0 | 371.0733 | 423.0996 | 41.094 | 41.094 | 423.0996 | 371.0733 | 97.44 | 97.44 |  | 28.9614
[7 empty] | 150 |  |  | 0.073025 | 371.0368 | 423.0996 | 41.094 | 41.094 | 423.0996 | 371.0003 | 97.4765 | 97.513 |  | 28.9979
FE | LEBT | ATT | 975 | FE_LEBT:ATT_D0975 | ATT1 | beam attenuator | 150 |  |  | 0 | 371.0003 | 423.0996 | 41.094 | 41.094 | 423.0996 | 371.0003 | 97.513 | 97.513 |  | 29.0344
[7 empty] | 150 |  |  | 0.1563 | 370.9222 | 423.0996 | 41.094 | 41.094 | 423.0996 | 370.844 | 97.5912 | 97.6693 |  | 29.1126
FE | LEBT | VD | 977 | FE_LEBT:VD_D0977 | V2 | viewer | 150 |  |  | 0 | 370.844 | 423.0996 | 41.094 | 41.094 | 423.0996 | 370.844 | 97.6693 | 97.6693 |  | 29.1907
[7 empty] | 150 |  |  | 0 | 370.844 | 423.0996 | 41.094 | 41.094 | 423.0996 | 370.844 | 97.6693 | 97.6693 |  | 29.1907
FE | LEBT | IP | 977 | FE_LEBT:IP_D0977 | IP_L | ion pump, 300 L/s | 150 |  |  | 0 | 370.844 | 423.0996 | 41.094 | 41.094 | 423.0996 | 370.844 | 97.6693 | 97.6693 |  | 29.1907
[7 empty] | 150 |  |  | 0 | 370.844 | 423.0996 | 41.094 | 41.094 | 423.0996 | 370.844 | 97.6693 | 97.6693 |  | 29.1907
FE | LEBT | FC | 977 | FE_LEBT:FC_D0977 | FC1 | faraday cup | 150 |  |  | 0 | 370.844 | 423.0996 | 41.094 | 41.094 | 423.0996 | 370.844 | 97.6693 | 97.6693 |  | 29.1907
[7 empty] | 150 |  |  | 0.12477 | 370.7816 | 423.0996 | 41.094 | 41.094 | 423.0996 | 370.7192 | 97.7317 | 97.7941 |  | 29.2531
FE | LEBT | DCHV | 979 | FE_LEBT:DCHV_D0979 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 370.624 | 423.0996 | 41.094 | 41.094 | 423.0996 | 370.5287 | 97.8893 | 97.9846 |  | 29.4107
[7 empty] | 150 |  |  | 0.0254 | 370.516 | 423.0996 | 41.094 | 41.094 | 423.0996 | 370.5033 | 97.9973 | 98.01 |  | 29.5187
FE | LEBT | SOLR | 982 | FE_LEBT:SOLR_D0982 | S4 | solenoid | 150 |  |  | 0.4 | 370.3033 | 423.0996 | 41.094 | 41.094 | 423.0996 | 370.1033 | 98.21 | 98.41 | 40cm effective length | 29.7314
[6 empty] | mhb box & bellows | 150 |  |  | 0.22 | 369.9933 | 423.0996 | 41.094 | 41.094 | 423.0996 | 369.8833 | 98.52 | 98.63 |  | 30.0414
[7 empty] | 150 |  |  | 0 | 369.8833 | 423.0996 | 41.094 | 41.094 | 423.0996 | 369.8833 | 98.63 | 98.63 |  | 30.1514
FE | LEBT | PM | 986 | FE_LEBT:PM_D0986 | PM1 | profile monitor (small fork) | 150 |  |  | 0 | 369.8833 | 423.0996 | 41.094 | 41.094 | 423.0996 | 369.8833 | 98.63 | 98.63 |  | 30.1514
[6 empty] | mhb box | 150 |  |  | 0.105 | 369.8308 | 423.0996 | 41.094 | 41.094 | 423.0996 | 369.7783 | 98.6825 | 98.735 |  | 30.2039
FE | LEBT | CAV | 987 | FE_LEBT:CAV_D0987 |  | multi harmonic buncher | 150 |  |  | 0 | 369.7783 | 423.0996 | 41.094 | 41.094 | 423.0996 | 369.7783 | 98.735 | 98.735 |  | 30.2564
[7 empty] | 150 |  |  | 0 | 369.7783 | 423.0996 | 41.094 | 41.094 | 423.0996 | 369.7783 | 98.735 | 98.735 |  | 30.2564
[6 empty] | mhb box | 150 |  |  | 0 | 369.7783 | 423.0996 | 41.094 | 41.094 | 423.0996 | 369.7783 | 98.735 | 98.735 |  | 30.2564
FE | LEBT | NEGP | 987 | FE_LEBT:NEGP_D0987 | NEG-0 | getter pump, 3500 l/s |  |  |  | 0 | 369.7783 | 423.0996 | 41.094 | 41.094 | 423.0996 | 369.7783 | 98.735 | 98.735 |  | 30.2564
[7 empty] | 150 |  |  | 0 | 369.7783 | 423.0996 | 41.094 | 41.094 | 423.0996 | 369.7783 | 98.735 | 98.735 |  | 30.2564
[7 empty] | 150 |  |  | 0 | 369.7783 | 423.0996 | 41.094 | 41.094 | 423.0996 | 369.7783 | 98.735 | 98.735 | Possible HV tube starts here | 30.2564
[7 empty] | 150 |  |  | 0.14649 | 369.7051 | 423.0996 | 41.094 | 41.094 | 423.0996 | 369.6318 | 98.8082 | 98.8815 |  | 30.3296
FE | LEBT | BCM | 989 | FE_LEBT:BCM_D0989 | BCM1 | AC current transformer | 150 | 50mm |  | 0 | 369.6318 | 423.0996 | 41.094 | 41.094 | 423.0996 | 369.6318 | 98.8815 | 98.8815 |  | 30.4029
[7 empty] | 150 |  |  | 0.22261 | 369.5205 | 423.0996 | 41.094 | 41.094 | 423.0996 | 369.4092 | 98.9928 | 99.1041 |  | 30.5142
FE | LEBT | DCHV | 992 | FE_LEBT:DCHV_D0992 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 369.314 | 423.0996 | 41.094 | 41.094 | 423.0996 | 369.2187 | 99.1993 | 99.2946 |  | 30.7207
[7 empty] | 150 |  |  | 0.0254 | 369.206 | 423.0996 | 41.094 | 41.094 | 423.0996 | 369.1933 | 99.3073 | 99.32 | Possible HV tube ends in solenoid | 30.8287
FE | LEBT | SOLR | 995 | FE_LEBT:SOLR_D0995 | S4b | solenoid | 150 |  |  | 0.4 | 368.9933 | 423.0996 | 41.094 | 41.094 | 423.0996 | 368.7933 | 99.52 | 99.72 | one side ISO flange connection | 31.0414
[7 empty] | 40 |  |  | 0.075524 | 368.7556 | 423.0996 | 41.094 | 41.094 | 423.0996 | 368.7178 | 99.7578 | 99.7955 |  | 31.2791
FE | LEBT | FC | 998 | FE_LEBT:FC_D0998 | FC1 | faraday cup | 40 |  |  | 0 | 368.7178 | 423.0996 | 41.094 | 41.094 | 423.0996 | 368.7178 | 99.7955 | 99.7955 | Not used currently | 31.3169
FE | LEBT | FFC | 998 | FE_LEBT:FFC_D0998 | LPM1 | fast faraday cup | 40 |  |  | 0 | 368.7178 | 423.0996 | 41.094 | 41.094 | 423.0996 | 368.7178 | 99.7955 | 99.7955 |  | 31.3169
[7 empty] | 40 |  |  | 0.023317 | 368.7062 | 423.0996 | 41.094 | 41.094 | 423.0996 | 368.6945 | 99.8072 | 99.8188 |  | 31.3286
FE | LEBT | PM | 998 | FE_LEBT:PM_D0998 | PM1 | profile monitor (small fork) | 40 |  |  | 0 | 368.6945 | 423.0996 | 41.094 | 41.094 | 423.0996 | 368.6945 | 99.8188 | 99.8188 |  | 31.3402
[7 empty] | 40 |  |  | 0.076159 | 368.6564 | 423.0996 | 41.094 | 41.094 | 423.0996 | 368.6183 | 99.8569 | 99.895 |  | 31.3783
FE | LEBT | CAV | 999 | FE_LEBT:CAV_D0999 |  | velocity equalizer resonator | 40 |  |  | 0 | 368.6183 | 423.0996 | 41.094 | 41.094 | 423.0996 | 368.6183 | 99.895 | 99.895 | port available | 31.4164
[7 empty] | 40 |  |  | 0.025336 | 368.6057 | 423.0996 | 41.094 | 41.094 | 423.0996 | 368.593 | 99.9077 | 99.9203 |  | 31.429
FE | LEBT | GV | 999 | FE_LEBT:GV_D0999 | GV_C | gate valve w/ special clamp | 40 |  |  | 0.0347 | 368.5757 | 423.0996 | 41.094 | 41.094 | 423.0996 | 368.5583 | 99.9377 | 99.955 | insertable gate valve | 31.4591
[10 empty] | 0.013215 | 368.5517 | 423.0996 | 41.094 | 41.094 | 423.0996 | 368.5451 | 99.9616 | 99.9682 |  | 31.483
[10 empty] | 0 | 368.5451 | 423.0996 | 41.094 | 41.094 | 423.0996 | 368.5451 | 99.9682 | 99.9682 | moved | 31.4896
FE | LEBT | end | [7 empty] | 0 | 368.5451 | 423.0996 | 41.094 | 41.094 | 423.0996 | 368.5451 | 99.9682 | 99.9682 |  | 31.4896
[6 empty] | RFQ end wall |  |  |  | 0.03175 | 368.5292 | 423.0996 | 41.094 | 41.094 | 423.0996 | 368.5133 | 99.9841 | 100 |  | 31.5055
[6 empty] | RFQ inn-wall (match point) |  |  |  | 0 | 368.5133 | 423.0996 | 41.094 | 41.094 | 423.0996 | 368.5133 | 100 | 100 |  | 31.5214
FE | RFQ | CAV | 1005 | FE_RFQ:CAV_D1005 |  | RFQ tank -1 |  |  |  | 1.0082 | 368.0093 | 423.0996 | 41.094 | 41.094 | 423.0996 | 367.5052 | 100.5041 | 101.0082 |  | 32.0254
FE | RFQ | IP | 1005 | FE_RFQ:IP_D1005 | NEG-I | getter-ion pump 2000-10 l/s |  |  |  | 0 | 367.5052 | 423.0996 | 41.094 | 41.094 | 423.0996 | 367.5052 | 101.0082 | 101.0082 |  | 32.5295
FE | RFQ | CAV | 1015 | FE_RFQ:CAV_D1015 |  | RFQ tank -2 |  |  |  | 1.0082 | 367.0011 | 423.0996 | 41.094 | 41.094 | 423.0996 | 366.497 | 101.5122 | 102.0163 |  | 33.0336
FE | RFQ | IP | 1015 | FE_RFQ:IP_D1015 | NEG-I | getter-ion pump 2000-10 l/s |  |  |  | 0 | 366.497 | 423.0996 | 41.094 | 41.094 | 423.0996 | 366.497 | 102.0163 | 102.0163 |  | 33.5377
FE | RFQ | CAV | 1025 | FE_RFQ:CAV_D1025 |  | RFQ tank -3 |  |  |  | 1.0082 | 365.993 | 423.0996 | 41.094 | 41.094 | 423.0996 | 365.4889 | 102.5204 | 103.0245 |  | 34.0418
FE | RFQ | TMP | 1025 | FE_RFQ:TMP_D1025 | TMP_M | pump, 700 l/s |  |  |  | 0 | 365.4889 | 423.0996 | 41.094 | 41.094 | 423.0996 | 365.4889 | 103.0245 | 103.0245 |  | 34.5458
FE | RFQ | CAV | 1035 | FE_RFQ:CAV_D1035 |  | RFQ tank -4 |  |  |  | 1.0082 | 364.9848 | 423.0996 | 41.094 | 41.094 | 423.0996 | 364.4807 | 103.5285 | 104.0326 |  | 35.0499
FE | RFQ | IP | 1035 | FE_RFQ:IP_D1035 | NEG-I | getter-ion pump 2000-10 l/s |  |  |  | 0 | 364.4807 | 423.0996 | 41.094 | 41.094 | 423.0996 | 364.4807 | 104.0326 | 104.0326 |  | 35.554
FE | RFQ | CAV | 1045 | FE_RFQ:CAV_D1045 |  | RFQ tank -5 |  |  |  | 1.0082 | 363.9766 | 423.0996 | 41.094 | 41.094 | 423.0996 | 363.4726 | 104.5367 | 105.0408 |  | 36.0581
FE | RFQ | IP | 1045 | FE_RFQ:IP_D1045 | NEG-I | getter-ion pump 2000-10 l/s |  |  |  | 0 | 363.4726 | 423.0996 | 41.094 | 41.094 | 423.0996 | 363.4726 | 105.0408 | 105.0408 |  | 36.5621
[6 empty] | RFQ inn-wall |  |  |  | 0 | 363.4726 | 423.0996 | 41.094 | 41.094 | 423.0996 | 363.4726 | 105.0408 | 105.0408 |  | 36.5621
[6 empty] | RFQ end wall |  |  |  | 0.03175 | 363.4567 | 423.0996 | 41.094 | 41.094 | 423.0996 | 363.4408 | 105.0567 | 105.0725 |  | 36.578
FE | MEBT | start | [7 empty] | 0 | 363.4408 | 423.0996 | 41.094 | 41.094 | 423.0996 | 363.4408 | 105.0725 | 105.0725 |  | 36.5939
[6 empty] | bellow | 30 |  |  | 0.057907 | 363.4119 | 423.0996 | 41.094 | 41.094 | 423.0996 | 363.3829 | 105.1015 | 105.1304 |  | 36.6228
FE | MEBT | GV | 1052 | FE_MEBT:GV_D1052 | GV_S | gate valve | 30 |  |  | 0.060706 | 363.3526 | 423.0996 | 41.094 | 41.094 | 423.0996 | 363.3222 | 105.1608 | 105.1911 |  | 36.6822
[10 empty] | 0.13118 | 363.2566 | 423.0996 | 41.094 | 41.094 | 423.0996 | 363.191 | 105.2567 | 105.3223 |  | 36.7781
FE | MEBT | PM | 1053 | FE_MEBT:PM_D1053 | PM1 | profile monitor (small fork) | 30 |  |  | 0 | 363.191 | 423.0996 | 41.094 | 41.094 | 423.0996 | 363.191 | 105.3223 | 105.3223 | this PM was removed per Gerlind K. 1/10/2017 | 36.8437
FE | MEBT | IP | 1053 | FE_MEBT:IP_D1053 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 363.191 | 423.0996 | 41.094 | 41.094 | 423.0996 | 363.191 | 105.3223 | 105.3223 |  | 36.8437
[7 empty] | 30 |  |  | 0.12291 | 363.1296 | 423.0996 | 41.094 | 41.094 | 423.0996 | 363.0681 | 105.3838 | 105.4452 |  | 36.9051
FE | MEBT | BCM | 1055 | FE_MEBT:BCM_D1055 | BCM1 | AC current transformer | 30 |  |  | 0.04 | 363.0481 | 423.0996 | 41.094 | 41.094 | 423.0996 | 363.0281 | 105.4652 | 105.4852 |  | 36.9866
[6 empty] | BPM bellow | 30 |  |  | 0.1035 | 362.9764 | 423.0996 | 41.094 | 41.094 | 423.0996 | 362.9246 | 105.537 | 105.5887 |  | 37.0584
FE | MEBT | BPM | 1056 | FE_MEBT:BPM_D1056 | BPM1 | position monitor | 40 |  |  | 0 | 362.9246 | 423.0996 | 41.094 | 41.094 | 423.0996 | 362.9246 | 105.5887 | 105.5887 |  | 37.1101
[7 empty] | 30 |  |  | 0.10663 | 362.8713 | 423.0996 | 41.094 | 41.094 | 423.0996 | 362.818 | 105.642 | 105.6954 |  | 37.1634
FE | MEBT | Q | 1057 | FE_MEBT:Q_D1057 | Q8 | quad, triplet 1-1 | 47.5 |  |  | 0.1 | 362.768 | 423.0996 | 41.094 | 41.094 | 423.0996 | 362.718 | 105.7454 | 105.7954 |  | 37.2667
[7 empty] | 47.5 |  |  | 0.1 | 362.668 | 423.0996 | 41.094 | 41.094 | 423.0996 | 362.618 | 105.8454 | 105.8954 |  | 37.3667
FE | MEBT | Q | 1060 | FE_MEBT:Q_D1060 | Q9 | quad, triplet 1-2 | 47.5 |  |  | 0.15 | 362.543 | 423.0996 | 41.094 | 41.094 | 423.0996 | 362.468 | 105.9704 | 106.0454 |  | 37.4917
[7 empty] | 47.5 |  |  | 0.1 | 362.418 | 423.0996 | 41.094 | 41.094 | 423.0996 | 362.368 | 106.0954 | 106.1454 |  | 37.6167
FE | MEBT | Q | 1062 | FE_MEBT:Q_D1062 | Q10 | quad+steerers, triplet 1-3 | 47.5 |  |  | 0.1 | 362.318 | 423.0996 | 41.094 | 41.094 | 423.0996 | 362.268 | 106.1954 | 106.2454 |  | 37.7167
FE | MEBT | DCHV | 1062 | FE_MEBT:DCHV_D1062 | Q10 | quad+steerers, triplet 1-3 | 47.5 |  |  | 0 | 362.268 | 423.0996 | 41.094 | 41.094 | 423.0996 | 362.268 | 106.2454 | 106.2454 |  | 37.7667
[6 empty] | bellow? | 47.5/30 |  |  | 0.24092 | 362.1475 | 423.0996 | 41.094 | 41.094 | 423.0996 | 362.0271 | 106.3658 | 106.4863 |  | 37.8872
FE | MEBT | CAV | 1066 | FE_MEBT:CAV_D1066 |  | rebuncher | 30 |  |  | 0.32 | 361.8671 | 423.0996 | 41.094 | 41.094 | 423.0996 | 361.7071 | 106.6463 | 106.8063 |  | 38.1676
FE | MEBT | IP | 1068 | FE_MEBT:IP_D1066 | IP_L | ion pump, 300 l/s |  |  |  | 0 | 361.7071 | 423.0996 | 41.094 | 41.094 | 423.0996 | 361.7071 | 106.8063 | 106.8063 |  | 38.3276
[6 empty] | bellow+tube ?? | 30 |  |  | 0.22222 | 361.596 | 423.0996 | 41.094 | 41.094 | 423.0996 | 361.4848 | 106.9174 | 107.0285 |  | 38.4388
[6 empty] | 6 way cross ?? | 30 |  |  | 0 | 361.4848 | 423.0996 | 41.094 | 41.094 | 423.0996 | 361.4848 | 107.0285 | 107.0285 |  | 38.5499
FE | MEBT | IP | 1070 | FE_MEBT:IP_D1070 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 361.4848 | 423.0996 | 41.094 | 41.094 | 423.0996 | 361.4848 | 107.0285 | 107.0285 |  | 38.5499
FE | MEBT | FAVS | 1070 | FE_MEBT:FAVS_D1070 | FVS | fast valve sensor |  |  |  | 0 | 361.4848 | 423.0996 | 41.094 | 41.094 | 423.0996 | 361.4848 | 107.0285 | 107.0285 |  | 38.5499
[10 empty] | 0.21444 | 361.3776 | 423.0996 | 41.094 | 41.094 | 423.0996 | 361.2704 | 107.1357 | 107.2429 |  | 38.6571
FE | MEBT | BPM | 1072 | FE_MEBT:BPM_D1072 | BPM1 | position monitor | 40 |  |  | 0 | 361.2704 | 423.0996 | 41.094 | 41.094 | 423.0996 | 361.2704 | 107.2429 | 107.2429 |  | 38.7643
[7 empty] | 30 |  |  | 0.10663 | 361.2171 | 423.0996 | 41.094 | 41.094 | 423.0996 | 361.1638 | 107.2963 | 107.3496 |  | 38.8176
FE | MEBT | Q | 1074 | FE_MEBT:Q_D1074 | Q8 | quad, triplet 2-1 | 47.5 |  |  | 0.1 | 361.1138 | 423.0996 | 41.094 | 41.094 | 423.0996 | 361.0638 | 107.3996 | 107.4496 |  | 38.9209
[7 empty] | 47.5 |  |  | 0.1 | 361.0138 | 423.0996 | 41.094 | 41.094 | 423.0996 | 360.9638 | 107.4996 | 107.5496 |  | 39.0209
FE | MEBT | Q | 1076 | FE_MEBT:Q_D1076 | Q9 | quad, triplet 2-2 | 47.5 |  |  | 0.15 | 360.8888 | 423.0996 | 41.094 | 41.094 | 423.0996 | 360.8138 | 107.6246 | 107.6996 |  | 39.1459
[7 empty] | 47.5 |  |  | 0.1 | 360.7638 | 423.0996 | 41.094 | 41.094 | 423.0996 | 360.7138 | 107.7496 | 107.7996 |  | 39.2709
FE | MEBT | Q | 1078 | FE_MEBT:Q_D1078 | Q10 | quad+steerers, triplet 2-3 | 47.5 |  |  | 0.1 | 360.6638 | 423.0996 | 41.094 | 41.094 | 423.0996 | 360.6138 | 107.8496 | 107.8996 |  | 39.3709
FE | MEBT | DCHV | 1078 | FE_MEBT:DCHV_D1078 | Q10 | quad+steerers, triplet 2-3 | 47.5 |  |  | 0 | 360.6138 | 423.0996 | 41.094 | 41.094 | 423.0996 | 360.6138 | 107.8996 | 107.8996 |  | 39.4209
[10 empty] | 0 | 360.6138 | 423.0996 | 41.094 | 41.094 | 423.0996 | 360.6138 | 107.8996 | 107.8996 |  | 39.4209
[10 empty] | 0.69723 | 360.2652 | 423.0996 | 41.094 | 41.094 | 423.0996 | 359.9165 | 108.2482 | 108.5968 |  | 39.7696
FE | MEBT | DH | 1086 | FE_MEBT:DH_D1086 | D6 | dipole-45deg, D6 (straight leg) | 40 |  |  | 0 | 359.9165 | 423.0996 | 41.094 | 41.094 | 423.0996 | 359.9165 | 108.5968 | 108.5968 |  | 40.1182 | FE | BTS | DH | 1089 | FE_BTS:DH_D1089 |  | dipole-45deg, D6 (bent leg) | 30 |  |  | 0.50921 | 359.5565 | 422.7395 | 41.094 | 41.094 | 422.7395 | 359.5565 | 108.8514 | 109.106 |  | 40.3932
[10 empty] | 0.562 | 359.6355 | 423.0996 | 41.094 | 41.094 | 423.0996 | 359.3545 | 108.8778 | 109.1588 |  | 40.3992 | [10 empty] | 0 | 359.5565 | 422.7395 | 41.094 | 41.094 | 422.7395 | 359.5565 | 109.106 | 109.106 |  | 40.3932
FE | MEBT | PM | 1092 | FE_MEBT:PM_D1092 | PM1 | profile monitor (small fork) | 30 ? |  |  | 0 | 359.3545 | 423.0996 | 41.094 | 41.094 | 423.0996 | 359.3545 | 109.1588 | 109.1588 |  | 40.6802 | FE | BTS | GV | 1091 | FE_BTS:GV_D1091 | GV_S | gate valve | 30 |  |  | 0.060706 | 359.535 | 422.7181 | 41.094 | 41.094 | 422.6966 | 359.5135 | 109.1364 | 109.1667 |  | 40.4539
FE | MEBT | IP | 1092 | FE_MEBT:IP_D1092 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 359.3545 | 423.0996 | 41.094 | 41.094 | 423.0996 | 359.3545 | 109.1588 | 109.1588 |  | 40.6802 | [10 empty] | 0.94143 | 359.1807 | 422.3638 | 41.094 | 41.094 | 422.0309 | 358.8479 | 109.6374 | 110.1081 |  | 41.3953
[6 empty] | bellow ? |  |  |  | 0.22641 | 359.2413 | 423.0996 | 41.094 | 41.094 | 423.0996 | 359.1281 | 109.272 | 109.3852 |  | 40.7934 | FE | BTS | SLT | 1101 | FE_BTS:SLT_D1101 |  | 2-jaw collimator | 30 |  |  | 0 | 358.8479 | 422.0309 | 41.094 | 41.094 | 422.0309 | 358.8479 | 110.1081 | 110.1081 |  | 41.3953
FE | MEBT | BPM | 1094 | FE_MEBT:BPM_D1094 | BPM1 | position monitor | 40 |  |  | 0 | 359.1281 | 423.0996 | 41.094 | 41.094 | 423.0996 | 359.1281 | 109.3852 | 109.3852 |  | 40.9066 | FE | BTS | VD | 1101 | FE_BTS:VD_D1101 | V2 | viewer | 30 |  |  | 0 | 358.8479 | 422.0309 | 41.094 | 41.094 | 422.0309 | 358.8479 | 110.1081 | 110.1081 |  | 41.3953
[7 empty] | 40 |  |  | 0.10663 | 359.0748 | 423.0996 | 41.094 | 41.094 | 423.0996 | 359.0215 | 109.4385 | 109.4918 |  | 40.9599 | FE | BTS | IP | 1101 | FE_BTS:IP_D1101 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 358.8479 | 422.0309 | 41.094 | 41.094 | 422.0309 | 358.8479 | 110.1081 | 110.1081 |  | 41.3953
[10 empty] | 0 | 359.0215 | 423.0996 | 41.094 | 41.094 | 423.0996 | 359.0215 | 109.4918 | 109.4918 |  | 41.0132 | [7 empty] | 40 |  |  | 0.18857 | 358.7812 | 421.9643 | 41.094 | 41.094 | 421.8976 | 358.7145 | 110.2024 | 110.2967 |  | 41.5839
FE | MEBT | Q | 1095 | FE_MEBT:Q_D1095 | Q8 | quad, triplet 3-1 | 47.5 |  |  | 0.1 | 358.9715 | 423.0996 | 41.094 | 41.094 | 423.0996 | 358.9215 | 109.5418 | 109.5918 |  | 41.0632 | FE | BTS | FC | 1104 | FE_BTS:FC_D1104 | FC1 | beam stopper & faraday cup | 40 |  |  | 0.21346 | 358.6391 | 421.8221 | 41.094 | 41.094 | 421.7466 | 358.5636 | 110.4034 | 110.5102 |  | 41.7973
[7 empty] | 47.5 |  |  | 0.1 | 358.8715 | 423.0996 | 41.094 | 41.094 | 423.0996 | 358.8215 | 109.6418 | 109.6918 |  | 41.1632
FE | MEBT | Q | 1098 | FE_MEBT:Q_D1098 | Q9 | quad, triplet 3-2 | 47.5 |  |  | 0.15 | 358.7465 | 423.0996 | 41.094 | 41.094 | 423.0996 | 358.6715 | 109.7668 | 109.8418 |  | 41.2882
[7 empty] | 47.5 |  |  | 0.1 | 358.6215 | 423.0996 | 41.094 | 41.094 | 423.0996 | 358.5715 | 109.8918 | 109.9418 |  | 41.4132
FE | MEBT | Q | 1100 | FE_MEBT:Q_D1100 | Q10 | quad+steerers, triplet 3-3 | 47.5 |  |  | 0.1 | 358.5215 | 423.0996 | 41.094 | 41.094 | 423.0996 | 358.4715 | 109.9918 | 110.0418 |  | 41.5132
FE | MEBT | DCHV | 1100 | FE_MEBT:DCHV_D1100 | Q10 | quad+steerers, triplet 3-3 | 47.5 |  |  | 0 | 358.4715 | 423.0996 | 41.094 | 41.094 | 423.0996 | 358.4715 | 110.0418 | 110.0418 |  | 41.5632
[10 empty] | 0.17064 | 358.3862 | 423.0996 | 41.094 | 41.094 | 423.0996 | 358.3009 | 110.1272 | 110.2125 |  | 41.6485
FE | MEBT | FC | 1102 | FE_MEBT:FC_D1102 | FC3 | beam stopper / faraday cup | 40 |  |  | 0 | 358.3009 | 423.0996 | 41.094 | 41.094 | 423.0996 | 358.3009 | 110.2125 | 110.2125 |  | 41.7339
FE | MEBT | FFC | 1102 | FE_MEBT:FFC_D1102 | LPM1 | fast faraday cup | 40 |  |  | 0 | 358.3009 | 423.0996 | 41.094 | 41.094 | 423.0996 | 358.3009 | 110.2125 | 110.2125 |  | 41.7339
FE | MEBT | IP | 1102 | FE_MEBT:IP_D1102 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 358.3009 | 423.0996 | 41.094 | 41.094 | 423.0996 | 358.3009 | 110.2125 | 110.2125 |  | 41.7339
[10 empty] | 0.11303 | 358.2443 | 423.0996 | 41.094 | 41.094 | 423.0996 | 358.1878 | 110.269 | 110.3255 |  | 41.7904
FE | MEBT | GV | 1104 | FE_MEBT:GV_D1104 | GV_S | gate valve | 30 ? |  |  | 0.060711 | 358.1575 | 423.0996 | 41.094 | 41.094 | 423.0996 | 358.1271 | 110.3559 | 110.3862 |  | 41.8772
[10 empty] | 0.18331 | 358.0355 | 423.0996 | 41.094 | 41.094 | 423.0996 | 357.9438 | 110.4779 | 110.5695 |  | 41.9992
FE | MEBT | CAV | 1107 | FE_MEBT:CAV_D1107 |  | rebuncher | 30 |  |  | 0.32 | 357.7838 | 423.0996 | 41.094 | 41.094 | 423.0996 | 357.6238 | 110.7295 | 110.8895 |  | 42.2509
FE | MEBT | IP | 1109 | FE_MEBT:IP_D1107 | IP_L | ion pump, 300 l/s |  |  |  | 0 | 357.6238 | 423.0996 | 41.094 | 41.094 | 423.0996 | 357.6238 | 110.8895 | 110.8895 |  | 42.4109
[6 empty] | bellow+tube | 40 |  |  | 0.21062 | 357.5185 | 423.0996 | 41.094 | 41.094 | 423.0996 | 357.4132 | 110.9948 | 111.1001 |  | 42.5162
FE | MEBT | BPM | 1111 | FE_MEBT:BPM_D1111 | BPM1 | position monitor | 40 |  |  | 0 | 357.4132 | 423.0996 | 41.094 | 41.094 | 423.0996 | 357.4132 | 111.1001 | 111.1001 |  | 42.6215
[7 empty] | 40 |  |  | 0.10663 | 357.3599 | 423.0996 | 41.094 | 41.094 | 423.0996 | 357.3066 | 111.1535 | 111.2068 |  | 42.6748
[10 empty] | 0 | 357.3066 | 423.0996 | 41.094 | 41.094 | 423.0996 | 357.3066 | 111.2068 | 111.2068 |  | 42.7281
FE | MEBT | Q | 1113 | FE_MEBT:Q_D1113 | Q8 | quad, triplet 4-1 | 47.5 |  |  | 0.1 | 357.2566 | 423.0996 | 41.094 | 41.094 | 423.0996 | 357.2066 | 111.2568 | 111.3068 |  | 42.7781
[7 empty] | 47.5 |  |  | 0.1 | 357.1566 | 423.0996 | 41.094 | 41.094 | 423.0996 | 357.1066 | 111.3568 | 111.4068 |  | 42.8781
FE | MEBT | Q | 1115 | FE_MEBT:Q_D1115 | Q9 | quad, triplet 4-2 | 47.5 |  |  | 0.15 | 357.0316 | 423.0996 | 41.094 | 41.094 | 423.0996 | 356.9566 | 111.4818 | 111.5568 |  | 43.0031
[7 empty] | 47.5 |  |  | 0.1 | 356.9066 | 423.0996 | 41.094 | 41.094 | 423.0996 | 356.8566 | 111.6068 | 111.6568 |  | 43.1281
FE | MEBT | Q | 1117 | FE_MEBT:Q_D1117 | Q10 | quad+steerers, triplet 4-3 | 47.5 |  |  | 0.1 | 356.8066 | 423.0996 | 41.094 | 41.094 | 423.0996 | 356.7566 | 111.7068 | 111.7568 |  | 43.2281
FE | MEBT | DCHV | 1117 | FE_MEBT:DCHV_D1117 | Q10 | quad+steerers, triplet 4-3 | 47.5 |  |  | 0 | 356.7566 | 423.0996 | 41.094 | 41.094 | 423.0996 | 356.7566 | 111.7568 | 111.7568 |  | 43.2781
[6 empty] | bellow | 40 |  |  | 0.05761 | 356.7278 | 423.0996 | 41.094 | 41.094 | 423.0996 | 356.699 | 111.7856 | 111.8144 |  | 43.307
FE | MEBT | FAV | 1118 | FE_MEBT:FAV_D1118 | FV | fast valve | 40 |  |  | 0.056602 | 356.6707 | 423.0996 | 41.094 | 41.094 | 423.0996 | 356.6424 | 111.8427 | 111.871 |  | 43.3641
[10 empty] | 0 | 356.6424 | 423.0996 | 41.094 | 41.094 | 423.0996 | 356.6424 | 111.871 | 111.871 |  | 43.3924
[10 empty] | 0.11748 | 356.5836 | 423.0996 | 41.094 | 41.094 | 423.0996 | 356.5249 | 111.9297 | 111.9885 |  | 43.4511
[10 empty] | 0 | 356.5249 | 423.0996 | 41.094 | 41.094 | 423.0996 | 356.5249 | 111.9885 | 111.9885 |  | 43.5098
FE | MEBT | BCM | 1120 | FE_MEBT:BCM_D1120 | BCM1 | AC current transformer | 40 |  |  | 0.04 | 356.5049 | 423.0996 | 41.094 | 41.094 | 423.0996 | 356.4849 | 112.0085 | 112.0285 |  | 43.5298
[10 empty] | 0 | 356.4849 | 423.0996 | 41.094 | 41.094 | 423.0996 | 356.4849 | 112.0285 | 112.0285 |  | 43.5498
[10 empty] | 0.11303 | 356.4284 | 423.0996 | 41.094 | 41.094 | 423.0996 | 356.3718 | 112.085 | 112.1415 |  | 43.6064
FE | MEBT | IP | 1121 | FE_MEBT:IP_D1121 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 356.3718 | 423.0996 | 41.094 | 41.094 | 423.0996 | 356.3718 | 112.1415 | 112.1415 |  | 43.6629
[7 empty] | 40 |  |  | 0.2305 | 356.2566 | 423.0996 | 41.094 | 41.094 | 423.0996 | 356.1413 | 112.2567 | 112.372 |  | 43.7781
[10 empty] | -2.63e-06 | 356.1413 | 423.0996 | 41.094 | 41.094 | 423.0996 | 356.1413 | 112.372 | 112.372 | adjustment for residuals | 43.8934
FE | MEBT | END | [7 empty] | 0 | 356.1413 | 423.0996 | 41.094 | 41.094 | 423.0996 | 356.1413 | 112.372 | 112.372 |  | 43.8934
SEGMENT 1 START (begin with b=0.041 QWR cryomodules) | [9 empty] | 0 | 356.1413 | 423.0996 | 41.094 | 41.094 | 423.0996 | 356.1413 | 112.372 | 112.372 |  | 43.8934
LS1 | CA01 | GV | 1124 | LS1_CA01:GV_D1124 |  | gate valve | 40 |  |  | 0.057404 | 356.1126 | 423.0996 | 41.094 | 41.094 | 423.0996 | 356.0839 | 112.4007 | 112.4294 |  | 43.9221
[6 empty] | bellows | 40 |  |  | 0.189 | 355.9894 | 423.0996 | 41.094 | 41.094 | 423.0996 | 355.8949 | 112.5239 | 112.6184 |  | 44.0453
LS1 | CA01 | CAV1 | 1127 | LS1_CA01:CAV1_D1127 |  | b04 resonator | 36 |  |  | 0.18336 | 355.8033 | 423.0996 | 41.094 | 41.094 | 423.0996 | 355.7116 | 112.7101 | 112.8018 |  | 44.2315
[10 empty] | 0 | 355.7116 | 423.0996 | 41.094 | 41.094 | 423.0996 | 355.7116 | 112.8018 | 112.8018 |  | 44.3231
LS1 | CA01 | BLM | 1128 | LS1_CA01:BLM_D1128 |  | outside loss monitor |  |  |  | 0 | 355.7116 | 423.0996 | 41.094 | 41.094 | 423.0996 | 355.7116 | 112.8018 | 112.8018 |  | 44.3231
[6 empty] | bellow | 40 |  |  | 0.09651 | 355.6633 | 423.0996 | 41.094 | 41.094 | 423.0996 | 355.6151 | 112.85 | 112.8983 |  | 44.3714
LS1 | CA01 | BPM | 1129 | LS1_CA01:BPM_D1129 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 355.6151 | 423.0996 | 41.094 | 41.094 | 423.0996 | 355.6151 | 112.8983 | 112.8983 |  | 44.4196
[6 empty] | BPM-box |  |  |  | 0.05588 | 355.5871 | 423.0996 | 41.094 | 41.094 | 423.0996 | 355.5592 | 112.9262 | 112.9542 |  | 44.4476
[6 empty] | solenoid-entry | 40 |  |  | 0.078206 | 355.5201 | 423.0996 | 41.094 | 41.094 | 423.0996 | 355.481 | 112.9933 | 113.0324 | 20170302 space adjusted to match Jlab design | 44.5146
LS1 | CA01 | SOL1 | 1132 | LS1_CA01:SOL1_D1132 | S1 | solenoid+correctors | 40 |  |  | 0.25 | 355.356 | 423.0996 | 41.094 | 41.094 | 423.0996 | 355.231 | 113.1574 | 113.2824 |  | 44.6787
[6 empty] | solenoid-exit | 40 |  |  | 0.078206 | 355.1919 | 423.0996 | 41.094 | 41.094 | 423.0996 | 355.1528 | 113.3215 | 113.3606 | 20170302 space adjusted to match Jlab design | 44.8428
[6 empty] | bellow | 40 |  |  | 0.099517 | 355.103 | 423.0996 | 41.094 | 41.094 | 423.0996 | 355.0533 | 113.4103 | 113.4601 |  | 44.9317
LS1 | CA01 | CAV2 | 1136 | LS1_CA01:CAV2_D1136 |  | b04 resonator | 36 |  |  | 0.18336 | 354.9616 | 423.0996 | 41.094 | 41.094 | 423.0996 | 354.8699 | 113.5518 | 113.6434 |  | 45.0731
[10 empty] | 0 | 354.8699 | 423.0996 | 41.094 | 41.094 | 423.0996 | 354.8699 | 113.6434 | 113.6434 |  | 45.1648
[10 empty] | 0 | 354.8699 | 423.0996 | 41.094 | 41.094 | 423.0996 | 354.8699 | 113.6434 | 113.6434 |  | 45.1648
[6 empty] | 2 bellows + tube | 36 |  |  | 0.45019 | 354.6448 | 423.0996 | 41.094 | 41.094 | 423.0996 | 354.4197 | 113.8685 | 114.0936 | 20170302 space adjusted to match Jlab design | 45.3899
[10 empty] | 0 | 354.4197 | 423.0996 | 41.094 | 41.094 | 423.0996 | 354.4197 | 114.0936 | 114.0936 |  | 45.615
LS1 | CA01 | CAV3 | 1142 | LS1_CA01:CAV3_D1142 |  | b04 resonator | 36 |  |  | 0.18336 | 354.328 | 423.0996 | 41.094 | 41.094 | 423.0996 | 354.2363 | 114.1853 | 114.277 |  | 45.7067
[10 empty] | 0 | 354.2363 | 423.0996 | 41.094 | 41.094 | 423.0996 | 354.2363 | 114.277 | 114.277 |  | 45.7984
LS1 | CA01 | BLM | 1143 | LS1_CA01:BLM_D1143 |  | outside loss monitor |  |  |  | 0 | 354.2363 | 423.0996 | 41.094 | 41.094 | 423.0996 | 354.2363 | 114.277 | 114.277 |  | 45.7984
[6 empty] | bellow |  |  |  | 0.09651 | 354.1881 | 423.0996 | 41.094 | 41.094 | 423.0996 | 354.1398 | 114.3253 | 114.3735 |  | 45.8466
LS1 | CA01 | BPM | 1144 | LS1_CA01:BPM_D1144 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 354.1398 | 423.0996 | 41.094 | 41.094 | 423.0996 | 354.1398 | 114.3735 | 114.3735 |  | 45.8949
[6 empty] | BPM-box |  |  |  | 0.05588 | 354.1119 | 423.0996 | 41.094 | 41.094 | 423.0996 | 354.0839 | 114.4015 | 114.4294 |  | 45.9228
[6 empty] | solenoid-entry |  |  |  | 0.078206 | 354.0448 | 423.0996 | 41.094 | 41.094 | 423.0996 | 354.0057 | 114.4685 | 114.5076 | 20170302 space adjusted to match Jlab design | 45.9899
LS1 | CA01 | SOL2 | 1146 | LS1_CA01:SOL2_D1146 | S1 | solenoid+correctors | 40 |  |  | 0.25 | 353.8807 | 423.0996 | 41.094 | 41.094 | 423.0996 | 353.7557 | 114.6326 | 114.7576 |  | 46.154
[6 empty] | solenoid-exit |  |  |  | 0.078206 | 353.7166 | 423.0996 | 41.094 | 41.094 | 423.0996 | 353.6775 | 114.7967 | 114.8358 | 20170302 space adjusted to match Jlab design | 46.3181
[6 empty] | bellow |  |  |  | 0.099517 | 353.6278 | 423.0996 | 41.094 | 41.094 | 423.0996 | 353.578 | 114.8856 | 114.9353 |  | 46.4069
LS1 | CA01 | CAV4 | 1150 | LS1_CA01:CAV4_D1150 |  | b04 resonator | 36 |  |  | 0.18336 | 353.4863 | 423.0996 | 41.094 | 41.094 | 423.0996 | 353.3947 | 115.027 | 115.1187 |  | 46.5484
[10 empty] | 0 | 353.3947 | 423.0996 | 41.094 | 41.094 | 423.0996 | 353.3947 | 115.1187 | 115.1187 |  | 46.6401
[10 empty] | 0 | 353.3947 | 423.0996 | 41.094 | 41.094 | 423.0996 | 353.3947 | 115.1187 | 115.1187 |  | 46.6401
[6 empty] | bellows |  |  |  | 0.189 | 353.3002 | 423.0996 | 41.094 | 41.094 | 423.0996 | 353.2057 | 115.2132 | 115.3077 |  | 46.7346
LS1 | CA01 | GV | 1153 | LS1_CA01:GV_D1153 |  | gate valve | 40 |  |  | 0.057404 | 353.177 | 423.0996 | 41.094 | 41.094 | 423.0996 | 353.1483 | 115.3364 | 115.3651 |  | 46.8578
[6 empty] | bellow | 40 |  |  | 0.1035 | 353.0965 | 423.0996 | 41.094 | 41.094 | 423.0996 | 353.0447 | 115.4168 | 115.4686 |  | 46.9382
LS1 | WA01 | BPM | 1155 | LS1_WA01:BPM_D1155 | BPM2 | position monitor |  |  |  | 0 | 353.0447 | 423.0996 | 41.094 | 41.094 | 423.0996 | 353.0447 | 115.4686 | 115.4686 |  | 46.99
[6 empty] | diagnostic box |  |  |  | 0.14528 | 352.9721 | 423.0996 | 41.094 | 41.094 | 423.0996 | 352.8995 | 115.5412 | 115.6139 |  | 47.0626
LS1 | WA01 | HMR | 1156 | LS1_WA01:HMR_D1156 | HR_L | Halo ring, large aperture | 33 |  |  | 0 | 352.8995 | 423.0996 | 41.094 | 41.094 | 423.0996 | 352.8995 | 115.6139 | 115.6139 |  | 47.1352
[6 empty] | diagnostic box |  |  |  | 0 | 352.8995 | 423.0996 | 41.094 | 41.094 | 423.0996 | 352.8995 | 115.6139 | 115.6139 |  | 47.1352
LS1 | WA01 | IP | 1156 | LS1_WA01:IP_D1156 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 352.8995 | 423.0996 | 41.094 | 41.094 | 423.0996 | 352.8995 | 115.6139 | 115.6139 |  | 47.1352
[6 empty] | bellow | 40 |  |  | 0.13122 | 352.8339 | 423.0996 | 41.094 | 41.094 | 423.0996 | 352.7683 | 115.6795 | 115.7451 |  | 47.2009
LS1 | CA02 | GV | 1158 | LS1_CA02:GV_D1158 |  | gate valve | 40 |  |  | 0.057404 | 352.7395 | 423.0996 | 41.094 | 41.094 | 423.0996 | 352.7108 | 115.7738 | 115.8025 |  | 47.2952
[6 empty] | bellows | 40 |  |  | 0.189 | 352.6163 | 423.0996 | 41.094 | 41.094 | 423.0996 | 352.5218 | 115.897 | 115.9915 |  | 47.4184
LS1 | CA02 | CAV1 | 1161 | LS1_CA02:CAV1_D1161 |  | b04 resonator | 36 |  |  | 0.18336 | 352.4302 | 423.0996 | 41.094 | 41.094 | 423.0996 | 352.3385 | 116.0832 | 116.1749 |  | 47.6045
[10 empty] | 0 | 352.3385 | 423.0996 | 41.094 | 41.094 | 423.0996 | 352.3385 | 116.1749 | 116.1749 |  | 47.6962
LS1 | CA02 | BLM | 1162 | LS1_CA02:BLM_D1162 |  | outside loss monitor |  |  |  | 0 | 352.3385 | 423.0996 | 41.094 | 41.094 | 423.0996 | 352.3385 | 116.1749 | 116.1749 |  | 47.6962
[6 empty] | bellow | 40 |  |  | 0.09651 | 352.2902 | 423.0996 | 41.094 | 41.094 | 423.0996 | 352.242 | 116.2231 | 116.2714 |  | 47.7445
LS1 | CA02 | BPM | 1163 | LS1_CA02:BPM_D1163 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 352.242 | 423.0996 | 41.094 | 41.094 | 423.0996 | 352.242 | 116.2714 | 116.2714 |  | 47.7927
[6 empty] | BPM-box |  |  |  | 0.05588 | 352.214 | 423.0996 | 41.094 | 41.094 | 423.0996 | 352.1861 | 116.2993 | 116.3272 |  | 47.8207
[6 empty] | solenoid-entry | 40 |  |  | 0.078206 | 352.147 | 423.0996 | 41.094 | 41.094 | 423.0996 | 352.1079 | 116.3664 | 116.4055 | 20170302 space adjusted to match Jlab design | 47.8877
LS1 | CA02 | SOL1 | 1165 | LS1_CA02:SOL1_D1165 | S1 | solenoid+correctors | 40 |  |  | 0.25 | 351.9829 | 423.0996 | 41.094 | 41.094 | 423.0996 | 351.8579 | 116.5305 | 116.6555 |  | 48.0518
[6 empty] | solenoid-exit | 40 |  |  | 0.078206 | 351.8188 | 423.0996 | 41.094 | 41.094 | 423.0996 | 351.7797 | 116.6946 | 116.7337 | 20170302 space adjusted to match Jlab design | 48.2159
[6 empty] | bellow | 40 |  |  | 0.099517 | 351.7299 | 423.0996 | 41.094 | 41.094 | 423.0996 | 351.6802 | 116.7834 | 116.8332 |  | 48.3048
LS1 | CA02 | CAV2 | 1169 | LS1_CA02:CAV2_D1169 |  | b04 resonator | 36 |  |  | 0.18336 | 351.5885 | 423.0996 | 41.094 | 41.094 | 423.0996 | 351.4968 | 116.9249 | 117.0165 |  | 48.4462
[10 empty] | 0 | 351.4968 | 423.0996 | 41.094 | 41.094 | 423.0996 | 351.4968 | 117.0165 | 117.0165 |  | 48.5379
[10 empty] | 0 | 351.4968 | 423.0996 | 41.094 | 41.094 | 423.0996 | 351.4968 | 117.0165 | 117.0165 |  | 48.5379
[6 empty] | 2 bellows + tube | 36 |  |  | 0.45019 | 351.2717 | 423.0996 | 41.094 | 41.094 | 423.0996 | 351.0466 | 117.2416 | 117.4667 | 20170302 space adjusted to match Jlab design | 48.763
[10 empty] | 0 | 351.0466 | 423.0996 | 41.094 | 41.094 | 423.0996 | 351.0466 | 117.4667 | 117.4667 |  | 48.9881
LS1 | CA02 | CAV3 | 1176 | LS1_CA02:CAV3_D1176 |  | b04 resonator | 36 |  |  | 0.18336 | 350.9549 | 423.0996 | 41.094 | 41.094 | 423.0996 | 350.8632 | 117.5584 | 117.6501 |  | 49.0798
[10 empty] | 0 | 350.8632 | 423.0996 | 41.094 | 41.094 | 423.0996 | 350.8632 | 117.6501 | 117.6501 |  | 49.1715
LS1 | CA02 | BLM | 1177 | LS1_CA02:BLM_D1177 |  | outside loss monitor |  |  |  | 0 | 350.8632 | 423.0996 | 41.094 | 41.094 | 423.0996 | 350.8632 | 117.6501 | 117.6501 |  | 49.1715
[6 empty] | bellow |  |  |  | 0.09651 | 350.815 | 423.0996 | 41.094 | 41.094 | 423.0996 | 350.7667 | 117.6984 | 117.7466 |  | 49.2197
LS1 | CA02 | BPM | 1177 | LS1_CA02:BPM_D1177 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 350.7667 | 423.0996 | 41.094 | 41.094 | 423.0996 | 350.7667 | 117.7466 | 117.7466 |  | 49.268
[6 empty] | BPM-box |  |  |  | 0.05588 | 350.7388 | 423.0996 | 41.094 | 41.094 | 423.0996 | 350.7109 | 117.7745 | 117.8025 |  | 49.2959
[6 empty] | solenoid-entry |  |  |  | 0.078206 | 350.6718 | 423.0996 | 41.094 | 41.094 | 423.0996 | 350.6326 | 117.8416 | 117.8807 | 20170302 space adjusted to match Jlab design | 49.363
LS1 | CA02 | SOL2 | 1180 | LS1_CA02:SOL2_D1180 | S1 | solenoid+correctors | 40 |  |  | 0.25 | 350.5076 | 423.0996 | 41.094 | 41.094 | 423.0996 | 350.3826 | 118.0057 | 118.1307 |  | 49.5271
[6 empty] | solenoid-exit |  |  |  | 0.078206 | 350.3435 | 423.0996 | 41.094 | 41.094 | 423.0996 | 350.3044 | 118.1698 | 118.2089 | 20170302 space adjusted to match Jlab design | 49.6912
[6 empty] | bellow |  |  |  | 0.099517 | 350.2547 | 423.0996 | 41.094 | 41.094 | 423.0996 | 350.2049 | 118.2587 | 118.3084 |  | 49.78
LS1 | CA02 | CAV4 | 1184 | LS1_CA02:CAV4_D1184 |  | b04 resonator | 36 |  |  | 0.18336 | 350.1132 | 423.0996 | 41.094 | 41.094 | 423.0996 | 350.0216 | 118.4001 | 118.4918 |  | 49.9215
[10 empty] | 0 | 350.0216 | 423.0996 | 41.094 | 41.094 | 423.0996 | 350.0216 | 118.4918 | 118.4918 |  | 50.0132
[10 empty] | 0 | 350.0216 | 423.0996 | 41.094 | 41.094 | 423.0996 | 350.0216 | 118.4918 | 118.4918 |  | 50.0132
[6 empty] | bellows |  |  |  | 0.189 | 349.9271 | 423.0996 | 41.094 | 41.094 | 423.0996 | 349.8326 | 118.5863 | 118.6808 |  | 50.1077
LS1 | CA02 | GV | 1187 | LS1_CA02:GV_D1187 |  | gate valve | 40 |  |  | 0.057404 | 349.8039 | 423.0996 | 41.094 | 41.094 | 423.0996 | 349.7752 | 118.7095 | 118.7382 |  | 50.2309
[6 empty] | bellow | 40 |  |  | 0.1035 | 349.7234 | 423.0996 | 41.094 | 41.094 | 423.0996 | 349.6717 | 118.7899 | 118.8417 |  | 50.3113
LS1 | WA02 | BPM | 1188 | LS1_WA02:BPM_D1188 | BPM2 | position monitor |  |  |  | 0 | 349.6717 | 423.0996 | 41.094 | 41.094 | 423.0996 | 349.6717 | 118.8417 | 118.8417 |  | 50.3631
[6 empty] | diagnostic box |  |  |  | 0.14528 | 349.599 | 423.0996 | 41.094 | 41.094 | 423.0996 | 349.5264 | 118.9143 | 118.987 |  | 50.4357
LS1 | WA02 | HMR | 1190 | LS1_WA02:HMR_D1190 | HR_L | Halo ring, large aperture | 33 |  |  | 0 | 349.5264 | 423.0996 | 41.094 | 41.094 | 423.0996 | 349.5264 | 118.987 | 118.987 |  | 50.5083
[6 empty] | diagnostic box |  |  |  | 0 | 349.5264 | 423.0996 | 41.094 | 41.094 | 423.0996 | 349.5264 | 118.987 | 118.987 |  | 50.5083
LS1 | WA02 | IP | 1190 | LS1_WA02:IP_D1190 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 349.5264 | 423.0996 | 41.094 | 41.094 | 423.0996 | 349.5264 | 118.987 | 118.987 |  | 50.5083
[6 empty] | bellow | 40 |  |  | 0.13122 | 349.4608 | 423.0996 | 41.094 | 41.094 | 423.0996 | 349.3952 | 119.0526 | 119.1182 |  | 50.5739
LS1 | CA03 | GV | 1191 | LS1_CA03:GV_D1191 |  | gate valve | 40 |  |  | 0.057404 | 349.3665 | 423.0996 | 41.094 | 41.094 | 423.0996 | 349.3378 | 119.1469 | 119.1756 |  | 50.6683
[6 empty] | bellows | 40 |  |  | 0.189 | 349.2433 | 423.0996 | 41.094 | 41.094 | 423.0996 | 349.1488 | 119.2701 | 119.3646 |  | 50.7915
LS1 | CA03 | CAV1 | 1195 | LS1_CA03:CAV1_D1195 |  | b04 resonator | 36 |  |  | 0.18336 | 349.0571 | 423.0996 | 41.094 | 41.094 | 423.0996 | 348.9654 | 119.4563 | 119.548 |  | 50.9776
[10 empty] | 0 | 348.9654 | 423.0996 | 41.094 | 41.094 | 423.0996 | 348.9654 | 119.548 | 119.548 |  | 51.0693
LS1 | CA03 | BLM | 1195 | LS1_CA03:BLM_D1195 |  | outside loss monitor |  |  |  | 0 | 348.9654 | 423.0996 | 41.094 | 41.094 | 423.0996 | 348.9654 | 119.548 | 119.548 |  | 51.0693
[6 empty] | bellow | 40 |  |  | 0.09651 | 348.9171 | 423.0996 | 41.094 | 41.094 | 423.0996 | 348.8689 | 119.5962 | 119.6445 |  | 51.1176
LS1 | CA03 | BPM | 1196 | LS1_CA03:BPM_D1196 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 348.8689 | 423.0996 | 41.094 | 41.094 | 423.0996 | 348.8689 | 119.6445 | 119.6445 |  | 51.1658
[6 empty] | BPM-box |  |  |  | 0.05588 | 348.8409 | 423.0996 | 41.094 | 41.094 | 423.0996 | 348.813 | 119.6724 | 119.7003 |  | 51.1938
[6 empty] | solenoid-entry | 40 |  |  | 0.078206 | 348.7739 | 423.0996 | 41.094 | 41.094 | 423.0996 | 348.7348 | 119.7394 | 119.7785 | 20170302 space adjusted to match Jlab design | 51.2608
LS1 | CA03 | SOL1 | 1199 | LS1_CA03:SOL1_D1199 | S1 | solenoid+correctors | 40 |  |  | 0.25 | 348.6098 | 423.0996 | 41.094 | 41.094 | 423.0996 | 348.4848 | 119.9035 | 120.0285 |  | 51.4249
[6 empty] | solenoid-exit | 40 |  |  | 0.078206 | 348.4457 | 423.0996 | 41.094 | 41.094 | 423.0996 | 348.4066 | 120.0677 | 120.1068 | 20170302 space adjusted to match Jlab design | 51.589
[6 empty] | bellow | 40 |  |  | 0.099517 | 348.3568 | 423.0996 | 41.094 | 41.094 | 423.0996 | 348.3071 | 120.1565 | 120.2063 |  | 51.6779
LS1 | CA03 | CAV2 | 1203 | LS1_CA03:CAV2_D1203 |  | b04 resonator | 36 |  |  | 0.18336 | 348.2154 | 423.0996 | 41.094 | 41.094 | 423.0996 | 348.1237 | 120.298 | 120.3896 |  | 51.8193
[10 empty] | 0 | 348.1237 | 423.0996 | 41.094 | 41.094 | 423.0996 | 348.1237 | 120.3896 | 120.3896 |  | 51.911
[10 empty] | 0 | 348.1237 | 423.0996 | 41.094 | 41.094 | 423.0996 | 348.1237 | 120.3896 | 120.3896 |  | 51.911
[6 empty] | 2 bellows + tube | 36 |  |  | 0.45019 | 347.8986 | 423.0996 | 41.094 | 41.094 | 423.0996 | 347.6735 | 120.6147 | 120.8398 | 20170302 space adjusted to match Jlab design | 52.1361
[10 empty] | 0 | 347.6735 | 423.0996 | 41.094 | 41.094 | 423.0996 | 347.6735 | 120.8398 | 120.8398 |  | 52.3612
LS1 | CA03 | CAV3 | 1209 | LS1_CA03:CAV3_D1209 |  | b04 resonator | 36 |  |  | 0.18336 | 347.5818 | 423.0996 | 41.094 | 41.094 | 423.0996 | 347.4901 | 120.9315 | 121.0232 |  | 52.4529
[10 empty] | 0 | 347.4901 | 423.0996 | 41.094 | 41.094 | 423.0996 | 347.4901 | 121.0232 | 121.0232 |  | 52.5446
LS1 | CA03 | BLM | 1210 | LS1_CA03:BLM_D1210 |  | outside loss monitor |  |  |  | 0 | 347.4901 | 423.0996 | 41.094 | 41.094 | 423.0996 | 347.4901 | 121.0232 | 121.0232 |  | 52.5446
[6 empty] | bellow |  |  |  | 0.09651 | 347.4419 | 423.0996 | 41.094 | 41.094 | 423.0996 | 347.3936 | 121.0714 | 121.1197 |  | 52.5928
LS1 | CA03 | BPM | 1211 | LS1_CA03:BPM_D1211 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 347.3936 | 423.0996 | 41.094 | 41.094 | 423.0996 | 347.3936 | 121.1197 | 121.1197 |  | 52.6411
[6 empty] | BPM-box |  |  |  | 0.05588 | 347.3657 | 423.0996 | 41.094 | 41.094 | 423.0996 | 347.3378 | 121.1476 | 121.1756 |  | 52.669
[6 empty] | solenoid-entry |  |  |  | 0.078206 | 347.2987 | 423.0996 | 41.094 | 41.094 | 423.0996 | 347.2596 | 121.2147 | 121.2538 | 20170302 space adjusted to match Jlab design | 52.7361
LS1 | CA03 | SOL2 | 1214 | LS1_CA03:SOL2_D1214 | S1 | solenoid+correctors | 40 |  |  | 0.25 | 347.1346 | 423.0996 | 41.094 | 41.094 | 423.0996 | 347.0096 | 121.3788 | 121.5038 |  | 52.9002
[6 empty] | solenoid-exit |  |  |  | 0.078206 | 346.9704 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.9313 | 121.5429 | 121.582 | 20170302 space adjusted to match Jlab design | 53.0643
[6 empty] | bellow |  |  |  | 0.099517 | 346.8816 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.8318 | 121.6318 | 121.6815 |  | 53.1531
LS1 | CA03 | CAV4 | 1218 | LS1_CA03:CAV4_D1218 |  | b04 resonator | 36 |  |  | 0.18336 | 346.7401 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.6485 | 121.7732 | 121.8649 |  | 53.2946
[10 empty] | 0 | 346.6485 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.6485 | 121.8649 | 121.8649 |  | 53.3862
[10 empty] | 0 | 346.6485 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.6485 | 121.8649 | 121.8649 |  | 53.3862
[6 empty] | bellows |  |  |  | 0.189 | 346.554 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.4595 | 121.9594 | 122.0539 |  | 53.4807 | exit valve of the last beta=0.041 module, diagnstic box, and D-station
LS1 | CA03 | GV | 1221 | LS1_CA03:GV_D1221 |  | gate valve | 40 |  |  | 0.057404 | 346.4308 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.4021 | 122.0826 | 122.1113 |  | 53.6039 | LS1 | CA03 | GV | 1221 | LS1_CA03:GV_D1221 |  | gate valve | 40 |  | M403_GV2 | 0.057404 | 346.4308 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.4021 | 122.0826 | 122.1113
[6 empty] | bellow | 40 |  |  | 0.1035 | 346.3503 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.2986 | 122.163 | 122.2148 |  | 53.6844 | [6 empty] | bellow | 40 |  |  | 0.1035 | 346.3503 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.2986 | 122.163 | 122.2148
LS1 | WA03 | BPM | 1222 | LS1_WA03:BPM_D1222 | BPM2 | position monitor |  |  |  | 0 | 346.2986 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.2986 | 122.2148 | 122.2148 |  | 53.7362 | LS1 | WA03 | BPM | 1222 | LS1_WA03:BPM_D1222 | BPM2 | position monitor |  |  | M403_BPM3 | 0 | 346.2986 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.2986 | 122.2148 | 122.2148
[6 empty] | diagnostic box |  |  |  | 0.14528 | 346.2259 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.1533 | 122.2874 | 122.3601 |  | 53.8088 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 346.2259 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.1533 | 122.2874 | 122.3601
LS1 | WA03 | HMR | 1224 | LS1_WA03:HMR_D1224 | HR_L | Halo ring, large aperture | 33 |  |  | 0 | 346.1533 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.1533 | 122.3601 | 122.3601 | 346.1533 | 53.8814 | LS1 | WA03 | PM | 1224 | LS1_WA03:PM_D1224 | PMx | profile monitor (not costed) | future PM |  |  | 0 | 346.1533 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.1533 | 122.3601 | 122.3601
LS1 | WA03 | BD | 1224 | LS1_WA03:BD_D1224 | dump | Beam dump with nibium plate |  |  |  | 0 | 346.1533 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.1533 | 122.3601 | 122.3601 |  | 53.8814 | [10 empty] | 0 | 346.1533 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.1533 | 122.3601 | 122.3601
[6 empty] | diagnostic box |  |  |  | 0 | 346.1533 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.1533 | 122.3601 | 122.3601 |  | 53.8814 | [6 empty] | diagnostic box |  |  |  | 0 | 346.1533 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.1533 | 122.3601 | 122.3601
LS1 | WA03 | IP | 1224 | LS1_WA03:IP_D1224 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 346.1533 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.1533 | 122.3601 | 122.3601 |  | 53.8814 | LS1 | WA03 | IP | 1224 | LS1_WA03:IP_D1224 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 346.1533 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.1533 | 122.3601 | 122.3601
[6 empty] | bellow | 40 |  |  | 0.13122 | 346.0877 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.0221 | 122.4257 | 122.4913 |  | 53.947 | [10 empty] | 0 | 346.1533 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.1533 | 122.3601 | 122.3601
D-station starts in "V" column | [30 empty] | 0 | 346.1533 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.1533 | 122.3601 | 122.3601
beta=0.085 QWR cryomodules START | [9 empty] | 0 | 346.0221 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.0221 | 122.4913 | 122.4913 |  | 54.0127 | [10 empty] | 0.036116 | 346.1352 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.1172 | 122.3781 | 122.3962
LS1 | CB01 | GV | 1225 | LS1_CB01:GV_D1225 |  | gate valve | 40 |  |  | 0.057404 | 345.9934 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.9647 | 122.52 | 122.5487 |  | 54.0414 | LS1 | WA03 | BLM | 1224 | LS1_WA03:BLM_D1224 | HR_L | Halo ring, large aperture | 33 |  | M403_HR | 0 | 346.1172 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.1172 | 122.3962 | 122.3962
[6 empty] | bellows |  |  |  | 0.19082 | 345.8692 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.7738 | 122.6441 | 122.7395 |  | 54.1655 | [6 empty] | bellow | 40 |  |  | 0.0951 | 346.0696 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.0221 | 122.4437 | 122.4913
LS1 | CB01 | CAV1 | 1229 | LS1_CB01:CAV1_D1229 |  | b08 resonator | 34 |  |  | 0.29817 | 345.6247 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.4757 | 122.8886 | 123.0377 |  | 54.41 | D-station start | [9 empty] | 0.028661 | 346.0077 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.9934 | 122.5056 | 122.5199
[10 empty] | 0 | 345.4757 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.4757 | 123.0377 | 123.0377 |  | 54.559 | LS1 | D | GV | 1225 | LS1_D:GV_D1225 | GV_M | gate valve |  |  | D_GV | 0 | 345.9934 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.9934 | 122.5199 | 122.5199 | borrowed from BDS
LS1 | CB01 | BLM | 1230 | LS1_CB01:BLM_D1230 |  | outside loss monitor |  |  |  | 0 | 345.4757 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.4757 | 123.0377 | 123.0377 |  | 54.559 | [10 empty] | 0.12629 | 345.9303 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.8671 | 122.5831 | 122.6462
[6 empty] | bellow | 40 |  |  | 0.09651 | 345.4274 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.3792 | 123.0859 | 123.1342 |  | 54.6073 | LS1 | D | FAV | 1226 | LS1_D:FAV_D1226 | FV | fast valve |  |  | D_FV | 0 | 345.8671 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.8671 | 122.6462 | 122.6462 | borrowed from BDS
LS1 | CB01 | BPM | 1231 | LS1_CB01:BPM_D1231 | BPM6 | position monitor, BPMc |  |  |  | 0 | 345.3792 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.3792 | 123.1342 | 123.1342 |  | 54.6556 | [10 empty] | 0.15123 | 345.7915 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.7159 | 122.7218 | 122.7975
[6 empty] | BPM-box | 40 |  |  | 0.05588 | 345.3512 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.3233 | 123.1621 | 123.1901 |  | 54.6835 | LS1 | D | LPM | 1228 | LS1_D:LPM_D1228 | LPM | longitudial profile monitor |  |  | D_LPM | 0 | 345.7159 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.7159 | 122.7975 | 122.7975 | Not installed
[6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 345.2852 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.2471 | 123.2281 | 123.2662 |  | 54.7495 | [10 empty] | 0.13376 | 345.649 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.5821 | 122.8643 | 122.9312
LS1 | CB01 | SOL1 | 1235 | LS1_CB01:SOL1_D1235 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 344.9971 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.7471 | 123.5162 | 123.7662 |  | 55.0376 | LS1 | D | BCM | 1229 | LS1_D:BCM_D1229 | BCM | beam current monitor |  |  | D_BCM1 | 0 | 345.5821 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.5821 | 122.9312 | 122.9312 | borrowed from BDS
[6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 344.7091 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.671 | 123.8043 | 123.8423 |  | 55.3256 | [10 empty] | 0.12349 | 345.5204 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.4586 | 122.993 | 123.0547
[6 empty] | bellow |  |  |  | 0.099517 | 344.6212 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.5715 | 123.8921 | 123.9419 |  | 55.4135 | [10 empty] | 0 | 345.4586 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.4586 | 123.0547 | 123.0547
LS1 | CB01 | CAV2 | 1241 | LS1_CB01:CAV2_D1241 |  | b08 resonator | 34 |  |  | 0.29817 | 344.4224 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.2733 | 124.0909 | 124.24 |  | 55.6123 | LS1 | D | BPM | 1231 | LS1_D:BPM_D1231 | BPM | beam position monitor |  |  | D_BPM1 | 0 | 345.4586 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.4586 | 123.0547 | 123.0547 | borrowed from BDS
[10 empty] | 0 | 344.2733 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.2733 | 124.24 | 124.24 |  | 55.7614 | [10 empty] | 0.14529 | 345.386 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.3133 | 123.1274 | 123.2
[10 empty] | 0 | 344.2733 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.2733 | 124.24 | 124.24 |  | 55.7614 | LS1 | D | PM | 1232 | LS1_D:PM_D1232 | PM | wire scanner |  |  | D_PM1 | 0 | 345.3133 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.3133 | 123.2 | 123.2 | borrowed from BDS
[6 empty] | bellow |  |  |  | 0.099517 | 344.2236 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.1738 | 124.2898 | 124.3395 |  | 55.8112 | LS1 | D | PM | 1232 | LS1_D:PM_D1232 | IP_S | ion pump, 75 l/s |  |  | D_IP | 0 | 345.3133 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.3133 | 123.2 | 123.2 | borrowed from BDS
LS1 | CB01 | CAV3 | 1245 | LS1_CB01:CAV3_D1245 |  | b08 resonator | 34 |  |  | 0.29817 | 344.0247 | 423.0996 | 41.094 | 41.094 | 423.0996 | 343.8756 | 124.4886 | 124.6377 |  | 56.01 | LS1 | D | CCG | 1232 | LS1_D:CCG_D1232 | CCG | cold cathde gauge |  |  | D_CCG1 | 0 | 345.3133 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.3133 | 123.2 | 123.2 | borrowed from BDS
[10 empty] | 0 | 343.8756 | 423.0996 | 41.094 | 41.094 | 423.0996 | 343.8756 | 124.6377 | 124.6377 |  | 56.1591 | [10 empty] | 0.37186 | 345.1274 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.9415 | 123.3859 | 123.5719
[10 empty] | 0 | 343.8756 | 423.0996 | 41.094 | 41.094 | 423.0996 | 343.8756 | 124.6377 | 124.6377 |  | 56.1591 | LS1 | D | SiD | 1236 | LS1_D:SiD_D1236 | SiD | silicon detector |  |  | D_SiD | 0 | 344.9415 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.9415 | 123.5719 | 123.5719 | borrowed from NSCL
[6 empty] | bellow |  |  |  | 0.099517 | 343.8259 | 423.0996 | 41.094 | 41.094 | 423.0996 | 343.7761 | 124.6875 | 124.7372 |  | 56.2088 | LS1 | D | TMP | 1236 | LS1_D:TMP_D1236 | TMP | turbo pump 500 l/s |  |  | D_TMP | 0 | 344.9415 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.9415 | 123.5719 | 123.5719 | borrowed from BDS
LS1 | CB01 | CAV4 | 1249 | LS1_CB01:CAV4_D1249 |  | b08 resonator | 34 |  |  | 0.29817 | 343.627 | 423.0996 | 41.094 | 41.094 | 423.0996 | 343.4779 | 124.8863 | 125.0354 |  | 56.4077 | [10 empty] | 0.23619 | 344.8234 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.7053 | 123.69 | 123.8081
[10 empty] | 0 | 343.4779 | 423.0996 | 41.094 | 41.094 | 423.0996 | 343.4779 | 125.0354 | 125.0354 |  | 56.5568 | LS1 | D | BPM | 1238 | LS1_D:BPM_D1238 | BPM | beam position monitor |  |  | D_BPM2 | 0 | 344.7053 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.7053 | 123.8081 | 123.8081 | borrowed from BDS
LS1 | CB01 | BLM | 1250 | LS1_CB01:BLM_D1250 |  | outside loss monitor |  |  |  | 0 | 343.4779 | 423.0996 | 41.094 | 41.094 | 423.0996 | 343.4779 | 125.0354 | 125.0354 |  | 56.5568 | [10 empty] | 0.14529 | 344.6326 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.56 | 123.8807 | 123.9533
[6 empty] | bellow |  |  |  | 0.09651 | 343.4297 | 423.0996 | 41.094 | 41.094 | 423.0996 | 343.3814 | 125.0837 | 125.1319 |  | 56.605 | LS1 | D | PM | 1240 | LS1_D:PM_D1240 | PM | wire scanner |  |  | D_PM2 | 0 | 344.56 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.56 | 123.9533 | 123.9533 | Not installed
LS1 | CB01 | BPM | 1251 | LS1_CB01:BPM_D1251 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 343.3814 | 423.0996 | 41.094 | 41.094 | 423.0996 | 343.3814 | 125.1319 | 125.1319 |  | 56.6533 | [10 empty] | 0.036144 | 344.5419 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.5239 | 123.9714 | 123.9895
[6 empty] | BPM-box |  |  |  | 0.05588 | 343.3535 | 423.0996 | 41.094 | 41.094 | 423.0996 | 343.3255 | 125.1599 | 125.1878 |  | 56.6812 | LS1 | D | BLM | 1240 | LS1_D:BLM_D1240 | HR_S | Halo ring, small aperture | 30 |  | D_HR | 0 | 344.5239 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.5239 | 123.9895 | 123.9895
[6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 343.2875 | 423.0996 | 41.094 | 41.094 | 423.0996 | 343.2494 | 125.2259 | 125.2639 |  | 56.7472 | [10 empty] | 0.11504 | 344.4663 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.4088 | 124.047 | 124.1045
LS1 | CB01 | SOL2 | 1255 | LS1_CB01:SOL2_D1255 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 342.9994 | 423.0996 | 41.094 | 41.094 | 423.0996 | 342.7494 | 125.5139 | 125.7639 |  | 57.0353 | LS1 | D | BCM | 1241 | LS1_D:BCM_D1241 | BCM | beam current monitor |  |  | D_BCM2 | 0 | 344.4088 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.4088 | 124.1045 | 124.1045 | borrowed from BDS
[6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 342.7113 | 423.0996 | 41.094 | 41.094 | 423.0996 | 342.6733 | 125.802 | 125.8401 |  | 57.3234 | [10 empty] | 0.26302 | 344.2773 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.1458 | 124.236 | 124.3675
[6 empty] | bellow |  |  |  | 0.099517 | 342.6235 | 423.0996 | 41.094 | 41.094 | 423.0996 | 342.5738 | 125.8898 | 125.9396 |  | 57.4112 | LS1 | D | FC | 1244 | LS1_D:FC_D1244 | FC | Faraday cup |  |  | D_FC | 0 | 344.1458 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.1458 | 124.3675 | 124.3675 | borrowed from FS2
LS1 | CB01 | CAV5 | 1261 | LS1_CB01:CAV5_D1261 |  | b08 resonator | 34 |  |  | 0.29817 | 342.4247 | 423.0996 | 41.094 | 41.094 | 423.0996 | 342.2756 | 126.0887 | 126.2378 |  | 57.61 | D-station end
[10 empty] | 0 | 342.2756 | 423.0996 | 41.094 | 41.094 | 423.0996 | 342.2756 | 126.2378 | 126.2378 |  | 57.7591
[10 empty] | 0 | 342.2756 | 423.0996 | 41.094 | 41.094 | 423.0996 | 342.2756 | 126.2378 | 126.2378 |  | 57.7591
[6 empty] | bellow |  |  |  | 0.099517 | 342.2258 | 423.0996 | 41.094 | 41.094 | 423.0996 | 342.1761 | 126.2875 | 126.3373 |  | 57.8089
LS1 | CB01 | CAV6 | 1265 | LS1_CB01:CAV6_D1265 |  | b08 resonator | 34 |  |  | 0.29817 | 342.027 | 423.0996 | 41.094 | 41.094 | 423.0996 | 341.8779 | 126.4864 | 126.6354 |  | 58.0077
[10 empty] | 0 | 341.8779 | 423.0996 | 41.094 | 41.094 | 423.0996 | 341.8779 | 126.6354 | 126.6354 |  | 58.1568
[10 empty] | 0 | 341.8779 | 423.0996 | 41.094 | 41.094 | 423.0996 | 341.8779 | 126.6354 | 126.6354 |  | 58.1568
[6 empty] | bellow |  |  |  | 0.099517 | 341.8281 | 423.0996 | 41.094 | 41.094 | 423.0996 | 341.7784 | 126.6852 | 126.735 |  | 58.2066
LS1 | CB01 | CAV7 | 1269 | LS1_CB01:CAV7_D1269 |  | b08 resonator | 34 |  |  | 0.29817 | 341.6293 | 423.0996 | 41.094 | 41.094 | 423.0996 | 341.4802 | 126.884 | 127.0331 |  | 58.4054
[10 empty] | 0 | 341.4802 | 423.0996 | 41.094 | 41.094 | 423.0996 | 341.4802 | 127.0331 | 127.0331 |  | 58.5545
LS1 | CB01 | BLM | 1270 | LS1_CB01:BLM_D1270 |  | outside loss monitor |  |  |  | 0 | 341.4802 | 423.0996 | 41.094 | 41.094 | 423.0996 | 341.4802 | 127.0331 | 127.0331 |  | 58.5545
[6 empty] | bellow |  |  |  | 0.09651 | 341.432 | 423.0996 | 41.094 | 41.094 | 423.0996 | 341.3837 | 127.0814 | 127.1296 |  | 58.6028
LS1 | CB01 | BPM | 1271 | LS1_CB01:BPM_D1271 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 341.3837 | 423.0996 | 41.094 | 41.094 | 423.0996 | 341.3837 | 127.1296 | 127.1296 |  | 58.651
[6 empty] | BPM-box |  |  |  | 0.05588 | 341.3558 | 423.0996 | 41.094 | 41.094 | 423.0996 | 341.3278 | 127.1576 | 127.1855 |  | 58.6789
[6 empty] | solenoid-entry |  |  |  | 0.076136 | 341.2898 | 423.0996 | 41.094 | 41.094 | 423.0996 | 341.2517 | 127.2236 | 127.2617 |  | 58.745
LS1 | CB01 | SOL3 | 1275 | LS1_CB01:SOL3_D1275 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 341.0017 | 423.0996 | 41.094 | 41.094 | 423.0996 | 340.7517 | 127.5117 | 127.7617 |  | 59.033
[6 empty] | solenoid-exit |  |  |  | 0.076136 | 340.7136 | 423.0996 | 41.094 | 41.094 | 423.0996 | 340.6756 | 127.7997 | 127.8378 |  | 59.3211
[6 empty] | bellow |  |  |  | 0.099517 | 340.6258 | 423.0996 | 41.094 | 41.094 | 423.0996 | 340.576 | 127.8875 | 127.9373 |  | 59.4089
LS1 | CB01 | CAV8 | 1281 | LS1_CB01:CAV8_D1281 |  | b08 resonator | 34 |  |  | 0.29817 | 340.4269 | 423.0996 | 41.094 | 41.094 | 423.0996 | 340.2779 | 128.0864 | 128.2355 |  | 59.6078
[10 empty] | 0 | 340.2779 | 423.0996 | 41.094 | 41.094 | 423.0996 | 340.2779 | 128.2355 | 128.2355 |  | 59.7568
[10 empty] | 0 | 340.2779 | 423.0996 | 41.094 | 41.094 | 423.0996 | 340.2779 | 128.2355 | 128.2355 |  | 59.7568
[6 empty] | bellows |  |  |  | 0.19082 | 340.1825 | 423.0996 | 41.094 | 41.094 | 423.0996 | 340.087 | 128.3309 | 128.4263 |  | 59.8523
LS1 | CB01 | GV | 1285 | LS1_CB01:GV_D1285 |  | gate valve | 40 |  |  | 0.057404 | 340.0583 | 423.0996 | 41.094 | 41.094 | 423.0996 | 340.0296 | 128.455 | 128.4837 |  | 59.9764
[6 empty] | bellow | 40 |  |  | 0.1035 | 339.9779 | 423.0996 | 41.094 | 41.094 | 423.0996 | 339.9261 | 128.5355 | 128.5872 |  | 60.0568
LS1 | WB01 | BPM | 1286 | LS1_WB01:BPM_D1286 | BPM2 | position monitor |  |  |  | 0 | 339.9261 | 423.0996 | 41.094 | 41.094 | 423.0996 | 339.9261 | 128.5872 | 128.5872 |  | 60.1086
[6 empty] | diagnostic box |  |  |  | 0.14528 | 339.8535 | 423.0996 | 41.094 | 41.094 | 423.0996 | 339.7809 | 128.6598 | 128.7325 |  | 60.1812
LS1 | WB01 | HMR | 1287 | LS1_WB01:HMR_D1287 | HR_L | Halo ring, large aperture | 33 |  |  | 0 | 339.7809 | 423.0996 | 41.094 | 41.094 | 423.0996 | 339.7809 | 128.7325 | 128.7325 |  | 60.2539
[6 empty] | diagnostic box |  |  |  | 0 | 339.7809 | 423.0996 | 41.094 | 41.094 | 423.0996 | 339.7809 | 128.7325 | 128.7325 |  | 60.2539
LS1 | WB01 | IP | 1287 | LS1_WB01:IP_D1287 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 339.7809 | 423.0996 | 41.094 | 41.094 | 423.0996 | 339.7809 | 128.7325 | 128.7325 |  | 60.2539
[6 empty] | bellow | 40 |  |  | 0.13122 | 339.7152 | 423.0996 | 41.094 | 41.094 | 423.0996 | 339.6496 | 128.7981 | 128.8637 |  | 60.3195
LS1 | CB02 | GV | 1289 | LS1_CB02:GV_D1289 |  | gate valve | 40 |  |  | 0.057404 | 339.6209 | 423.0996 | 41.094 | 41.094 | 423.0996 | 339.5922 | 128.8924 | 128.9211 |  | 60.4138
[6 empty] | bellows |  |  |  | 0.19082 | 339.4968 | 423.0996 | 41.094 | 41.094 | 423.0996 | 339.4014 | 129.0165 | 129.1119 |  | 60.5379
LS1 | CB02 | CAV1 | 1293 | LS1_CB02:CAV1_D1293 |  | b08 resonator | 34 |  |  | 0.29817 | 339.2523 | 423.0996 | 41.094 | 41.094 | 423.0996 | 339.1032 | 129.261 | 129.4101 |  | 60.7824
[10 empty] | 0 | 339.1032 | 423.0996 | 41.094 | 41.094 | 423.0996 | 339.1032 | 129.4101 | 129.4101 |  | 60.9315
LS1 | CB02 | BLM | 1294 | LS1_CB02:BLM_D1294 |  | outside loss monitor |  |  |  | 0 | 339.1032 | 423.0996 | 41.094 | 41.094 | 423.0996 | 339.1032 | 129.4101 | 129.4101 |  | 60.9315
[6 empty] | bellow | 40 |  |  | 0.09651 | 339.055 | 423.0996 | 41.094 | 41.094 | 423.0996 | 339.0067 | 129.4584 | 129.5066 |  | 60.9797
LS1 | CB02 | BPM | 1295 | LS1_CB02:BPM_D1295 | BPM6 | position monitor, BPMc |  |  |  | 0 | 339.0067 | 423.0996 | 41.094 | 41.094 | 423.0996 | 339.0067 | 129.5066 | 129.5066 |  | 61.028
[6 empty] | BPM-box | 40 |  |  | 0.05588 | 338.9788 | 423.0996 | 41.094 | 41.094 | 423.0996 | 338.9508 | 129.5346 | 129.5625 |  | 61.0559
[6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 338.9128 | 423.0996 | 41.094 | 41.094 | 423.0996 | 338.8747 | 129.6006 | 129.6386 |  | 61.1219
LS1 | CB02 | SOL1 | 1299 | LS1_CB02:SOL1_D1299 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 338.6247 | 423.0996 | 41.094 | 41.094 | 423.0996 | 338.3747 | 129.8886 | 130.1386 |  | 61.41
[6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 338.3366 | 423.0996 | 41.094 | 41.094 | 423.0996 | 338.2986 | 130.1767 | 130.2148 |  | 61.6981
[6 empty] | bellow |  |  |  | 0.099517 | 338.2488 | 423.0996 | 41.094 | 41.094 | 423.0996 | 338.1991 | 130.2645 | 130.3143 |  | 61.7859
LS1 | CB02 | CAV2 | 1305 | LS1_CB02:CAV2_D1305 |  | b08 resonator | 34 |  |  | 0.29817 | 338.05 | 423.0996 | 41.094 | 41.094 | 423.0996 | 337.9009 | 130.4634 | 130.6125 |  | 61.9847
[10 empty] | 0 | 337.9009 | 423.0996 | 41.094 | 41.094 | 423.0996 | 337.9009 | 130.6125 | 130.6125 |  | 62.1338
[10 empty] | 0 | 337.9009 | 423.0996 | 41.094 | 41.094 | 423.0996 | 337.9009 | 130.6125 | 130.6125 |  | 62.1338
[6 empty] | bellow |  |  |  | 0.099517 | 337.8511 | 423.0996 | 41.094 | 41.094 | 423.0996 | 337.8014 | 130.6622 | 130.712 |  | 62.1836
LS1 | CB02 | CAV3 | 1309 | LS1_CB02:CAV3_D1309 |  | b08 resonator | 34 |  |  | 0.29817 | 337.6523 | 423.0996 | 41.094 | 41.094 | 423.0996 | 337.5032 | 130.8611 | 131.0101 |  | 62.3824
[10 empty] | 0 | 337.5032 | 423.0996 | 41.094 | 41.094 | 423.0996 | 337.5032 | 131.0101 | 131.0101 |  | 62.5315
[10 empty] | 0 | 337.5032 | 423.0996 | 41.094 | 41.094 | 423.0996 | 337.5032 | 131.0101 | 131.0101 |  | 62.5315
[6 empty] | bellow |  |  |  | 0.099517 | 337.4534 | 423.0996 | 41.094 | 41.094 | 423.0996 | 337.4037 | 131.0599 | 131.1097 |  | 62.5813
LS1 | CB02 | CAV4 | 1313 | LS1_CB02:CAV4_D1313 |  | b08 resonator | 34 |  |  | 0.29817 | 337.2546 | 423.0996 | 41.094 | 41.094 | 423.0996 | 337.1055 | 131.2587 | 131.4078 |  | 62.7801
[10 empty] | 0 | 337.1055 | 423.0996 | 41.094 | 41.094 | 423.0996 | 337.1055 | 131.4078 | 131.4078 |  | 62.9292
LS1 | CB02 | BLM | 1314 | LS1_CB02:BLM_D1314 |  | outside loss monitor |  |  |  | 0 | 337.1055 | 423.0996 | 41.094 | 41.094 | 423.0996 | 337.1055 | 131.4078 | 131.4078 |  | 62.9292
[6 empty] | bellow |  |  |  | 0.09651 | 337.0573 | 423.0996 | 41.094 | 41.094 | 423.0996 | 337.009 | 131.4561 | 131.5043 |  | 62.9775
LS1 | CB02 | BPM | 1315 | LS1_CB02:BPM_D1315 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 337.009 | 423.0996 | 41.094 | 41.094 | 423.0996 | 337.009 | 131.5043 | 131.5043 |  | 63.0257
[6 empty] | BPM-box |  |  |  | 0.05588 | 336.9811 | 423.0996 | 41.094 | 41.094 | 423.0996 | 336.9531 | 131.5323 | 131.5602 |  | 63.0536
[6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 336.9151 | 423.0996 | 41.094 | 41.094 | 423.0996 | 336.877 | 131.5983 | 131.6364 |  | 63.1197
LS1 | CB02 | SOL2 | 1319 | LS1_CB02:SOL2_D1319 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 336.627 | 423.0996 | 41.094 | 41.094 | 423.0996 | 336.377 | 131.8864 | 132.1364 |  | 63.4077
[6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 336.3389 | 423.0996 | 41.094 | 41.094 | 423.0996 | 336.3009 | 132.1744 | 132.2125 |  | 63.6958
[6 empty] | bellow |  |  |  | 0.099517 | 336.2511 | 423.0996 | 41.094 | 41.094 | 423.0996 | 336.2013 | 132.2622 | 132.312 |  | 63.7836
LS1 | CB02 | CAV5 | 1325 | LS1_CB02:CAV5_D1325 |  | b08 resonator | 34 |  |  | 0.29817 | 336.0522 | 423.0996 | 41.094 | 41.094 | 423.0996 | 335.9032 | 132.4611 | 132.6102 |  | 63.9825
[10 empty] | 0 | 335.9032 | 423.0996 | 41.094 | 41.094 | 423.0996 | 335.9032 | 132.6102 | 132.6102 |  | 64.1315
[10 empty] | 0 | 335.9032 | 423.0996 | 41.094 | 41.094 | 423.0996 | 335.9032 | 132.6102 | 132.6102 |  | 64.1315
[6 empty] | bellow |  |  |  | 0.099517 | 335.8534 | 423.0996 | 41.094 | 41.094 | 423.0996 | 335.8036 | 132.6599 | 132.7097 |  | 64.1813
LS1 | CB02 | CAV6 | 1329 | LS1_CB02:CAV6_D1329 |  | b08 resonator | 34 |  |  | 0.29817 | 335.6546 | 423.0996 | 41.094 | 41.094 | 423.0996 | 335.5055 | 132.8588 | 133.0079 |  | 64.3802
[10 empty] | 0 | 335.5055 | 423.0996 | 41.094 | 41.094 | 423.0996 | 335.5055 | 133.0079 | 133.0079 |  | 64.5292
[10 empty] | 0 | 335.5055 | 423.0996 | 41.094 | 41.094 | 423.0996 | 335.5055 | 133.0079 | 133.0079 |  | 64.5292
[6 empty] | bellow |  |  |  | 0.099517 | 335.4557 | 423.0996 | 41.094 | 41.094 | 423.0996 | 335.406 | 133.0576 | 133.1074 |  | 64.579
LS1 | CB02 | CAV7 | 1333 | LS1_CB02:CAV7_D1333 |  | b08 resonator | 34 |  |  | 0.29817 | 335.2569 | 423.0996 | 41.094 | 41.094 | 423.0996 | 335.1078 | 133.2565 | 133.4056 |  | 64.7778
[10 empty] | 0 | 335.1078 | 423.0996 | 41.094 | 41.094 | 423.0996 | 335.1078 | 133.4056 | 133.4056 |  | 64.9269
LS1 | CB02 | BLM | 1334 | LS1_CB02:BLM_D1334 |  | outside loss monitor |  |  |  | 0 | 335.1078 | 423.0996 | 41.094 | 41.094 | 423.0996 | 335.1078 | 133.4056 | 133.4056 |  | 64.9269
[6 empty] | bellow |  |  |  | 0.09651 | 335.0595 | 423.0996 | 41.094 | 41.094 | 423.0996 | 335.0113 | 133.4538 | 133.5021 |  | 64.9752
LS1 | CB02 | BPM | 1335 | LS1_CB02:BPM_D1335 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 335.0113 | 423.0996 | 41.094 | 41.094 | 423.0996 | 335.0113 | 133.5021 | 133.5021 |  | 65.0234
[6 empty] | BPM-box |  |  |  | 0.05588 | 334.9833 | 423.0996 | 41.094 | 41.094 | 423.0996 | 334.9554 | 133.53 | 133.5579 |  | 65.0514
[6 empty] | solenoid-entry |  |  |  | 0.076136 | 334.9173 | 423.0996 | 41.094 | 41.094 | 423.0996 | 334.8793 | 133.596 | 133.6341 |  | 65.1174
LS1 | CB02 | SOL3 | 1339 | LS1_CB02:SOL3_D1339 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 334.6293 | 423.0996 | 41.094 | 41.094 | 423.0996 | 334.3793 | 133.8841 | 134.1341 |  | 65.4054
[6 empty] | solenoid-exit |  |  |  | 0.076136 | 334.3412 | 423.0996 | 41.094 | 41.094 | 423.0996 | 334.3031 | 134.1721 | 134.2102 |  | 65.6935
[6 empty] | bellow |  |  |  | 0.099517 | 334.2534 | 423.0996 | 41.094 | 41.094 | 423.0996 | 334.2036 | 134.26 | 134.3097 |  | 65.7813
LS1 | CB02 | CAV8 | 1345 | LS1_CB02:CAV8_D1345 |  | b08 resonator | 34 |  |  | 0.29817 | 334.0545 | 423.0996 | 41.094 | 41.094 | 423.0996 | 333.9054 | 134.4588 | 134.6079 |  | 65.9802
[10 empty] | 0 | 333.9054 | 423.0996 | 41.094 | 41.094 | 423.0996 | 333.9054 | 134.6079 | 134.6079 |  | 66.1293
[10 empty] | 0 | 333.9054 | 423.0996 | 41.094 | 41.094 | 423.0996 | 333.9054 | 134.6079 | 134.6079 |  | 66.1293
[6 empty] | bellows |  |  |  | 0.19082 | 333.81 | 423.0996 | 41.094 | 41.094 | 423.0996 | 333.7146 | 134.7033 | 134.7987 |  | 66.2247
LS1 | CB02 | GV | 1348 | LS1_CB02:GV_D1348 |  | gate valve | 40 |  |  | 0.057404 | 333.6859 | 423.0996 | 41.094 | 41.094 | 423.0996 | 333.6572 | 134.8274 | 134.8561 |  | 66.3488
[6 empty] | bellow | 40 |  |  | 0.1035 | 333.6055 | 423.0996 | 41.094 | 41.094 | 423.0996 | 333.5537 | 134.9079 | 134.9596 |  | 66.4293
LS1 | WB02 | BPM | 1350 | LS1_WB02:BPM_D1350 | BPM2 | position monitor |  |  |  | 0 | 333.5537 | 423.0996 | 41.094 | 41.094 | 423.0996 | 333.5537 | 134.9596 | 134.9596 |  | 66.481
[6 empty] | diagnostic box |  |  |  | 0.14528 | 333.4811 | 423.0996 | 41.094 | 41.094 | 423.0996 | 333.4084 | 135.0323 | 135.1049 |  | 66.5536
LS1 | WB02 | HMR | 1351 | LS1_WB02:HMR_D1351 | HR_L | Halo ring, large aperture | 33 |  |  | 0 | 333.4084 | 423.0996 | 41.094 | 41.094 | 423.0996 | 333.4084 | 135.1049 | 135.1049 |  | 66.6263
[6 empty] | diagnostic box |  |  |  | 0 | 333.4084 | 423.0996 | 41.094 | 41.094 | 423.0996 | 333.4084 | 135.1049 | 135.1049 |  | 66.6263
LS1 | WB02 | IP | 1351 | LS1_WB02:IP_D1351 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 333.4084 | 423.0996 | 41.094 | 41.094 | 423.0996 | 333.4084 | 135.1049 | 135.1049 |  | 66.6263
[6 empty] | bellow | 40 |  |  | 0.13122 | 333.3428 | 423.0996 | 41.094 | 41.094 | 423.0996 | 333.2772 | 135.1705 | 135.2361 |  | 66.6919
LS1 | CB03 | GV | 1353 | LS1_CB03:GV_D1353 |  | gate valve | 40 |  |  | 0.057404 | 333.2485 | 423.0996 | 41.094 | 41.094 | 423.0996 | 333.2198 | 135.2648 | 135.2935 |  | 66.7862
[6 empty] | bellows |  |  |  | 0.19082 | 333.1244 | 423.0996 | 41.094 | 41.094 | 423.0996 | 333.029 | 135.3889 | 135.4844 |  | 66.9103
LS1 | CB03 | CAV1 | 1356 | LS1_CB03:CAV1_D1356 |  | b08 resonator | 34 |  |  | 0.29817 | 332.8799 | 423.0996 | 41.094 | 41.094 | 423.0996 | 332.7308 | 135.6334 | 135.7825 |  | 67.1548
[10 empty] | 0 | 332.7308 | 423.0996 | 41.094 | 41.094 | 423.0996 | 332.7308 | 135.7825 | 135.7825 |  | 67.3039
LS1 | CB03 | BLM | 1358 | LS1_CB03:BLM_D1358 |  | outside loss monitor |  |  |  | 0 | 332.7308 | 423.0996 | 41.094 | 41.094 | 423.0996 | 332.7308 | 135.7825 | 135.7825 |  | 67.3039
[6 empty] | bellow | 40 |  |  | 0.09651 | 332.6826 | 423.0996 | 41.094 | 41.094 | 423.0996 | 332.6343 | 135.8308 | 135.879 |  | 67.3522
LS1 | CB03 | BPM | 1359 | LS1_CB03:BPM_D1359 | BPM6 | position monitor, BPMc |  |  |  | 0 | 332.6343 | 423.0996 | 41.094 | 41.094 | 423.0996 | 332.6343 | 135.879 | 135.879 |  | 67.4004
[6 empty] | BPM-box | 40 |  |  | 0.05588 | 332.6064 | 423.0996 | 41.094 | 41.094 | 423.0996 | 332.5784 | 135.907 | 135.9349 |  | 67.4283
[6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 332.5404 | 423.0996 | 41.094 | 41.094 | 423.0996 | 332.5023 | 135.973 | 136.0111 |  | 67.4944
LS1 | CB03 | SOL1 | 1363 | LS1_CB03:SOL1_D1363 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 332.2523 | 423.0996 | 41.094 | 41.094 | 423.0996 | 332.0023 | 136.2611 | 136.5111 |  | 67.7824
[6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 331.9642 | 423.0996 | 41.094 | 41.094 | 423.0996 | 331.9262 | 136.5491 | 136.5872 |  | 68.0705
[6 empty] | bellow |  |  |  | 0.099517 | 331.8764 | 423.0996 | 41.094 | 41.094 | 423.0996 | 331.8266 | 136.6369 | 136.6867 |  | 68.1583
LS1 | CB03 | CAV2 | 1368 | LS1_CB03:CAV2_D1368 |  | b08 resonator | 34 |  |  | 0.29817 | 331.6776 | 423.0996 | 41.094 | 41.094 | 423.0996 | 331.5285 | 136.8358 | 136.9849 |  | 68.3572
[10 empty] | 0 | 331.5285 | 423.0996 | 41.094 | 41.094 | 423.0996 | 331.5285 | 136.9849 | 136.9849 |  | 68.5062
[10 empty] | 0 | 331.5285 | 423.0996 | 41.094 | 41.094 | 423.0996 | 331.5285 | 136.9849 | 136.9849 |  | 68.5062
[6 empty] | bellow |  |  |  | 0.099517 | 331.4787 | 423.0996 | 41.094 | 41.094 | 423.0996 | 331.4289 | 137.0346 | 137.0844 |  | 68.556
LS1 | CB03 | CAV3 | 1372 | LS1_CB03:CAV3_D1372 |  | b08 resonator | 34 |  |  | 0.29817 | 331.2799 | 423.0996 | 41.094 | 41.094 | 423.0996 | 331.1308 | 137.2335 | 137.3826 |  | 68.7548
[10 empty] | 0 | 331.1308 | 423.0996 | 41.094 | 41.094 | 423.0996 | 331.1308 | 137.3826 | 137.3826 |  | 68.9039
[10 empty] | 0 | 331.1308 | 423.0996 | 41.094 | 41.094 | 423.0996 | 331.1308 | 137.3826 | 137.3826 |  | 68.9039
[6 empty] | bellow |  |  |  | 0.099517 | 331.081 | 423.0996 | 41.094 | 41.094 | 423.0996 | 331.0313 | 137.4323 | 137.4821 |  | 68.9537
LS1 | CB03 | CAV4 | 1376 | LS1_CB03:CAV4_D1376 |  | b08 resonator | 34 |  |  | 0.29817 | 330.8822 | 423.0996 | 41.094 | 41.094 | 423.0996 | 330.7331 | 137.6312 | 137.7803 |  | 69.1525
[10 empty] | 0 | 330.7331 | 423.0996 | 41.094 | 41.094 | 423.0996 | 330.7331 | 137.7803 | 137.7803 |  | 69.3016
LS1 | CB03 | BLM | 1378 | LS1_CB03:BLM_D1378 |  | outside loss monitor |  |  |  | 0 | 330.7331 | 423.0996 | 41.094 | 41.094 | 423.0996 | 330.7331 | 137.7803 | 137.7803 |  | 69.3016
[6 empty] | bellow |  |  |  | 0.09651 | 330.6848 | 423.0996 | 41.094 | 41.094 | 423.0996 | 330.6366 | 137.8285 | 137.8768 |  | 69.3499
LS1 | CB03 | BPM | 1379 | LS1_CB03:BPM_D1379 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 330.6366 | 423.0996 | 41.094 | 41.094 | 423.0996 | 330.6366 | 137.8768 | 137.8768 |  | 69.3981
[6 empty] | BPM-box |  |  |  | 0.05588 | 330.6086 | 423.0996 | 41.094 | 41.094 | 423.0996 | 330.5807 | 137.9047 | 137.9326 |  | 69.4261
[6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 330.5426 | 423.0996 | 41.094 | 41.094 | 423.0996 | 330.5046 | 137.9707 | 138.0088 |  | 69.4921
LS1 | CB03 | SOL2 | 1383 | LS1_CB03:SOL2_D1383 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 330.2546 | 423.0996 | 41.094 | 41.094 | 423.0996 | 330.0046 | 138.2588 | 138.5088 |  | 69.7801
[6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 329.9665 | 423.0996 | 41.094 | 41.094 | 423.0996 | 329.9284 | 138.5468 | 138.5849 |  | 70.0682
[6 empty] | bellow |  |  |  | 0.099517 | 329.8787 | 423.0996 | 41.094 | 41.094 | 423.0996 | 329.8289 | 138.6347 | 138.6844 |  | 70.156
LS1 | CB03 | CAV5 | 1388 | LS1_CB03:CAV5_D1388 |  | b08 resonator | 34 |  |  | 0.29817 | 329.6798 | 423.0996 | 41.094 | 41.094 | 423.0996 | 329.5307 | 138.8335 | 138.9826 |  | 70.3549
[10 empty] | 0 | 329.5307 | 423.0996 | 41.094 | 41.094 | 423.0996 | 329.5307 | 138.9826 | 138.9826 |  | 70.504
[10 empty] | 0 | 329.5307 | 423.0996 | 41.094 | 41.094 | 423.0996 | 329.5307 | 138.9826 | 138.9826 |  | 70.504
[6 empty] | bellow |  |  |  | 0.099517 | 329.481 | 423.0996 | 41.094 | 41.094 | 423.0996 | 329.4312 | 139.0324 | 139.0821 |  | 70.5537
LS1 | CB03 | CAV6 | 1392 | LS1_CB03:CAV6_D1392 |  | b08 resonator | 34 |  |  | 0.29817 | 329.2821 | 423.0996 | 41.094 | 41.094 | 423.0996 | 329.1331 | 139.2312 | 139.3803 |  | 70.7526
[10 empty] | 0 | 329.1331 | 423.0996 | 41.094 | 41.094 | 423.0996 | 329.1331 | 139.3803 | 139.3803 |  | 70.9017
[10 empty] | 0 | 329.1331 | 423.0996 | 41.094 | 41.094 | 423.0996 | 329.1331 | 139.3803 | 139.3803 |  | 70.9017
[6 empty] | bellow |  |  |  | 0.099517 | 329.0833 | 423.0996 | 41.094 | 41.094 | 423.0996 | 329.0335 | 139.43 | 139.4798 |  | 70.9514
LS1 | CB03 | CAV7 | 1396 | LS1_CB03:CAV7_D1396 |  | b08 resonator | 34 |  |  | 0.29817 | 328.8844 | 423.0996 | 41.094 | 41.094 | 423.0996 | 328.7354 | 139.6289 | 139.778 |  | 71.1503
[10 empty] | 0 | 328.7354 | 423.0996 | 41.094 | 41.094 | 423.0996 | 328.7354 | 139.778 | 139.778 |  | 71.2993
LS1 | CB03 | BLM | 1398 | LS1_CB03:BLM_D1398 |  | outside loss monitor |  |  |  | 0 | 328.7354 | 423.0996 | 41.094 | 41.094 | 423.0996 | 328.7354 | 139.778 | 139.778 |  | 71.2993
[6 empty] | bellow |  |  |  | 0.09651 | 328.6871 | 423.0996 | 41.094 | 41.094 | 423.0996 | 328.6389 | 139.8262 | 139.8745 |  | 71.3476
LS1 | CB03 | BPM | 1399 | LS1_CB03:BPM_D1399 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 328.6389 | 423.0996 | 41.094 | 41.094 | 423.0996 | 328.6389 | 139.8745 | 139.8745 |  | 71.3959
[6 empty] | BPM-box |  |  |  | 0.05588 | 328.6109 | 423.0996 | 41.094 | 41.094 | 423.0996 | 328.583 | 139.9024 | 139.9304 |  | 71.4238
[6 empty] | solenoid-entry |  |  |  | 0.076136 | 328.5449 | 423.0996 | 41.094 | 41.094 | 423.0996 | 328.5068 | 139.9684 | 140.0065 |  | 71.4898
LS1 | CB03 | SOL3 | 1403 | LS1_CB03:SOL3_D1403 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 328.2568 | 423.0996 | 41.094 | 41.094 | 423.0996 | 328.0068 | 140.2565 | 140.5065 |  | 71.7779
[6 empty] | solenoid-exit |  |  |  | 0.076136 | 327.9688 | 423.0996 | 41.094 | 41.094 | 423.0996 | 327.9307 | 140.5446 | 140.5826 |  | 72.0659
[6 empty] | bellow |  |  |  | 0.099517 | 327.8809 | 423.0996 | 41.094 | 41.094 | 423.0996 | 327.8312 | 140.6324 | 140.6822 |  | 72.1538
LS1 | CB03 | CAV8 | 1408 | LS1_CB03:CAV8_D1408 |  | b08 resonator | 34 |  |  | 0.29817 | 327.6821 | 423.0996 | 41.094 | 41.094 | 423.0996 | 327.533 | 140.8312 | 140.9803 |  | 72.3526
[10 empty] | 0 | 327.533 | 423.0996 | 41.094 | 41.094 | 423.0996 | 327.533 | 140.9803 | 140.9803 |  | 72.5017
[10 empty] | 0 | 327.533 | 423.0996 | 41.094 | 41.094 | 423.0996 | 327.533 | 140.9803 | 140.9803 |  | 72.5017
[6 empty] | bellows |  |  |  | 0.19082 | 327.4376 | 423.0996 | 41.094 | 41.094 | 423.0996 | 327.3422 | 141.0757 | 141.1712 |  | 72.5971
LS1 | CB03 | GV | 1412 | LS1_CB03:GV_D1412 |  | gate valve | 40 |  |  | 0.057404 | 327.3135 | 423.0996 | 41.094 | 41.094 | 423.0996 | 327.2848 | 141.1999 | 141.2286 |  | 72.7212
[6 empty] | bellow | 40 |  |  | 0.1035 | 327.233 | 423.0996 | 41.094 | 41.094 | 423.0996 | 327.1813 | 141.2803 | 141.3321 |  | 72.8017
LS1 | WB03 | BPM | 1413 | LS1_WB03:BPM_D1413 | BPM2 | position monitor |  |  |  | 0 | 327.1813 | 423.0996 | 41.094 | 41.094 | 423.0996 | 327.1813 | 141.3321 | 141.3321 |  | 72.8534
[6 empty] | diagnostic box |  |  |  | 0.14528 | 327.1086 | 423.0996 | 41.094 | 41.094 | 423.0996 | 327.036 | 141.4047 | 141.4773 |  | 72.9261
LS1 | WB03 | HMR | 1415 | LS1_WB03:HMR_D1415 | HR_L | Halo ring, large aperture | 33 |  |  | 0 | 327.036 | 423.0996 | 41.094 | 41.094 | 423.0996 | 327.036 | 141.4773 | 141.4773 |  | 72.9987
[6 empty] | diagnostic box |  |  |  | 0 | 327.036 | 423.0996 | 41.094 | 41.094 | 423.0996 | 327.036 | 141.4773 | 141.4773 |  | 72.9987
LS1 | WB03 | IP | 1415 | LS1_WB03:IP_D1415 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 327.036 | 423.0996 | 41.094 | 41.094 | 423.0996 | 327.036 | 141.4773 | 141.4773 |  | 72.9987
[6 empty] | bellow | 40 |  |  | 0.13122 | 326.9704 | 423.0996 | 41.094 | 41.094 | 423.0996 | 326.9048 | 141.5429 | 141.6086 |  | 73.0643
LS1 | CB04 | GV | 1416 | LS1_CB04:GV_D1416 |  | gate valve | 40 |  |  | 0.057404 | 326.8761 | 423.0996 | 41.094 | 41.094 | 423.0996 | 326.8474 | 141.6373 | 141.666 |  | 73.1586
[6 empty] | bellows |  |  |  | 0.19082 | 326.752 | 423.0996 | 41.094 | 41.094 | 423.0996 | 326.6566 | 141.7614 | 141.8568 |  | 73.2827
LS1 | CB04 | CAV1 | 1420 | LS1_CB04:CAV1_D1420 |  | b08 resonator | 34 |  |  | 0.29817 | 326.5075 | 423.0996 | 41.094 | 41.094 | 423.0996 | 326.3584 | 142.0059 | 142.155 |  | 73.5272
[10 empty] | 0 | 326.3584 | 423.0996 | 41.094 | 41.094 | 423.0996 | 326.3584 | 142.155 | 142.155 |  | 73.6763
LS1 | CB04 | BLM | 1422 | LS1_CB04:BLM_D1422 |  | outside loss monitor |  |  |  | 0 | 326.3584 | 423.0996 | 41.094 | 41.094 | 423.0996 | 326.3584 | 142.155 | 142.155 |  | 73.6763
[6 empty] | bellow | 40 |  |  | 0.09651 | 326.3101 | 423.0996 | 41.094 | 41.094 | 423.0996 | 326.2619 | 142.2032 | 142.2515 |  | 73.7246
LS1 | CB04 | BPM | 1423 | LS1_CB04:BPM_D1423 | BPM6 | position monitor, BPMc |  |  |  | 0 | 326.2619 | 423.0996 | 41.094 | 41.094 | 423.0996 | 326.2619 | 142.2515 | 142.2515 |  | 73.7728
[6 empty] | BPM-box | 40 |  |  | 0.05588 | 326.2339 | 423.0996 | 41.094 | 41.094 | 423.0996 | 326.206 | 142.2794 | 142.3073 |  | 73.8008
[6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 326.1679 | 423.0996 | 41.094 | 41.094 | 423.0996 | 326.1299 | 142.3454 | 142.3835 |  | 73.8668
LS1 | CB04 | SOL1 | 1426 | LS1_CB04:SOL1_D1426 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 325.8799 | 423.0996 | 41.094 | 41.094 | 423.0996 | 325.6299 | 142.6335 | 142.8835 |  | 74.1548
[6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 325.5918 | 423.0996 | 41.094 | 41.094 | 423.0996 | 325.5537 | 142.9215 | 142.9596 |  | 74.4429
[6 empty] | bellow |  |  |  | 0.099517 | 325.504 | 423.0996 | 41.094 | 41.094 | 423.0996 | 325.4542 | 143.0094 | 143.0591 |  | 74.5307
LS1 | CB04 | CAV2 | 1432 | LS1_CB04:CAV2_D1432 |  | b08 resonator | 34 |  |  | 0.29817 | 325.3051 | 423.0996 | 41.094 | 41.094 | 423.0996 | 325.156 | 143.2082 | 143.3573 |  | 74.7296
[10 empty] | 0 | 325.156 | 423.0996 | 41.094 | 41.094 | 423.0996 | 325.156 | 143.3573 | 143.3573 |  | 74.8787
[10 empty] | 0 | 325.156 | 423.0996 | 41.094 | 41.094 | 423.0996 | 325.156 | 143.3573 | 143.3573 |  | 74.8787
[6 empty] | bellow |  |  |  | 0.099517 | 325.1063 | 423.0996 | 41.094 | 41.094 | 423.0996 | 325.0565 | 143.4071 | 143.4568 |  | 74.9284
LS1 | CB04 | CAV3 | 1436 | LS1_CB04:CAV3_D1436 |  | b08 resonator | 34 |  |  | 0.29817 | 324.9074 | 423.0996 | 41.094 | 41.094 | 423.0996 | 324.7584 | 143.6059 | 143.755 |  | 75.1273
[10 empty] | 0 | 324.7584 | 423.0996 | 41.094 | 41.094 | 423.0996 | 324.7584 | 143.755 | 143.755 |  | 75.2764
[10 empty] | 0 | 324.7584 | 423.0996 | 41.094 | 41.094 | 423.0996 | 324.7584 | 143.755 | 143.755 |  | 75.2764
[6 empty] | bellow |  |  |  | 0.099517 | 324.7086 | 423.0996 | 41.094 | 41.094 | 423.0996 | 324.6588 | 143.8047 | 143.8545 |  | 75.3261
LS1 | CB04 | CAV4 | 1440 | LS1_CB04:CAV4_D1440 |  | b08 resonator | 34 |  |  | 0.29817 | 324.5097 | 423.0996 | 41.094 | 41.094 | 423.0996 | 324.3607 | 144.0036 | 144.1527 |  | 75.525
[10 empty] | 0 | 324.3607 | 423.0996 | 41.094 | 41.094 | 423.0996 | 324.3607 | 144.1527 | 144.1527 |  | 75.674
LS1 | CB04 | BLM | 1442 | LS1_CB04:BLM_D1442 |  | outside loss monitor |  |  |  | 0 | 324.3607 | 423.0996 | 41.094 | 41.094 | 423.0996 | 324.3607 | 144.1527 | 144.1527 |  | 75.674
[6 empty] | bellow |  |  |  | 0.09651 | 324.3124 | 423.0996 | 41.094 | 41.094 | 423.0996 | 324.2642 | 144.2009 | 144.2492 |  | 75.7223
LS1 | CB04 | BPM | 1442 | LS1_CB04:BPM_D1442 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 324.2642 | 423.0996 | 41.094 | 41.094 | 423.0996 | 324.2642 | 144.2492 | 144.2492 |  | 75.7706
[6 empty] | BPM-box |  |  |  | 0.05588 | 324.2362 | 423.0996 | 41.094 | 41.094 | 423.0996 | 324.2083 | 144.2771 | 144.3051 |  | 75.7985
[6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 324.1702 | 423.0996 | 41.094 | 41.094 | 423.0996 | 324.1321 | 144.3431 | 144.3812 |  | 75.8645
LS1 | CB04 | SOL2 | 1446 | LS1_CB04:SOL2_D1446 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 323.8821 | 423.0996 | 41.094 | 41.094 | 423.0996 | 323.6321 | 144.6312 | 144.8812 |  | 76.1526
[6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 323.5941 | 423.0996 | 41.094 | 41.094 | 423.0996 | 323.556 | 144.9193 | 144.9573 |  | 76.4406
[6 empty] | bellow |  |  |  | 0.099517 | 323.5062 | 423.0996 | 41.094 | 41.094 | 423.0996 | 323.4565 | 145.0071 | 145.0569 |  | 76.5285
LS1 | CB04 | CAV5 | 1452 | LS1_CB04:CAV5_D1452 |  | b08 resonator | 34 |  |  | 0.29817 | 323.3074 | 423.0996 | 41.094 | 41.094 | 423.0996 | 323.1583 | 145.2059 | 145.355 |  | 76.7273
[10 empty] | 0 | 323.1583 | 423.0996 | 41.094 | 41.094 | 423.0996 | 323.1583 | 145.355 | 145.355 |  | 76.8764
[10 empty] | 0 | 323.1583 | 423.0996 | 41.094 | 41.094 | 423.0996 | 323.1583 | 145.355 | 145.355 |  | 76.8764
[6 empty] | bellow |  |  |  | 0.099517 | 323.1086 | 423.0996 | 41.094 | 41.094 | 423.0996 | 323.0588 | 145.4048 | 145.4545 |  | 76.9262
LS1 | CB04 | CAV6 | 1456 | LS1_CB04:CAV6_D1456 |  | b08 resonator | 34 |  |  | 0.29817 | 322.9097 | 423.0996 | 41.094 | 41.094 | 423.0996 | 322.7606 | 145.6036 | 145.7527 |  | 77.125
[10 empty] | 0 | 322.7606 | 423.0996 | 41.094 | 41.094 | 423.0996 | 322.7606 | 145.7527 | 145.7527 |  | 77.2741
[10 empty] | 0 | 322.7606 | 423.0996 | 41.094 | 41.094 | 423.0996 | 322.7606 | 145.7527 | 145.7527 |  | 77.2741
[6 empty] | bellow |  |  |  | 0.099517 | 322.7109 | 423.0996 | 41.094 | 41.094 | 423.0996 | 322.6611 | 145.8025 | 145.8522 |  | 77.3238
LS1 | CB04 | CAV7 | 1460 | LS1_CB04:CAV7_D1460 |  | b08 resonator | 34 |  |  | 0.29817 | 322.512 | 423.0996 | 41.094 | 41.094 | 423.0996 | 322.3629 | 146.0013 | 146.1504 |  | 77.5227
[10 empty] | 0 | 322.3629 | 423.0996 | 41.094 | 41.094 | 423.0996 | 322.3629 | 146.1504 | 146.1504 |  | 77.6718
LS1 | CB04 | BLM | 1462 | LS1_CB04:BLM_D1462 |  | outside loss monitor |  |  |  | 0 | 322.3629 | 423.0996 | 41.094 | 41.094 | 423.0996 | 322.3629 | 146.1504 | 146.1504 |  | 77.6718
[6 empty] | bellow |  |  |  | 0.09651 | 322.3147 | 423.0996 | 41.094 | 41.094 | 423.0996 | 322.2664 | 146.1987 | 146.2469 |  | 77.72
LS1 | CB04 | BPM | 1462 | LS1_CB04:BPM_D1462 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 322.2664 | 423.0996 | 41.094 | 41.094 | 423.0996 | 322.2664 | 146.2469 | 146.2469 |  | 77.7683
[6 empty] | BPM-box |  |  |  | 0.05588 | 322.2385 | 423.0996 | 41.094 | 41.094 | 423.0996 | 322.2105 | 146.2749 | 146.3028 |  | 77.7962
[6 empty] | solenoid-entry |  |  |  | 0.076136 | 322.1725 | 423.0996 | 41.094 | 41.094 | 423.0996 | 322.1344 | 146.3409 | 146.3789 |  | 77.8622
LS1 | CB04 | SOL3 | 1466 | LS1_CB04:SOL3_D1466 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 321.8844 | 423.0996 | 41.094 | 41.094 | 423.0996 | 321.6344 | 146.6289 | 146.8789 |  | 78.1503
[6 empty] | solenoid-exit |  |  |  | 0.076136 | 321.5963 | 423.0996 | 41.094 | 41.094 | 423.0996 | 321.5583 | 146.917 | 146.9551 |  | 78.4384
[6 empty] | bellow |  |  |  | 0.099517 | 321.5085 | 423.0996 | 41.094 | 41.094 | 423.0996 | 321.4588 | 147.0048 | 147.0546 |  | 78.5262
LS1 | CB04 | CAV8 | 1472 | LS1_CB04:CAV8_D1472 |  | b08 resonator | 34 |  |  | 0.29817 | 321.3097 | 423.0996 | 41.094 | 41.094 | 423.0996 | 321.1606 | 147.2037 | 147.3528 |  | 78.725
[10 empty] | 0 | 321.1606 | 423.0996 | 41.094 | 41.094 | 423.0996 | 321.1606 | 147.3528 | 147.3528 |  | 78.8741
[10 empty] | 0 | 321.1606 | 423.0996 | 41.094 | 41.094 | 423.0996 | 321.1606 | 147.3528 | 147.3528 |  | 78.8741
[6 empty] | bellows |  |  |  | 0.19082 | 321.0652 | 423.0996 | 41.094 | 41.094 | 423.0996 | 320.9698 | 147.4482 | 147.5436 |  | 78.9695
LS1 | CB04 | GV | 1476 | LS1_CB04:GV_D1476 |  | gate valve | 40 |  |  | 0.057404 | 320.9411 | 423.0996 | 41.094 | 41.094 | 423.0996 | 320.9124 | 147.5723 | 147.601 |  | 79.0936
[6 empty] | bellow | 40 |  |  | 0.1035 | 320.8606 | 423.0996 | 41.094 | 41.094 | 423.0996 | 320.8089 | 147.6527 | 147.7045 |  | 79.1741
LS1 | WB04 | BPM | 1477 | LS1_WB04:BPM_D1477 | BPM2 | position monitor |  |  |  | 0 | 320.8089 | 423.0996 | 41.094 | 41.094 | 423.0996 | 320.8089 | 147.7045 | 147.7045 |  | 79.2259
[6 empty] | diagnostic box |  |  |  | 0.14528 | 320.7362 | 423.0996 | 41.094 | 41.094 | 423.0996 | 320.6636 | 147.7771 | 147.8498 |  | 79.2985
LS1 | WB04 | HMR | 1478 | LS1_WB04:HMR_D1478 | HR_L | Halo ring, large aperture | 33 |  |  | 0 | 320.6636 | 423.0996 | 41.094 | 41.094 | 423.0996 | 320.6636 | 147.8498 | 147.8498 |  | 79.3711
[6 empty] | diagnostic box |  |  |  | 0 | 320.6636 | 423.0996 | 41.094 | 41.094 | 423.0996 | 320.6636 | 147.8498 | 147.8498 |  | 79.3711
LS1 | WB04 | IP | 1478 | LS1_WB04:IP_D1478 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 320.6636 | 423.0996 | 41.094 | 41.094 | 423.0996 | 320.6636 | 147.8498 | 147.8498 |  | 79.3711
[6 empty] | bellow | 40 |  |  | 0.13122 | 320.598 | 423.0996 | 41.094 | 41.094 | 423.0996 | 320.5324 | 147.9154 | 147.981 |  | 79.4367
LS1 | CB05 | GV | 1480 | LS1_CB05:GV_D1480 |  | gate valve | 40 |  |  | 0.057404 | 320.5037 | 423.0996 | 41.094 | 41.094 | 423.0996 | 320.475 | 148.0097 | 148.0384 |  | 79.5311
[6 empty] | bellows |  |  |  | 0.19082 | 320.3795 | 423.0996 | 41.094 | 41.094 | 423.0996 | 320.2841 | 148.1338 | 148.2292 |  | 79.6552
LS1 | CB05 | CAV1 | 1484 | LS1_CB05:CAV1_D1484 |  | b08 resonator | 34 |  |  | 0.29817 | 320.1351 | 423.0996 | 41.094 | 41.094 | 423.0996 | 319.986 | 148.3783 | 148.5274 |  | 79.8997
[10 empty] | 0 | 319.986 | 423.0996 | 41.094 | 41.094 | 423.0996 | 319.986 | 148.5274 | 148.5274 |  | 80.0487
LS1 | CB05 | BLM | 1485 | LS1_CB05:BLM_D1485 |  | outside loss monitor |  |  |  | 0 | 319.986 | 423.0996 | 41.094 | 41.094 | 423.0996 | 319.986 | 148.5274 | 148.5274 |  | 80.0487
[6 empty] | bellow | 40 |  |  | 0.09651 | 319.9377 | 423.0996 | 41.094 | 41.094 | 423.0996 | 319.8895 | 148.5756 | 148.6239 |  | 80.097
LS1 | CB05 | BPM | 1486 | LS1_CB05:BPM_D1486 | BPM6 | position monitor, BPMc |  |  |  | 0 | 319.8895 | 423.0996 | 41.094 | 41.094 | 423.0996 | 319.8895 | 148.6239 | 148.6239 |  | 80.1453
[6 empty] | BPM-box | 40 |  |  | 0.05588 | 319.8615 | 423.0996 | 41.094 | 41.094 | 423.0996 | 319.8336 | 148.6518 | 148.6798 |  | 80.1732
[6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 319.7955 | 423.0996 | 41.094 | 41.094 | 423.0996 | 319.7574 | 148.7178 | 148.7559 |  | 80.2392
LS1 | CB05 | SOL1 | 1490 | LS1_CB05:SOL1_D1490 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 319.5074 | 423.0996 | 41.094 | 41.094 | 423.0996 | 319.2574 | 149.0059 | 149.2559 |  | 80.5273
[6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 319.2194 | 423.0996 | 41.094 | 41.094 | 423.0996 | 319.1813 | 149.294 | 149.332 |  | 80.8153
[6 empty] | bellow |  |  |  | 0.099517 | 319.1315 | 423.0996 | 41.094 | 41.094 | 423.0996 | 319.0818 | 149.3818 | 149.4316 |  | 80.9032
LS1 | CB05 | CAV2 | 1496 | LS1_CB05:CAV2_D1496 |  | b08 resonator | 34 |  |  | 0.29817 | 318.9327 | 423.0996 | 41.094 | 41.094 | 423.0996 | 318.7836 | 149.5806 | 149.7297 |  | 81.102
[10 empty] | 0 | 318.7836 | 423.0996 | 41.094 | 41.094 | 423.0996 | 318.7836 | 149.7297 | 149.7297 |  | 81.2511
[10 empty] | 0 | 318.7836 | 423.0996 | 41.094 | 41.094 | 423.0996 | 318.7836 | 149.7297 | 149.7297 |  | 81.2511
[6 empty] | bellow |  |  |  | 0.099517 | 318.7339 | 423.0996 | 41.094 | 41.094 | 423.0996 | 318.6841 | 149.7795 | 149.8292 |  | 81.3009
LS1 | CB05 | CAV3 | 1500 | LS1_CB05:CAV3_D1500 |  | b08 resonator | 34 |  |  | 0.29817 | 318.535 | 423.0996 | 41.094 | 41.094 | 423.0996 | 318.3859 | 149.9783 | 150.1274 |  | 81.4997
[10 empty] | 0 | 318.3859 | 423.0996 | 41.094 | 41.094 | 423.0996 | 318.3859 | 150.1274 | 150.1274 |  | 81.6488
[10 empty] | 0 | 318.3859 | 423.0996 | 41.094 | 41.094 | 423.0996 | 318.3859 | 150.1274 | 150.1274 |  | 81.6488
[6 empty] | bellow |  |  |  | 0.099517 | 318.3362 | 423.0996 | 41.094 | 41.094 | 423.0996 | 318.2864 | 150.1772 | 150.2269 |  | 81.6985
LS1 | CB05 | CAV4 | 1504 | LS1_CB05:CAV4_D1504 |  | b08 resonator | 34 |  |  | 0.29817 | 318.1373 | 423.0996 | 41.094 | 41.094 | 423.0996 | 317.9882 | 150.376 | 150.5251 |  | 81.8974
[10 empty] | 0 | 317.9882 | 423.0996 | 41.094 | 41.094 | 423.0996 | 317.9882 | 150.5251 | 150.5251 |  | 82.0465
LS1 | CB05 | BLM | 1505 | LS1_CB05:BLM_D1505 |  | outside loss monitor |  |  |  | 0 | 317.9882 | 423.0996 | 41.094 | 41.094 | 423.0996 | 317.9882 | 150.5251 | 150.5251 |  | 82.0465
[6 empty] | bellow |  |  |  | 0.09651 | 317.94 | 423.0996 | 41.094 | 41.094 | 423.0996 | 317.8917 | 150.5734 | 150.6216 |  | 82.0947
LS1 | CB05 | BPM | 1506 | LS1_CB05:BPM_D1506 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 317.8917 | 423.0996 | 41.094 | 41.094 | 423.0996 | 317.8917 | 150.6216 | 150.6216 |  | 82.143
[6 empty] | BPM-box |  |  |  | 0.05588 | 317.8638 | 423.0996 | 41.094 | 41.094 | 423.0996 | 317.8359 | 150.6496 | 150.6775 |  | 82.1709
[6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 317.7978 | 423.0996 | 41.094 | 41.094 | 423.0996 | 317.7597 | 150.7156 | 150.7536 |  | 82.2369
LS1 | CB05 | SOL2 | 1510 | LS1_CB05:SOL2_D1510 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 317.5097 | 423.0996 | 41.094 | 41.094 | 423.0996 | 317.2597 | 151.0036 | 151.2536 |  | 82.525
[6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 317.2216 | 423.0996 | 41.094 | 41.094 | 423.0996 | 317.1836 | 151.2917 | 151.3298 |  | 82.8131
[6 empty] | bellow |  |  |  | 0.099517 | 317.1338 | 423.0996 | 41.094 | 41.094 | 423.0996 | 317.0841 | 151.3795 | 151.4293 |  | 82.9009
LS1 | CB05 | CAV5 | 1516 | LS1_CB05:CAV5_D1516 |  | b08 resonator | 34 |  |  | 0.29817 | 316.935 | 423.0996 | 41.094 | 41.094 | 423.0996 | 316.7859 | 151.5784 | 151.7275 |  | 83.0997
[10 empty] | 0 | 316.7859 | 423.0996 | 41.094 | 41.094 | 423.0996 | 316.7859 | 151.7275 | 151.7275 |  | 83.2488
[10 empty] | 0 | 316.7859 | 423.0996 | 41.094 | 41.094 | 423.0996 | 316.7859 | 151.7275 | 151.7275 |  | 83.2488
[6 empty] | bellow |  |  |  | 0.099517 | 316.7361 | 423.0996 | 41.094 | 41.094 | 423.0996 | 316.6864 | 151.7772 | 151.827 |  | 83.2986
LS1 | CB05 | CAV6 | 1520 | LS1_CB05:CAV6_D1520 |  | b08 resonator | 34 |  |  | 0.29817 | 316.5373 | 423.0996 | 41.094 | 41.094 | 423.0996 | 316.3882 | 151.9761 | 152.1251 |  | 83.4974
[10 empty] | 0 | 316.3882 | 423.0996 | 41.094 | 41.094 | 423.0996 | 316.3882 | 152.1251 | 152.1251 |  | 83.6465
[10 empty] | 0 | 316.3882 | 423.0996 | 41.094 | 41.094 | 423.0996 | 316.3882 | 152.1251 | 152.1251 |  | 83.6465
[6 empty] | bellow |  |  |  | 0.099517 | 316.3384 | 423.0996 | 41.094 | 41.094 | 423.0996 | 316.2887 | 152.1749 | 152.2247 |  | 83.6963
LS1 | CB05 | CAV7 | 1524 | LS1_CB05:CAV7_D1524 |  | b08 resonator | 34 |  |  | 0.29817 | 316.1396 | 423.0996 | 41.094 | 41.094 | 423.0996 | 315.9905 | 152.3737 | 152.5228 |  | 83.8951
[10 empty] | 0 | 315.9905 | 423.0996 | 41.094 | 41.094 | 423.0996 | 315.9905 | 152.5228 | 152.5228 |  | 84.0442
LS1 | CB05 | BLM | 1525 | LS1_CB05:BLM_D1525 |  | outside loss monitor |  |  |  | 0 | 315.9905 | 423.0996 | 41.094 | 41.094 | 423.0996 | 315.9905 | 152.5228 | 152.5228 |  | 84.0442
[6 empty] | bellow |  |  |  | 0.09651 | 315.9423 | 423.0996 | 41.094 | 41.094 | 423.0996 | 315.894 | 152.5711 | 152.6193 |  | 84.0925
LS1 | CB05 | BPM | 1526 | LS1_CB05:BPM_D1526 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 315.894 | 423.0996 | 41.094 | 41.094 | 423.0996 | 315.894 | 152.6193 | 152.6193 |  | 84.1407
[6 empty] | BPM-box |  |  |  | 0.05588 | 315.8661 | 423.0996 | 41.094 | 41.094 | 423.0996 | 315.8381 | 152.6473 | 152.6752 |  | 84.1686
[6 empty] | solenoid-entry |  |  |  | 0.076136 | 315.8001 | 423.0996 | 41.094 | 41.094 | 423.0996 | 315.762 | 152.7133 | 152.7514 |  | 84.2347
LS1 | CB05 | SOL3 | 1530 | LS1_CB05:SOL3_D1530 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 315.512 | 423.0996 | 41.094 | 41.094 | 423.0996 | 315.262 | 153.0014 | 153.2514 |  | 84.5227
[6 empty] | solenoid-exit |  |  |  | 0.076136 | 315.2239 | 423.0996 | 41.094 | 41.094 | 423.0996 | 315.1859 | 153.2894 | 153.3275 |  | 84.8108
[6 empty] | bellow |  |  |  | 0.099517 | 315.1361 | 423.0996 | 41.094 | 41.094 | 423.0996 | 315.0863 | 153.3772 | 153.427 |  | 84.8986
LS1 | CB05 | CAV8 | 1536 | LS1_CB05:CAV8_D1536 |  | b08 resonator | 34 |  |  | 0.29817 | 314.9372 | 423.0996 | 41.094 | 41.094 | 423.0996 | 314.7882 | 153.5761 | 153.7252 |  | 85.0975
[10 empty] | 0 | 314.7882 | 423.0996 | 41.094 | 41.094 | 423.0996 | 314.7882 | 153.7252 | 153.7252 |  | 85.2465
[10 empty] | 0 | 314.7882 | 423.0996 | 41.094 | 41.094 | 423.0996 | 314.7882 | 153.7252 | 153.7252 |  | 85.2465
[6 empty] | bellows |  |  |  | 0.19082 | 314.6928 | 423.0996 | 41.094 | 41.094 | 423.0996 | 314.5973 | 153.8206 | 153.916 |  | 85.342
LS1 | CB05 | GV | 1539 | LS1_CB05:GV_D1539 |  | gate valve | 40 |  |  | 0.057404 | 314.5686 | 423.0996 | 41.094 | 41.094 | 423.0996 | 314.5399 | 153.9447 | 153.9734 |  | 85.4661
[6 empty] | bellow | 40 |  |  | 0.1035 | 314.4882 | 423.0996 | 41.094 | 41.094 | 423.0996 | 314.4364 | 154.0252 | 154.0769 |  | 85.5465
LS1 | WB05 | BPM | 1541 | LS1_WB05:BPM_D1541 | BPM2 | position monitor |  |  |  | 0 | 314.4364 | 423.0996 | 41.094 | 41.094 | 423.0996 | 314.4364 | 154.0769 | 154.0769 |  | 85.5983
[6 empty] | diagnostic box |  |  |  | 0.14528 | 314.3638 | 423.0996 | 41.094 | 41.094 | 423.0996 | 314.2912 | 154.1495 | 154.2222 |  | 85.6709
LS1 | WB05 | HMR | 1542 | LS1_WB05:HMR_D1542 | HR_L | Halo ring, large aperture | 33 |  |  | 0 | 314.2912 | 423.0996 | 41.094 | 41.094 | 423.0996 | 314.2912 | 154.2222 | 154.2222 |  | 85.7436
[6 empty] | diagnostic box |  |  |  | 0 | 314.2912 | 423.0996 | 41.094 | 41.094 | 423.0996 | 314.2912 | 154.2222 | 154.2222 |  | 85.7436
LS1 | WB05 | IP | 1542 | LS1_WB05:IP_D1542 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 314.2912 | 423.0996 | 41.094 | 41.094 | 423.0996 | 314.2912 | 154.2222 | 154.2222 |  | 85.7436
[6 empty] | bellow | 40 |  |  | 0.13122 | 314.2255 | 423.0996 | 41.094 | 41.094 | 423.0996 | 314.1599 | 154.2878 | 154.3534 |  | 85.8092
LS1 | CB06 | GV | 1544 | LS1_CB06:GV_D1544 |  | gate valve | 40 |  |  | 0.057404 | 314.1312 | 423.0996 | 41.094 | 41.094 | 423.0996 | 314.1025 | 154.3821 | 154.4108 |  | 85.9035
[6 empty] | bellows |  |  |  | 0.19082 | 314.0071 | 423.0996 | 41.094 | 41.094 | 423.0996 | 313.9117 | 154.5062 | 154.6016 |  | 86.0276
LS1 | CB06 | CAV1 | 1548 | LS1_CB06:CAV1_D1548 |  | b08 resonator | 34 |  |  | 0.29817 | 313.7626 | 423.0996 | 41.094 | 41.094 | 423.0996 | 313.6135 | 154.7507 | 154.8998 |  | 86.2721
[10 empty] | 0 | 313.6135 | 423.0996 | 41.094 | 41.094 | 423.0996 | 313.6135 | 154.8998 | 154.8998 |  | 86.4212
LS1 | CB06 | BLM | 1549 | LS1_CB06:BLM_D1549 |  | outside loss monitor |  |  |  | 0 | 313.6135 | 423.0996 | 41.094 | 41.094 | 423.0996 | 313.6135 | 154.8998 | 154.8998 |  | 86.4212
[6 empty] | bellow | 40 |  |  | 0.09651 | 313.5653 | 423.0996 | 41.094 | 41.094 | 423.0996 | 313.517 | 154.9481 | 154.9963 |  | 86.4694
LS1 | CB06 | BPM | 1550 | LS1_CB06:BPM_D1550 | BPM6 | position monitor, BPMc |  |  |  | 0 | 313.517 | 423.0996 | 41.094 | 41.094 | 423.0996 | 313.517 | 154.9963 | 154.9963 |  | 86.5177
[6 empty] | BPM-box | 40 |  |  | 0.05588 | 313.4891 | 423.0996 | 41.094 | 41.094 | 423.0996 | 313.4612 | 155.0242 | 155.0522 |  | 86.5456
[6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 313.4231 | 423.0996 | 41.094 | 41.094 | 423.0996 | 313.385 | 155.0903 | 155.1283 |  | 86.6116
LS1 | CB06 | SOL1 | 1554 | LS1_CB06:SOL1_D1554 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 313.135 | 423.0996 | 41.094 | 41.094 | 423.0996 | 312.885 | 155.3783 | 155.6283 |  | 86.8997
[6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 312.8469 | 423.0996 | 41.094 | 41.094 | 423.0996 | 312.8089 | 155.6664 | 155.7045 |  | 87.1878
[6 empty] | bellow |  |  |  | 0.099517 | 312.7591 | 423.0996 | 41.094 | 41.094 | 423.0996 | 312.7094 | 155.7542 | 155.804 |  | 87.2756
LS1 | CB06 | CAV2 | 1560 | LS1_CB06:CAV2_D1560 |  | b08 resonator | 34 |  |  | 0.29817 | 312.5603 | 423.0996 | 41.094 | 41.094 | 423.0996 | 312.4112 | 155.9531 | 156.1021 |  | 87.4744
[10 empty] | 0 | 312.4112 | 423.0996 | 41.094 | 41.094 | 423.0996 | 312.4112 | 156.1021 | 156.1021 |  | 87.6235
[10 empty] | 0 | 312.4112 | 423.0996 | 41.094 | 41.094 | 423.0996 | 312.4112 | 156.1021 | 156.1021 |  | 87.6235
[6 empty] | bellow |  |  |  | 0.099517 | 312.3614 | 423.0996 | 41.094 | 41.094 | 423.0996 | 312.3117 | 156.1519 | 156.2017 |  | 87.6733
LS1 | CB06 | CAV3 | 1564 | LS1_CB06:CAV3_D1564 |  | b08 resonator | 34 |  |  | 0.29817 | 312.1626 | 423.0996 | 41.094 | 41.094 | 423.0996 | 312.0135 | 156.3508 | 156.4998 |  | 87.8721
[10 empty] | 0 | 312.0135 | 423.0996 | 41.094 | 41.094 | 423.0996 | 312.0135 | 156.4998 | 156.4998 |  | 88.0212
[10 empty] | 0 | 312.0135 | 423.0996 | 41.094 | 41.094 | 423.0996 | 312.0135 | 156.4998 | 156.4998 |  | 88.0212
[6 empty] | bellow |  |  |  | 0.099517 | 311.9637 | 423.0996 | 41.094 | 41.094 | 423.0996 | 311.914 | 156.5496 | 156.5994 |  | 88.071
LS1 | CB06 | CAV4 | 1567 | LS1_CB06:CAV4_D1567 |  | b08 resonator | 34 |  |  | 0.29817 | 311.7649 | 423.0996 | 41.094 | 41.094 | 423.0996 | 311.6158 | 156.7484 | 156.8975 |  | 88.2698
[10 empty] | 0 | 311.6158 | 423.0996 | 41.094 | 41.094 | 423.0996 | 311.6158 | 156.8975 | 156.8975 |  | 88.4189
LS1 | CB06 | BLM | 1569 | LS1_CB06:BLM_D1569 |  | outside loss monitor |  |  |  | 0 | 311.6158 | 423.0996 | 41.094 | 41.094 | 423.0996 | 311.6158 | 156.8975 | 156.8975 |  | 88.4189
[6 empty] | bellow |  |  |  | 0.09651 | 311.5676 | 423.0996 | 41.094 | 41.094 | 423.0996 | 311.5193 | 156.9458 | 156.994 |  | 88.4672
LS1 | CB06 | BPM | 1570 | LS1_CB06:BPM_D1570 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 311.5193 | 423.0996 | 41.094 | 41.094 | 423.0996 | 311.5193 | 156.994 | 156.994 |  | 88.5154
[6 empty] | BPM-box |  |  |  | 0.05588 | 311.4914 | 423.0996 | 41.094 | 41.094 | 423.0996 | 311.4634 | 157.022 | 157.0499 |  | 88.5433
[6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 311.4254 | 423.0996 | 41.094 | 41.094 | 423.0996 | 311.3873 | 157.088 | 157.1261 |  | 88.6094
LS1 | CB06 | SOL2 | 1574 | LS1_CB06:SOL2_D1574 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 311.1373 | 423.0996 | 41.094 | 41.094 | 423.0996 | 310.8873 | 157.3761 | 157.6261 |  | 88.8974
[6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 310.8492 | 423.0996 | 41.094 | 41.094 | 423.0996 | 310.8112 | 157.6641 | 157.7022 |  | 89.1855
[6 empty] | bellow |  |  |  | 0.099517 | 310.7614 | 423.0996 | 41.094 | 41.094 | 423.0996 | 310.7116 | 157.7519 | 157.8017 |  | 89.2733
LS1 | CB06 | CAV5 | 1580 | LS1_CB06:CAV5_D1580 |  | b08 resonator | 34 |  |  | 0.29817 | 310.5626 | 423.0996 | 41.094 | 41.094 | 423.0996 | 310.4135 | 157.9508 | 158.0999 |  | 89.4722
[10 empty] | 0 | 310.4135 | 423.0996 | 41.094 | 41.094 | 423.0996 | 310.4135 | 158.0999 | 158.0999 |  | 89.6212
[10 empty] | 0 | 310.4135 | 423.0996 | 41.094 | 41.094 | 423.0996 | 310.4135 | 158.0999 | 158.0999 |  | 89.6212
[6 empty] | bellow |  |  |  | 0.099517 | 310.3637 | 423.0996 | 41.094 | 41.094 | 423.0996 | 310.3139 | 158.1496 | 158.1994 |  | 89.671
LS1 | CB06 | CAV6 | 1583 | LS1_CB06:CAV6_D1583 |  | b08 resonator | 34 |  |  | 0.29817 | 310.1649 | 423.0996 | 41.094 | 41.094 | 423.0996 | 310.0158 | 158.3485 | 158.4976 |  | 89.8698
[10 empty] | 0 | 310.0158 | 423.0996 | 41.094 | 41.094 | 423.0996 | 310.0158 | 158.4976 | 158.4976 |  | 90.0189
[10 empty] | 0 | 310.0158 | 423.0996 | 41.094 | 41.094 | 423.0996 | 310.0158 | 158.4976 | 158.4976 |  | 90.0189
[6 empty] | bellow |  |  |  | 0.099517 | 309.966 | 423.0996 | 41.094 | 41.094 | 423.0996 | 309.9163 | 158.5473 | 158.5971 |  | 90.0687
LS1 | CB06 | CAV7 | 1587 | LS1_CB06:CAV7_D1587 |  | b08 resonator | 34 |  |  | 0.29817 | 309.7672 | 423.0996 | 41.094 | 41.094 | 423.0996 | 309.6181 | 158.7462 | 158.8953 |  | 90.2675
[10 empty] | 0 | 309.6181 | 423.0996 | 41.094 | 41.094 | 423.0996 | 309.6181 | 158.8953 | 158.8953 |  | 90.4166
LS1 | CB06 | BLM | 1589 | LS1_CB06:BLM_D1589 |  | outside loss monitor |  |  |  | 0 | 309.6181 | 423.0996 | 41.094 | 41.094 | 423.0996 | 309.6181 | 158.8953 | 158.8953 |  | 90.4166
[6 empty] | bellow |  |  |  | 0.09651 | 309.5698 | 423.0996 | 41.094 | 41.094 | 423.0996 | 309.5216 | 158.9435 | 158.9918 |  | 90.4649
LS1 | CB06 | BPM | 1590 | LS1_CB06:BPM_D1590 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 309.5216 | 423.0996 | 41.094 | 41.094 | 423.0996 | 309.5216 | 158.9918 | 158.9918 |  | 90.5131
[6 empty] | BPM-box |  |  |  | 0.05588 | 309.4936 | 423.0996 | 41.094 | 41.094 | 423.0996 | 309.4657 | 159.0197 | 159.0476 |  | 90.5411
[6 empty] | solenoid-entry |  |  |  | 0.076136 | 309.4276 | 423.0996 | 41.094 | 41.094 | 423.0996 | 309.3896 | 159.0857 | 159.1238 |  | 90.6071
LS1 | CB06 | SOL3 | 1594 | LS1_CB06:SOL3_D1594 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 309.1396 | 423.0996 | 41.094 | 41.094 | 423.0996 | 308.8896 | 159.3738 | 159.6238 |  | 90.8951
[6 empty] | solenoid-exit |  |  |  | 0.076136 | 308.8515 | 423.0996 | 41.094 | 41.094 | 423.0996 | 308.8134 | 159.6618 | 159.6999 |  | 91.1832
[6 empty] | bellow |  |  |  | 0.099517 | 308.7637 | 423.0996 | 41.094 | 41.094 | 423.0996 | 308.7139 | 159.7497 | 159.7994 |  | 91.271
LS1 | CB06 | CAV8 | 1599 | LS1_CB06:CAV8_D1599 |  | b08 resonator | 34 |  |  | 0.29817 | 308.5648 | 423.0996 | 41.094 | 41.094 | 423.0996 | 308.4157 | 159.9485 | 160.0976 |  | 91.4699
[10 empty] | 0 | 308.4157 | 423.0996 | 41.094 | 41.094 | 423.0996 | 308.4157 | 160.0976 | 160.0976 |  | 91.619
[10 empty] | 0 | 308.4157 | 423.0996 | 41.094 | 41.094 | 423.0996 | 308.4157 | 160.0976 | 160.0976 |  | 91.619
[6 empty] | bellows |  |  |  | 0.19082 | 308.3203 | 423.0996 | 41.094 | 41.094 | 423.0996 | 308.2249 | 160.193 | 160.2884 |  | 91.7144
LS1 | CB06 | GV | 1603 | LS1_CB06:GV_D1603 |  | gate valve | 40 |  |  | 0.057404 | 308.1962 | 423.0996 | 41.094 | 41.094 | 423.0996 | 308.1675 | 160.3171 | 160.3458 |  | 91.8385
[6 empty] | bellow | 40 |  |  | 0.1035 | 308.1158 | 423.0996 | 41.094 | 41.094 | 423.0996 | 308.064 | 160.3976 | 160.4493 |  | 91.9189
LS1 | WB06 | BPM | 1604 | LS1_WB06:BPM_D1604 | BPM2 | position monitor |  |  |  | 0 | 308.064 | 423.0996 | 41.094 | 41.094 | 423.0996 | 308.064 | 160.4493 | 160.4493 |  | 91.9707
[6 empty] | diagnostic box |  |  |  | 0.14528 | 307.9914 | 423.0996 | 41.094 | 41.094 | 423.0996 | 307.9187 | 160.522 | 160.5946 |  | 92.0433
LS1 | WB06 | HMR | 1606 | LS1_WB06:HMR_D1606 | HR_L | Halo ring, large aperture | 33 |  |  | 0 | 307.9187 | 423.0996 | 41.094 | 41.094 | 423.0996 | 307.9187 | 160.5946 | 160.5946 |  | 92.116
[6 empty] | diagnostic box |  |  |  | 0 | 307.9187 | 423.0996 | 41.094 | 41.094 | 423.0996 | 307.9187 | 160.5946 | 160.5946 |  | 92.116
LS1 | WB06 | IP | 1606 | LS1_WB06:IP_D1606 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 307.9187 | 423.0996 | 41.094 | 41.094 | 423.0996 | 307.9187 | 160.5946 | 160.5946 |  | 92.116
[6 empty] | bellow | 40 |  |  | 0.13122 | 307.8531 | 423.0996 | 41.094 | 41.094 | 423.0996 | 307.7875 | 160.6602 | 160.7258 |  | 92.1816
LS1 | CB07 | GV | 1608 | LS1_CB07:GV_D1608 |  | gate valve | 40 |  |  | 0.057404 | 307.7588 | 423.0996 | 41.094 | 41.094 | 423.0996 | 307.7301 | 160.7545 | 160.7832 |  | 92.2759
[6 empty] | bellows |  |  |  | 0.19082 | 307.6347 | 423.0996 | 41.094 | 41.094 | 423.0996 | 307.5393 | 160.8786 | 160.9741 |  | 92.4
LS1 | CB07 | CAV1 | 1611 | LS1_CB07:CAV1_D1611 |  | b08 resonator | 34 |  |  | 0.29817 | 307.3902 | 423.0996 | 41.094 | 41.094 | 423.0996 | 307.2411 | 161.1231 | 161.2722 |  | 92.6445
[10 empty] | 0 | 307.2411 | 423.0996 | 41.094 | 41.094 | 423.0996 | 307.2411 | 161.2722 | 161.2722 |  | 92.7936
LS1 | CB07 | BLM | 1613 | LS1_CB07:BLM_D1613 |  | outside loss monitor |  |  |  | 0 | 307.2411 | 423.0996 | 41.094 | 41.094 | 423.0996 | 307.2411 | 161.2722 | 161.2722 |  | 92.7936
[6 empty] | bellow | 40 |  |  | 0.09651 | 307.1929 | 423.0996 | 41.094 | 41.094 | 423.0996 | 307.1446 | 161.3205 | 161.3687 |  | 92.8418
LS1 | CB07 | BPM | 1614 | LS1_CB07:BPM_D1614 | BPM6 | position monitor, BPMc |  |  |  | 0 | 307.1446 | 423.0996 | 41.094 | 41.094 | 423.0996 | 307.1446 | 161.3687 | 161.3687 |  | 92.8901
[6 empty] | BPM-box | 40 |  |  | 0.05588 | 307.1167 | 423.0996 | 41.094 | 41.094 | 423.0996 | 307.0887 | 161.3967 | 161.4246 |  | 92.918
[6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 307.0507 | 423.0996 | 41.094 | 41.094 | 423.0996 | 307.0126 | 161.4627 | 161.5008 |  | 92.9841
LS1 | CB07 | SOL1 | 1618 | LS1_CB07:SOL1_D1618 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 306.7626 | 423.0996 | 41.094 | 41.094 | 423.0996 | 306.5126 | 161.7508 | 162.0008 |  | 93.2721
[6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 306.4745 | 423.0996 | 41.094 | 41.094 | 423.0996 | 306.4365 | 162.0388 | 162.0769 |  | 93.5602
[6 empty] | bellow |  |  |  | 0.099517 | 306.3867 | 423.0996 | 41.094 | 41.094 | 423.0996 | 306.3369 | 162.1266 | 162.1764 |  | 93.648
LS1 | CB07 | CAV2 | 1623 | LS1_CB07:CAV2_D1623 |  | b08 resonator | 34 |  |  | 0.29817 | 306.1879 | 423.0996 | 41.094 | 41.094 | 423.0996 | 306.0388 | 162.3255 | 162.4746 |  | 93.8469
[10 empty] | 0 | 306.0388 | 423.0996 | 41.094 | 41.094 | 423.0996 | 306.0388 | 162.4746 | 162.4746 |  | 93.9959
[10 empty] | 0 | 306.0388 | 423.0996 | 41.094 | 41.094 | 423.0996 | 306.0388 | 162.4746 | 162.4746 |  | 93.9959
[6 empty] | bellow |  |  |  | 0.099517 | 305.989 | 423.0996 | 41.094 | 41.094 | 423.0996 | 305.9392 | 162.5243 | 162.5741 |  | 94.0457
LS1 | CB07 | CAV3 | 1627 | LS1_CB07:CAV3_D1627 |  | b08 resonator | 34 |  |  | 0.29817 | 305.7902 | 423.0996 | 41.094 | 41.094 | 423.0996 | 305.6411 | 162.7232 | 162.8723 |  | 94.2445
[10 empty] | 0 | 305.6411 | 423.0996 | 41.094 | 41.094 | 423.0996 | 305.6411 | 162.8723 | 162.8723 |  | 94.3936
[10 empty] | 0 | 305.6411 | 423.0996 | 41.094 | 41.094 | 423.0996 | 305.6411 | 162.8723 | 162.8723 |  | 94.3936
[6 empty] | bellow |  |  |  | 0.099517 | 305.5913 | 423.0996 | 41.094 | 41.094 | 423.0996 | 305.5416 | 162.922 | 162.9718 |  | 94.4434
LS1 | CB07 | CAV4 | 1631 | LS1_CB07:CAV4_D1631 |  | b08 resonator | 34 |  |  | 0.29817 | 305.3925 | 423.0996 | 41.094 | 41.094 | 423.0996 | 305.2434 | 163.1209 | 163.27 |  | 94.6422
[10 empty] | 0 | 305.2434 | 423.0996 | 41.094 | 41.094 | 423.0996 | 305.2434 | 163.27 | 163.27 |  | 94.7913
LS1 | CB07 | BLM | 1633 | LS1_CB07:BLM_D1633 |  | outside loss monitor |  |  |  | 0 | 305.2434 | 423.0996 | 41.094 | 41.094 | 423.0996 | 305.2434 | 163.27 | 163.27 |  | 94.7913
[6 empty] | bellow |  |  |  | 0.09651 | 305.1951 | 423.0996 | 41.094 | 41.094 | 423.0996 | 305.1469 | 163.3182 | 163.3665 |  | 94.8396
LS1 | CB07 | BPM | 1634 | LS1_CB07:BPM_D1634 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 305.1469 | 423.0996 | 41.094 | 41.094 | 423.0996 | 305.1469 | 163.3665 | 163.3665 |  | 94.8878
[6 empty] | BPM-box |  |  |  | 0.05588 | 305.1189 | 423.0996 | 41.094 | 41.094 | 423.0996 | 305.091 | 163.3944 | 163.4223 |  | 94.9158
[6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 305.0529 | 423.0996 | 41.094 | 41.094 | 423.0996 | 305.0149 | 163.4604 | 163.4985 |  | 94.9818
LS1 | CB07 | SOL2 | 1637 | LS1_CB07:SOL2_D1637 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 304.7649 | 423.0996 | 41.094 | 41.094 | 423.0996 | 304.5149 | 163.7485 | 163.9985 |  | 95.2698
[6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 304.4768 | 423.0996 | 41.094 | 41.094 | 423.0996 | 304.4387 | 164.0365 | 164.0746 |  | 95.5579
[6 empty] | bellow |  |  |  | 0.099517 | 304.389 | 423.0996 | 41.094 | 41.094 | 423.0996 | 304.3392 | 164.1244 | 164.1741 |  | 95.6457
LS1 | CB07 | CAV5 | 1643 | LS1_CB07:CAV5_D1643 |  | b08 resonator | 34 |  |  | 0.29817 | 304.1901 | 423.0996 | 41.094 | 41.094 | 423.0996 | 304.041 | 164.3232 | 164.4723 |  | 95.8446
[10 empty] | 0 | 304.041 | 423.0996 | 41.094 | 41.094 | 423.0996 | 304.041 | 164.4723 | 164.4723 |  | 95.9937
[10 empty] | 0 | 304.041 | 423.0996 | 41.094 | 41.094 | 423.0996 | 304.041 | 164.4723 | 164.4723 |  | 95.9937
[6 empty] | bellow |  |  |  | 0.099517 | 303.9913 | 423.0996 | 41.094 | 41.094 | 423.0996 | 303.9415 | 164.5221 | 164.5718 |  | 96.0434
LS1 | CB07 | CAV6 | 1647 | LS1_CB07:CAV6_D1647 |  | b08 resonator | 34 |  |  | 0.29817 | 303.7924 | 423.0996 | 41.094 | 41.094 | 423.0996 | 303.6434 | 164.7209 | 164.87 |  | 96.2423
[10 empty] | 0 | 303.6434 | 423.0996 | 41.094 | 41.094 | 423.0996 | 303.6434 | 164.87 | 164.87 |  | 96.3914
[10 empty] | 0 | 303.6434 | 423.0996 | 41.094 | 41.094 | 423.0996 | 303.6434 | 164.87 | 164.87 |  | 96.3914
[6 empty] | bellow |  |  |  | 0.099517 | 303.5936 | 423.0996 | 41.094 | 41.094 | 423.0996 | 303.5438 | 164.9197 | 164.9695 |  | 96.4411
LS1 | CB07 | CAV7 | 1651 | LS1_CB07:CAV7_D1651 |  | b08 resonator | 34 |  |  | 0.29817 | 303.3948 | 423.0996 | 41.094 | 41.094 | 423.0996 | 303.2457 | 165.1186 | 165.2677 |  | 96.64
[10 empty] | 0 | 303.2457 | 423.0996 | 41.094 | 41.094 | 423.0996 | 303.2457 | 165.2677 | 165.2677 |  | 96.789
LS1 | CB07 | BLM | 1653 | LS1_CB07:BLM_D1653 |  | outside loss monitor |  |  |  | 0 | 303.2457 | 423.0996 | 41.094 | 41.094 | 423.0996 | 303.2457 | 165.2677 | 165.2677 |  | 96.789
[6 empty] | bellow |  |  |  | 0.09651 | 303.1974 | 423.0996 | 41.094 | 41.094 | 423.0996 | 303.1492 | 165.3159 | 165.3642 |  | 96.8373
LS1 | CB07 | BPM | 1654 | LS1_CB07:BPM_D1654 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 303.1492 | 423.0996 | 41.094 | 41.094 | 423.0996 | 303.1492 | 165.3642 | 165.3642 |  | 96.8856
[6 empty] | BPM-box |  |  |  | 0.05588 | 303.1212 | 423.0996 | 41.094 | 41.094 | 423.0996 | 303.0933 | 165.3921 | 165.4201 |  | 96.9135
[6 empty] | solenoid-entry |  |  |  | 0.076136 | 303.0552 | 423.0996 | 41.094 | 41.094 | 423.0996 | 303.0171 | 165.4581 | 165.4962 |  | 96.9795
LS1 | CB07 | SOL3 | 1657 | LS1_CB07:SOL3_D1657 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 302.7671 | 423.0996 | 41.094 | 41.094 | 423.0996 | 302.5171 | 165.7462 | 165.9962 |  | 97.2676
[6 empty] | solenoid-exit |  |  |  | 0.076136 | 302.4791 | 423.0996 | 41.094 | 41.094 | 423.0996 | 302.441 | 166.0343 | 166.0723 |  | 97.5556
[6 empty] | bellow |  |  |  | 0.099517 | 302.3912 | 423.0996 | 41.094 | 41.094 | 423.0996 | 302.3415 | 166.1221 | 166.1719 |  | 97.6435
LS1 | CB07 | CAV8 | 1663 | LS1_CB07:CAV8_D1663 |  | b08 resonator | 34 |  |  | 0.29817 | 302.1924 | 423.0996 | 41.094 | 41.094 | 423.0996 | 302.0433 | 166.3209 | 166.47 |  | 97.8423
[10 empty] | 0 | 302.0433 | 423.0996 | 41.094 | 41.094 | 423.0996 | 302.0433 | 166.47 | 166.47 |  | 97.9914
[10 empty] | 0 | 302.0433 | 423.0996 | 41.094 | 41.094 | 423.0996 | 302.0433 | 166.47 | 166.47 |  | 97.9914
[6 empty] | bellows |  |  |  | 0.19082 | 301.9479 | 423.0996 | 41.094 | 41.094 | 423.0996 | 301.8525 | 166.5654 | 166.6608 |  | 98.0868
LS1 | CB07 | GV | 1667 | LS1_CB07:GV_D1667 |  | gate valve | 40 |  |  | 0.057404 | 301.8238 | 423.0996 | 41.094 | 41.094 | 423.0996 | 301.7951 | 166.6896 | 166.7183 |  | 98.2109
[6 empty] | bellow | 40 |  |  | 0.1035 | 301.7433 | 423.0996 | 41.094 | 41.094 | 423.0996 | 301.6916 | 166.77 | 166.8218 |  | 98.2914
LS1 | WB07 | BPM | 1668 | LS1_WB07:BPM_D1668 | BPM2 | position monitor |  |  |  | 0 | 301.6916 | 423.0996 | 41.094 | 41.094 | 423.0996 | 301.6916 | 166.8218 | 166.8218 |  | 98.3431
[6 empty] | diagnostic box |  |  |  | 0.14528 | 301.6189 | 423.0996 | 41.094 | 41.094 | 423.0996 | 301.5463 | 166.8944 | 166.967 |  | 98.4158
LS1 | WB07 | HMR | 1670 | LS1_WB07:HMR_D1670 | HR_L | Halo ring, large aperture | 33 |  |  | 0 | 301.5463 | 423.0996 | 41.094 | 41.094 | 423.0996 | 301.5463 | 166.967 | 166.967 |  | 98.4884
[6 empty] | diagnostic box |  |  |  | 0 | 301.5463 | 423.0996 | 41.094 | 41.094 | 423.0996 | 301.5463 | 166.967 | 166.967 |  | 98.4884
LS1 | WB07 | IP | 1670 | LS1_WB07:IP_D1670 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 301.5463 | 423.0996 | 41.094 | 41.094 | 423.0996 | 301.5463 | 166.967 | 166.967 |  | 98.4884
[6 empty] | bellow | 40 |  |  | 0.13122 | 301.4807 | 423.0996 | 41.094 | 41.094 | 423.0996 | 301.4151 | 167.0326 | 167.0983 |  | 98.554
LS1 | CB08 | GV | 1671 | LS1_CB08:GV_D1671 |  | gate valve | 40 |  |  | 0.057404 | 301.3864 | 423.0996 | 41.094 | 41.094 | 423.0996 | 301.3577 | 167.127 | 167.1557 |  | 98.6483
[6 empty] | bellows |  |  |  | 0.19082 | 301.2623 | 423.0996 | 41.094 | 41.094 | 423.0996 | 301.1669 | 167.2511 | 167.3465 |  | 98.7724
LS1 | CB08 | CAV1 | 1675 | LS1_CB08:CAV1_D1675 |  | b08 resonator | 34 |  |  | 0.29817 | 301.0178 | 423.0996 | 41.094 | 41.094 | 423.0996 | 300.8687 | 167.4956 | 167.6446 |  | 99.0169
[10 empty] | 0 | 300.8687 | 423.0996 | 41.094 | 41.094 | 423.0996 | 300.8687 | 167.6446 | 167.6446 |  | 99.166
LS1 | CB08 | BLM | 1676 | LS1_CB08:BLM_D1676 |  | outside loss monitor |  |  |  | 0 | 300.8687 | 423.0996 | 41.094 | 41.094 | 423.0996 | 300.8687 | 167.6446 | 167.6446 |  | 99.166
[6 empty] | bellow | 40 |  |  | 0.09651 | 300.8204 | 423.0996 | 41.094 | 41.094 | 423.0996 | 300.7722 | 167.6929 | 167.7412 |  | 99.2143
LS1 | CB08 | BPM | 1677 | LS1_CB08:BPM_D1677 | BPM6 | position monitor, BPMc |  |  |  | 0 | 300.7722 | 423.0996 | 41.094 | 41.094 | 423.0996 | 300.7722 | 167.7412 | 167.7412 |  | 99.2625
[6 empty] | BPM-box | 40 |  |  | 0.05588 | 300.7442 | 423.0996 | 41.094 | 41.094 | 423.0996 | 300.7163 | 167.7691 | 167.797 |  | 99.2905
[6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 300.6782 | 423.0996 | 41.094 | 41.094 | 423.0996 | 300.6402 | 167.8351 | 167.8732 |  | 99.3565
LS1 | CB08 | SOL1 | 1681 | LS1_CB08:SOL1_D1681 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 300.3902 | 423.0996 | 41.094 | 41.094 | 423.0996 | 300.1402 | 168.1232 | 168.3732 |  | 99.6445
[6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 300.1021 | 423.0996 | 41.094 | 41.094 | 423.0996 | 300.064 | 168.4112 | 168.4493 |  | 99.9326
[6 empty] | bellow |  |  |  | 0.099517 | 300.0143 | 423.0996 | 41.094 | 41.094 | 423.0996 | 299.9645 | 168.4991 | 168.5488 |  | 100.0204
LS1 | CB08 | CAV2 | 1687 | LS1_CB08:CAV2_D1687 |  | b08 resonator | 34 |  |  | 0.29817 | 299.8154 | 423.0996 | 41.094 | 41.094 | 423.0996 | 299.6663 | 168.6979 | 168.847 |  | 100.2193
[10 empty] | 0 | 299.6663 | 423.0996 | 41.094 | 41.094 | 423.0996 | 299.6663 | 168.847 | 168.847 |  | 100.3684
[10 empty] | 0 | 299.6663 | 423.0996 | 41.094 | 41.094 | 423.0996 | 299.6663 | 168.847 | 168.847 |  | 100.3684
[6 empty] | bellow |  |  |  | 0.099517 | 299.6166 | 423.0996 | 41.094 | 41.094 | 423.0996 | 299.5668 | 168.8968 | 168.9465 |  | 100.4181
LS1 | CB08 | CAV3 | 1691 | LS1_CB08:CAV3_D1691 |  | b08 resonator | 34 |  |  | 0.29817 | 299.4177 | 423.0996 | 41.094 | 41.094 | 423.0996 | 299.2687 | 169.0956 | 169.2447 |  | 100.617
[10 empty] | 0 | 299.2687 | 423.0996 | 41.094 | 41.094 | 423.0996 | 299.2687 | 169.2447 | 169.2447 |  | 100.7661
[10 empty] | 0 | 299.2687 | 423.0996 | 41.094 | 41.094 | 423.0996 | 299.2687 | 169.2447 | 169.2447 |  | 100.7661
[6 empty] | bellow |  |  |  | 0.099517 | 299.2189 | 423.0996 | 41.094 | 41.094 | 423.0996 | 299.1691 | 169.2944 | 169.3442 |  | 100.8158
LS1 | CB08 | CAV4 | 1695 | LS1_CB08:CAV4_D1695 |  | b08 resonator | 34 |  |  | 0.29817 | 299.0201 | 423.0996 | 41.094 | 41.094 | 423.0996 | 298.871 | 169.4933 | 169.6424 |  | 101.0147
[10 empty] | 0 | 298.871 | 423.0996 | 41.094 | 41.094 | 423.0996 | 298.871 | 169.6424 | 169.6424 |  | 101.1637
LS1 | CB08 | BLM | 1696 | LS1_CB08:BLM_D1696 |  | outside loss monitor |  |  |  | 0 | 298.871 | 423.0996 | 41.094 | 41.094 | 423.0996 | 298.871 | 169.6424 | 169.6424 |  | 101.1637
[6 empty] | bellow |  |  |  | 0.09651 | 298.8227 | 423.0996 | 41.094 | 41.094 | 423.0996 | 298.7745 | 169.6906 | 169.7389 |  | 101.212
LS1 | CB08 | BPM | 1697 | LS1_CB08:BPM_D1697 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 298.7745 | 423.0996 | 41.094 | 41.094 | 423.0996 | 298.7745 | 169.7389 | 169.7389 |  | 101.2603
[6 empty] | BPM-box |  |  |  | 0.05588 | 298.7465 | 423.0996 | 41.094 | 41.094 | 423.0996 | 298.7186 | 169.7668 | 169.7948 |  | 101.2882
[6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 298.6805 | 423.0996 | 41.094 | 41.094 | 423.0996 | 298.6424 | 169.8328 | 169.8709 |  | 101.3542
LS1 | CB08 | SOL2 | 1701 | LS1_CB08:SOL2_D1701 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 298.3924 | 423.0996 | 41.094 | 41.094 | 423.0996 | 298.1424 | 170.1209 | 170.3709 |  | 101.6423
[6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 298.1044 | 423.0996 | 41.094 | 41.094 | 423.0996 | 298.0663 | 170.409 | 170.447 |  | 101.9303
[6 empty] | bellow |  |  |  | 0.099517 | 298.0165 | 423.0996 | 41.094 | 41.094 | 423.0996 | 297.9668 | 170.4968 | 170.5466 |  | 102.0182
LS1 | CB08 | CAV5 | 1707 | LS1_CB08:CAV5_D1707 |  | b08 resonator | 34 |  |  | 0.29817 | 297.8177 | 423.0996 | 41.094 | 41.094 | 423.0996 | 297.6686 | 170.6956 | 170.8447 |  | 102.217
[10 empty] | 0 | 297.6686 | 423.0996 | 41.094 | 41.094 | 423.0996 | 297.6686 | 170.8447 | 170.8447 |  | 102.3661
[10 empty] | 0 | 297.6686 | 423.0996 | 41.094 | 41.094 | 423.0996 | 297.6686 | 170.8447 | 170.8447 |  | 102.3661
[6 empty] | bellow |  |  |  | 0.099517 | 297.6189 | 423.0996 | 41.094 | 41.094 | 423.0996 | 297.5691 | 170.8945 | 170.9442 |  | 102.4159
LS1 | CB08 | CAV6 | 1711 | LS1_CB08:CAV6_D1711 |  | b08 resonator | 34 |  |  | 0.29817 | 297.42 | 423.0996 | 41.094 | 41.094 | 423.0996 | 297.2709 | 171.0933 | 171.2424 |  | 102.6147
[10 empty] | 0 | 297.2709 | 423.0996 | 41.094 | 41.094 | 423.0996 | 297.2709 | 171.2424 | 171.2424 |  | 102.7638
[10 empty] | 0 | 297.2709 | 423.0996 | 41.094 | 41.094 | 423.0996 | 297.2709 | 171.2424 | 171.2424 |  | 102.7638
[6 empty] | bellow |  |  |  | 0.099517 | 297.2212 | 423.0996 | 41.094 | 41.094 | 423.0996 | 297.1714 | 171.2922 | 171.3419 |  | 102.8135
LS1 | CB08 | CAV7 | 1715 | LS1_CB08:CAV7_D1715 |  | b08 resonator | 34 |  |  | 0.29817 | 297.0223 | 423.0996 | 41.094 | 41.094 | 423.0996 | 296.8732 | 171.491 | 171.6401 |  | 103.0124
[10 empty] | 0 | 296.8732 | 423.0996 | 41.094 | 41.094 | 423.0996 | 296.8732 | 171.6401 | 171.6401 |  | 103.1615
LS1 | CB08 | BLM | 1716 | LS1_CB08:BLM_D1716 |  | outside loss monitor |  |  |  | 0 | 296.8732 | 423.0996 | 41.094 | 41.094 | 423.0996 | 296.8732 | 171.6401 | 171.6401 |  | 103.1615
[6 empty] | bellow |  |  |  | 0.09651 | 296.825 | 423.0996 | 41.094 | 41.094 | 423.0996 | 296.7767 | 171.6884 | 171.7366 |  | 103.2097
LS1 | CB08 | BPM | 1717 | LS1_CB08:BPM_D1717 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 296.7767 | 423.0996 | 41.094 | 41.094 | 423.0996 | 296.7767 | 171.7366 | 171.7366 |  | 103.258
[6 empty] | BPM-box |  |  |  | 0.05588 | 296.7488 | 423.0996 | 41.094 | 41.094 | 423.0996 | 296.7209 | 171.7646 | 171.7925 |  | 103.2859
[6 empty] | solenoid-entry |  |  |  | 0.076136 | 296.6828 | 423.0996 | 41.094 | 41.094 | 423.0996 | 296.6447 | 171.8306 | 171.8686 |  | 103.3519
LS1 | CB08 | SOL3 | 1721 | LS1_CB08:SOL3_D1721 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 296.3947 | 423.0996 | 41.094 | 41.094 | 423.0996 | 296.1447 | 172.1186 | 172.3686 |  | 103.64
[6 empty] | solenoid-exit |  |  |  | 0.076136 | 296.1066 | 423.0996 | 41.094 | 41.094 | 423.0996 | 296.0686 | 172.4067 | 172.4448 |  | 103.9281
[6 empty] | bellow |  |  |  | 0.099517 | 296.0188 | 423.0996 | 41.094 | 41.094 | 423.0996 | 295.9691 | 172.4945 | 172.5443 |  | 104.0159
LS1 | CB08 | CAV8 | 1727 | LS1_CB08:CAV8_D1727 |  | b08 resonator | 34 |  |  | 0.29817 | 295.82 | 423.0996 | 41.094 | 41.094 | 423.0996 | 295.6709 | 172.6934 | 172.8425 |  | 104.2147
[10 empty] | 0 | 295.6709 | 423.0996 | 41.094 | 41.094 | 423.0996 | 295.6709 | 172.8425 | 172.8425 |  | 104.3638
[10 empty] | 0 | 295.6709 | 423.0996 | 41.094 | 41.094 | 423.0996 | 295.6709 | 172.8425 | 172.8425 |  | 104.3638
[6 empty] | bellows |  |  |  | 0.19082 | 295.5755 | 423.0996 | 41.094 | 41.094 | 423.0996 | 295.4801 | 172.9379 | 173.0333 |  | 104.4592
LS1 | CB08 | GV | 1731 | LS1_CB08:GV_D1731 |  | gate valve | 40 |  |  | 0.057404 | 295.4514 | 423.0996 | 41.094 | 41.094 | 423.0996 | 295.4227 | 173.062 | 173.0907 |  | 104.5833
[6 empty] | bellow | 40 |  |  | 0.1035 | 295.3709 | 423.0996 | 41.094 | 41.094 | 423.0996 | 295.3192 | 173.1424 | 173.1942 |  | 104.6638
LS1 | WB08 | BPM | 1732 | LS1_WB08:BPM_D1732 | BPM2 | position monitor |  |  |  | 0 | 295.3192 | 423.0996 | 41.094 | 41.094 | 423.0996 | 295.3192 | 173.1942 | 173.1942 |  | 104.7155
[6 empty] | diagnostic box |  |  |  | 0.14528 | 295.2465 | 423.0996 | 41.094 | 41.094 | 423.0996 | 295.1739 | 173.2668 | 173.3395 |  | 104.7882
LS1 | WB08 | HMR | 1733 | LS1_WB08:HMR_D1733 | HR_L | Halo ring, large aperture | 33 |  |  | 0 | 295.1739 | 423.0996 | 41.094 | 41.094 | 423.0996 | 295.1739 | 173.3395 | 173.3395 |  | 104.8608
[6 empty] | diagnostic box |  |  |  | 0 | 295.1739 | 423.0996 | 41.094 | 41.094 | 423.0996 | 295.1739 | 173.3395 | 173.3395 |  | 104.8608
LS1 | WB08 | IP | 1733 | LS1_WB08:IP_D1733 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 295.1739 | 423.0996 | 41.094 | 41.094 | 423.0996 | 295.1739 | 173.3395 | 173.3395 |  | 104.8608
[6 empty] | bellow | 40 |  |  | 0.13122 | 295.1083 | 423.0996 | 41.094 | 41.094 | 423.0996 | 295.0427 | 173.4051 | 173.4707 |  | 104.9264
LS1 | CB09 | GV | 1735 | LS1_CB09:GV_D1735 |  | gate valve | 40 |  |  | 0.057404 | 295.014 | 423.0996 | 41.094 | 41.094 | 423.0996 | 294.9853 | 173.4994 | 173.5281 |  | 105.0207
[6 empty] | bellows |  |  |  | 0.19082 | 294.8898 | 423.0996 | 41.094 | 41.094 | 423.0996 | 294.7944 | 173.6235 | 173.7189 |  | 105.1449
LS1 | CB09 | CAV1 | 1739 | LS1_CB09:CAV1_D1739 |  | b08 resonator | 34 |  |  | 0.29817 | 294.6454 | 423.0996 | 41.094 | 41.094 | 423.0996 | 294.4963 | 173.868 | 174.0171 |  | 105.3894
[10 empty] | 0 | 294.4963 | 423.0996 | 41.094 | 41.094 | 423.0996 | 294.4963 | 174.0171 | 174.0171 |  | 105.5384
LS1 | CB09 | BLM | 1740 | LS1_CB09:BLM_D1740 |  | outside loss monitor |  |  |  | 0 | 294.4963 | 423.0996 | 41.094 | 41.094 | 423.0996 | 294.4963 | 174.0171 | 174.0171 |  | 105.5384
[6 empty] | bellow | 40 |  |  | 0.09651 | 294.448 | 423.0996 | 41.094 | 41.094 | 423.0996 | 294.3998 | 174.0653 | 174.1136 |  | 105.5867
LS1 | CB09 | BPM | 1741 | LS1_CB09:BPM_D1741 | BPM6 | position monitor, BPMc |  |  |  | 0 | 294.3998 | 423.0996 | 41.094 | 41.094 | 423.0996 | 294.3998 | 174.1136 | 174.1136 |  | 105.635
[6 empty] | BPM-box | 40 |  |  | 0.05588 | 294.3718 | 423.0996 | 41.094 | 41.094 | 423.0996 | 294.3439 | 174.1415 | 174.1695 |  | 105.6629
[6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 294.3058 | 423.0996 | 41.094 | 41.094 | 423.0996 | 294.2677 | 174.2075 | 174.2456 |  | 105.7289
LS1 | CB09 | SOL1 | 1745 | LS1_CB09:SOL1_D1745 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 294.0177 | 423.0996 | 41.094 | 41.094 | 423.0996 | 293.7677 | 174.4956 | 174.7456 |  | 106.017
[6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 293.7297 | 423.0996 | 41.094 | 41.094 | 423.0996 | 293.6916 | 174.7837 | 174.8217 |  | 106.305
[6 empty] | bellow |  |  |  | 0.099517 | 293.6418 | 423.0996 | 41.094 | 41.094 | 423.0996 | 293.5921 | 174.8715 | 174.9213 |  | 106.3929
LS1 | CB09 | CAV2 | 1751 | LS1_CB09:CAV2_D1751 |  | b08 resonator | 34 |  |  | 0.29817 | 293.443 | 423.0996 | 41.094 | 41.094 | 423.0996 | 293.2939 | 175.0703 | 175.2194 |  | 106.5917
[10 empty] | 0 | 293.2939 | 423.0996 | 41.094 | 41.094 | 423.0996 | 293.2939 | 175.2194 | 175.2194 |  | 106.7408
[10 empty] | 0 | 293.2939 | 423.0996 | 41.094 | 41.094 | 423.0996 | 293.2939 | 175.2194 | 175.2194 |  | 106.7408
[6 empty] | bellow |  |  |  | 0.099517 | 293.2442 | 423.0996 | 41.094 | 41.094 | 423.0996 | 293.1944 | 175.2692 | 175.3189 |  | 106.7906
LS1 | CB09 | CAV3 | 1755 | LS1_CB09:CAV3_D1755 |  | b08 resonator | 34 |  |  | 0.29817 | 293.0453 | 423.0996 | 41.094 | 41.094 | 423.0996 | 292.8962 | 175.468 | 175.6171 |  | 106.9894
[10 empty] | 0 | 292.8962 | 423.0996 | 41.094 | 41.094 | 423.0996 | 292.8962 | 175.6171 | 175.6171 |  | 107.1385
[10 empty] | 0 | 292.8962 | 423.0996 | 41.094 | 41.094 | 423.0996 | 292.8962 | 175.6171 | 175.6171 |  | 107.1385
[6 empty] | bellow |  |  |  | 0.099517 | 292.8465 | 423.0996 | 41.094 | 41.094 | 423.0996 | 292.7967 | 175.6669 | 175.7166 |  | 107.1882
LS1 | CB09 | CAV4 | 1759 | LS1_CB09:CAV4_D1759 |  | b08 resonator | 34 |  |  | 0.29817 | 292.6476 | 423.0996 | 41.094 | 41.094 | 423.0996 | 292.4985 | 175.8657 | 176.0148 |  | 107.3871
[10 empty] | 0 | 292.4985 | 423.0996 | 41.094 | 41.094 | 423.0996 | 292.4985 | 176.0148 | 176.0148 |  | 107.5362
LS1 | CB09 | BLM | 1760 | LS1_CB09:BLM_D1760 |  | outside loss monitor |  |  |  | 0 | 292.4985 | 423.0996 | 41.094 | 41.094 | 423.0996 | 292.4985 | 176.0148 | 176.0148 |  | 107.5362
[6 empty] | bellow |  |  |  | 0.09651 | 292.4503 | 423.0996 | 41.094 | 41.094 | 423.0996 | 292.402 | 176.0631 | 176.1113 |  | 107.5844
LS1 | CB09 | BPM | 1761 | LS1_CB09:BPM_D1761 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 292.402 | 423.0996 | 41.094 | 41.094 | 423.0996 | 292.402 | 176.1113 | 176.1113 |  | 107.6327
[6 empty] | BPM-box |  |  |  | 0.05588 | 292.3741 | 423.0996 | 41.094 | 41.094 | 423.0996 | 292.3462 | 176.1392 | 176.1672 |  | 107.6606
[6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 292.3081 | 423.0996 | 41.094 | 41.094 | 423.0996 | 292.27 | 176.2053 | 176.2433 |  | 107.7266
LS1 | CB09 | SOL2 | 1765 | LS1_CB09:SOL2_D1765 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 292.02 | 423.0996 | 41.094 | 41.094 | 423.0996 | 291.77 | 176.4933 | 176.7433 |  | 108.0147
[6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 291.7319 | 423.0996 | 41.094 | 41.094 | 423.0996 | 291.6939 | 176.7814 | 176.8195 |  | 108.3028
[6 empty] | bellow |  |  |  | 0.099517 | 291.6441 | 423.0996 | 41.094 | 41.094 | 423.0996 | 291.5944 | 176.8692 | 176.919 |  | 108.3906
LS1 | CB09 | CAV5 | 1771 | LS1_CB09:CAV5_D1771 |  | b08 resonator | 34 |  |  | 0.29817 | 291.4453 | 423.0996 | 41.094 | 41.094 | 423.0996 | 291.2962 | 177.0681 | 177.2171 |  | 108.5894
[10 empty] | 0 | 291.2962 | 423.0996 | 41.094 | 41.094 | 423.0996 | 291.2962 | 177.2171 | 177.2171 |  | 108.7385
[10 empty] | 0 | 291.2962 | 423.0996 | 41.094 | 41.094 | 423.0996 | 291.2962 | 177.2171 | 177.2171 |  | 108.7385
[6 empty] | bellow |  |  |  | 0.099517 | 291.2464 | 423.0996 | 41.094 | 41.094 | 423.0996 | 291.1967 | 177.2669 | 177.3167 |  | 108.7883
LS1 | CB09 | CAV6 | 1775 | LS1_CB09:CAV6_D1775 |  | b08 resonator | 34 |  |  | 0.29817 | 291.0476 | 423.0996 | 41.094 | 41.094 | 423.0996 | 290.8985 | 177.4658 | 177.6148 |  | 108.9871
[10 empty] | 0 | 290.8985 | 423.0996 | 41.094 | 41.094 | 423.0996 | 290.8985 | 177.6148 | 177.6148 |  | 109.1362
[10 empty] | 0 | 290.8985 | 423.0996 | 41.094 | 41.094 | 423.0996 | 290.8985 | 177.6148 | 177.6148 |  | 109.1362
[6 empty] | bellow |  |  |  | 0.099517 | 290.8487 | 423.0996 | 41.094 | 41.094 | 423.0996 | 290.799 | 177.6646 | 177.7144 |  | 109.186
LS1 | CB09 | CAV7 | 1779 | LS1_CB09:CAV7_D1779 |  | b08 resonator | 34 |  |  | 0.29817 | 290.6499 | 423.0996 | 41.094 | 41.094 | 423.0996 | 290.5008 | 177.8634 | 178.0125 |  | 109.3848
[10 empty] | 0 | 290.5008 | 423.0996 | 41.094 | 41.094 | 423.0996 | 290.5008 | 178.0125 | 178.0125 |  | 109.5339
LS1 | CB09 | BLM | 1780 | LS1_CB09:BLM_D1780 |  | outside loss monitor |  |  |  | 0 | 290.5008 | 423.0996 | 41.094 | 41.094 | 423.0996 | 290.5008 | 178.0125 | 178.0125 |  | 109.5339
[6 empty] | bellow |  |  |  | 0.09651 | 290.4526 | 423.0996 | 41.094 | 41.094 | 423.0996 | 290.4043 | 178.0608 | 178.109 |  | 109.5822
LS1 | CB09 | BPM | 1781 | LS1_CB09:BPM_D1781 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 290.4043 | 423.0996 | 41.094 | 41.094 | 423.0996 | 290.4043 | 178.109 | 178.109 |  | 109.6304
[6 empty] | BPM-box |  |  |  | 0.05588 | 290.3764 | 423.0996 | 41.094 | 41.094 | 423.0996 | 290.3484 | 178.137 | 178.1649 |  | 109.6583
[6 empty] | solenoid-entry |  |  |  | 0.076136 | 290.3104 | 423.0996 | 41.094 | 41.094 | 423.0996 | 290.2723 | 178.203 | 178.2411 |  | 109.7244
LS1 | CB09 | SOL3 | 1785 | LS1_CB09:SOL3_D1785 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 290.0223 | 423.0996 | 41.094 | 41.094 | 423.0996 | 289.7723 | 178.4911 | 178.7411 |  | 110.0124
[6 empty] | solenoid-exit |  |  |  | 0.076136 | 289.7342 | 423.0996 | 41.094 | 41.094 | 423.0996 | 289.6962 | 178.7791 | 178.8172 |  | 110.3005
[6 empty] | bellow |  |  |  | 0.099517 | 289.6464 | 423.0996 | 41.094 | 41.094 | 423.0996 | 289.5966 | 178.8669 | 178.9167 |  | 110.3883
LS1 | CB09 | CAV8 | 1791 | LS1_CB09:CAV8_D1791 |  | b08 resonator | 34 |  |  | 0.29817 | 289.4476 | 423.0996 | 41.094 | 41.094 | 423.0996 | 289.2985 | 179.0658 | 179.2149 |  | 110.5872
[10 empty] | 0 | 289.2985 | 423.0996 | 41.094 | 41.094 | 423.0996 | 289.2985 | 179.2149 | 179.2149 |  | 110.7362
[10 empty] | 0 | 289.2985 | 423.0996 | 41.094 | 41.094 | 423.0996 | 289.2985 | 179.2149 | 179.2149 |  | 110.7362
[6 empty] | bellows |  |  |  | 0.19082 | 289.2031 | 423.0996 | 41.094 | 41.094 | 423.0996 | 289.1076 | 179.3103 | 179.4057 |  | 110.8317
LS1 | CB09 | GV | 1794 | LS1_CB09:GV_D1794 |  | gate valve | 40 |  |  | 0.057404 | 289.0789 | 423.0996 | 41.094 | 41.094 | 423.0996 | 289.0502 | 179.4344 | 179.4631 |  | 110.9558
[6 empty] | bellow | 40 |  |  | 0.1035 | 288.9985 | 423.0996 | 41.094 | 41.094 | 423.0996 | 288.9467 | 179.5149 | 179.5666 |  | 111.0362
LS1 | WB09 | BPM | 1796 | LS1_WB09:BPM_D1796 | BPM2 | position monitor |  |  |  | 0 | 288.9467 | 423.0996 | 41.094 | 41.094 | 423.0996 | 288.9467 | 179.5666 | 179.5666 |  | 111.088
[6 empty] | diagnostic box |  |  |  | 0.14528 | 288.8741 | 423.0996 | 41.094 | 41.094 | 423.0996 | 288.8015 | 179.6392 | 179.7119 |  | 111.1606
LS1 | WB09 | HMR | 1797 | LS1_WB09:HMR_D1797 | HR_L | Halo ring, large aperture | 33 |  |  | 0 | 288.8015 | 423.0996 | 41.094 | 41.094 | 423.0996 | 288.8015 | 179.7119 | 179.7119 |  | 111.2333
[6 empty] | diagnostic box |  |  |  | 0 | 288.8015 | 423.0996 | 41.094 | 41.094 | 423.0996 | 288.8015 | 179.7119 | 179.7119 |  | 111.2333
LS1 | WB09 | IP | 1797 | LS1_WB09:IP_D1797 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 288.8015 | 423.0996 | 41.094 | 41.094 | 423.0996 | 288.8015 | 179.7119 | 179.7119 |  | 111.2333
[6 empty] | bellow | 40 |  |  | 0.13122 | 288.7358 | 423.0996 | 41.094 | 41.094 | 423.0996 | 288.6702 | 179.7775 | 179.8431 |  | 111.2989
LS1 | CB10 | GV | 1799 | LS1_CB10:GV_D1799 |  | gate valve | 40 |  |  | 0.057404 | 288.6415 | 423.0996 | 41.094 | 41.094 | 423.0996 | 288.6128 | 179.8718 | 179.9005 |  | 111.3932
[6 empty] | bellows |  |  |  | 0.19082 | 288.5174 | 423.0996 | 41.094 | 41.094 | 423.0996 | 288.422 | 179.9959 | 180.0913 |  | 111.5173
LS1 | CB10 | CAV1 | 1802 | LS1_CB10:CAV1_D1802 |  | b08 resonator | 34 |  |  | 0.29817 | 288.2729 | 423.0996 | 41.094 | 41.094 | 423.0996 | 288.1238 | 180.2404 | 180.3895 |  | 111.7618
[10 empty] | 0 | 288.1238 | 423.0996 | 41.094 | 41.094 | 423.0996 | 288.1238 | 180.3895 | 180.3895 |  | 111.9109
LS1 | CB10 | BLM | 1804 | LS1_CB10:BLM_D1804 |  | outside loss monitor |  |  |  | 0 | 288.1238 | 423.0996 | 41.094 | 41.094 | 423.0996 | 288.1238 | 180.3895 | 180.3895 |  | 111.9109
[6 empty] | bellow | 40 |  |  | 0.09651 | 288.0756 | 423.0996 | 41.094 | 41.094 | 423.0996 | 288.0273 | 180.4378 | 180.486 |  | 111.9591
LS1 | CB10 | BPM | 1805 | LS1_CB10:BPM_D1805 | BPM6 | position monitor, BPMc |  |  |  | 0 | 288.0273 | 423.0996 | 41.094 | 41.094 | 423.0996 | 288.0273 | 180.486 | 180.486 |  | 112.0074
[6 empty] | BPM-box | 40 |  |  | 0.05588 | 287.9994 | 423.0996 | 41.094 | 41.094 | 423.0996 | 287.9715 | 180.5139 | 180.5419 |  | 112.0353
[6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 287.9334 | 423.0996 | 41.094 | 41.094 | 423.0996 | 287.8953 | 180.58 | 180.618 |  | 112.1013
LS1 | CB10 | SOL1 | 1809 | LS1_CB10:SOL1_D1809 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 287.6453 | 423.0996 | 41.094 | 41.094 | 423.0996 | 287.3953 | 180.868 | 181.118 |  | 112.3894
[6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 287.3572 | 423.0996 | 41.094 | 41.094 | 423.0996 | 287.3192 | 181.1561 | 181.1942 |  | 112.6775
[6 empty] | bellow |  |  |  | 0.099517 | 287.2694 | 423.0996 | 41.094 | 41.094 | 423.0996 | 287.2197 | 181.2439 | 181.2937 |  | 112.7653
LS1 | CB10 | CAV2 | 1814 | LS1_CB10:CAV2_D1814 |  | b08 resonator | 34 |  |  | 0.29817 | 287.0706 | 423.0996 | 41.094 | 41.094 | 423.0996 | 286.9215 | 181.4428 | 181.5918 |  | 112.9641
[10 empty] | 0 | 286.9215 | 423.0996 | 41.094 | 41.094 | 423.0996 | 286.9215 | 181.5918 | 181.5918 |  | 113.1132
[10 empty] | 0 | 286.9215 | 423.0996 | 41.094 | 41.094 | 423.0996 | 286.9215 | 181.5918 | 181.5918 |  | 113.1132
[6 empty] | bellow |  |  |  | 0.099517 | 286.8717 | 423.0996 | 41.094 | 41.094 | 423.0996 | 286.822 | 181.6416 | 181.6914 |  | 113.163
LS1 | CB10 | CAV3 | 1818 | LS1_CB10:CAV3_D1818 |  | b08 resonator | 34 |  |  | 0.29817 | 286.6729 | 423.0996 | 41.094 | 41.094 | 423.0996 | 286.5238 | 181.8405 | 181.9895 |  | 113.3618
[10 empty] | 0 | 286.5238 | 423.0996 | 41.094 | 41.094 | 423.0996 | 286.5238 | 181.9895 | 181.9895 |  | 113.5109
[10 empty] | 0 | 286.5238 | 423.0996 | 41.094 | 41.094 | 423.0996 | 286.5238 | 181.9895 | 181.9895 |  | 113.5109
[6 empty] | bellow |  |  |  | 0.099517 | 286.474 | 423.0996 | 41.094 | 41.094 | 423.0996 | 286.4243 | 182.0393 | 182.0891 |  | 113.5607
LS1 | CB10 | CAV4 | 1822 | LS1_CB10:CAV4_D1822 |  | b08 resonator | 34 |  |  | 0.29817 | 286.2752 | 423.0996 | 41.094 | 41.094 | 423.0996 | 286.1261 | 182.2381 | 182.3872 |  | 113.7595
[10 empty] | 0 | 286.1261 | 423.0996 | 41.094 | 41.094 | 423.0996 | 286.1261 | 182.3872 | 182.3872 |  | 113.9086
LS1 | CB10 | BLM | 1824 | LS1_CB10:BLM_D1824 |  | outside loss monitor |  |  |  | 0 | 286.1261 | 423.0996 | 41.094 | 41.094 | 423.0996 | 286.1261 | 182.3872 | 182.3872 |  | 113.9086
[6 empty] | bellow |  |  |  | 0.09651 | 286.0779 | 423.0996 | 41.094 | 41.094 | 423.0996 | 286.0296 | 182.4355 | 182.4837 |  | 113.9568
LS1 | CB10 | BPM | 1825 | LS1_CB10:BPM_D1825 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 286.0296 | 423.0996 | 41.094 | 41.094 | 423.0996 | 286.0296 | 182.4837 | 182.4837 |  | 114.0051
[6 empty] | BPM-box |  |  |  | 0.05588 | 286.0017 | 423.0996 | 41.094 | 41.094 | 423.0996 | 285.9737 | 182.5117 | 182.5396 |  | 114.033
[6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 285.9357 | 423.0996 | 41.094 | 41.094 | 423.0996 | 285.8976 | 182.5777 | 182.6157 |  | 114.0991
LS1 | CB10 | SOL2 | 1829 | LS1_CB10:SOL2_D1829 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 285.6476 | 423.0996 | 41.094 | 41.094 | 423.0996 | 285.3976 | 182.8657 | 183.1157 |  | 114.3871
[6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 285.3595 | 423.0996 | 41.094 | 41.094 | 423.0996 | 285.3215 | 183.1538 | 183.1919 |  | 114.6752
[6 empty] | bellow |  |  |  | 0.099517 | 285.2717 | 423.0996 | 41.094 | 41.094 | 423.0996 | 285.2219 | 183.2416 | 183.2914 |  | 114.763
LS1 | CB10 | CAV5 | 1834 | LS1_CB10:CAV5_D1834 |  | b08 resonator | 34 |  |  | 0.29817 | 285.0729 | 423.0996 | 41.094 | 41.094 | 423.0996 | 284.9238 | 183.4405 | 183.5896 |  | 114.9619
[10 empty] | 0 | 284.9238 | 423.0996 | 41.094 | 41.094 | 423.0996 | 284.9238 | 183.5896 | 183.5896 |  | 115.1109
[10 empty] | 0 | 284.9238 | 423.0996 | 41.094 | 41.094 | 423.0996 | 284.9238 | 183.5896 | 183.5896 |  | 115.1109
[6 empty] | bellow |  |  |  | 0.099517 | 284.874 | 423.0996 | 41.094 | 41.094 | 423.0996 | 284.8243 | 183.6393 | 183.6891 |  | 115.1607
LS1 | CB10 | CAV6 | 1838 | LS1_CB10:CAV6_D1838 |  | b08 resonator | 34 |  |  | 0.29817 | 284.6752 | 423.0996 | 41.094 | 41.094 | 423.0996 | 284.5261 | 183.8382 | 183.9873 |  | 115.3595
[10 empty] | 0 | 284.5261 | 423.0996 | 41.094 | 41.094 | 423.0996 | 284.5261 | 183.9873 | 183.9873 |  | 115.5086
[10 empty] | 0 | 284.5261 | 423.0996 | 41.094 | 41.094 | 423.0996 | 284.5261 | 183.9873 | 183.9873 |  | 115.5086
[6 empty] | bellow |  |  |  | 0.099517 | 284.4763 | 423.0996 | 41.094 | 41.094 | 423.0996 | 284.4266 | 184.037 | 184.0868 |  | 115.5584
LS1 | CB10 | CAV7 | 1842 | LS1_CB10:CAV7_D1842 |  | b08 resonator | 34 |  |  | 0.29817 | 284.2775 | 423.0996 | 41.094 | 41.094 | 423.0996 | 284.1284 | 184.2359 | 184.3849 |  | 115.7572
[10 empty] | 0 | 284.1284 | 423.0996 | 41.094 | 41.094 | 423.0996 | 284.1284 | 184.3849 | 184.3849 |  | 115.9063
LS1 | CB10 | BLM | 1844 | LS1_CB10:BLM_D1844 |  | outside loss monitor |  |  |  | 0 | 284.1284 | 423.0996 | 41.094 | 41.094 | 423.0996 | 284.1284 | 184.3849 | 184.3849 |  | 115.9063
[6 empty] | bellow |  |  |  | 0.09651 | 284.0801 | 423.0996 | 41.094 | 41.094 | 423.0996 | 284.0319 | 184.4332 | 184.4815 |  | 115.9546
LS1 | CB10 | BPM | 1845 | LS1_CB10:BPM_D1845 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 284.0319 | 423.0996 | 41.094 | 41.094 | 423.0996 | 284.0319 | 184.4815 | 184.4815 |  | 116.0028
[6 empty] | BPM-box |  |  |  | 0.05588 | 284.0039 | 423.0996 | 41.094 | 41.094 | 423.0996 | 283.976 | 184.5094 | 184.5373 |  | 116.0308
[6 empty] | solenoid-entry |  |  |  | 0.076136 | 283.9379 | 423.0996 | 41.094 | 41.094 | 423.0996 | 283.8999 | 184.5754 | 184.6135 |  | 116.0968
LS1 | CB10 | SOL3 | 1849 | LS1_CB10:SOL3_D1849 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 283.6499 | 423.0996 | 41.094 | 41.094 | 423.0996 | 283.3999 | 184.8635 | 185.1135 |  | 116.3848
[6 empty] | solenoid-exit |  |  |  | 0.076136 | 283.3618 | 423.0996 | 41.094 | 41.094 | 423.0996 | 283.3237 | 185.1515 | 185.1896 |  | 116.6729
[6 empty] | bellow |  |  |  | 0.099517 | 283.274 | 423.0996 | 41.094 | 41.094 | 423.0996 | 283.2242 | 185.2394 | 185.2891 |  | 116.7607
LS1 | CB10 | CAV8 | 1854 | LS1_CB10:CAV8_D1854 |  | b08 resonator | 34 |  |  | 0.29817 | 283.0751 | 423.0996 | 41.094 | 41.094 | 423.0996 | 282.926 | 185.4382 | 185.5873 |  | 116.9596
[10 empty] | 0 | 282.926 | 423.0996 | 41.094 | 41.094 | 423.0996 | 282.926 | 185.5873 | 185.5873 |  | 117.1087
[10 empty] | 0 | 282.926 | 423.0996 | 41.094 | 41.094 | 423.0996 | 282.926 | 185.5873 | 185.5873 |  | 117.1087
[6 empty] | bellows |  |  |  | 0.19082 | 282.8306 | 423.0996 | 41.094 | 41.094 | 423.0996 | 282.7352 | 185.6827 | 185.7781 |  | 117.2041
LS1 | CB10 | GV | 1858 | LS1_CB10:GV_D1858 |  | gate valve | 40 |  |  | 0.057404 | 282.7065 | 423.0996 | 41.094 | 41.094 | 423.0996 | 282.6778 | 185.8068 | 185.8355 |  | 117.3282
[6 empty] | bellow | 40 |  |  | 0.1035 | 282.6261 | 423.0996 | 41.094 | 41.094 | 423.0996 | 282.5743 | 185.8873 | 185.939 |  | 117.4086
LS1 | WB10 | BPM | 1859 | LS1_WB10:BPM_D1859 | BPM2 | position monitor |  |  |  | 0 | 282.5743 | 423.0996 | 41.094 | 41.094 | 423.0996 | 282.5743 | 185.939 | 185.939 |  | 117.4604
[6 empty] | diagnostic box |  |  |  | 0.14528 | 282.5017 | 423.0996 | 41.094 | 41.094 | 423.0996 | 282.429 | 186.0117 | 186.0843 |  | 117.533
LS1 | WB10 | HMR | 1861 | LS1_WB10:HMR_D1861 | HR_L | Halo ring, large aperture | 33 |  |  | 0 | 282.429 | 423.0996 | 41.094 | 41.094 | 423.0996 | 282.429 | 186.0843 | 186.0843 |  | 117.6057
[6 empty] | diagnostic box |  |  |  | 0 | 282.429 | 423.0996 | 41.094 | 41.094 | 423.0996 | 282.429 | 186.0843 | 186.0843 |  | 117.6057
LS1 | WB10 | IP | 1861 | LS1_WB10:IP_D1861 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 282.429 | 423.0996 | 41.094 | 41.094 | 423.0996 | 282.429 | 186.0843 | 186.0843 |  | 117.6057
[6 empty] | bellow | 40 |  |  | 0.13122 | 282.3634 | 423.0996 | 41.094 | 41.094 | 423.0996 | 282.2978 | 186.1499 | 186.2155 |  | 117.6713
LS1 | CB11 | GV | 1862 | LS1_CB11:GV_D1862 |  | gate valve | 40 |  |  | 0.057404 | 282.2691 | 423.0996 | 41.094 | 41.094 | 423.0996 | 282.2404 | 186.2442 | 186.2729 |  | 117.7656
[6 empty] | bellows |  |  |  | 0.19082 | 282.145 | 423.0996 | 41.094 | 41.094 | 423.0996 | 282.0496 | 186.3683 | 186.4638 |  | 117.8897
LS1 | CB11 | CAV1 | 1866 | LS1_CB11:CAV1_D1866 |  | b08 resonator | 34 |  |  | 0.29817 | 281.9005 | 423.0996 | 41.094 | 41.094 | 423.0996 | 281.7514 | 186.6128 | 186.7619 |  | 118.1342
[10 empty] | 0 | 281.7514 | 423.0996 | 41.094 | 41.094 | 423.0996 | 281.7514 | 186.7619 | 186.7619 |  | 118.2833
LS1 | CB11 | BLM | 1868 | LS1_CB11:BLM_D1868 |  | outside loss monitor |  |  |  | 0 | 281.7514 | 423.0996 | 41.094 | 41.094 | 423.0996 | 281.7514 | 186.7619 | 186.7619 |  | 118.2833
[6 empty] | bellow | 40 |  |  | 0.09651 | 281.7032 | 423.0996 | 41.094 | 41.094 | 423.0996 | 281.6549 | 186.8102 | 186.8584 |  | 118.3315
LS1 | CB11 | BPM | 1869 | LS1_CB11:BPM_D1869 | BPM6 | position monitor, BPMc |  |  |  | 0 | 281.6549 | 423.0996 | 41.094 | 41.094 | 423.0996 | 281.6549 | 186.8584 | 186.8584 |  | 118.3798
[6 empty] | BPM-box | 40 |  |  | 0.05588 | 281.627 | 423.0996 | 41.094 | 41.094 | 423.0996 | 281.599 | 186.8864 | 186.9143 |  | 118.4077
[6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 281.561 | 423.0996 | 41.094 | 41.094 | 423.0996 | 281.5229 | 186.9524 | 186.9904 |  | 118.4738
LS1 | CB11 | SOL1 | 1872 | LS1_CB11:SOL1_D1872 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 281.2729 | 423.0996 | 41.094 | 41.094 | 423.0996 | 281.0229 | 187.2404 | 187.4904 |  | 118.7618
[6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 280.9848 | 423.0996 | 41.094 | 41.094 | 423.0996 | 280.9468 | 187.5285 | 187.5666 |  | 119.0499
[6 empty] | bellow |  |  |  | 0.099517 | 280.897 | 423.0996 | 41.094 | 41.094 | 423.0996 | 280.8472 | 187.6163 | 187.6661 |  | 119.1377
LS1 | CB11 | CAV2 | 1878 | LS1_CB11:CAV2_D1878 |  | b08 resonator | 34 |  |  | 0.29817 | 280.6982 | 423.0996 | 41.094 | 41.094 | 423.0996 | 280.5491 | 187.8152 | 187.9643 |  | 119.3366
[10 empty] | 0 | 280.5491 | 423.0996 | 41.094 | 41.094 | 423.0996 | 280.5491 | 187.9643 | 187.9643 |  | 119.4856
[10 empty] | 0 | 280.5491 | 423.0996 | 41.094 | 41.094 | 423.0996 | 280.5491 | 187.9643 | 187.9643 |  | 119.4856
[6 empty] | bellow |  |  |  | 0.099517 | 280.4993 | 423.0996 | 41.094 | 41.094 | 423.0996 | 280.4496 | 188.014 | 188.0638 |  | 119.5354
LS1 | CB11 | CAV3 | 1882 | LS1_CB11:CAV3_D1882 |  | b08 resonator | 34 |  |  | 0.29817 | 280.3005 | 423.0996 | 41.094 | 41.094 | 423.0996 | 280.1514 | 188.2129 | 188.362 |  | 119.7342
[10 empty] | 0 | 280.1514 | 423.0996 | 41.094 | 41.094 | 423.0996 | 280.1514 | 188.362 | 188.362 |  | 119.8833
[10 empty] | 0 | 280.1514 | 423.0996 | 41.094 | 41.094 | 423.0996 | 280.1514 | 188.362 | 188.362 |  | 119.8833
[6 empty] | bellow |  |  |  | 0.099517 | 280.1016 | 423.0996 | 41.094 | 41.094 | 423.0996 | 280.0519 | 188.4117 | 188.4615 |  | 119.9331
LS1 | CB11 | CAV4 | 1886 | LS1_CB11:CAV4_D1886 |  | b08 resonator | 34 |  |  | 0.29817 | 279.9028 | 423.0996 | 41.094 | 41.094 | 423.0996 | 279.7537 | 188.6106 | 188.7596 |  | 120.1319
[10 empty] | 0 | 279.7537 | 423.0996 | 41.094 | 41.094 | 423.0996 | 279.7537 | 188.7596 | 188.7596 |  | 120.281
LS1 | CB11 | BLM | 1888 | LS1_CB11:BLM_D1888 |  | outside loss monitor |  |  |  | 0 | 279.7537 | 423.0996 | 41.094 | 41.094 | 423.0996 | 279.7537 | 188.7596 | 188.7596 |  | 120.281
[6 empty] | bellow |  |  |  | 0.09651 | 279.7054 | 423.0996 | 41.094 | 41.094 | 423.0996 | 279.6572 | 188.8079 | 188.8562 |  | 120.3293
LS1 | CB11 | BPM | 1889 | LS1_CB11:BPM_D1889 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 279.6572 | 423.0996 | 41.094 | 41.094 | 423.0996 | 279.6572 | 188.8562 | 188.8562 |  | 120.3775
[6 empty] | BPM-box |  |  |  | 0.05588 | 279.6292 | 423.0996 | 41.094 | 41.094 | 423.0996 | 279.6013 | 188.8841 | 188.912 |  | 120.4055
[6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 279.5632 | 423.0996 | 41.094 | 41.094 | 423.0996 | 279.5252 | 188.9501 | 188.9882 |  | 120.4715
LS1 | CB11 | SOL2 | 1892 | LS1_CB11:SOL2_D1892 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 279.2752 | 423.0996 | 41.094 | 41.094 | 423.0996 | 279.0252 | 189.2382 | 189.4882 |  | 120.7595
[6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 278.9871 | 423.0996 | 41.094 | 41.094 | 423.0996 | 278.949 | 189.5262 | 189.5643 |  | 121.0476
[6 empty] | bellow |  |  |  | 0.099517 | 278.8993 | 423.0996 | 41.094 | 41.094 | 423.0996 | 278.8495 | 189.6141 | 189.6638 |  | 121.1354
LS1 | CB11 | CAV5 | 1898 | LS1_CB11:CAV5_D1898 |  | b08 resonator | 34 |  |  | 0.29817 | 278.7004 | 423.0996 | 41.094 | 41.094 | 423.0996 | 278.5513 | 189.8129 | 189.962 |  | 121.3343
[10 empty] | 0 | 278.5513 | 423.0996 | 41.094 | 41.094 | 423.0996 | 278.5513 | 189.962 | 189.962 |  | 121.4834
[10 empty] | 0 | 278.5513 | 423.0996 | 41.094 | 41.094 | 423.0996 | 278.5513 | 189.962 | 189.962 |  | 121.4834
[6 empty] | bellow |  |  |  | 0.099517 | 278.5016 | 423.0996 | 41.094 | 41.094 | 423.0996 | 278.4518 | 190.0118 | 190.0615 |  | 121.5331
LS1 | CB11 | CAV6 | 1902 | LS1_CB11:CAV6_D1902 |  | b08 resonator | 34 |  |  | 0.29817 | 278.3027 | 423.0996 | 41.094 | 41.094 | 423.0996 | 278.1537 | 190.2106 | 190.3597 |  | 121.732
[10 empty] | 0 | 278.1537 | 423.0996 | 41.094 | 41.094 | 423.0996 | 278.1537 | 190.3597 | 190.3597 |  | 121.8811
[10 empty] | 0 | 278.1537 | 423.0996 | 41.094 | 41.094 | 423.0996 | 278.1537 | 190.3597 | 190.3597 |  | 121.8811
[6 empty] | bellow |  |  |  | 0.099517 | 278.1039 | 423.0996 | 41.094 | 41.094 | 423.0996 | 278.0541 | 190.4094 | 190.4592 |  | 121.9308
LS1 | CB11 | CAV7 | 1906 | LS1_CB11:CAV7_D1906 |  | b08 resonator | 34 |  |  | 0.29817 | 277.9051 | 423.0996 | 41.094 | 41.094 | 423.0996 | 277.756 | 190.6083 | 190.7574 |  | 122.1297
[10 empty] | 0 | 277.756 | 423.0996 | 41.094 | 41.094 | 423.0996 | 277.756 | 190.7574 | 190.7574 |  | 122.2787
LS1 | CB11 | BLM | 1908 | LS1_CB11:BLM_D1908 |  | outside loss monitor |  |  |  | 0 | 277.756 | 423.0996 | 41.094 | 41.094 | 423.0996 | 277.756 | 190.7574 | 190.7574 |  | 122.2787
[6 empty] | bellow |  |  |  | 0.09651 | 277.7077 | 423.0996 | 41.094 | 41.094 | 423.0996 | 277.6595 | 190.8056 | 190.8539 |  | 122.327
LS1 | CB11 | BPM | 1909 | LS1_CB11:BPM_D1909 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 277.6595 | 423.0996 | 41.094 | 41.094 | 423.0996 | 277.6595 | 190.8539 | 190.8539 |  | 122.3753
[6 empty] | BPM-box |  |  |  | 0.05588 | 277.6315 | 423.0996 | 41.094 | 41.094 | 423.0996 | 277.6036 | 190.8818 | 190.9098 |  | 122.4032
[6 empty] | solenoid-entry |  |  |  | 0.076136 | 277.5655 | 423.0996 | 41.094 | 41.094 | 423.0996 | 277.5274 | 190.9478 | 190.9859 |  | 122.4692
LS1 | CB11 | SOL3 | 1912 | LS1_CB11:SOL3_D1912 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 277.2774 | 423.0996 | 41.094 | 41.094 | 423.0996 | 277.0274 | 191.2359 | 191.4859 |  | 122.7573
[6 empty] | solenoid-exit |  |  |  | 0.076136 | 276.9894 | 423.0996 | 41.094 | 41.094 | 423.0996 | 276.9513 | 191.524 | 191.562 |  | 123.0453
[6 empty] | bellow |  |  |  | 0.099517 | 276.9015 | 423.0996 | 41.094 | 41.094 | 423.0996 | 276.8518 | 191.6118 | 191.6616 |  | 123.1332
LS1 | CB11 | CAV8 | 1918 | LS1_CB11:CAV8_D1918 |  | b08 resonator | 34 |  |  | 0.29817 | 276.7027 | 423.0996 | 41.094 | 41.094 | 423.0996 | 276.5536 | 191.8106 | 191.9597 |  | 123.332
[10 empty] | 0 | 276.5536 | 423.0996 | 41.094 | 41.094 | 423.0996 | 276.5536 | 191.9597 | 191.9597 |  | 123.4811
[10 empty] | 0 | 276.5536 | 423.0996 | 41.094 | 41.094 | 423.0996 | 276.5536 | 191.9597 | 191.9597 |  | 123.4811
[6 empty] | bellows |  |  |  | 0.19082 | 276.4582 | 423.0996 | 41.094 | 41.094 | 423.0996 | 276.3628 | 192.0551 | 192.1505 |  | 123.5765
LS1 | CB11 | GV | 1922 | LS1_CB11:GV_D1922 |  | gate valve | 40 |  |  | 0.057404 | 276.3341 | 423.0996 | 41.094 | 41.094 | 423.0996 | 276.3054 | 192.1792 | 192.208 |  | 123.7006
[6 empty] | bellow | 40 |  |  | 0.1035 | 276.2536 | 423.0996 | 41.094 | 41.094 | 423.0996 | 276.2019 | 192.2597 | 192.3115 |  | 123.7811
LS1 | WB11 | BPM | 1923 | LS1_WB11:BPM_D1923 | BPM2 | position monitor |  |  |  | 0 | 276.2019 | 423.0996 | 41.094 | 41.094 | 423.0996 | 276.2019 | 192.3115 | 192.3115 |  | 123.8328
[6 empty] | diagnostic box |  |  |  | 0.14528 | 276.1292 | 423.0996 | 41.094 | 41.094 | 423.0996 | 276.0566 | 192.3841 | 192.4567 |  | 123.9055
LS1 | WB11 | HMR | 1925 | LS1_WB11:HMR_D1925 | HR_L | Halo ring, large aperture | 33 |  |  | 0 | 276.0566 | 423.0996 | 41.094 | 41.094 | 423.0996 | 276.0566 | 192.4567 | 192.4567 |  | 123.9781
[6 empty] | diagnostic box |  |  |  | 0 | 276.0566 | 423.0996 | 41.094 | 41.094 | 423.0996 | 276.0566 | 192.4567 | 192.4567 |  | 123.9781
LS1 | WB11 | IP | 1925 | LS1_WB11:IP_D1925 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 276.0566 | 423.0996 | 41.094 | 41.094 | 423.0996 | 276.0566 | 192.4567 | 192.4567 |  | 123.9781
[6 empty] | bellow | 40 |  |  | 0.13122 | 275.991 | 423.0996 | 41.094 | 41.094 | 423.0996 | 275.9254 | 192.5223 | 192.588 |  | 124.0437
SEGMENT 1 END | [9 empty] | -9.6e-07 | 275.9254 | 423.0996 | 41.094 | 41.094 | 423.0996 | 275.9254 | 192.5879 | 192.5879 | adjustment for residuals | 124.1093
[10 empty] | 1.05 | 275.4004 | 423.0996 | 41.094 | 41.094 | 423.0996 | 274.8754 | 193.1129 | 193.6379 |  | 124.6343
LS1 | BTS | DCHV | 1937 | LS1_BTS:DCHV_D1937 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 274.7754 | 423.0996 | 41.094 | 41.094 | 423.0996 | 274.6754 | 193.7379 | 193.8379 |  | 125.2593
[10 empty] | 0.25 | 274.5504 | 423.0996 | 41.094 | 41.094 | 423.0996 | 274.4254 | 193.9629 | 194.0879 |  | 125.4843
LS1 | BTS | QH | 1942 | LS1_BTS:QH_D1942 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 274.3004 | 423.0996 | 41.094 | 41.094 | 423.0996 | 274.1754 | 194.2129 | 194.3379 |  | 125.7343
[10 empty] | 0.29595 | 274.0274 | 423.0996 | 41.094 | 41.094 | 423.0996 | 273.8794 | 194.4859 | 194.6339 |  | 126.0073
LS1 | BTS | GV | 1947 | LS1_BTS:GV_D1947 | GV_M | gate valve | 63 | 1.9924e-06 | 273.8419 | 0.075 | 273.8419 | 423.0996 | 41.094 | 41.094 | 423.0996 | 273.8044 | 194.6714 | 194.7089 |  | 126.1928
[10 empty] | 0.12905 | 273.7399 | 423.0996 | 41.094 | 41.094 | 423.0996 | 273.6754 | 194.7734 | 194.8379 |  | 126.2948
LS1 | BTS | QV | 1950 | LS1_BTS:QV_D1950 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 273.5504 | 423.0996 | 41.094 | 41.094 | 423.0996 | 273.4254 | 194.9629 | 195.0879 |  | 126.4843
[10 empty] | 1.198 | 272.8264 | 423.0996 | 41.094 | 41.094 | 423.0996 | 272.2274 | 195.6869 | 196.2859 |  | 127.2083
LS1 | BTS | DCHV | 1964 | LS1_BTS:DCHV_D1964 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 272.1274 | 423.0996 | 41.094 | 41.094 | 423.0996 | 272.0274 | 196.3859 | 196.4859 |  | 127.9073
[10 empty] | 0 | 272.0274 | 423.0996 | 41.094 | 41.094 | 423.0996 | 272.0274 | 196.4859 | 196.4859 |  | 128.0073
[8 empty] | 2.9924e-06 |  | 0.19214 | 271.9313 | 423.0996 | 41.094 | 41.094 | 423.0996 | 271.8352 | 196.582 | 196.6781 |  | 128.1034
LS1 | BTS | BPM | 1967 | LS1_BTS:BPM_D1967 | BPM3 | position monitor | 47.5 | 271.8353 | 271.8352 | 0 | 271.8352 | 423.0996 | 41.094 | 41.094 | 423.0996 | 271.8352 | 196.6781 | 196.6781 | 0.137 | 128.1995
[8 empty] | inconsisent with next one |  | 0.057856 | 271.8063 | 423.0996 | 41.094 | 41.094 | 423.0996 | 271.7774 | 196.707 | 196.7359 |  | 128.2284
LS1 | BTS | QH | 1969 | LS1_BTS:QH_D1969 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 271.6524 | 423.0996 | 41.094 | 41.094 | 423.0996 | 271.5274 | 196.8609 | 196.9859 |  | 128.3823
[10 empty] | 0.5 | 271.2774 | 423.0996 | 41.094 | 41.094 | 423.0996 | 271.0274 | 197.2359 | 197.4859 |  | 128.7573
LS1 | BTS | QV | 1976 | LS1_BTS:QV_D1976 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 270.9024 | 423.0996 | 41.094 | 41.094 | 423.0996 | 270.7774 | 197.6109 | 197.7359 |  | 129.1323
[10 empty] | 0.1325 | 270.7111 | 423.0996 | 41.094 | 41.094 | 423.0996 | 270.6449 | 197.8022 | 197.8685 |  | 129.3236
[6 empty] | bellow |  |  |  | 0.1035 | 270.5931 | 423.0996 | 41.094 | 41.094 | 423.0996 | 270.5414 | 197.9202 | 197.972 |  | 129.4416
LS1 | BTS | BPM | 1980 | LS1_BTS:BPM_D1980 | BPM2 | position monitor | 47.5 | replace by a nipple |  | 0 | 270.5414 | 423.0996 | 41.094 | 41.094 | 423.0996 | 270.5414 | 197.972 | 197.972 | Should be BMP3a, 4/24/2017 | 129.4933
[6 empty] | diagnostic box |  |  |  | 0.14528 | 270.4687 | 423.0996 | 41.094 | 41.094 | 423.0996 | 270.3961 | 198.0446 | 198.1172 |  | 129.566
LS1 | BTS | PORT | 1981 | LS1_BTS:PORT_D1981 |  | port for future diagnostics |  |  |  | 0 | 270.3961 | 423.0996 | 41.094 | 41.094 | 423.0996 | 270.3961 | 198.1172 | 198.1172 |  | 129.6386
[6 empty] | diagnostic box |  |  |  | 0 | 270.3961 | 423.0996 | 41.094 | 41.094 | 423.0996 | 270.3961 | 198.1172 | 198.1172 |  | 129.6386
LS1 | BTS | IP | 1981 | LS1_BTS:IP_D1981 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 270.3961 | 423.0996 | 41.094 | 41.094 | 423.0996 | 270.3961 | 198.1172 | 198.1172 |  | 129.6386
[6 empty] | bellow |  |  |  | 0.13122 | 270.3305 | 423.0996 | 41.094 | 41.094 | 423.0996 | 270.2649 | 198.1828 | 198.2485 |  | 129.7042
[10 empty] | 0 | 270.2649 | 423.0996 | 41.094 | 41.094 | 423.0996 | 270.2649 | 198.2485 | 198.2485 |  | 129.7698
[10 empty] | 0.059497 | 270.2351 | 423.0996 | 41.094 | 41.094 | 423.0996 | 270.2054 | 198.2782 | 198.3079 |  | 129.7996
[10 empty] | 0.052097 | 270.1793 | 423.0996 | 41.094 | 41.094 | 423.0996 | 270.1533 | 198.334 | 198.36 |  | 129.8554
[10 empty] | 0.32364 | 269.9915 | 423.0996 | 41.094 | 41.094 | 423.0996 | 269.8297 | 198.5219 | 198.6837 |  | 130.0432
[10 empty] | 0 | 269.8297 | 423.0996 | 41.094 | 41.094 | 423.0996 | 269.8297 | 198.6837 | 198.6837 |  | 130.2051
[10 empty] | 0 | 269.8297 | 423.0996 | 41.094 | 41.094 | 423.0996 | 269.8297 | 198.6837 | 198.6837 |  | 130.2051
[10 empty] | 0.099517 | 269.7799 | 423.0996 | 41.094 | 41.094 | 423.0996 | 269.7301 | 198.7334 | 198.7832 |  | 130.2548
[10 empty] | 0.32364 | 269.5683 | 423.0996 | 41.094 | 41.094 | 423.0996 | 269.4065 | 198.945 | 199.1069 |  | 130.4664
[10 empty] | 0 | 269.4065 | 423.0996 | 41.094 | 41.094 | 423.0996 | 269.4065 | 199.1069 | 199.1069 |  | 130.6282
[10 empty] | 0 | 269.4065 | 423.0996 | 41.094 | 41.094 | 423.0996 | 269.4065 | 199.1069 | 199.1069 |  | 130.6282
[10 empty] | 0.052097 | 269.3804 | 423.0996 | 41.094 | 41.094 | 423.0996 | 269.3544 | 199.1329 | 199.1589 |  | 130.6543
[10 empty] | 0.072 | 269.3184 | 423.0996 | 41.094 | 41.094 | 423.0996 | 269.2824 | 199.1949 | 199.2309 |  | 130.7163
[10 empty] | 0 | 269.2824 | 423.0996 | 41.094 | 41.094 | 423.0996 | 269.2824 | 199.2309 | 199.2309 |  | 130.7523
[6 empty] | bellow |  |  |  | 0.094006 | 269.2354 | 423.0996 | 41.094 | 41.094 | 423.0996 | 269.1884 | 199.278 | 199.325 |  | 130.7993
[6 empty] | vacuum box |  |  |  | 0 | 269.1884 | 423.0996 | 41.094 | 41.094 | 423.0996 | 269.1884 | 199.325 | 199.325 |  | 130.8463
LS1 | BTS | PORT | 1993 | LS1_BTS:PORT_D1993 |  | future diagnostic port |  |  |  | 0 | 269.1884 | 423.0996 | 41.094 | 41.094 | 423.0996 | 269.1884 | 199.325 | 199.325 |  | 130.8463
[6 empty] | vacuum box |  | 2.9924e-06 |  | 0 | 269.1884 | 423.0996 | 41.094 | 41.094 | 423.0996 | 269.1884 | 199.325 | 199.325 |  | 130.8463
LS1 | BTS | IP | 1993 | LS1_BTS:IP_D1993 | IP_S | ion pump, 75 l/s |  | 269.1884 | 269.1884 | 0 | 269.1884 | 423.0996 | 41.094 | 41.094 | 423.0996 | 269.1884 | 199.325 | 199.325 |  | 130.8463
[6 empty] | bellow |  |  |  | 0.18999 | 269.0934 | 423.0996 | 41.094 | 41.094 | 423.0996 | 268.9984 | 199.42 | 199.5149 |  | 130.9413
[10 empty] | 0.106 | 268.9454 | 423.0996 | 41.094 | 41.094 | 423.0996 | 268.8924 | 199.5679 | 199.6209 |  | 131.0893
LS1 | BTS | DCHV | 1997 | LS1_BTS:DCHV_D1997 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 268.7924 | 423.0996 | 41.094 | 41.094 | 423.0996 | 268.6924 | 199.7209 | 199.8209 |  | 131.2423
[10 empty] | 0.25 | 268.5674 | 423.0996 | 41.094 | 41.094 | 423.0996 | 268.4424 | 199.9459 | 200.0709 |  | 131.4673
LS1 | BTS | QH | 2002 | LS1_BTS:QH_D2002 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 268.3174 | 423.0996 | 41.094 | 41.094 | 423.0996 | 268.1924 | 200.1959 | 200.3209 |  | 131.7173
[10 empty] | 0.88621 | 267.7493 | 423.0996 | 41.094 | 41.094 | 423.0996 | 267.3062 | 200.7641 | 201.2072 |  | 132.2854
LS1 | BTS | QV | 2013 | LS1_BTS:QV_D2013 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 267.1812 | 423.0996 | 41.094 | 41.094 | 423.0996 | 267.0562 | 201.3322 | 201.4572 |  | 132.8535
[10 empty] | 0.87 | 266.6212 | 423.0996 | 41.094 | 41.094 | 423.0996 | 266.1862 | 201.8922 | 202.3272 |  | 133.4135
LS1 | BTS | DCHV | 2024 | LS1_BTS:DCHV_D2024 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 266.0862 | 423.0996 | 41.094 | 41.094 | 423.0996 | 265.9862 | 202.4272 | 202.5272 |  | 133.9485
[8 empty] | 2.9924e-06 |  | 0.1909 | 265.8907 | 423.0996 | 41.094 | 41.094 | 423.0996 | 265.7953 | 202.6226 | 202.7181 | 0.137 | 134.144
LS1 | BTS | BPM | 2027 | LS1_BTS:BPM_D2027 | BPM3 | position monitor | 47.5 | 265.7953 | 265.7953 | 0 | 265.7953 | 423.0996 | 41.094 | 41.094 | 423.0996 | 265.7953 | 202.7181 | 202.7181 |  | 134.2394
[8 empty] | inconsisent with previous one |  | 0 | 265.7953 | 423.0996 | 41.094 | 41.094 | 423.0996 | 265.7953 | 202.7181 | 202.7181 |  | 134.2394
[10 empty] | 0.059095 | 265.7657 | 423.0996 | 41.094 | 41.094 | 423.0996 | 265.7362 | 202.7476 | 202.7772 |  | 134.269
LS1 | BTS | QH | 2029 | LS1_BTS:QH_D2029 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 265.6112 | 423.0996 | 41.094 | 41.094 | 423.0996 | 265.4862 | 202.9022 | 203.0272 |  | 134.4235
[10 empty] | 1.0862 | 264.9431 | 423.0996 | 41.094 | 41.094 | 423.0996 | 264.4 | 203.5703 | 204.1134 |  | 135.0916
LS1 | BTS | QV | 2042 | LS1_BTS:QV_D2042 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 264.275 | 423.0996 | 41.094 | 41.094 | 423.0996 | 264.15 | 204.2384 | 204.3634 |  | 135.7597
[10 empty] | 0.97033 | 263.6648 | 423.0996 | 41.094 | 41.094 | 423.0996 | 263.1796 | 204.8485 | 205.3337 |  | 136.3699
[6 empty] | bellow |  | 2.9924e-06 |  | 0.1035 | 263.1279 | 423.0996 | 41.094 | 41.094 | 423.0996 | 263.0761 | 205.3855 | 205.4372 |  | 136.9068
LS1 | BTS | BPM | 2054 | LS1_BTS:BPM_D2054 | BPM3 | position monitor | 47.5 | 263.0761 | 263.0761 | 0 | 263.0761 | 423.0996 | 41.094 | 41.094 | 423.0996 | 263.0761 | 205.4372 | 205.4372 |  | 136.9586
[6 empty] | diagnostic box |  |  |  | 0.14528 | 263.0035 | 423.0996 | 41.094 | 41.094 | 423.0996 | 262.9309 | 205.5098 | 205.5825 |  | 137.0312
LS1 | BTS | PM | 2056 | LS1_BTS:PM_D2056 | PM3 | profile monitor |  |  |  | 0 | 262.9309 | 423.0996 | 41.094 | 41.094 | 423.0996 | 262.9309 | 205.5825 | 205.5825 |  | 137.1039
[6 empty] | diagnostic box |  |  |  | 0 | 262.9309 | 423.0996 | 41.094 | 41.094 | 423.0996 | 262.9309 | 205.5825 | 205.5825 |  | 137.1039
LS1 | BTS | IP | 2056 | LS1_BTS:IP_D2056 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 262.9309 | 423.0996 | 41.094 | 41.094 | 423.0996 | 262.9309 | 205.5825 | 205.5825 |  | 137.1039
[6 empty] | bellow |  |  |  | 0.13122 | 262.8652 | 423.0996 | 41.094 | 41.094 | 423.0996 | 262.7996 | 205.6481 | 205.7137 |  | 137.1695
[10 empty] | 0 | 262.7996 | 423.0996 | 41.094 | 41.094 | 423.0996 | 262.7996 | 205.7137 | 205.7137 |  | 137.2351
[10 empty] | 0.26878 | 262.6652 | 423.0996 | 41.094 | 41.094 | 423.0996 | 262.5309 | 205.8481 | 205.9825 | 2012-03-09 why this odd length ? Make it say 0.78 | 137.3695
LS1 | BTS | DCHV | 2061 | LS1_BTS:DCHV_D2061 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 262.4309 | 423.0996 | 41.094 | 41.094 | 423.0996 | 262.3309 | 206.0825 | 206.1825 |  | 137.6039
[10 empty] | 0.25 | 262.2059 | 423.0996 | 41.094 | 41.094 | 423.0996 | 262.0809 | 206.3075 | 206.4325 |  | 137.8289
LS1 | BTS | QV | 2066 | LS1_BTS:QV_D2066 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 261.9559 | 423.0996 | 41.094 | 41.094 | 423.0996 | 261.8309 | 206.5575 | 206.6825 |  | 138.0789
[10 empty] | 0.5 | 261.5809 | 423.0996 | 41.094 | 41.094 | 423.0996 | 261.3309 | 206.9325 | 207.1825 |  | 138.4539
LS1 | BTS | QH | 2073 | LS1_BTS:QH_D2073 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 261.2059 | 423.0996 | 41.094 | 41.094 | 423.0996 | 261.0809 | 207.3075 | 207.4325 |  | 138.8289
[10 empty] | 0.7 | 260.7309 | 423.0996 | 41.094 | 41.094 | 423.0996 | 260.3809 | 207.7825 | 208.1325 |  | 139.3039
[10 empty] | 0 | 260.3809 | 423.0996 | 41.094 | 41.094 | 423.0996 | 260.3809 | 208.1325 | 208.1325 |  | 139.6539
LS1 | CH01 | GV | 2082 | LS1_CH01:GV_D2082 |  | gate valve | 40 | 0.057404 |  | 0.057404 | 260.3522 | 423.0996 | 41.094 | 41.094 | 423.0996 | 260.3235 | 208.1612 | 208.1899 |  | 139.6826
[6 empty] | bellows |  |  |  | 0.19049 | 260.2282 | 423.0996 | 41.094 | 41.094 | 423.0996 | 260.133 | 208.2851 | 208.3804 |  | 139.8065
LS1 | CH01 | CAV1 | 2085 | LS1_CH01:CAV1_D2085 |  | b08 resonator | 34 | 0.29817 |  | 0.29817 | 259.9839 | 423.0996 | 41.094 | 41.094 | 423.0996 | 259.8348 | 208.5295 | 208.6785 |  | 140.0508
[6 empty] | bellows |  | 0.099517 |  | 0.099517 | 259.785 | 423.0996 | 41.094 | 41.094 | 423.0996 | 259.7353 | 208.7283 | 208.7781 |  | 140.2497
LS1 | CH01 | CAV2 | 2089 | LS1_CH01:CAV2_D2089 |  | b08 resonator | 34 | 0.29817 |  | 0.29817 | 259.5862 | 423.0996 | 41.094 | 41.094 | 423.0996 | 259.4371 | 208.9272 | 209.0762 |  | 140.4485
[6 empty] | bellows |  |  |  | 0.099517 | 259.3873 | 423.0996 | 41.094 | 41.094 | 423.0996 | 259.3376 | 209.126 | 209.1758 |  | 140.6474
[10 empty] | 0.19798 | 259.2386 | 423.0996 | 41.094 | 41.094 | 423.0996 | 259.1396 | 209.2747 | 209.3737 |  | 140.7961
[6 empty] | bellows |  |  |  | 0.099517 | 259.0898 | 423.0996 | 41.094 | 41.094 | 423.0996 | 259.0401 | 209.4235 | 209.4733 |  | 140.9449
LS1 | CH01 | CAV3 | 2096 | LS1_CH01:CAV3_D2096 |  | b08 resonator | 34 | 0.29817 |  | 0.29817 | 258.891 | 423.0996 | 41.094 | 41.094 | 423.0996 | 258.7419 | 209.6223 | 209.7714 |  | 141.1437
[6 empty] | bellows |  | 0.099517 |  | 0.099517 | 258.6922 | 423.0996 | 41.094 | 41.094 | 423.0996 | 258.6424 | 209.8212 | 209.8709 |  | 141.3426
LS1 | CH01 | CAV4 | 2100 | LS1_CH01:CAV4_D2100 |  | b08 resonator | 34 | 0.29817 |  | 0.29817 | 258.4933 | 423.0996 | 41.094 | 41.094 | 423.0996 | 258.3442 | 210.02 | 210.1691 |  | 141.5414
[6 empty] | bellows |  |  |  | 0.19049 | 258.249 | 423.0996 | 41.094 | 41.094 | 423.0996 | 258.1537 | 210.2644 | 210.3596 |  | 141.7857
LS1 | CH01 | GV | 2104 | LS1_CH01:GV_D2104 |  | gate valve | 40 | 0.057404 |  | 0.057404 | 258.125 | 423.0996 | 41.094 | 41.094 | 423.0996 | 258.0963 | 210.3883 | 210.417 |  | 141.9097
[10 empty] | 0.058708 | 258.067 | 423.0996 | 41.094 | 41.094 | 423.0996 | 258.0376 | 210.4464 | 210.4757 |  | 141.9677
LS1 | BTS | GV | 2105 | LS1_BTS:GV_D2105 | GV_M | gate valve | 63 | 1.9924e-06 | 258.0001 | 0.075 | 258.0001 | 423.0996 | 41.094 | 41.094 | 423.0996 | 257.9626 | 210.5132 | 210.5507 |  | 142.0346
[10 empty] | 0.087596 | 257.9188 | 423.0996 | 41.094 | 41.094 | 423.0996 | 257.875 | 210.5945 | 210.6383 |  | 142.1159
LS1 | BTS | IP | 2106 | LS1_BTS:IP_D2106 | IP_S | ion pump, 75 l/s |  | 257.875 | 258.0376 | 0 | 257.875 | 423.0996 | 41.094 | 41.094 | 423.0996 | 257.875 | 210.6383 | 210.6383 |  | 142.1597
[6 empty] | vacuum box |  |  |  | 0.51924 | 257.6154 | 423.0996 | 41.094 | 41.094 | 423.0996 | 257.3558 | 210.8979 | 211.1576 |  | 142.4193
[10 empty] | 0 | 257.3558 | 423.0996 | 41.094 | 41.094 | 423.0996 | 257.3558 | 211.1576 | 211.1576 |  | 142.6789
[10 empty] | 0.1 | 257.3058 | 423.0996 | 41.094 | 41.094 | 423.0996 | 257.2558 | 211.2076 | 211.2576 |  | 142.7289
LS1 | BTS | DCHV | 2114 | LS1_BTS:DCHV_D2114 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 257.1558 | 423.0996 | 41.094 | 41.094 | 423.0996 | 257.0558 | 211.3576 | 211.4576 |  | 142.8789
[10 empty] | 0 | 257.0558 | 423.0996 | 41.094 | 41.094 | 423.0996 | 257.0558 | 211.4576 | 211.4576 |  | 142.9789
[8 empty] | 2.9924e-06 |  | 0.19218 | 256.9597 | 423.0996 | 41.094 | 41.094 | 423.0996 | 256.8636 | 211.5536 | 211.6497 |  | 143.075
LS1 | BTS | BPM | 2116 | LS1_BTS:BPM_D2116 | BPM3 | position monitor | 47.5 | 256.8636 | 256.8636 | 0 | 256.8636 | 423.0996 | 41.094 | 41.094 | 423.0996 | 256.8636 | 211.6497 | 211.6497 | 0.137 | 143.1711
[10 empty] | 0.057824 | 256.8347 | 423.0996 | 41.094 | 41.094 | 423.0996 | 256.8058 | 211.6786 | 211.7076 |  | 143.2
LS1 | BTS | QV | 2118 | LS1_BTS:QV_D2118 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 256.6808 | 423.0996 | 41.094 | 41.094 | 423.0996 | 256.5558 | 211.8326 | 211.9576 |  | 143.3539
[10 empty] | 0.5 | 256.3058 | 423.0996 | 41.094 | 41.094 | 423.0996 | 256.0558 | 212.2076 | 212.4576 |  | 143.7289
LS1 | BTS | QH | 2126 | LS1_BTS:QH_D2126 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 255.9308 | 423.0996 | 41.094 | 41.094 | 423.0996 | 255.8058 | 212.5826 | 212.7076 |  | 144.1039
[10 empty] | 0.16165 | 255.725 | 423.0996 | 41.094 | 41.094 | 423.0996 | 255.6441 | 212.7884 | 212.8692 |  | 144.3097
[6 empty] | bellow |  |  |  | 0.1035 | 255.5924 | 423.0996 | 41.094 | 41.094 | 423.0996 | 255.5406 | 212.921 | 212.9727 |  | 144.4423
LS1 | BTS | BPM | 2130 | LS1_BTS:BPM_D2130 | BPM2 | position monitor | 40 | 1.9924e-06 | 255.5406 | 0 | 255.5406 | 423.0996 | 41.094 | 41.094 | 423.0996 | 255.5406 | 212.9727 | 212.9727 |  | 144.4941
[6 empty] | diagnostic box |  |  |  | 0.14528 | 255.468 | 423.0996 | 41.094 | 41.094 | 423.0996 | 255.3954 | 213.0453 | 213.118 |  | 144.5667
LS1 | BTS | PM | 2131 | LS1_BTS:PM_D2131 | PM2 | profile monitor | defer installation |  |  | 0 | 255.3954 | 423.0996 | 41.094 | 41.094 | 423.0996 | 255.3954 | 213.118 | 213.118 |  | 144.6394
[6 empty] | diagnostic box |  | 1.9924e-06 | 255.3954 | 0 | 255.3954 | 423.0996 | 41.094 | 41.094 | 423.0996 | 255.3954 | 213.118 | 213.118 |  | 144.6394
LS1 | BTS | IP | 2131 | LS1_BTS:IP_D2131 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 255.3954 | 423.0996 | 41.094 | 41.094 | 423.0996 | 255.3954 | 213.118 | 213.118 |  | 144.6394
[6 empty] | bellow |  | 0.16514 |  | 0.13122 | 255.3297 | 423.0996 | 41.094 | 41.094 | 423.0996 | 255.2641 | 213.1836 | 213.2492 |  | 144.705
[8 empty] | 1.9924e-06 | 255.1755 | 0.088633 | 255.2198 | 423.0996 | 41.094 | 41.094 | 423.0996 | 255.1755 | 213.2935 | 213.3378 |  | 144.8149
FS1 | CH01 | GV | 2134 | FS1_CH02:GV_D2134 |  | gate valve | 40 | 0.057404 | 2.2845 | 0.057404 | 255.1468 | 423.0996 | 41.094 | 41.094 | 423.0996 | 255.1181 | 213.3665 | 213.3952 | b85 buncher moved 0.165123 m upstream | 144.8879
[6 empty] | bellows |  |  |  | 0.19049 | 255.0229 | 423.0996 | 41.094 | 41.094 | 423.0996 | 254.9276 | 213.4905 | 213.5857 |  | 145.0119
FS1 | CH01 | CAV1 | 2137 | FS1_CH02:CAV1_D2137 |  | b08 resonator | 34 | 0.29817 |  | 0.29817 | 254.7785 | 423.0996 | 41.094 | 41.094 | 423.0996 | 254.6294 | 213.7348 | 213.8839 |  | 145.2562
[6 empty] | bellows |  | 0.099517 |  | 0.099517 | 254.5797 | 423.0996 | 41.094 | 41.094 | 423.0996 | 254.5299 | 213.9337 | 213.9834 |  | 145.455
FS1 | CH01 | CAV2 | 2141 | FS1_CH02:CAV2_D2141 |  | b08 resonator | 34 | 0.29817 |  | 0.29817 | 254.3808 | 423.0996 | 41.094 | 41.094 | 423.0996 | 254.2318 | 214.1325 | 214.2816 |  | 145.6539
[6 empty] | bellows |  |  |  | 0.099517 | 254.182 | 423.0996 | 41.094 | 41.094 | 423.0996 | 254.1322 | 214.3313 | 214.3811 |  | 145.8527
[10 empty] | 0.19798 | 254.0332 | 423.0996 | 41.094 | 41.094 | 423.0996 | 253.9343 | 214.4801 | 214.5791 |  | 146.0015
[6 empty] | bellows |  |  |  | 0.099517 | 253.8845 | 423.0996 | 41.094 | 41.094 | 423.0996 | 253.8347 | 214.6288 | 214.6786 |  | 146.1502
FS1 | CH01 | CAV3 | 2148 | FS1_CH02:CAV3_D2148 |  | b08 resonator | 34 | 0.29817 |  | 0.29817 | 253.6856 | 423.0996 | 41.094 | 41.094 | 423.0996 | 253.5366 | 214.8277 | 214.9768 |  | 146.3491
[6 empty] | bellows |  | 0.099517 |  | 0.099517 | 253.4868 | 423.0996 | 41.094 | 41.094 | 423.0996 | 253.437 | 215.0265 | 215.0763 |  | 146.5479
FS1 | CH01 | CAV4 | 2152 | FS1_CH02:CAV4_D2152 |  | b08 resonator | 34 | 0.29817 |  | 0.29817 | 253.288 | 423.0996 | 41.094 | 41.094 | 423.0996 | 253.1389 | 215.2254 | 215.3745 |  | 146.7468
[6 empty] | bellows |  |  |  | 0.19049 | 253.0436 | 423.0996 | 41.094 | 41.094 | 423.0996 | 252.9484 | 215.4697 | 215.565 |  | 146.9911
FS1 | CH01 | GV | 2156 | FS1_CH02:GV_D2156 |  | gate valve | 40 | 0.057404 |  | 0.057404 | 252.9197 | 423.0996 | 41.094 | 41.094 | 423.0996 | 252.891 | 215.5937 | 215.6224 |  | 147.115
[8 empty] | -7.6277e-09 | 252.891 | 0.16514 | 252.8084 | 423.0996 | 41.094 | 41.094 | 423.0996 | 252.7258 | 215.7049 | 215.7875 |  | 147.2263
[6 empty] | bellow |  |  |  | 0.011054 | 252.7203 | 423.0996 | 41.094 | 41.094 | 423.0996 | 252.7148 | 215.793 | 215.7986 |  | 147.3144
[6 empty] | vacuum box |  |  |  | 0 | 252.7148 | 423.0996 | 41.094 | 41.094 | 423.0996 | 252.7148 | 215.7986 | 215.7986 |  | 147.3199
FS1 | CSS | FAVS | 2158 | FS1_CSS:FAVS_D2158 | FVS | fast valve sensor |  | 0.031522 |  | 0 | 252.7148 | 423.0996 | 41.094 | 41.094 | 423.0996 | 252.7148 | 215.7986 | 215.7986 |  | 147.3199
[6 empty] | vacuum box |  | 2.1992e-05 | 252.7148 | 0 | 252.7148 | 423.0996 | 41.094 | 41.094 | 423.0996 | 252.7148 | 215.7986 | 215.7986 |  | 147.3199
FS1 | CSS | IP | 2158 | FS1_CSS:IP_D2158 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 252.7148 | 423.0996 | 41.094 | 41.094 | 423.0996 | 252.7148 | 215.7986 | 215.7986 |  | 147.3199
[6 empty] | bellow |  | 0.027624 | 0.076642 | 0.16868 | 252.6304 | 423.0996 | 41.094 | 41.094 | 423.0996 | 252.5461 | 215.8829 | 215.9672 |  | 147.4043
FS1 | CSS | FAV | 2160 | FS1_CSS:FAV_D2160 | FV | fast valve | 40 | 2.1992e-05 | 252.5461 | 0 | 252.5461 | 423.0996 | 41.094 | 41.094 | 423.0996 | 252.5461 | 215.9672 | 215.9672 | 0.067196 | 147.4886
[10 empty] | 0.20307 | 252.4446 | 423.0996 | 41.094 | 41.094 | 423.0996 | 252.343 | 216.0688 | 216.1703 |  | 147.5901
FS1 | CSS | DH | 2163 | FS1_CSS:DH_D2163 | D1 | dipole-5deg, D1-1, C1 | 40 x 40 |  |  | 0.3 | 252.1932 | 423.0931 | 41.094 | 41.094 | 423.0865 | 252.0434 | 216.3203 | 216.4703 |  | 147.8417
[10 empty] | 0.25 | 251.9189 | 423.0756 | 41.094 | 41.094 | 423.0647 | 251.7944 | 216.5953 | 216.7203 |  | 148.1167
FS1 | CSS | DH | 2169 | FS1_CSS:DH_D2169 | D1 | dipole-5deg, D1-2, C1 | 40 x 40 |  |  | 0.3 | 251.6446 | 423.0582 | 41.094 | 41.094 | 423.0517 | 251.4947 | 216.8703 | 217.0203 |  | 148.3917
[10 empty] | 0.091402 | 251.449 | 423.0517 | 41.094 | 41.094 | 423.0517 | 251.4033 | 217.066 | 217.1117 |  | 148.5874
[10 empty] | 0.067196 | 251.3697 | 423.0517 | 41.094 | 41.094 | 423.0517 | 251.3362 | 217.1453 | 217.1789 | 20170214 moved upstream | 148.6667
[10 empty] | 0.091402 | 251.2905 | 423.0517 | 41.094 | 41.094 | 423.0517 | 251.2447 | 217.2246 | 217.2703 |  | 148.746
FS1 | CSS | DH | 2174 | FS1_CSS:DH_D2174 | D1 | dipole-5deg, D1-3, C1 | 40 x 40 |  |  | 0.3 | 251.0949 | 423.0582 | 41.094 | 41.094 | 423.0647 | 250.9451 | 217.4203 | 217.5703 |  | 148.9417
[10 empty] | 0.25 | 250.8206 | 423.0756 | 41.094 | 41.094 | 423.0865 | 250.6961 | 217.6953 | 217.8203 |  | 149.2167
FS1 | CSS | DH | 2180 | FS1_CSS:DH_D2180 | D1 | dipole-5deg, D1-4, C1 | 40 x 40 |  |  | 0.3 | 250.5463 | 423.0931 | 41.094 | 41.094 | 423.0996 | 250.3965 | 217.9703 | 218.1203 |  | 149.4917
[6 empty] | coil-out (assumed) |  |  |  | 0.06 | 250.3665 | 423.0996 | 41.094 | 41.094 | 423.0996 | 250.3365 | 218.1503 | 218.1803 |  | 149.6717
[8 empty] | 2.2834e-05 |  | 0.11723 | 250.2778 | 423.0996 | 41.094 | 41.094 | 423.0996 | 250.2192 | 218.2389 | 218.2975 |  | 149.7603
FS1 | CSS | BCM | 2183 | FS1_CSS:BCM_D2183 | BCM2 | current monitor |  | 250.1693 | 250.1692 | 0.1 | 250.1692 | 423.0996 | 41.094 | 41.094 | 423.0996 | 250.1192 | 218.3475 | 218.3975 |  | 149.8689
[10 empty] | 0.17244 | 250.033 | 423.0996 | 41.094 | 41.094 | 423.0996 | 249.9468 | 218.4838 | 218.57 |  | 150.0051
FS1 | CSS | TMP | 2186 | FS1_CSS:TMP_D2186 | TMP_L | turbo pump 2050 l/s |  | 2.233e-05 | 249.9468 | 0 | 249.9468 | 423.0996 | 41.094 | 41.094 | 423.0996 | 249.9468 | 218.57 | 218.57 |  | 150.0913
[10 empty] | 0.13933 | 249.8771 | 423.0996 | 41.094 | 41.094 | 423.0996 | 249.8075 | 218.6396 | 218.7093 |  | 150.161
[6 empty] | bellow |  |  |  | 0.1 | 249.7575 | 423.0996 | 41.094 | 41.094 | 423.0996 | 249.7075 | 218.7593 | 218.8093 |  | 150.2807
[6 empty] | coil-out |  |  |  | 0.036 | 249.6895 | 423.0996 | 41.094 | 41.094 | 423.0996 | 249.6715 | 218.8273 | 218.8453 |  | 150.3487
FS1 | CSS | DCHV | 2189 | FS1_CSS:DCHV_D2189 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 249.5715 | 423.0996 | 41.094 | 41.094 | 423.0996 | 249.4715 | 218.9453 | 219.0453 |  | 150.4667
[10 empty] | 0.25 | 249.3465 | 423.0996 | 41.094 | 41.094 | 423.0996 | 249.2215 | 219.1703 | 219.2953 |  | 150.6917
FS1 | CSS | QH | 2194 | FS1_CSS:QH_D2194 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 249.0965 | 423.0996 | 41.094 | 41.094 | 423.0996 | 248.9715 | 219.4203 | 219.5453 |  | 150.9417
[10 empty] | 0.0475 | 248.9477 | 423.0996 | 41.094 | 41.094 | 423.0996 | 248.924 | 219.5691 | 219.5928 |  | 151.0904
[10 empty] | 0 | 248.924 | 423.0996 | 41.094 | 41.094 | 423.0996 | 248.924 | 219.5928 | 219.5928 |  | 151.1142
[6 empty] | diagnostic box |  | 2.2834e-05 |  | 0.19515 | 248.8264 | 423.0996 | 41.094 | 41.094 | 423.0996 | 248.7288 | 219.6904 | 219.788 |  | 151.2117
FS1 | CSS | LPM | 2198 | FS1_CSS:LPM_D2198 | LPMx | bunch length monitor (no costed) | future LPM3 | 248.7288 | 248.7288 | 0 | 248.7288 | 423.0996 | 41.094 | 41.094 | 423.0996 | 248.7288 | 219.788 | 219.788 |  | 151.3093
FS1 | CSS | PM | 2198 | FS1_CSS:PM_D2198 | PM3 | profile monitor | defer installation |  |  | 0 | 248.7288 | 423.0996 | 41.094 | 41.094 | 423.0996 | 248.7288 | 219.788 | 219.788 |  | 151.3093
[10 empty] | 0.12238 | 248.6676 | 423.0996 | 41.094 | 41.094 | 423.0996 | 248.6064 | 219.8491 | 219.9103 |  | 151.3705
FS1 | CSS | GV | 2199 | FS1_CSS:GV_D2199 | GV_M | gate valve | 63 | 2.1834e-05 | 248.569 | 0.075 | 248.5689 | 423.0996 | 41.094 | 41.094 | 423.0996 | 248.5314 | 219.9478 | 219.9853 |  | 151.4692
[10 empty] | 0.097816 | 248.4825 | 423.0996 | 41.094 | 41.094 | 423.0996 | 248.4336 | 220.0342 | 220.0832 |  | 151.5556
FS1 | CSS | QV | 2202 | FS1_CSS:QV_D2202 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 248.3086 | 423.0996 | 41.094 | 41.094 | 423.0996 | 248.1836 | 220.2082 | 220.3332 |  | 151.7295
[10 empty] | 0.13552 | 248.1159 | 423.0996 | 41.094 | 41.094 | 423.0996 | 248.0481 | 220.4009 | 220.4687 |  | 151.9223
[10 empty] | 0.15102 | 247.9726 | 423.0996 | 41.094 | 41.094 | 423.0996 | 247.8971 | 220.5442 | 220.6197 |  | 152.0656
FS1 | CSS | TMP | 2206 | FS1_CSS:TMP_D2206 | TMP_L | turbo pump 2050 l/s |  | 2.3334e-05 |  | 0 | 247.8971 | 423.0996 | 41.094 | 41.094 | 423.0996 | 247.8971 | 220.6197 | 220.6197 |  | 152.1411
FS1 | CSS | FAVS | 2206 | FS1_CSS:FAVS_D2206 | FVS | fast valve sensor |  | 247.8971 | 247.8971 | 0 | 247.8971 | 423.0996 | 41.094 | 41.094 | 423.0996 | 247.8971 | 220.6197 | 220.6197 |  | 152.1411
[10 empty] | 0.26346 | 247.7653 | 423.0996 | 41.094 | 41.094 | 423.0996 | 247.6336 | 220.7514 | 220.8832 |  | 152.2728
FS1 | CSS | DCHV | 2210 | FS1_CSS:DCHV_D2210 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 247.5336 | 423.0996 | 41.094 | 41.094 | 423.0996 | 247.4336 | 220.9832 | 221.0832 |  | 152.5045
[8 empty] | 2.2834e-05 |  | 0.027106 | 247.4201 | 423.0996 | 41.094 | 41.094 | 423.0996 | 247.4065 | 221.0967 | 221.1103 |  | 152.6181
FS1 | CSS | BPM | 2212 | FS1_CSS:BPM_D2212 | BPM3 | position monitor | 47.5 | 247.338 | 247.338 | 0.137 | 247.338 | 423.0996 | 41.094 | 41.094 | 423.0996 | 247.2695 | 221.1788 | 221.2473 |  | 152.7001
[10 empty] | 0.085894 | 247.2266 | 423.0996 | 41.094 | 41.094 | 423.0996 | 247.1836 | 221.2902 | 221.3332 |  | 152.8116
FS1 | CSS | QH | 2215 | FS1_CSS:QH_D2215 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 247.0586 | 423.0996 | 41.094 | 41.094 | 423.0996 | 246.9336 | 221.4582 | 221.5832 |  | 152.9795
[10 empty] | 0.25 | 246.8086 | 423.0996 | 41.094 | 41.094 | 423.0996 | 246.6836 | 221.7082 | 221.8332 |  | 153.2295
FS1 | CSS | QV | 2220 | FS1_CSS:QV_D2220 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 246.5586 | 423.0996 | 41.094 | 41.094 | 423.0996 | 246.4336 | 221.9582 | 222.0832 |  | 153.4795
[10 empty] | 0 | 246.4336 | 423.0996 | 41.094 | 41.094 | 423.0996 | 246.4336 | 222.0832 | 222.0832 |  | 153.6045
[8 empty] | 2.2834e-05 |  | 0.074011 | 246.3966 | 423.0996 | 41.094 | 41.094 | 423.0996 | 246.3596 | 222.1202 | 222.1572 |  | 153.6415
FS1 | CSS | GV | 2222 | FS1_CSS:GV_D2222 | GV_M | gate valve | 63 | 246.3221 | 246.3221 | 0.075 | 246.3221 | 423.0996 | 41.094 | 41.094 | 423.0996 | 246.2846 | 222.1947 | 222.2322 |  | 153.716
[9 empty] | 0.045989 | 0 | 246.2846 | 423.0996 | 41.094 | 41.094 | 423.0996 | 246.2846 | 222.2322 | 222.2322 |  | 153.7535
[6 empty] | bellow |  | 0.054792 |  | 0.094699 | 246.2373 | 423.0996 | 41.094 | 41.094 | 423.0996 | 246.1899 | 222.2795 | 222.3269 | bellow compressed by 0.008803 m, 4/24/2017 | 153.8009
FS1 | CSS | BPM | 2223 | FS1_CSS:BPM_D2223 | BPM3 | position monitor |  | 2.1834e-05 | 246.1899 | 0 | 246.1899 | 423.0996 | 41.094 | 41.094 | 423.0996 | 246.1899 | 222.3269 | 222.3269 |  | 153.8482
[6 empty] | diagnostic box |  |  |  | 0.14528 | 246.1173 | 423.0996 | 41.094 | 41.094 | 423.0996 | 246.0446 | 222.3995 | 222.4721 |  | 153.9209
FS1 | CSS | PM | 2225 | FS1_CSS:PM_D2225 | PM6 | profile monitor (small) |  |  |  | 0 | 246.0446 | 423.0996 | 41.094 | 41.094 | 423.0996 | 246.0446 | 222.4721 | 222.4721 |  | 153.9935
[6 empty] | diagnostic box |  |  |  | 0 | 246.0446 | 423.0996 | 41.094 | 41.094 | 423.0996 | 246.0446 | 222.4721 | 222.4721 |  | 153.9935
FS1 | CSS | IP | 2225 | FS1_CSS:IP_D2225 | IP_S | ion pump, 75 l/s |  | 2.1834e-05 | 246.0446 | 0 | 246.0446 | 423.0996 | 41.094 | 41.094 | 423.0996 | 246.0446 | 222.4721 | 222.4721 |  | 153.9935
[10 empty] | 0 | 246.0446 | 423.0996 | 41.094 | 41.094 | 423.0996 | 246.0446 | 222.4721 | 222.4721 |  | 153.9935
[6 empty] | bellow |  |  |  | 0.13122 | 245.979 | 423.0996 | 41.094 | 41.094 | 423.0996 | 245.9134 | 222.5378 | 222.6034 |  | 154.0591
[10 empty] | 0.054792 | 245.886 | 423.0996 | 41.094 | 41.094 | 423.0996 | 245.8586 | 222.6308 | 222.6582 |  | 154.1521
FS1 | STRL | GV | 2227 | FS1_STRL:GV_D2227 |  | gate valve |  |  |  | 0.072 | 245.8226 | 423.0996 | 41.094 | 41.094 | 423.0996 | 245.7866 | 222.6942 | 222.7302 | stripper module length: 1.975607 | 154.2155
[6 empty] | stripper module |  |  |  | 0.528 | 245.5226 | 423.0996 | 41.094 | 41.094 | 423.0996 | 245.2586 | 222.9942 | 223.2582 |  | 154.5155
FS1 | STRL | CP | 2233 | FS1_STRL:CP_D2233 | CP_L | cryo pump, 1500 l/s |  |  |  | 0 | 245.2586 | 423.0996 | 41.094 | 41.094 | 423.0996 | 245.2586 | 223.2582 | 223.2582 |  | 154.7795
[10 empty] | 0.4 | 245.0586 | 423.0996 | 41.094 | 41.094 | 423.0996 | 244.8586 | 223.4582 | 223.6582 |  | 154.9795
FS1 | STRL | STRIP | 2237 | FS1_STRL:STRIP_D2237 |  | lithium film stripper |  |  |  | 0 | 244.8586 | 423.0996 | 41.094 | 41.094 | 423.0996 | 244.8586 | 223.6582 | 223.6582 |  | 155.1795
[10 empty] | 0.4 | 244.6586 | 423.0996 | 41.094 | 41.094 | 423.0996 | 244.4586 | 223.8582 | 224.0582 |  | 155.3795
FS1 | STRL | CP | 2241 | FS1_STRL:CP_D2241 | CP_L | cryo pump, 1500 l/s |  |  |  | 0 | 244.4586 | 423.0996 | 41.094 | 41.094 | 423.0996 | 244.4586 | 224.0582 | 224.0582 |  | 155.5795
[6 empty] | stripper module |  |  |  | 0.528 | 244.1946 | 423.0996 | 41.094 | 41.094 | 423.0996 | 243.9306 | 224.3222 | 224.5862 |  | 155.8435
FS1 | STRL | GV | 2246 | FS1_STRL:GV_D2246 |  | gate valve |  |  |  | 0.072 | 243.8946 | 423.0996 | 41.094 | 41.094 | 423.0996 | 243.8586 | 224.6222 | 224.6582 |  | 156.1435
[10 empty] | 0.008836 | 243.8542 | 423.0996 | 41.094 | 41.094 | 423.0996 | 243.8498 | 224.6626 | 224.667 | stripper module increased by 0.008836 m, 4/24/2017 | 156.1839
[6 empty] | bellow |  | -0.008836 |  | 0.1035 | 243.798 | 423.0996 | 41.094 | 41.094 | 423.0996 | 243.7463 | 224.7187 | 224.7705 |  | 156.2401
FS1 | CSS | BPM | 2248 | FS1_CSS:BPM_D2248 | BPM3 | position monitor |  | 2.1834e-05 | 243.7463 | 0 | 243.7463 | 423.0996 | 41.094 | 41.094 | 423.0996 | 243.7463 | 224.7705 | 224.7705 |  | 156.2919
[6 empty] | diagnostic box |  |  |  | 0.14528 | 243.6736 | 423.0996 | 41.094 | 41.094 | 423.0996 | 243.601 | 224.8431 | 224.9158 |  | 156.3645
FS1 | CSS | PM | 2249 | FS1_CSS:PM_D2249 | PM6 | profile monitor | defer installation |  |  | 0 | 243.601 | 423.0996 | 41.094 | 41.094 | 423.0996 | 243.601 | 224.9158 | 224.9158 |  | 156.4371
[6 empty] | diagnostic box |  |  |  | 0 | 243.601 | 423.0996 | 41.094 | 41.094 | 423.0996 | 243.601 | 224.9158 | 224.9158 |  | 156.4371
FS1 | CSS | IP | 2249 | FS1_CSS:IP_D2249 | IP_S | ion pump, 75 l/s |  | 2.1834e-05 | 243.601 | 0 | 243.601 | 423.0996 | 41.094 | 41.094 | 423.0996 | 243.601 | 224.9158 | 224.9158 |  | 156.4371
[10 empty] | 0 | 243.601 | 423.0996 | 41.094 | 41.094 | 423.0996 | 243.601 | 224.9158 | 224.9158 |  | 156.4371
[6 empty] | bellow |  |  |  | 0.12238 | 243.5398 | 423.0996 | 41.094 | 41.094 | 423.0996 | 243.4786 | 224.977 | 225.0382 | bellow compressed by 0.008836 m, 4/24/2017 | 156.4983
FS1 | CSS | GV | 2251 | FS1_CSS:GV_D2251 | GV_M | gate valve | 63 |  |  | 0.075 | 243.4411 | 423.0996 | 41.094 | 41.094 | 423.0996 | 243.4036 | 225.0757 | 225.1132 |  | 156.597
[10 empty] | 0.12 | 243.3436 | 423.0996 | 41.094 | 41.094 | 423.0996 | 243.2836 | 225.1732 | 225.2332 |  | 156.6945
FS1 | CSS | QV | 2254 | FS1_CSS:QV_D2254 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 243.1586 | 423.0996 | 41.094 | 41.094 | 423.0996 | 243.0336 | 225.3582 | 225.4832 |  | 156.8795
[8 empty] | 2.2834e-05 |  | 0.092506 | 242.9874 | 423.0996 | 41.094 | 41.094 | 423.0996 | 242.9411 | 225.5294 | 225.5757 |  | 157.0508
FS1 | CSS | DCHV | 2257 | FS1_CSS:DCHV_D2257 | C2 | correctors-X,Y | 50 | 242.8411 | 242.8411 | 0.2 | 242.8411 | 423.0996 | 41.094 | 41.094 | 423.0996 | 242.7411 | 225.6757 | 225.7757 |  | 157.197
[10 empty] | 0 | 242.7411 | 423.0996 | 41.094 | 41.094 | 423.0996 | 242.7411 | 225.7757 | 225.7757 |  | 157.297
[10 empty] | 0.10749 | 242.6874 | 423.0996 | 41.094 | 41.094 | 423.0996 | 242.6336 | 225.8294 | 225.8832 |  | 157.3508
[10 empty] | 0 | 242.6336 | 423.0996 | 41.094 | 41.094 | 423.0996 | 242.6336 | 225.8832 | 225.8832 |  | 157.4045
FS1 | CSS | QH | 2260 | FS1_CSS:QH_D2260 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 242.5086 | 423.0996 | 41.094 | 41.094 | 423.0996 | 242.3836 | 226.0082 | 226.1332 |  | 157.5295
[6 empty] | coil out + leads |  |  |  | 0.098 | 242.3346 | 423.0996 | 41.094 | 41.094 | 423.0996 | 242.2856 | 226.1822 | 226.2312 |  | 157.7035
[8 empty] | 2.2834e-05 |  | 0.084251 | 242.2435 | 423.0996 | 41.094 | 41.094 | 423.0996 | 242.2014 | 226.2733 | 226.3154 |  | 157.7946
FS1 | CSS | LPM | 2263 | FS1_CSS:LPM_D2263 | LPM3 | bunch length monitor |  | 242.2014 | 242.2014 | 0 | 242.2014 | 423.0996 | 41.094 | 41.094 | 423.0996 | 242.2014 | 226.3154 | 226.3154 |  | 157.8368
[8 empty] | 2.2834e-05 |  | 0.08375 | 242.1595 | 423.0996 | 41.094 | 41.094 | 423.0996 | 242.1176 | 226.3573 | 226.3992 |  | 157.8787
FS1 | CSS | BCM | 2264 | FS1_CSS:BCM_D2264 | BCM4 | current monitor |  | 242.0676 | 242.0676 | 0.1 | 242.0676 | 423.0996 | 41.094 | 41.094 | 423.0996 | 242.0176 | 226.4492 | 226.4992 |  | 157.9705
[10 empty] | 0.20355 | 241.9158 | 423.0996 | 41.094 | 41.094 | 423.0996 | 241.8141 | 226.6009 | 226.7027 |  | 158.1223
FS1 | CSS | TMP | 2267 | FS1_CSS:TMP_D2267 | TMP_L | turbo pump 2050 l/s |  | 2.2834e-05 |  | 0 | 241.8141 | 423.0996 | 41.094 | 41.094 | 423.0996 | 241.8141 | 226.7027 | 226.7027 |  | 158.2241
FS1 | CSS | FAVS | 2267 | FS1_CSS:FAVS_D2267 | FVS | fast valve sensor |  | 241.8141 | 241.8141 | 0 | 241.8141 | 423.0996 | 41.094 | 41.094 | 423.0996 | 241.8141 | 226.7027 | 226.7027 |  | 158.2241
[7 empty] | 0.69455 |  |  | 0 | 241.8141 | 423.0996 | 41.094 | 41.094 | 423.0996 | 241.8141 | 226.7027 | 226.7027 |  | 158.2241
[6 empty] | collimation flange |  |  |  | 0.28245 | 241.6728 | 423.0996 | 41.094 | 41.094 | 423.0996 | 241.5316 | 226.8439 | 226.9852 |  | 158.3653
[6 empty] | coil out + leads |  |  |  | 0.098 | 241.4826 | 423.0996 | 41.094 | 41.094 | 423.0996 | 241.4336 | 227.0342 | 227.0832 |  | 158.5555
FS1 | CSS | QV | 2272 | FS1_CSS:QV_D2272 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 241.3086 | 423.0996 | 41.094 | 41.094 | 423.0996 | 241.1836 | 227.2082 | 227.3332 |  | 158.7295
[8 empty] | 2.2834e-05 | -0.050806 | 0.13866 | 241.1143 | 423.0996 | 41.094 | 41.094 | 423.0996 | 241.045 | 227.4025 | 227.4718 |  | 158.9239
FS1 | CSS | DCHV | 2276 | FS1_CSS:DCHV_D2276 | C2 | correctors-X,Y | 50 | 240.945 | 240.945 | 0.2 | 240.945 | 423.0996 | 41.094 | 41.094 | 423.0996 | 240.845 | 227.5718 | 227.6718 |  | 159.0932
[8 empty] | 2.1834e-05 | -0.052701 | 0.13939 | 240.7753 | 423.0996 | 41.094 | 41.094 | 423.0996 | 240.7056 | 227.7415 | 227.8112 |  | 159.2629
FS1 | CSS | BPM | 2278 | FS1_CSS:BPM_D2278 | BPM3 | position monitor | 47.5 | 240.7056 | 240.7056 | 0 | 240.7056 | 423.0996 | 41.094 | 41.094 | 423.0996 | 240.7056 | 227.8112 | 227.8112 | 0.137 | 159.3326
[8 empty] | 0.068501 | -0.008701 | 0.0598 | 240.6757 | 423.0996 | 41.094 | 41.094 | 423.0996 | 240.6458 | 227.8411 | 227.871 |  | 159.3625
FS1 | CSS | QH | 2280 | FS1_CSS:QH_D2280 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 240.5208 | 423.0996 | 41.094 | 41.094 | 423.0996 | 240.3958 | 227.996 | 228.121 |  | 159.5174
[10 empty] | 0.28 | 240.2558 | 423.0996 | 41.094 | 41.094 | 423.0996 | 240.1158 | 228.261 | 228.401 |  | 159.7824
[8 empty] | 2.2834e-05 |  | 0.082672 | 240.0744 | 423.0996 | 41.094 | 41.094 | 423.0996 | 240.0331 | 228.4423 | 228.4837 |  | 159.9637
FS1 | CSS | TMP | 2285 | FS1_CSS:TMP_D2285 | TMP_L | turbo pump 2050 l/s |  | 240.0331 | 240.0331 | 0 | 240.0331 | 423.0996 | 41.094 | 41.094 | 423.0996 | 240.0331 | 228.4837 | 228.4837 |  | 160.005
[10 empty] | 0.2199 | 239.9231 | 423.0996 | 41.094 | 41.094 | 423.0996 | 239.8132 | 228.5936 | 228.7036 |  | 160.115
[10 empty] | 0 | 239.8132 | 423.0996 | 41.094 | 41.094 | 423.0996 | 239.8132 | 228.7036 | 228.7036 |  | 160.2249
[8 empty] | -0.017571 |  | 0 | 239.8132 | 423.0996 | 41.094 | 41.094 | 423.0996 | 239.8132 | 228.7036 | 228.7036 |  | 160.2249
[6 empty] | collimation flange |  | 2.1834e-05 | 239.8132 | 0 | 239.8132 | 423.0996 | 41.094 | 41.094 | 423.0996 | 239.8132 | 228.7036 | 228.7036 |  | 160.2249
[10 empty] | 0.18302 | 239.7217 | 423.0996 | 41.094 | 41.094 | 423.0996 | 239.6302 | 228.7951 | 228.8866 |  | 160.3165
[8 empty] | 2.3521e-05 | 0.2294 | 0 | 239.6302 | 423.0996 | 41.094 | 41.094 | 423.0996 | 239.6302 | 228.8866 | 228.8866 | 20170215 D1-1 & D1-2 moved 0.229406 m upstream | 160.408
FS1 | CSS | DH | 2290 | FS1_CSS:DH_D2290 | D1 | dipole-5deg, D1-1, C2 | 40 x 40 | 239.4804 | 239.4804 | 0.3 | 239.4804 | 423.0931 | 41.094 | 41.094 | 423.0865 | 239.3305 | 229.0366 | 229.1866 |  | 160.558
[8 empty] | 2.3568e-05 |  | 0.25 | 239.206 | 423.0756 | 41.094 | 41.094 | 423.0647 | 239.0815 | 229.3116 | 229.4366 |  | 160.833
FS1 | CSS | DH | 2296 | FS1_CSS:DH_D2296 | D1 | dipole-5deg, D1-1, C2 | 40 x 40 | 238.9317 | 238.9317 | 0.3 | 238.9317 | 423.0582 | 41.094 | 41.094 | 423.0517 | 238.7819 | 229.5866 | 229.7366 |  | 161.108
[10 empty] | 0.082804 | 238.7405 | 423.0517 | 41.094 | 41.094 | 423.0517 | 238.6991 | 229.778 | 229.8194 |  | 161.2994
[6 empty] | 2nd chicane not installed initially, replaced by a pipe |  |  |  | 0.1 | 238.6491 | 423.0517 | 41.094 | 41.094 | 423.0517 | 238.5991 | 229.8694 | 229.9194 |  | 161.3908
[10 empty] | 0.067196 | 238.5655 | 423.0517 | 41.094 | 41.094 | 423.0517 | 238.5319 | 229.953 | 229.9866 |  | 161.4744
[8 empty] | 2.3942e-05 |  | 0.1 | 238.4819 | 423.0517 | 41.094 | 41.094 | 423.0517 | 238.4319 | 230.0366 | 230.0866 |  | 161.558
FS1 | CSS | DH | 2302 | FS1_CSS:DH_D2302 | D1 | dipole-5deg, D1-1, C2 | 40 x 40 | 238.2821 | 238.2821 | 0.3 | 238.2821 | 423.0582 | 41.094 | 41.094 | 423.0647 | 238.1323 | 230.2366 | 230.3866 |  | 161.758
[8 empty] | 2.3989e-05 |  | 0.25 | 238.0077 | 423.0756 | 41.094 | 41.094 | 423.0865 | 237.8832 | 230.5116 | 230.6366 | 20170215 added 0.2294 m | 162.033
FS1 | CSS | DH | 2308 | FS1_CSS:DH_D2308 | D1 | dipole-5deg, D1-1, C2 | 40 x 40 | 237.7334 | 237.7334 | 0.3 | 237.7334 | 423.0931 | 41.094 | 41.094 | 423.0996 | 237.5836 | 230.7866 | 230.9366 |  | 162.308
[7 empty] | 237.3497 | 2.3676e-05 | -0.19191 | 0.19645 | 237.4854 | 423.0996 | 41.094 | 41.094 | 423.0996 | 237.3871 | 231.0348 | 231.133 |  | 162.5562
FS1 | CSS | FAV | 2312 | FS1_CSS:FAV_D2312 | FV | fast valve | 40 | 237.3536 | 237.3535 | 0.067196 | 237.3535 | 423.0996 | 41.094 | 41.094 | 423.0996 | 237.32 | 231.1666 | 231.2002 |  | 162.688
[7 empty] | FV still required ? |  |  | 0.003898 | 237.318 | 423.0996 | 41.094 | 41.094 | 423.0996 | 237.3161 | 231.2022 | 231.2041 |  | 162.7236
[6 empty] | bellow | 40 | 2.3676e-05 | 0.0070924 | 0.1035 | 237.2643 | 423.0996 | 41.094 | 41.094 | 423.0996 | 237.2126 | 231.2559 | 231.3076 |  | 162.7773
FS1 | CSS | BPM | 2313 | FS1_CSS:BPM_D2313 | BPM2 | position monitor | 237.3497 | 237.2126 | 237.2126 | 0 | 237.2126 | 423.0996 | 41.094 | 41.094 | 423.0996 | 237.2126 | 231.3076 | 231.3076 |  | 162.829
[6 empty] | diagnostic box |  |  |  | 0.14528 | 237.1399 | 423.0996 | 41.094 | 41.094 | 423.0996 | 237.0673 | 231.3803 | 231.4529 | 0.39757 | 162.9017
FS1 | CSS | PM | 2315 | FS1_CSS:PM_D2315 | PMx | profile monitor (not costed) | future PM |  |  | 0 | 237.0673 | 423.0996 | 41.094 | 41.094 | 423.0996 | 237.0673 | 231.4529 | 231.4529 | 2.2 | 162.9743
[6 empty] | diagnostic box |  | 2.3676e-05 |  | 0 | 237.0673 | 423.0996 | 41.094 | 41.094 | 423.0996 | 237.0673 | 231.4529 | 231.4529 | 1.4 | 162.9743
FS1 | CSS | IP | 2315 | FS1_CSS:IP_D2315 | IP_S | pump, 75 l/s |  | 237.0673 | 237.0673 | 0 | 237.0673 | 423.0996 | 41.094 | 41.094 | 423.0996 | 237.0673 | 231.4529 | 231.4529 | 0.253 | 162.9743
FS1 | CSS | FAVS | 2315 | FS1_CSS:FAVS_D2315 | FVS | fast valve sensor |  |  |  | 0 | 237.0673 | 423.0996 | 41.094 | 41.094 | 423.0996 | 237.0673 | 231.4529 | 231.4529 |  | 162.9743
[6 empty] | bellow | 40 |  |  | 0.13122 | 237.0017 | 423.0996 | 41.094 | 41.094 | 423.0996 | 236.9361 | 231.5185 | 231.5841 |  | 163.0399
[10 empty] | 0.017429 | 236.9273 | 423.0996 | 41.094 | 41.094 | 423.0996 | 236.9186 | 231.5929 | 231.6016 | 20170216 reduced 0.022604 m | 163.1142
[6 empty] | collimation flange |  | 2.2676e-05 | 236.9186 | 0 | 236.9186 | 423.0996 | 41.094 | 41.094 | 423.0996 | 236.9186 | 231.6016 | 231.6016 |  | 163.1229
[10 empty] | 0.096757 | 236.8702 | 423.0996 | 41.094 | 41.094 | 423.0996 | 236.8219 | 231.6499 | 231.6983 |  | 163.1713
[9 empty] | enough or not ? | 0.1 | 236.7719 | 423.0996 | 41.094 | 41.094 | 423.0996 | 236.7219 | 231.7483 | 231.7983 |  | 163.2697
FS1 | MGB01 | GV | 2318 | FS1_MGB01:GV_D2318 |  | gate valve | 40 | -0.34819 | 236.345 | 0.057404 | 236.6932 | 423.0996 | 41.094 | 41.094 | 423.0996 | 236.6645 | 231.827 | 231.8557 |  | 163.3484
[10 empty] | 0 | 236.6645 | 423.0996 | 41.094 | 41.094 | 423.0996 | 236.6645 | 231.8557 | 231.8557 |  | 163.3771
FS1 | MGB01 | CAV | 2320 | FS1_MGB01:CAV_D2320 |  | b19 resonator 1 1/4 | 36 |  |  | 0.31833 | 236.5053 | 423.0996 | 41.094 | 41.094 | 423.0996 | 236.3461 | 232.0149 | 232.1741 |  | 163.5363
FS1 | MGB01 | IP | 2322 | FS1_MGB01:IP_D2322 | IP_L | ion pump, ? l/s |  |  |  | 0 | 236.3461 | 423.0996 | 41.094 | 41.094 | 423.0996 | 236.3461 | 232.1741 | 232.1741 |  | 163.6954
FS1 | MGB01 | CAV | 2323 | FS1_MGB01:CAV_D2323 |  | b19 resonator 1 2/4 | 36 |  |  | 0.31833 | 236.187 | 423.0996 | 41.094 | 41.094 | 423.0996 | 236.0278 | 232.3332 | 232.4924 |  | 163.8546
FS1 | MGB01 | CAV | 2325 | FS1_MGB01:CAV_D2325 |  | rf coupler |  | -0.34819 | 235.6796 | 0 | 236.0278 | 423.0996 | 41.094 | 41.094 | 423.0996 | 236.0278 | 232.4924 | 232.4924 |  | 164.0137
[6 empty] | This buncher will not be installed initially by replacing a pipe |  |  |  | 0 | 236.0278 | 423.0996 | 41.094 | 41.094 | 423.0996 | 236.0278 | 232.4924 | 232.4924 |  | 164.0137
FS1 | MGB01 | CAV | 2327 | FS1_MGB01:CAV_D2327 |  | b19 resonator 1 3/4 | 36 |  |  | 0.31833 | 235.8686 | 423.0996 | 41.094 | 41.094 | 423.0996 | 235.7095 | 232.6515 | 232.8107 |  | 164.1729
FS1 | MGB01 | TMP | 2328 | FS1_MGB01:TMP_D2330 | TMP_L | turbo pump, ? l/s |  |  |  | 0 | 235.7095 | 423.0996 | 41.094 | 41.094 | 423.0996 | 235.7095 | 232.8107 | 232.8107 |  | 164.3321
FS1 | MGB01 | CAV | 2330 | FS1_MGB01:CAV_D2330 |  | b19 resonator 1 4/4 | 36 |  |  | 0.31833 | 235.5503 | 423.0996 | 41.094 | 41.094 | 423.0996 | 235.3912 | 232.9699 | 233.129 |  | 164.4912
[10 empty] | 0 | 235.3912 | 423.0996 | 41.094 | 41.094 | 423.0996 | 235.3912 | 233.129 | 233.129 |  | 164.6504
FS1 | MGB01 | GV | 2332 | FS1_MGB01:GV_D2332 |  | gate valve | 40 | -0.34819 | 235.0143 | 0.057404 | 235.3625 | 423.0996 | 41.094 | 41.094 | 423.0996 | 235.3338 | 233.1577 | 233.1864 |  | 164.6791
[10 empty] | 0.91761 | 234.875 | 423.0996 | 41.094 | 41.094 | 423.0996 | 234.4162 | 233.6452 | 234.104 | diag bx moved down by 0.12119 m, 4/24/2017 | 165.1666
[6 empty] | bellow | 40 |  | 0.13122 | 0.152 | 234.3402 | 423.0996 | 41.094 | 41.094 | 423.0996 | 234.2642 | 234.18 | 234.256 |  | 165.7014
[6 empty] | vacuum box |  |  |  | 0 | 234.2642 | 423.0996 | 41.094 | 41.094 | 423.0996 | 234.2642 | 234.256 | 234.256 |  | 165.7774
FS1 | CSS | FAVS | 2343 | FS1_CSS:FAVS_D2343 | FVS | fast valve sensor |  |  |  | 0 | 234.2642 | 423.0996 | 41.094 | 41.094 | 423.0996 | 234.2642 | 234.256 | 234.256 |  | 165.7774
[6 empty] | vacuum box |  |  |  | 0 | 234.2642 | 423.0996 | 41.094 | 41.094 | 423.0996 | 234.2642 | 234.256 | 234.256 |  | 165.7774
FS1 | CSS | IP | 2343 | FS1_CSS:IP_D2343 | IP_S | ion pump, 75 l/s |  | 2.2676e-05 | 234.2642 | 0 | 234.2642 | 423.0996 | 41.094 | 41.094 | 423.0996 | 234.2642 | 234.256 | 234.256 |  | 165.7774
[6 empty] | bellow | 40 |  |  | 0.132 | 234.1982 | 423.0996 | 41.094 | 41.094 | 423.0996 | 234.1322 | 234.322 | 234.388 |  | 165.8434
[10 empty] | 0.001052 | 234.1316 | 423.0996 | 41.094 | 41.094 | 423.0996 | 234.1311 | 234.3886 | 234.3891 |  | 165.9099
[8 empty] | -0.038013 |  | 0.015594 | 234.1233 | 423.0996 | 41.094 | 41.094 | 423.0996 | 234.1155 | 234.3969 | 234.4047 |  | 165.9183
[6 empty] | collimation flange |  | 2.2676e-05 | 234.1155 | 0 | 234.1155 | 423.0996 | 41.094 | 41.094 | 423.0996 | 234.1155 | 234.4047 | 234.4047 |  | 165.9261
[10 empty] | 0.30958 | 233.9607 | 423.0996 | 41.094 | 41.094 | 423.0996 | 233.8059 | 234.5595 | 234.7143 |  | 166.0808
[10 empty] | 0 | 233.8059 | 423.0996 | 41.094 | 41.094 | 423.0996 | 233.8059 | 234.7143 | 234.7143 |  | 166.2356
[10 empty] | 0 | 233.8059 | 423.0996 | 41.094 | 41.094 | 423.0996 | 233.8059 | 234.7143 | 234.7143 |  | 166.2356
[7 empty] | FV still required ? |  |  | 0.069596 | 233.7711 | 423.0996 | 41.094 | 41.094 | 423.0996 | 233.7363 | 234.7491 | 234.7839 |  | 166.2704
FS1 | CSS | FAV | 2348 | FS1_CSS:FAV_D2348 | FV | fast valve | 40 | 0.067196 |  | 0.067196 | 233.7027 | 423.0996 | 41.094 | 41.094 | 423.0996 | 233.6691 | 234.8175 | 234.8511 |  | 166.3388
[10 empty] | 0.14724 | 233.5955 | 423.0996 | 41.094 | 41.094 | 423.0996 | 233.5219 | 234.9247 | 234.9983 |  | 166.4461
FS1 | CSS | DCHV | 2351 | FS1_CSS:DCHV_D2351 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 233.4219 | 423.0996 | 41.094 | 41.094 | 423.0996 | 233.3219 | 235.0983 | 235.1983 |  | 166.6197
[10 empty] | 0.1067 | 233.2685 | 423.0996 | 41.094 | 41.094 | 423.0996 | 233.2152 | 235.2517 | 235.305 |  | 166.773
FS1 | CSS | BCM | 2353 | FS1_CSS:BCM_D2353 | BCM4 | current monitor |  | 2.2676e-05 | 233.2152 | 0 | 233.2152 | 423.0996 | 41.094 | 41.094 | 423.0996 | 233.2152 | 235.305 | 235.305 |  | 166.8264
[10 empty] | 0.1433 | 233.1435 | 423.0996 | 41.094 | 41.094 | 423.0996 | 233.0719 | 235.3767 | 235.4483 |  | 166.898
FS1 | CSS | QH | 2356 | FS1_CSS:QH_D2356 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 232.9469 | 423.0996 | 41.094 | 41.094 | 423.0996 | 232.8219 | 235.5733 | 235.6983 |  | 167.0947
[6 empty] | coil out |  |  |  | 0.047 | 232.7984 | 423.0996 | 41.094 | 41.094 | 423.0996 | 232.7749 | 235.7218 | 235.7453 |  | 167.2432
[10 empty] | 0.19117 | 232.6793 | 423.0996 | 41.094 | 41.094 | 423.0996 | 232.5837 | 235.8409 | 235.9365 | 20170216 space adjusted | 167.3623
[6 empty] | collimation flange |  | 2.2676e-05 | 232.5837 | 0 | 232.5837 | 423.0996 | 41.094 | 41.094 | 423.0996 | 232.5837 | 235.9365 | 235.9365 |  | 167.4578
[10 empty] | 0.038832 | 232.5643 | 423.0996 | 41.094 | 41.094 | 423.0996 | 232.5449 | 235.9559 | 235.9753 |  | 167.4773
[6 empty] | coil out + leads |  |  |  | 0.098 | 232.4959 | 423.0996 | 41.094 | 41.094 | 423.0996 | 232.4469 | 236.0243 | 236.0733 |  | 167.5457
FS1 | CSS | QH | 2362 | FS1_CSS:QH_D2362 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 232.3219 | 423.0996 | 41.094 | 41.094 | 423.0996 | 232.1969 | 236.1983 | 236.3233 |  | 167.7197
[10 empty] | 0 | 232.1969 | 423.0996 | 41.094 | 41.094 | 423.0996 | 232.1969 | 236.3233 | 236.3233 |  | 167.8447
[10 empty] | 0 | 232.1969 | 423.0996 | 41.094 | 41.094 | 423.0996 | 232.1969 | 236.3233 | 236.3233 |  | 167.8447
[10 empty] | 0.3 | 232.0469 | 423.0996 | 41.094 | 41.094 | 423.0996 | 231.8969 | 236.4733 | 236.6233 |  | 167.9947
FS1 | CSS | DCHV | 2367 | FS1_CSS:DCHV_D2367 | C2 | correctors-X,Y | 50 |  | 0.15635 | 0.2 | 231.7969 | 423.0996 | 41.094 | 41.094 | 423.0996 | 231.6969 | 236.7233 | 236.8233 |  | 168.2447
[8 empty] | 2.3676e-05 | 0.08115 | -0.01215 | 231.703 | 423.0996 | 41.094 | 41.094 | 423.0996 | 231.709 | 236.8172 | 236.8112 |  | 168.3386
FS1 | CSS | BPM | 2369 | FS1_CSS:BPM_D2369 | BPM3 | position monitor | 47.5 | 231.6406 | 231.6405 | 0.137 | 231.6405 | 423.0996 | 41.094 | 41.094 | 423.0996 | 231.572 | 236.8797 | 236.9482 |  | 168.401
[10 empty] | 0.12515 | 231.5095 | 423.0996 | 41.094 | 41.094 | 423.0996 | 231.4469 | 237.0107 | 237.0733 | 20170216 space adjusted | 168.5321
FS1 | CSS | QV | 2372 | FS1_CSS:QV_D2372 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 231.3219 | 423.0996 | 41.094 | 41.094 | 423.0996 | 231.1969 | 237.1983 | 237.3233 |  | 168.7197
[10 empty] | 0.25 | 231.0719 | 423.0996 | 41.094 | 41.094 | 423.0996 | 230.9469 | 237.4483 | 237.5733 |  | 168.9697
FS1 | CSS | QH | 2377 | FS1_CSS:QH_D2377 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 230.8219 | 423.0996 | 41.094 | 41.094 | 423.0996 | 230.6969 | 237.6983 | 237.8233 |  | 169.2197
[10 empty] | 0.145 | 230.6244 | 423.0996 | 41.094 | 41.094 | 423.0996 | 230.5519 | 237.8958 | 237.9683 |  | 169.4172
FS1 | CSS | DCHV | 2381 | FS1_CSS:DCHV_D2381 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 230.4519 | 423.0996 | 41.094 | 41.094 | 423.0996 | 230.3519 | 238.0683 | 238.1683 |  | 169.5897
[10 empty] | 0.059488 | 230.3221 | 423.0996 | 41.094 | 41.094 | 423.0996 | 230.2924 | 238.198 | 238.2278 | 20170216 space adjusted | 169.7194
[6 empty] | bellow | 40 | 2.3676e-05 | 0.033167 | 0.1035 | 230.2407 | 423.0996 | 41.094 | 41.094 | 423.0996 | 230.1889 | 238.2795 | 238.3313 |  | 169.8009
FS1 | CSS | BPM | 2383 | FS1_CSS:BPM_D2383 | BPM3 | position monitor |  | 230.1889 | 230.1889 | 0 | 230.1889 | 423.0996 | 41.094 | 41.094 | 423.0996 | 230.1889 | 238.3313 | 238.3313 |  | 169.8527
[6 empty] | diagnostic box |  |  |  | 0.14528 | 230.1163 | 423.0996 | 41.094 | 41.094 | 423.0996 | 230.0436 | 238.4039 | 238.4766 |  | 169.9253
FS1 | CSS | PM | 2385 | FS1_CSS:PM_D2385 | PM3 | profile monitor |  |  |  | 0 | 230.0436 | 423.0996 | 41.094 | 41.094 | 423.0996 | 230.0436 | 238.4766 | 238.4766 |  | 169.9979
[6 empty] | diagnostic box |  | 2.4676e-05 |  | 0 | 230.0436 | 423.0996 | 41.094 | 41.094 | 423.0996 | 230.0436 | 238.4766 | 238.4766 |  | 169.9979
FS1 | CSS | FAVS | 2385 | FS1_CSS:FAVS_D2385 | FVS | fast valve sensor |  | 230.0436 | 230.0436 | 0 | 230.0436 | 423.0996 | 41.094 | 41.094 | 423.0996 | 230.0436 | 238.4766 | 238.4766 |  | 169.9979
FS1 | CSS | IP | 2385 | FS1_CSS:IP_D2384 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 230.0436 | 423.0996 | 41.094 | 41.094 | 423.0996 | 230.0436 | 238.4766 | 238.4766 |  | 169.9979
FS1 | CSS | IP | 2385 | FS1_CSS:IP_D2385 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 230.0436 | 423.0996 | 41.094 | 41.094 | 423.0996 | 230.0436 | 238.4766 | 238.4766 |  | 169.9979
[6 empty] | bellow | 40 | 2.3676e-05 | 0.008831 | 0.12238 | 229.9824 | 423.0996 | 41.094 | 41.094 | 423.0996 | 229.9212 | 238.5378 | 238.599 | 20170216 space adjusted | 170.0591
FS1 | CSS | GV | 2386 | FS1_CSS:GV_D2386 | GV_M | gate valve | 63 | 229.8838 | 229.8837 | 0.075 | 229.8837 | 423.0996 | 41.094 | 41.094 | 423.0996 | 229.8462 | 238.6365 | 238.674 |  | 170.1578 | [19 empty] | straight line
First beam dump line starts in "V" column | [9 empty] | 0.18422 | 229.7541 | 423.0996 | 41.094 | 41.094 | 423.0996 | 229.662 | 238.7661 | 238.8582 | adjustment for residuals | 170.2874 | dump1 line start (FS1 straight line dump) | [13 empty] | 41.094 | 423.0996 | 229.662 |  |  | 2013/06/20 length updated from 5.892 to 5.922 Nathan Bultman
FS1 | BBS | DH | 2394 | FS1_BBS:DH_D2394 | D2 | dipole-45deg, D2-1 | 40 x 100 |  |  | 1.0407 | 229.1936 | 422.9056 | 41.094 | 41.094 | 422.7115 | 228.7251 | 239.3785 | 239.8988 |  | 170.8999 | FS1 | BTS | DH | 2394 | FS1_BTS:DH_D2394 | D2 | dipole-45deg, D2-1 | 40 x 100 | 1.83e-06 |  | 1.0974 | 228.5646 | 423.0996 | 41.094 | 41.094 | 423.0996 | 228.5646 | 239.4069 | 239.9556
[10 empty] | 0.63 | 228.5024 | 422.4888 | 41.094 | 41.094 | 422.266 | 228.2796 | 240.2138 | 240.5288 |  | 171.7352 | FS1 | BTS | GV | 2400 | FS1_BTS:GV_D2400 | GV_M | gate valve | 63 | 228.5271 | 228.5271 | 0.075 | 228.5271 | 423.0996 | 41.094 | 41.094 | 423.0996 | 228.4896 | 239.9931 | 240.0306
FS1 | BBS | SLH | 2405 | FS1_BBS:SLH_D2405 |  | charge selection slit high power |  |  |  | 0 | 228.2796 | 422.266 | 41.094 | 41.094 | 422.266 | 228.2796 | 240.5288 | 240.5288 |  | 172.0502 | [10 empty] | 0.7826 | 228.0983 | 423.0996 | 41.094 | 41.094 | 423.0996 | 227.707 | 240.4219 | 240.8132
[10 empty] | 0 | 228.2796 | 422.266 | 41.094 | 41.094 | 422.266 | 228.2796 | 240.5288 | 240.5288 |  | 172.0502 | [8 empty] | 1.83e-06 |  | 0 | 227.707 | 423.0996 | 41.094 | 41.094 | 423.0996 | 227.707 | 240.8132 | 240.8132
FS1 | BBS | TMP | 2405 | FS1_BBS:TMP_D2405 | TMP_L | turbo pump 2050 l/s |  |  |  | 0 | 228.2796 | 422.266 | 41.094 | 41.094 | 422.266 | 228.2796 | 240.5288 | 240.5288 |  | 172.0502 | FS1 | BTS | Q | 2409 | FS1_BTS:Q_D2409 | Q1 | quadrupole-25cm |  | 227.582 | 227.582 | 0.25 | 227.582 | 423.0996 | 41.094 | 41.094 | 423.0996 | 227.457 | 240.9382 | 241.0632
[10 empty] | 0.58 | 228.0746 | 422.061 | 41.094 | 41.094 | 421.8559 | 227.8695 | 240.8188 | 241.1088 |  | 172.3402 | [8 empty] | 1.83e-06 |  | 0.09 | 227.412 | 423.0996 | 41.094 | 41.094 | 423.0996 | 227.367 | 241.1082 | 241.1532
FS1 | BBS | DCHV | 2412 | FS1_BBS:DCHV_D2412 | C1 | correctors-X,Y | 140 |  |  | 0.2 | 227.7988 | 421.7852 | 41.094 | 41.094 | 421.7145 | 227.7281 | 241.2088 | 241.3088 |  | 172.7302 | FS1 | BTS | Q | 2413 | FS1_BTS:Q_D2413 | Q1 | quadrupole-25cm |  | 227.242 | 227.242 | 0.25 | 227.242 | 423.0996 | 41.094 | 41.094 | 423.0996 | 227.117 | 241.2782 | 241.4032
[10 empty] | 0.15 | 227.675 | 421.6615 | 41.094 | 41.094 | 421.6084 | 227.622 | 241.3838 | 241.4588 |  | 172.9052 | [10 empty] | 0.58999 | 226.822 | 423.0996 | 41.094 | 41.094 | 423.0996 | 226.527 | 241.6982 | 241.9932
FS1 | BBS | QH | 2416 | FS1_BBS:QH_D2416 | Q6 | quadrupole-X, B1 | 40 x 140 |  |  | 0.25 | 227.5336 | 421.52 | 41.094 | 41.094 | 421.4317 | 227.4452 | 241.5838 | 241.7088 |  | 173.1052 | FS1 | BTS | Q | 2421 | FS1_BTS:Q_D2421 | Q2 | quadrupole-26cm |  |  |  | 0.26 | 226.397 | 423.0996 | 41.094 | 41.094 | 423.0996 | 226.267 | 242.1232 | 242.2532
[10 empty] | 0 | 227.4452 | 421.4317 | 41.094 | 41.094 | 421.4317 | 227.4452 | 241.7088 | 241.7088 |  | 173.2302 | [10 empty] | 1.2438 | 225.6451 | 423.0996 | 41.094 | 41.094 | 423.0996 | 225.0232 | 242.8751 | 243.4969
[10 empty] | 0.14 | 227.3957 | 421.3822 | 41.094 | 41.094 | 421.3327 | 227.3462 | 241.7788 | 241.8488 |  | 173.3002 | [6 empty] | bellow |  | 1.83e-06 |  | 0.1035 | 224.9715 | 423.0996 | 41.094 | 41.094 | 423.0996 | 224.9197 | 243.5487 | 243.6004
FS1 | BBS | S | 2419 | FS1_BBS:S_D2419 | H3 | hexapole-X+ | 140 |  |  | 0.15 | 227.2932 | 421.2796 | 41.094 | 41.094 | 421.2266 | 227.2402 | 241.9238 | 241.9988 |  | 173.4452 | FS1 | BTS | BPM | 2436 | FS1_BTS:BPM_D2436 | BPM4 | position monitor | 100 | 224.9197 | 224.9197 | 0 | 224.9197 | 423.0996 | 41.094 | 41.094 | 423.0996 | 224.9197 | 243.6004 | 243.6004 | 20170427 adjusted based on ME
[8 empty] | -3.5182e-06 |  | 0.10618 | 227.2026 | 421.1891 | 41.094 | 41.094 | 421.1515 | 227.1651 | 242.0519 | 242.105 |  | 173.5733 | [6 empty] | diagnostic box |  |  |  | 0.2966 | 224.7714 | 423.0996 | 41.094 | 41.094 | 423.0996 | 224.6231 | 243.7487 | 243.8971
FS1 | BBS | BPM | 2421 | FS1_BBS:BPM_D2421 | BPM5 | position monitor 142x42 | 133.6 | 227.1353 | 227.1353 | 0.084138 | 227.1353 | 421.1218 | 41.094 | 41.094 | 421.092 | 227.1056 | 242.1471 | 242.1891 | 20170216 move up 0.03175 m | 173.6684 | FS1 | BTS | PM | 2439 | FS1_BTS:PM_D2439 | PM5 | profile monitor | 100 | defer installation |  | 0 | 224.6231 | 423.0996 | 41.094 | 41.094 | 423.0996 | 224.6231 | 243.8971 | 243.8971
[10 empty] | 0.11968 | 227.0633 | 421.0497 | 41.094 | 41.094 | 421.0074 | 227.021 | 242.249 | 242.3088 |  | 173.7704 | [6 empty] | diagnostic box |  | 1.83e-06 | 0.2966 | 0 | 224.6231 | 423.0996 | 41.094 | 41.094 | 423.0996 | 224.6231 | 243.8971 | 243.8971
FS1 | BBS | QV | 2424 | FS1_BBS:QV_D2424 | Q6 | quadrupole-Y, B1 | 40 x 140 |  |  | 0.25 | 226.9326 | 420.919 | 41.094 | 41.094 | 420.8306 | 226.8442 | 242.4338 | 242.5588 |  | 173.9552 | FS1 | BTS | IP | 2439 | FS1_BTS:IP_D2439 | IP_L | ion pump 300 l/s |  | 224.6231 | 224.6231 | 0 | 224.6231 | 423.0996 | 41.094 | 41.094 | 423.0996 | 224.6231 | 243.8971 | 243.8971
[10 empty] | 0.40322 | 226.7016 | 420.6881 | 41.094 | 41.094 | 420.5455 | 226.5591 | 242.7604 | 242.9621 |  | 174.2818 | [6 empty] | bellow |  |  |  | 0.13122 | 224.5575 | 423.0996 | 41.094 | 41.094 | 423.0996 | 224.4919 | 243.9627 | 244.0283 | 20170427 adjusted based on ME
[10 empty] | 0 | 226.5591 | 420.5455 | 41.094 | 41.094 | 420.5455 | 226.5591 | 242.9621 | 242.9621 |  | 174.4834 | [10 empty] | 0.75191 | 224.116 | 423.0996 | 41.094 | 41.094 | 423.0996 | 223.74 | 244.4042 | 244.7802
[10 empty] | 0 | 226.5591 | 420.5455 | 41.094 | 41.094 | 420.5455 | 226.5591 | 242.9621 | 242.9621 |  | 174.4834 | FS1 | BTS | BD | 2448 | FS1_BTS:BD_D2448 |  | dump1 |  |  |  | 0 | 223.74 | 423.0996 | 41.094 | 41.094 | 423.0996 | 223.74 | 244.7802 | 244.7802
[10 empty] | 0 | 226.5591 | 420.5455 | 41.094 | 41.094 | 420.5455 | 226.5591 | 242.9621 | 242.9621 |  | 174.4834 | dump1 line end (FS1 straight line dump) | [18 empty] | 45-deg line
Second beam dump line starts in "V" column | [9 empty] | 0 | 226.5591 | 420.5455 | 41.094 | 41.094 | 420.5455 | 226.5591 | 242.9621 | 242.9621 |  | 174.4834 | dump2 line start (FS1 45-deg line dump) | [13 empty] | 41.094 | 420.5455 | 226.5591 |  |  | 2013/06/20 length updated from 5.222 to 4.922 Nathan Bultman
FS1 | BBS | DH | 2435 | FS1_BBS:DH_D2435 | D2 | dipole-45deg, D2-2 | 40 x 100 |  |  | 1.0407 | 226.365 | 420.077 | 41.094 | 41.094 | 419.6086 | 226.171 | 243.4824 | 244.0027 |  | 175.0037 | FS1 | BTS | DH | 2435 | FS1_BTS:DH_D2435 | D2 | dipole-45deg, D2-2 | 40 x 100 |  |  | 1.0974 | 225.7831 | 419.7695 | 41.094 | 41.094 | 419.7695 | 225.7831 | 243.5108 | 244.0595
[10 empty] | 0 | 226.171 | 419.6086 | 41.094 | 41.094 | 419.6086 | 226.171 | 244.0027 | 244.0027 |  | 175.5241 | FS1 | BTS | GV | 2441 | FS1_BTS:GV_D2441 | GV_M | gate valve | 63 |  |  | 0.075 | 225.7566 | 419.743 | 41.094 | 41.094 | 419.7165 | 225.7301 | 244.097 | 244.1345
[10 empty] | 0.38904 | 226.171 | 419.4141 | 41.094 | 41.094 | 419.2195 | 226.171 | 244.1972 | 244.3917 |  | 175.7186 | [10 empty] | 0.26759 | 225.6354 | 419.6219 | 41.094 | 41.094 | 419.5273 | 225.5408 | 244.2683 | 244.4021
FS1 | BBS | PM | 2444 | FS1_BBS:PM_D2444 | PM4 | profile monitor | 150 | 0.00092926 |  | 0 | 226.171 | 419.2195 | 41.094 | 41.094 | 419.2195 | 226.171 | 244.3917 | 244.3917 | 20170216 moved dwon 0.004971 m | 175.9131 | FS1 | BTS | Q | 2445 | FS1_BTS:Q_D2445 | Q1 | quadrupole-25cm | 50 |  |  | 0.25 | 225.4525 | 419.4389 | 41.094 | 41.094 | 419.3505 | 225.3641 | 244.5271 | 244.6521
[10 empty] | 0 | 226.171 | 419.2195 | 41.094 | 41.094 | 419.2195 | 226.171 | 244.3917 | 244.3917 |  | 175.9131 | [10 empty] | 0.25 | 225.2757 | 419.2621 | 41.094 | 41.094 | 419.1737 | 225.1873 | 244.7771 | 244.9021
[10 empty] | 0 | 226.171 | 419.2195 | 41.094 | 41.094 | 419.2195 | 226.171 | 244.3917 | 244.3917 |  | 175.9131 |  |  | BCM | 2449 | FS1_BTS:BCM_D2449 | BCM? | temporary BCM until ARR04 | 75 |  |  | 0 | 225.1873 | 419.1737 | 41.094 | 41.094 | 419.1737 | 225.1873 | 244.9021 | 244.9021
FS1 | BBS | IP | 2444 | FS1_BBS:IP_D2444 | IP_L | ion pump, 300 l/s |  | 0 | 419.2195 | 0 | 226.171 | 419.2195 | 41.094 | 41.094 | 419.2195 | 226.171 | 244.3917 | 244.3917 |  | 175.9131 | [10 empty] | 0.26 | 225.0954 | 419.0818 | 41.094 | 41.094 | 418.9899 | 225.0034 | 245.0321 | 245.1621
FS1 | BBS | FAVS | 2444 | FS1_BBS:FAVS_D2444 | FVS | fast valve sensor |  |  |  | 0 | 226.171 | 419.2195 | 41.094 | 41.094 | 419.2195 | 226.171 | 244.3917 | 244.3917 |  | 175.9131 | FS1 | BTS | Q | 2453 | FS1_BTS:Q_D2453 | Q2 | quadrupole-26cm | 75 |  |  | 0.26 | 224.9115 | 418.8979 | 41.094 | 41.094 | 418.806 | 224.8196 | 245.2921 | 245.4221
[10 empty] | 0.38096 | 226.171 | 419.0291 | 41.094 | 41.094 | 418.8386 | 226.171 | 244.5822 | 244.7727 |  | 176.1036 | [10 empty] | 1.1978 | 224.3961 | 418.3825 | 41.094 | 41.094 | 417.959 | 223.9726 | 246.021 | 246.6199
[10 empty] | 0 | 226.171 | 418.8386 | 41.094 | 41.094 | 418.8386 | 226.171 | 244.7727 | 244.7727 |  | 176.2941 | [6 empty] | bellow |  |  |  | 0.1035 | 223.936 | 417.9224 | 41.094 | 41.094 | 417.8858 | 223.8994 | 246.6716 | 246.7234
FS1 | BBS | DH | 2453 | FS1_BBS:DH_D2453 | D2 | dipole-45deg, D2-3 | 40 x 100 |  |  | 1.0407 | 226.365 | 418.3701 | 41.094 | 41.094 | 417.9017 | 226.5591 | 245.293 | 245.8134 |  | 176.8144 | FS1 | BTS | BPM | 2467 | FS1_BTS:BPM_D2467 | BPM4 | position monitor | 100 |  |  | 0 | 223.8994 | 417.8858 | 41.094 | 41.094 | 417.8858 | 223.8994 | 246.7234 | 246.7234 | 20170427 adjusted based on ME
[10 empty] | 0.1873 | 226.6253 | 417.8354 | 41.094 | 41.094 | 417.7692 | 226.6915 | 245.907 | 246.0007 |  | 177.4284 | [6 empty] | diagnostic box |  |  |  | 0.29661 | 223.7946 | 417.781 | 41.094 | 41.094 | 417.6761 | 223.6897 | 246.8717 | 247.02
[6 empty] | collimation flange |  | 1.4378e-06 | 226.6915 | 0 | 226.6915 | 417.7692 | 41.094 | 41.094 | 417.7692 | 226.6915 | 246.0007 | 246.0007 |  | 177.522 | [10 empty] | 0 | 223.6897 | 417.6761 | 41.094 | 41.094 | 417.6761 | 223.6897 | 247.02 | 247.02
[8 empty] | 0.358 | 0.34093 | 0.21593 | 226.7679 | 417.6929 | 41.094 | 41.094 | 417.6165 | 226.8442 | 246.1086 | 246.2166 |  | 177.63 | FS1 | BTS | PM | 2470 | FS1_BTS:PM_D2470 | PM5 | profile monitor | 100 | defer installation |  | 0 | 223.6897 | 417.6761 | 41.094 | 41.094 | 417.6761 | 223.6897 | 247.02 | 247.02
FS1 | BBS | QV | 2463 | FS1_BBS:QV_D2463 | Q6 | quadrupole-Y, B1 | 40 x 140 |  | 0.125 | 0.25 | 226.9326 | 417.5282 | 41.094 | 41.094 | 417.4398 | 227.021 | 246.3416 | 246.4666 |  | 177.8629 | [6 empty] | diagnostic box |  |  |  | 0 | 223.6897 | 417.6761 | 41.094 | 41.094 | 417.6761 | 223.6897 | 247.02 | 247.02
[10 empty] | 0.11968 | 227.0633 | 417.3975 | 41.094 | 41.094 | 417.3551 | 227.1056 | 246.5264 | 246.5863 |  | 178.0478 | FS1 | BTS | IP | 2470 | FS1_BTS:IP_D2470 | IP_L | ion pump 300 l/s |  |  |  | 0 | 223.6897 | 417.6761 | 41.094 | 41.094 | 417.6761 | 223.6897 | 247.02 | 247.02
FS1 | BBS | BPM | 2466 | FS1_BBS:BPM_D2466 | BPM5 | position monitor R=75mm | 133.6 |  |  | 0.084138 | 227.1353 | 417.3254 | 41.094 | 41.094 | 417.2956 | 227.1651 | 246.6283 | 246.6704 |  | 178.1497 | [6 empty] | bellow |  |  |  | 0.13122 | 223.6433 | 417.6297 | 41.094 | 41.094 | 417.5833 | 223.5969 | 247.0856 | 247.1512 | 20170427 adjusted based on ME
[6 empty] | collimation flange |  |  |  | 0.10618 | 227.2026 | 417.2581 | 41.094 | 41.094 | 417.2206 | 227.2402 | 246.7235 | 246.7766 |  | 178.2449 | [10 empty] | 0.73286 | 223.3378 | 417.3242 | 41.094 | 41.094 | 417.0651 | 223.0787 | 247.5176 | 247.8841
FS1 | BBS | S | 2469 | FS1_BBS:S_D2469 | H3 | hexapole-X+ | 140 |  |  | 0.15 | 227.2932 | 417.1675 | 41.094 | 41.094 | 417.1145 | 227.3462 | 246.8516 | 246.9266 |  | 178.3729 | FS1 | BTS | BD | 2479 | FS1_BTS:BD_D2479 |  | dump2 |  |  |  | 0 | 223.0787 | 417.0651 | 41.094 | 41.094 | 417.0651 | 223.0787 | 247.8841 | 247.8841
[10 empty] | 0.14 | 227.3957 | 417.065 | 41.094 | 41.094 | 417.0155 | 227.4452 | 246.9966 | 247.0666 |  | 178.5179 | dump2 line end (FS1 45-deg line dump)
FS1 | BBS | QH | 2472 | FS1_BBS:QH_D2472 | Q6 | quadrupole-Y, B1 | 40 x 140 |  |  | 0.25 | 227.5336 | 416.9271 | 41.094 | 41.094 | 416.8387 | 227.622 | 247.1916 | 247.3166 |  | 178.7129
[10 empty] | 0.15 | 227.675 | 416.7857 | 41.094 | 41.094 | 416.7327 | 227.7281 | 247.3916 | 247.4666 |  | 178.9129
FS1 | BBS | DCHV | 2476 | FS1_BBS:DCHV_D2476 | C1 | correctors-X,Y | 140 |  |  | 0.2 | 227.7988 | 416.6619 | 41.094 | 41.094 | 416.5912 | 227.8695 | 247.5666 | 247.6666 |  | 179.0879
[10 empty] | 0.58 | 228.0746 | 416.3862 | 41.094 | 41.094 | 416.1811 | 228.2796 | 247.9566 | 248.2466 |  | 179.4779
FS1 | BBS | PM | 2482 | FS1_BBS:PM_D2482 | PM4 | profile monitor | 150 |  |  | 0 | 228.2796 | 416.1811 | 41.094 | 41.094 | 416.1811 | 228.2796 | 248.2466 | 248.2466 |  | 179.7679
FS1 | BBS | IP | 2482 | FS1_BBS:IP_D2482 | IP_L | ion pump, 300 l/s |  |  |  | 0 | 228.2796 | 416.1811 | 41.094 | 41.094 | 416.1811 | 228.2796 | 248.2466 | 248.2466 |  | 179.7679
[10 empty] | 0.63 | 228.5024 | 415.9584 | 41.094 | 41.094 | 415.7356 | 228.7251 | 248.5616 | 248.8766 |  | 180.0829
FS1 | BBS | DH | 2494 | FS1_BBS:DH_D2494 | D2 | dipole-45deg, D2-4 | 40 x 100 |  |  | 1.0407 | 229.1936 | 415.5416 | 41.094 | 41.094 | 415.3476 | 229.662 | 249.3969 | 249.9172 |  | 180.9183
[8 empty] | 3.176e-06 | 0.016423 | 0.15866 | 229.7413 | 415.3476 | 41.094 | 41.094 | 415.3476 | 229.8207 | 249.9966 | 250.0759 |  | 181.5179
FS1 | BMS | GV | 2501 | FS1_BMS:GV_D2501 | GV_M | gate valve | 63 | 229.8582 | 229.8582 | 0.075 | 229.8582 | 415.3476 | 41.094 | 41.094 | 415.3476 | 229.8957 | 250.1134 | 250.1509 | 20170216 seems X of GV incorrect | 181.6348
[10 empty] | 0 | 229.8957 | 415.3476 | 41.094 | 41.094 | 415.3476 | 229.8957 | 250.1509 | 250.1509 |  | 181.6723
[6 empty] | bellow | 40 | -0.008801 |  | 0.094701 | 229.943 | 415.3476 | 41.094 | 41.094 | 415.3476 | 229.9904 | 250.1982 | 250.2456 | bellow compressed by 0.008801 m, 4/24/2017 | 181.7196
FS1 | BMS | BPM | 2502 | FS1_BMS:BPM_D2502 | BPM3 | position monitor | 50 | 2.176e-06 | 229.9904 | 0 | 229.9904 | 415.3476 | 41.094 | 41.094 | 415.3476 | 229.9904 | 250.2456 | 250.2456 |  | 181.767
[6 empty] | diagnostic box |  |  |  | 0.14528 | 230.063 | 415.3476 | 41.094 | 41.094 | 415.3476 | 230.1357 | 250.3182 | 250.3909 |  | 181.8396
FS1 | BMS | PM | 2504 | FS1_BMS:PM_D2504 | PM3 | profile monitor | 50 | defer installation |  | 0 | 230.1357 | 415.3476 | 41.094 | 41.094 | 415.3476 | 230.1357 | 250.3909 | 250.3909 |  | 181.9122
[6 empty] | diagnostic box |  |  |  | 0 | 230.1357 | 415.3476 | 41.094 | 41.094 | 415.3476 | 230.1357 | 250.3909 | 250.3909 |  | 181.9122
FS1 | BMS | FAVS | 2504 | FS1_BMS:FAVS_D2504 | FVS | fast valve sensor |  |  |  | 0 | 230.1357 | 415.3476 | 41.094 | 41.094 | 415.3476 | 230.1357 | 250.3909 | 250.3909 |  | 181.9122
FS1 | BMS | IP | 2504 | FS1_BMS:IP_D2504 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 230.1357 | 415.3476 | 41.094 | 41.094 | 415.3476 | 230.1357 | 250.3909 | 250.3909 |  | 181.9122
[6 empty] | bellow | 40 | 0.085138 |  | 0.13122 | 230.2013 | 415.3476 | 41.094 | 41.094 | 415.3476 | 230.2669 | 250.4565 | 250.5221 |  | 181.9779
[10 empty] | 0.01743 | 230.2756 | 415.3476 | 41.094 | 41.094 | 415.3476 | 230.2843 | 250.5308 | 250.5395 |  | 182.0522
[6 empty] | collimation flange |  | 2.176e-06 | 230.2843 | 0 | 230.2843 | 415.3476 | 41.094 | 41.094 | 415.3476 | 230.2843 | 250.5395 | 250.5395 | 20170216 collimation flange moved from upstream | 182.0609
[10 empty] | 0.067708 | 230.3182 | 415.3476 | 41.094 | 41.094 | 415.3476 | 230.352 | 250.5734 | 250.6072 |  | 182.0947
FS1 | BMS | DCHV | 2507 | FS1_BMS:DCHV_D2507 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 230.452 | 415.3476 | 41.094 | 41.094 | 415.3476 | 230.552 | 250.7072 | 250.8072 |  | 182.2286
[10 empty] | 0.145 | 230.6245 | 415.3476 | 41.094 | 41.094 | 415.3476 | 230.697 | 250.8797 | 250.9522 |  | 182.4011
FS1 | BMS | QV | 2511 | FS1_BMS:QV_D2511 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 230.822 | 415.3476 | 41.094 | 41.094 | 415.3476 | 230.947 | 251.0772 | 251.2022 |  | 182.5986
[6 empty] | collimation flange moved?? |  |  |  | 0.2 | 231.047 | 415.3476 | 41.094 | 41.094 | 415.3476 | 231.147 | 251.3022 | 251.4022 |  | 182.8236
FS1 | BMS | QH | 2515 | FS1_BMS:QH_D2515 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 231.272 | 415.3476 | 41.094 | 41.094 | 415.3476 | 231.397 | 251.5272 | 251.6522 |  | 183.0486
[8 empty] | 3.176e-06 | 0.20683 | 0.23183 | 231.5129 | 415.3476 | 41.094 | 41.094 | 415.3476 | 231.6288 | 251.7681 | 251.8841 | 20170216 space adjusted | 183.2895
FS1 | BMS | BCM | 2519 | FS1_BMS:BCM_D2519 | BCM3 | current monitor |  | 231.6788 | 231.6788 | 0.1 | 231.6788 | 415.3476 | 41.094 | 41.094 | 415.3476 | 231.7288 | 251.9341 | 251.9841 |  | 183.4554
[10 empty] | 1.1932 | 232.3254 | 415.3476 | 41.094 | 41.094 | 415.3476 | 232.922 | 252.5806 | 253.1772 | 20170216 space adjusted | 184.102
[6 empty] | collimation flange moved?? |  |  |  | 0.1 | 232.972 | 415.3476 | 41.094 | 41.094 | 415.3476 | 233.022 | 253.2272 | 253.2772 |  | 184.7486
FS1 | BMS | DCHV | 2534 | FS1_BMS:DCHV_D2534 | C2 | correctors-X,Y | 50 |  | 0.29242 | 0.2 | 233.122 | 415.3476 | 41.094 | 41.094 | 415.3476 | 233.222 | 253.3772 | 253.4772 |  | 184.8986
[8 empty] | 3.176e-06 | 0.054921 | 0.12392 | 233.284 | 415.3476 | 41.094 | 41.094 | 415.3476 | 233.3459 | 253.5392 | 253.6012 | 20170216 space adjusted | 185.0606
FS1 | BMS | BPM | 2537 | FS1_BMS:BPM_D2537 | BPM3 | position monitor | 47.5 | 233.4144 | 233.4144 | 0.137 | 233.4144 | 415.3476 | 41.094 | 41.094 | 415.3476 | 233.4829 | 253.6697 | 253.7382 |  | 185.191
[9 empty] | 0.18258 | -0.010921 | 233.4775 | 415.3476 | 41.094 | 41.094 | 415.3476 | 233.472 | 253.7327 | 253.7272 | 20170216 space adjusted | 185.2541
FS1 | BMS | QV | 2539 | FS1_BMS:QV_D2539 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 233.597 | 415.3476 | 41.094 | 41.094 | 415.3476 | 233.722 | 253.8522 | 253.9772 |  | 185.3736
[10 empty] | 1.2 | 234.322 | 415.3476 | 41.094 | 41.094 | 415.3476 | 234.922 | 254.5772 | 255.1772 |  | 186.0986
FS1 | BMS | IP | 2552 | FS1_BMS:IP_D2552 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 234.922 | 415.3476 | 41.094 | 41.094 | 415.3476 | 234.922 | 255.1772 | 255.1772 |  | 186.6986
[10 empty] | 1 | 235.422 | 415.3476 | 41.094 | 41.094 | 415.3476 | 235.922 | 255.6772 | 256.1772 |  | 187.1986
FS1 | BMS | QH | 2563 | FS1_BMS:QH_D2563 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 236.047 | 415.3476 | 41.094 | 41.094 | 415.3476 | 236.172 | 256.3022 | 256.4272 |  | 187.8236
[10 empty] | 0 | 236.172 | 415.3476 | 41.094 | 41.094 | 415.3476 | 236.172 | 256.4272 | 256.4272 |  | 187.9486
[6 empty] | collimation flange |  |  |  | 0.1 | 236.222 | 415.3476 | 41.094 | 41.094 | 415.3476 | 236.272 | 256.4772 | 256.5272 |  | 187.9986
[8 empty] | 3.176e-06 |  | 0.004193 | 236.2741 | 415.3476 | 41.094 | 41.094 | 415.3476 | 236.2762 | 256.5293 | 256.5314 | 20170216 space adjusted | 188.0507
FS1 | BMS | GV | 2566 | FS1_BMS:GV_D2566 | GV_M | gate valve | 63 | 236.3137 | 236.3137 | 0.075 | 236.3137 | 415.3476 | 41.094 | 41.094 | 415.3476 | 236.3512 | 256.5689 | 256.6064 | 20170216 GV inserted | 188.0903
[8 empty] | 3.076e-06 |  | 1.7184 | 237.2104 | 415.3476 | 41.094 | 41.094 | 415.3476 | 238.0696 | 257.4656 | 258.3248 | 20170216 space adjusted | 188.987
FS1 | BMS | DCHV | 2584 | FS1_BMS:DCHV_D2584 | C2 | correctors-X,Y | 50 | 238.1696 | 238.1696 | 0.2 | 238.1696 | 415.3476 | 41.094 | 41.094 | 415.3476 | 238.2696 | 258.4248 | 258.5248 |  | 189.9462
[8 empty] | 3.076e-06 |  | 0.073805 | 238.3065 | 415.3476 | 41.094 | 41.094 | 415.3476 | 238.3434 | 258.5617 | 258.5986 | 20170216 space adjusted | 190.0831
FS1 | BMS | BPM | 2587 | FS1_BMS:BPM_D2587 | BPM3 | position monitor | 47.5 | 238.4119 | 238.4119 | 0.137 | 238.4119 | 415.3476 | 41.094 | 41.094 | 415.3476 | 238.4804 | 258.6671 | 258.7356 |  | 190.1885
[10 empty] | 0.090002 | 238.5254 | 415.3476 | 41.094 | 41.094 | 415.3476 | 238.5704 | 258.7806 | 258.8256 | 20170216 space adjusted | 190.302
FS1 | BMS | QV | 2590 | FS1_BMS:QV_D2590 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 238.6954 | 415.3476 | 41.094 | 41.094 | 415.3476 | 238.8204 | 258.9506 | 259.0756 |  | 190.472
[8 empty] | 3.076e-06 | -0.25408 | 0.24592 | 238.9434 | 415.3476 | 41.094 | 41.094 | 415.3476 | 239.0663 | 259.1986 | 259.3216 | 20170216 space adjusted | 190.72
FS1 | BMS | FAV | 2594 | FS1_BMS:FAV_D2594 | FV | fast valve | 50 | 239.0999 | 239.0999 | 0.067196 | 239.0999 | 415.3476 | 41.094 | 41.094 | 415.3476 | 239.1335 | 259.3552 | 259.3888 | 20170216 fast valve moved from downstream | 190.8765
[7 empty] | Is it the good place for FV ? |  |  | 0.18688 | 239.227 | 415.3476 | 41.094 | 41.094 | 415.3476 | 239.3204 | 259.4822 | 259.5756 | 20170216 space adjusted | 191.0036
FS1 | BMS | QH | 2597 | FS1_BMS:QH_D2597 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 239.4454 | 415.3476 | 41.094 | 41.094 | 415.3476 | 239.5704 | 259.7006 | 259.8256 |  | 191.222
[10 empty] | 0.10898 | 239.6249 | 415.3476 | 41.094 | 41.094 | 415.3476 | 239.6794 | 259.8801 | 259.9346 | 20170216 space adjusted | 191.4015
[6 empty] | bellow | 50 | 3.076e-06 |  | 0.1035 | 239.7312 | 415.3476 | 41.094 | 41.094 | 415.3476 | 239.7829 | 259.9864 | 260.0381 |  | 191.5077
FS1 | BMS | BPM | 2600 | FS1_BMS:BPM_D2600 | BPM2 | position monitor |  | 239.7829 | 239.7829 | 0 | 239.7829 | 415.3476 | 41.094 | 41.094 | 415.3476 | 239.7829 | 260.0381 | 260.0381 |  | 191.5595
[6 empty] | diagnostic box |  |  |  | 0.14528 | 239.8555 | 415.3476 | 41.094 | 41.094 | 415.3476 | 239.9282 | 260.1108 | 260.1834 |  | 191.6321
FS1 | BMS | PM | 2602 | FS1_BMS:PM_D2602 | PM2 | profile monitor |  | 3.076e-06 |  | 0 | 239.9282 | 415.3476 | 41.094 | 41.094 | 415.3476 | 239.9282 | 260.1834 | 260.1834 |  | 191.7048
[6 empty] | diagnostic box |  | 239.9282 | 239.9282 | 0 | 239.9282 | 415.3476 | 41.094 | 41.094 | 415.3476 | 239.9282 | 260.1834 | 260.1834 |  | 191.7048
FS1 | BMS | FAVS | 2602 | FS1_BMS:FAVS_D2602 | FVS | fast valve sensor |  |  |  | 0 | 239.9282 | 415.3476 | 41.094 | 41.094 | 415.3476 | 239.9282 | 260.1834 | 260.1834 |  | 191.7048
FS1 | BMS | IP | 2602 | FS1_BMS:IP_D2602 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 239.9282 | 415.3476 | 41.094 | 41.094 | 415.3476 | 239.9282 | 260.1834 | 260.1834 |  | 191.7048
[6 empty] | bellow | 50 | 0 | 0.019014 | 0.13122 | 239.9938 | 415.3476 | 41.094 | 41.094 | 415.3476 | 240.0594 | 260.249 | 260.3146 |  | 191.7704
[8 empty] | 0.5776 | 0.5776 | 0.088621 | 240.1037 | 415.3476 | 41.094 | 41.094 | 415.3476 | 240.148 | 260.3589 | 260.4032 | 20170216 space adjusted | 191.8803
FS1 | MGB02 | GV | 2604 | FS1_MGB02:GV_D2604 |  | gate valve | 40 | -240.1767 |  | 0.057404 | 240.1767 | 415.3476 | 41.094 | 41.094 | 415.3476 | 240.2054 | 260.4319 | 260.4606 |  | 191.9533
[10 empty] | 0 | 240.2054 | 415.3476 | 41.094 | 41.094 | 415.3476 | 240.2054 | 260.4606 | 260.4606 |  | 191.982
FS1 | MGB02 | CAV | 2606 | FS1_MGB02:CAV_D2606 |  | b19 resonator 1 1/4 | 36 |  |  | 0.31833 | 240.3646 | 415.3476 | 41.094 | 41.094 | 415.3476 | 240.5238 | 260.6198 | 260.779 |  | 192.1412
FS1 | MGB02 | IP | 2608 | FS1_MGB02:CAV_D2608 | IP_L | ion pump, ? l/s |  |  |  | 0 | 240.5238 | 415.3476 | 41.094 | 41.094 | 415.3476 | 240.5238 | 260.779 | 260.779 |  | 192.3003
FS1 | MGB02 | CAV | 2609 | FS1_MGB02:CAV_D2609 |  | b19 resonator 1 2/4 | 36 |  |  | 0.31833 | 240.6829 | 415.3476 | 41.094 | 41.094 | 415.3476 | 240.8421 | 260.9381 | 261.0973 |  | 192.4595
|  |  | 2611 | FS1_MGB02:CAV_D2611 |  | rf coupler |  | -240.8421 |  | 0 | 240.8421 | 415.3476 | 41.094 | 41.094 | 415.3476 | 240.8421 | 261.0973 | 261.0973 |  | 192.6187
[10 empty] | 0 | 240.8421 | 415.3476 | 41.094 | 41.094 | 415.3476 | 240.8421 | 261.0973 | 261.0973 |  | 192.6187
FS1 | MGB02 | CAV | 2613 | FS1_MGB02:CAV_D2613 |  | b19 resonator 1 3/4 | 36 |  |  | 0.31833 | 241.0012 | 415.3476 | 41.094 | 41.094 | 415.3476 | 241.1604 | 261.2565 | 261.4156 |  | 192.7778
FS1 | MGB02 | TMP | 2614 | FS1_MGB02:CAV_D2614 | TMP_L | turbo pump, ? l/s |  |  |  | 0 | 241.1604 | 415.3476 | 41.094 | 41.094 | 415.3476 | 241.1604 | 261.4156 | 261.4156 |  | 192.937
FS1 | MGB02 | CAV | 2616 | FS1_MGB02:CAV_D2616 |  | b19 resonator 1 4/4 | 36 |  |  | 0.31833 | 241.3196 | 415.3476 | 41.094 | 41.094 | 415.3476 | 241.4787 | 261.5748 | 261.7339 |  | 193.0962
[10 empty] | 0 | 241.4787 | 415.3476 | 41.094 | 41.094 | 415.3476 | 241.4787 | 261.7339 | 261.7339 |  | 193.2553
FS1 | MGB02 | GV | 2618 | FS1_MGB02:GV_D2618 |  | gate valve | 40 | -241.5074 |  | 0.057404 | 241.5074 | 415.3476 | 41.094 | 41.094 | 415.3476 | 241.5361 | 261.7626 | 261.7914 |  | 193.284
[10 empty] | 0.89642 | 241.9843 | 415.3476 | 41.094 | 41.094 | 415.3476 | 242.4326 | 262.2396 | 262.6878 |  | 193.7609
[6 empty] | bellow | 40 |  |  | 0 | 242.4326 | 415.3476 | 41.094 | 41.094 | 415.3476 | 242.4326 | 262.6878 | 262.6878 |  | 194.2091
[10 empty] | 0.24879 | 242.5569 | 415.3476 | 41.094 | 41.094 | 415.3476 | 242.6813 | 262.8122 | 262.9366 | 20170216 space adjusted | 194.3335
[6 empty] | vacuum box |  | 3.076e-06 | 0.15479 | 0 | 242.6813 | 415.3476 | 41.094 | 41.094 | 415.3476 | 242.6813 | 262.9366 | 262.9366 |  | 194.4579
FS1 | BMS | FAVS | 2629 | FS1_BMS:FAVS_D2629 | FVS | fast valve sensor |  | 242.6813 | 242.6813 | 0 | 242.6813 | 415.3476 | 41.094 | 41.094 | 415.3476 | 242.6813 | 262.9366 | 262.9366 |  | 194.4579
[6 empty] | vacuum box |  |  |  | 0 | 242.6813 | 415.3476 | 41.094 | 41.094 | 415.3476 | 242.6813 | 262.9366 | 262.9366 |  | 194.4579
FS1 | BMS | IP | 2629 | FS1_BMS:IP_D2629 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 242.6813 | 415.3476 | 41.094 | 41.094 | 415.3476 | 242.6813 | 262.9366 | 262.9366 |  | 194.4579
[6 empty] | bellow | 40 |  |  | 0.035206 | 242.6989 | 415.3476 | 41.094 | 41.094 | 415.3476 | 242.7166 | 262.9542 | 262.9718 | 20170216 space adjusted | 194.4755
[10 empty] | 0.001052 | 242.7171 | 415.3476 | 41.094 | 41.094 | 415.3476 | 242.7176 | 262.9723 | 262.9728 |  | 194.4937
[10 empty] | 0.072 | 242.7536 | 415.3476 | 41.094 | 41.094 | 415.3476 | 242.7896 | 263.0088 | 263.0448 | 20170216 moved from down stream | 194.5302
[10 empty] | 0.37429 | 242.9768 | 415.3476 | 41.094 | 41.094 | 415.3476 | 243.1639 | 263.232 | 263.4191 |  | 194.7533
FS1 | BMS | BD | 2634 | FS1_BMS:BD_D2634 | dump | dump 25W no current reading |  | -2.076e-06 | 243.1639 | 0 | 243.1639 | 415.3476 | 41.094 | 41.094 | 415.3476 | 243.1639 | 263.4191 | 263.4191 |  | 194.9405
[8 empty] | 0.64996 |  | 0.27567 | 243.3017 | 415.3476 | 41.094 | 41.094 | 415.3476 | 243.4396 | 263.5569 | 263.6948 |  | 195.0783
[10 empty] | 0.072 | 243.4756 | 415.3476 | 41.094 | 41.094 | 415.3476 | 243.5116 | 263.7308 | 263.7668 |  | 195.2522
[10 empty] | 0.1775 | 243.6003 | 415.3476 | 41.094 | 41.094 | 415.3476 | 243.6891 | 263.8555 | 263.9443 |  | 195.3769
FS1 | BMS | DCHV | 2640 | FS1_BMS:DCHV_D2640 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 243.7891 | 415.3476 | 41.094 | 41.094 | 415.3476 | 243.8891 | 264.0443 | 264.1443 |  | 195.5657
[10 empty] | 0.25 | 244.0141 | 415.3476 | 41.094 | 41.094 | 415.3476 | 244.1391 | 264.2693 | 264.3943 |  | 195.7907
FS1 | BMS | QH | 2645 | FS1_BMS:QH_D2645 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 244.2641 | 415.3476 | 41.094 | 41.094 | 415.3476 | 244.3891 | 264.5193 | 264.6443 |  | 196.0407
[10 empty] | 0.41601 | 244.5971 | 415.3476 | 41.094 | 41.094 | 415.3476 | 244.8051 | 264.8523 | 265.0603 |  | 196.3737
FS1 | BMS | GV | 2651 | FS1_BMS:GV_D2651 | GV_M | gate valve | 63 | 2.076e-06 | 244.8426 | 0.075 | 244.8426 | 415.3476 | 41.094 | 41.094 | 415.3476 | 244.8801 | 265.0978 | 265.1353 |  | 196.6192
[10 empty] | 0.13139 | 244.9458 | 415.3476 | 41.094 | 41.094 | 415.3476 | 245.0115 | 265.201 | 265.2667 |  | 196.7224
FS1 | BMS | QV | 2654 | FS1_BMS:QV_D2654 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 245.1365 | 415.3476 | 41.094 | 41.094 | 415.3476 | 245.2615 | 265.3917 | 265.5167 |  | 196.9131
[10 empty] | 0.55 | 245.5365 | 415.3476 | 41.094 | 41.094 | 415.3476 | 245.8115 | 265.7917 | 266.0667 |  | 197.3131
FS1 | BMS | DCHV | 2662 | FS1_BMS:DCHV_D2662 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 245.9115 | 415.3476 | 41.094 | 41.094 | 415.3476 | 246.0115 | 266.1667 | 266.2667 |  | 197.6881
[8 empty] | 2.776e-06 | 0.054921 | 0.12392 | 246.0734 | 415.3476 | 41.094 | 41.094 | 415.3476 | 246.1354 | 266.3286 | 266.3906 |  | 197.85
FS1 | BMS | BPM | 2665 | FS1_BMS:BPM_D2665 | BPM3 | position monitor | 47.5 | 246.2039 | 246.2039 | 0.137 | 246.2039 | 415.3476 | 41.094 | 41.094 | 415.3476 | 246.2724 | 266.4591 | 266.5276 |  | 197.9805
[10 empty] | -0.010921 | 246.2669 | 415.3476 | 41.094 | 41.094 | 415.3476 | 246.2615 | 266.5221 | 266.5167 |  | 198.0435
FS1 | BMS | QH | 2666 | FS1_BMS:QH_D2666 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 246.3865 | 415.3476 | 41.094 | 41.094 | 415.3476 | 246.5115 | 266.6417 | 266.7667 |  | 198.1631
[10 empty] | 0.49027 | 246.7566 | 415.3476 | 41.094 | 41.094 | 415.3476 | 247.0017 | 267.0118 | 267.257 | 20170216 space adjusted | 198.5332
FS1 | BMS | LPM | 2673 | FS1_BMS:LPM_D2673 | LPMx | bunch length monitor (not costed) | future LPM2 | 2.976e-06 | -0.034726 | 0 | 247.0017 | 415.3476 | 41.094 | 41.094 | 415.3476 | 247.0017 | 267.257 | 267.257 |  | 198.7783
FS1 | BMS | IP | 2673 | FS1_BMS:IP_D2673 | IP_S | ion pump, 75 l/s |  | 247.0017 | 247.0017 | 0 | 247.0017 | 415.3476 | 41.094 | 41.094 | 415.3476 | 247.0017 | 267.257 | 267.257 |  | 198.7783
[8 empty] | 2.976e-06 |  | 0.14619 | 247.0748 | 415.3476 | 41.094 | 41.094 | 415.3476 | 247.1479 | 267.3301 | 267.4031 | 20170216 space adjusted | 198.8514
FS1 | BMS | BCM | 2675 | FS1_BMS:BCM_D2675 | BCM2 | current monitor |  | 247.2479 | 247.2479 | 0.2 | 247.2479 | 415.3476 | 41.094 | 41.094 | 415.3476 | 247.3479 | 267.5031 | 267.6031 |  | 199.0245
[10 empty] | 0.21354 | 247.4547 | 415.3476 | 41.094 | 41.094 | 415.3476 | 247.5615 | 267.7099 | 267.8167 | 20170216 space adjusted | 199.2313
FS1 | BMS | QH | 2679 | FS1_BMS:QH_D2679 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 247.6865 | 415.3476 | 41.094 | 41.094 | 415.3476 | 247.8115 | 267.9417 | 268.0667 |  | 199.4631
[10 empty] | 0.3495 | 247.9862 | 415.3476 | 41.094 | 41.094 | 415.3476 | 248.161 | 268.2414 | 268.4162 |  | 199.7628
[10 empty] | 0 | 248.161 | 415.3476 | 41.094 | 41.094 | 415.3476 | 248.161 | 268.4162 | 268.4162 |  | 199.9376
[8 empty] | 2.976e-06 | -0.058407 | 0.29109 | 248.3065 | 415.3476 | 41.094 | 41.094 | 415.3476 | 248.4521 | 268.5617 | 268.7073 | 20170216 space adjusted | 200.0831
FS1 | BMS | DCHV | 2688 | FS1_BMS:DCHV_D2688 | C2 | correctors-X,Y | 50 | 248.5521 | 248.5521 | 0.2 | 248.5521 | 415.3476 | 41.094 | 41.094 | 415.3476 | 248.6521 | 268.8073 | 268.9073 |  | 200.3286
[8 empty] | 2.976e-06 | -0.13856 | -0.011153 | 248.6465 | 415.3476 | 41.094 | 41.094 | 415.3476 | 248.6409 | 268.9017 | 268.8961 | 20170216 space adjusted | 200.4231
FS1 | BMS | BPM | 2690 | FS1_BMS:BPM_D2690 | BPM3 | position monitor | 47.5 | 248.7094 | 248.7094 | 0.137 | 248.7094 | 415.3476 | 41.094 | 41.094 | 415.3476 | 248.7779 | 268.9646 | 269.0331 |  | 200.486
[10 empty] | 0.18256 | 248.8692 | 415.3476 | 41.094 | 41.094 | 415.3476 | 248.9605 | 269.1244 | 269.2157 | 20170216 space adjusted | 200.6458
FS1 | BMS | QV | 2693 | FS1_BMS:QV_D2693 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 249.0855 | 415.3476 | 41.094 | 41.094 | 415.3476 | 249.2105 | 269.3407 | 269.4657 |  | 200.8621
[10 empty] | 0.25 | 249.3355 | 415.3476 | 41.094 | 41.094 | 415.3476 | 249.4605 | 269.5907 | 269.7157 |  | 201.1121
FS1 | BMS | QH | 2698 | FS1_BMS:QH_D2698 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 249.5855 | 415.3476 | 41.094 | 41.094 | 415.3476 | 249.7105 | 269.8407 | 269.9657 |  | 201.3621
[10 empty] | 0.12 | 249.7705 | 415.3476 | 41.094 | 41.094 | 415.3476 | 249.8305 | 270.0257 | 270.0857 |  | 201.5471
[6 empty] | bellow | 40 | -0.058788 |  | 0.1035 | 249.8822 | 415.3476 | 41.094 | 41.094 | 415.3476 | 249.934 | 270.1374 | 270.1892 |  | 201.6588
FS1 | BMS | BPM | 2702 | FS1_BMS:BPM_D2702 | BPM2 | position monitor |  | 249.8752 | 249.934 | 0 | 249.934 | 415.3476 | 41.094 | 41.094 | 415.3476 | 249.934 | 270.1892 | 270.1892 |  | 201.7106
[6 empty] | diagnostic box |  |  |  | 0.14528 | 250.0066 | 415.3476 | 41.094 | 41.094 | 415.3476 | 250.0793 | 270.2618 | 270.3345 |  | 201.7832
FS1 | BMS | PM | 2703 | FS1_BMS:PM_D2703 | PM2 | profile monitor | TO BE CHECKED | defer installation |  | 0 | 250.0793 | 415.3476 | 41.094 | 41.094 | 415.3476 | 250.0793 | 270.3345 | 270.3345 |  | 201.8558
[6 empty] | diagnostic box |  | 250.0205 | 250.0793 | 0 | 250.0793 | 415.3476 | 41.094 | 41.094 | 415.3476 | 250.0793 | 270.3345 | 270.3345 |  | 201.8558
FS1 | BMS | IP | 2703 | FS1_BMS:IP_D2703 | IP_S | ion pump, 75 l/s |  | -0.058788 | 250.0793 | 0 | 250.0793 | 415.3476 | 41.094 | 41.094 | 415.3476 | 250.0793 | 270.3345 | 270.3345 |  | 201.8558
[10 empty] | 0 | 250.0793 | 415.3476 | 41.094 | 41.094 | 415.3476 | 250.0793 | 270.3345 | 270.3345 |  | 201.8558
[6 empty] | bellow | 40 |  |  | 0.13122 | 250.1449 | 415.3476 | 41.094 | 41.094 | 415.3476 | 250.2105 | 270.4001 | 270.4657 | adjustment for residuals | 201.9214
SEGMENT 2 START (12 b29 CMs +12 b54 CMs) | [9 empty] | -2e-07 | 250.2105 | 415.3476 | 41.094 | 41.094 | 415.3476 | 250.2105 | 270.4657 | 270.4657 |  | 201.9871
LS2 | CC01 | GV | 2705 | LS2_CC01:GV_D2705 |  | gate valve | 40 |  |  | 0.057404 | 250.2392 | 415.3476 | 41.094 | 41.094 | 415.3476 | 250.2679 | 270.4944 | 270.5231 |  | 202.0158
[6 empty] | bellows |  |  |  | 0.1677 | 250.3517 | 415.3476 | 41.094 | 41.094 | 415.3476 | 250.4356 | 270.6069 | 270.6908 |  | 202.1283
LS2 | CC01 | CAV1 | 2706 | LS2_CC01:CAV1_D2706 |  | b29 resonator | 40 | 0.32364 |  | 0.32364 | 250.5974 | 415.3476 | 41.094 | 41.094 | 415.3476 | 250.7592 | 270.8526 | 271.0144 |  | 202.374
LS2 | CC01 | BLM | 2709 | LS2_CC01:BLM_D2709 |  | outside loss monitor |  |  |  | 0 | 250.7592 | 415.3476 | 41.094 | 41.094 | 415.3476 | 250.7592 | 271.0144 | 271.0144 |  | 202.5358
[10 empty] | 0 | 250.7592 | 415.3476 | 41.094 | 41.094 | 415.3476 | 250.7592 | 271.0144 | 271.0144 |  | 202.5358
[6 empty] | bellow |  |  |  | 0.099517 | 250.809 | 415.3476 | 41.094 | 41.094 | 415.3476 | 250.8587 | 271.0642 | 271.114 |  | 202.5856
LS2 | CC01 | CAV2 | 2711 | LS2_CC01:CAV2_D2711 |  | b29 resonator | 40 |  |  | 0.32364 | 251.0206 | 415.3476 | 41.094 | 41.094 | 415.3476 | 251.1824 | 271.2758 | 271.4376 |  | 202.7971
[10 empty] | 0 | 251.1824 | 415.3476 | 41.094 | 41.094 | 415.3476 | 251.1824 | 271.4376 | 271.4376 |  | 202.959
[10 empty] | 0 | 251.1824 | 415.3476 | 41.094 | 41.094 | 415.3476 | 251.1824 | 271.4376 | 271.4376 |  | 202.959
[6 empty] | bellow |  |  |  | 0.099517 | 251.2321 | 415.3476 | 41.094 | 41.094 | 415.3476 | 251.2819 | 271.4874 | 271.5371 |  | 203.0087
LS2 | CC01 | CAV3 | 2715 | LS2_CC01:CAV3_D2715 |  | b29 resonator | 40 |  |  | 0.32364 | 251.4437 | 415.3476 | 41.094 | 41.094 | 415.3476 | 251.6055 | 271.6989 | 271.8608 |  | 203.2203
LS2 | CC01 | BLM | 2717 | LS2_CC01:BLM_D2717 |  | outside loss monitor |  |  |  | 0 | 251.6055 | 415.3476 | 41.094 | 41.094 | 415.3476 | 251.6055 | 271.8608 | 271.8608 |  | 203.3821
[10 empty] | 0 | 251.6055 | 415.3476 | 41.094 | 41.094 | 415.3476 | 251.6055 | 271.8608 | 271.8608 |  | 203.3821
[6 empty] | bellow+tube |  |  |  | 0.15525 | 251.6832 | 415.3476 | 41.094 | 41.094 | 415.3476 | 251.7608 | 271.9384 | 272.016 |  | 203.4598
LS2 | CC01 | SOL1 | 2723 | LS2_CC01:SOL1_D2723 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 252.0108 | 415.3476 | 41.094 | 41.094 | 415.3476 | 252.2608 | 272.266 | 272.516 |  | 203.7874
[6 empty] | bellow+tube |  |  |  | 0.15525 | 252.3384 | 415.3476 | 41.094 | 41.094 | 415.3476 | 252.416 | 272.5936 | 272.6713 |  | 204.115
LS2 | CC01 | CAV4 | 2726 | LS2_CC01:CAV4_D2726 |  | b29 resonator | 40 |  |  | 0.32364 | 252.5779 | 415.3476 | 41.094 | 41.094 | 415.3476 | 252.7397 | 272.8331 | 272.9949 |  | 204.3545
[10 empty] | 0 | 252.7397 | 415.3476 | 41.094 | 41.094 | 415.3476 | 252.7397 | 272.9949 | 272.9949 |  | 204.5163
[10 empty] | 0 | 252.7397 | 415.3476 | 41.094 | 41.094 | 415.3476 | 252.7397 | 272.9949 | 272.9949 |  | 204.5163
[10 empty] | 0.099517 | 252.7894 | 415.3476 | 41.094 | 41.094 | 415.3476 | 252.8392 | 273.0447 | 273.0944 |  | 204.566
LS2 | CC01 | CAV5 | 2730 | LS2_CC01:CAV5_D2730 |  | b29 resonator | 40 |  |  | 0.32364 | 253.001 | 415.3476 | 41.094 | 41.094 | 415.3476 | 253.1628 | 273.2562 | 273.4181 |  | 204.7776
[10 empty] | 0 | 253.1628 | 415.3476 | 41.094 | 41.094 | 415.3476 | 253.1628 | 273.4181 | 273.4181 |  | 204.9394
[10 empty] | 0 | 253.1628 | 415.3476 | 41.094 | 41.094 | 415.3476 | 253.1628 | 273.4181 | 273.4181 |  | 204.9394
[10 empty] | 0.099517 | 253.2126 | 415.3476 | 41.094 | 41.094 | 415.3476 | 253.2624 | 273.4678 | 273.5176 |  | 204.9892
LS2 | CC01 | CAV6 | 2735 | LS2_CC01:CAV6_D2735 |  | b29 resonator | 40 |  |  | 0.32364 | 253.4242 | 415.3476 | 41.094 | 41.094 | 415.3476 | 253.586 | 273.6794 | 273.8412 |  | 205.2008
LS2 | CC01 | BLM | 2737 | LS2_CC01:BLM_D2737 |  | outside loss monitor |  |  |  | 0 | 253.586 | 415.3476 | 41.094 | 41.094 | 415.3476 | 253.586 | 273.8412 | 273.8412 |  | 205.3626
[10 empty] | 0 | 253.586 | 415.3476 | 41.094 | 41.094 | 415.3476 | 253.586 | 273.8412 | 273.8412 |  | 205.3626
[6 empty] | bellows |  |  |  | 0.1677 | 253.6699 | 415.3476 | 41.094 | 41.094 | 415.3476 | 253.7537 | 273.9251 | 274.0089 |  | 205.4464
LS2 | CC01 | GV | 2740 | LS2_CC01:GV_D2740 |  | gate valve | 40 |  |  | 0.057404 | 253.7824 | 415.3476 | 41.094 | 41.094 | 415.3476 | 253.8111 | 274.0376 | 274.0663 |  | 205.559
[6 empty] | bellow | 40 |  |  | 0.1035 | 253.8629 | 415.3476 | 41.094 | 41.094 | 415.3476 | 253.9146 | 274.1181 | 274.1698 |  | 205.6395
LS2 | WC01 | BPM | 2742 | LS2_WC01:BPM_D2742 | BPM2 | position monitor |  |  |  | 0 | 253.9146 | 415.3476 | 41.094 | 41.094 | 415.3476 | 253.9146 | 274.1698 | 274.1698 |  | 205.6912
[6 empty] | diagnostic box |  |  |  | 0.14528 | 253.9873 | 415.3476 | 41.094 | 41.094 | 415.3476 | 254.0599 | 274.2425 | 274.3151 |  | 205.7638
LS2 | WC01 | PORT | 2743 | LS2_WC01:PORT_D2743 |  | port for future diagnostics |  |  |  | 0 | 254.0599 | 415.3476 | 41.094 | 41.094 | 415.3476 | 254.0599 | 274.3151 | 274.3151 |  | 205.8365
[6 empty] | diagnostic box |  |  |  | 0 | 254.0599 | 415.3476 | 41.094 | 41.094 | 415.3476 | 254.0599 | 274.3151 | 274.3151 |  | 205.8365
LS2 | WC01 | IP | 2743 | LS2_WC01:IP_D2743 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 254.0599 | 415.3476 | 41.094 | 41.094 | 415.3476 | 254.0599 | 274.3151 | 274.3151 |  | 205.8365
[6 empty] | bellow | 40 |  |  | 0.13122 | 254.1255 | 415.3476 | 41.094 | 41.094 | 415.3476 | 254.1911 | 274.3807 | 274.4463 |  | 205.9021
LS2 | CC02 | GV | 2745 | LS2_CC02:GV_D2745 |  | gate valve | 40 |  |  | 0.057404 | 254.2198 | 415.3476 | 41.094 | 41.094 | 415.3476 | 254.2485 | 274.475 | 274.5037 |  | 205.9964
[6 empty] | bellows |  |  |  | 0.1677 | 254.3324 | 415.3476 | 41.094 | 41.094 | 415.3476 | 254.4162 | 274.5876 | 274.6714 |  | 206.109
LS2 | CC02 | CAV1 | 2746 | LS2_CC02:CAV1_D2746 |  | b29 resonator | 40 | 0.32364 |  | 0.32364 | 254.578 | 415.3476 | 41.094 | 41.094 | 415.3476 | 254.7399 | 274.8333 | 274.9951 |  | 206.3546
LS2 | CC02 | BLM | 2748 | LS2_CC02:BLM_D2748 |  | outside loss monitor |  |  |  | 0 | 254.7399 | 415.3476 | 41.094 | 41.094 | 415.3476 | 254.7399 | 274.9951 | 274.9951 |  | 206.5165
[10 empty] | 0 | 254.7399 | 415.3476 | 41.094 | 41.094 | 415.3476 | 254.7399 | 274.9951 | 274.9951 |  | 206.5165
[6 empty] | bellow |  |  |  | 0.099517 | 254.7896 | 415.3476 | 41.094 | 41.094 | 415.3476 | 254.8394 | 275.0448 | 275.0946 |  | 206.5662
LS2 | CC02 | CAV2 | 2750 | LS2_CC02:CAV2_D2750 |  | b29 resonator | 40 |  |  | 0.32364 | 255.0012 | 415.3476 | 41.094 | 41.094 | 415.3476 | 255.163 | 275.2564 | 275.4182 |  | 206.7778
[10 empty] | 0 | 255.163 | 415.3476 | 41.094 | 41.094 | 415.3476 | 255.163 | 275.4182 | 275.4182 |  | 206.9396
[10 empty] | 0 | 255.163 | 415.3476 | 41.094 | 41.094 | 415.3476 | 255.163 | 275.4182 | 275.4182 |  | 206.9396
[6 empty] | bellow |  |  |  | 0.099517 | 255.2128 | 415.3476 | 41.094 | 41.094 | 415.3476 | 255.2625 | 275.468 | 275.5178 |  | 206.9894
LS2 | CC02 | CAV3 | 2755 | LS2_CC02:CAV3_D2755 |  | b29 resonator | 40 |  |  | 0.32364 | 255.4244 | 415.3476 | 41.094 | 41.094 | 415.3476 | 255.5862 | 275.6796 | 275.8414 |  | 207.201
LS2 | CC02 | BLM | 2757 | LS2_CC02:BLM_D2757 |  | outside loss monitor |  |  |  | 0 | 255.5862 | 415.3476 | 41.094 | 41.094 | 415.3476 | 255.5862 | 275.8414 | 275.8414 |  | 207.3628
[10 empty] | 0 | 255.5862 | 415.3476 | 41.094 | 41.094 | 415.3476 | 255.5862 | 275.8414 | 275.8414 |  | 207.3628
[6 empty] | bellow+tube |  |  |  | 0.15525 | 255.6638 | 415.3476 | 41.094 | 41.094 | 415.3476 | 255.7414 | 275.919 | 275.9967 |  | 207.4404
LS2 | CC02 | SOL1 | 2762 | LS2_CC02:SOL1_D2762 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 255.9914 | 415.3476 | 41.094 | 41.094 | 415.3476 | 256.2414 | 276.2467 | 276.4967 |  | 207.768
[6 empty] | bellow+tube |  |  |  | 0.15525 | 256.3191 | 415.3476 | 41.094 | 41.094 | 415.3476 | 256.3967 | 276.5743 | 276.6519 |  | 208.0957
LS2 | CC02 | CAV4 | 2766 | LS2_CC02:CAV4_D2766 |  | b29 resonator | 40 |  |  | 0.32364 | 256.5585 | 415.3476 | 41.094 | 41.094 | 415.3476 | 256.7203 | 276.8137 | 276.9756 |  | 208.3351
[10 empty] | 0 | 256.7203 | 415.3476 | 41.094 | 41.094 | 415.3476 | 256.7203 | 276.9756 | 276.9756 |  | 208.4969
[10 empty] | 0 | 256.7203 | 415.3476 | 41.094 | 41.094 | 415.3476 | 256.7203 | 276.9756 | 276.9756 |  | 208.4969
[10 empty] | 0.099517 | 256.7701 | 415.3476 | 41.094 | 41.094 | 415.3476 | 256.8199 | 277.0253 | 277.0751 |  | 208.5467
LS2 | CC02 | CAV5 | 2770 | LS2_CC02:CAV5_D2770 |  | b29 resonator | 40 |  |  | 0.32364 | 256.9817 | 415.3476 | 41.094 | 41.094 | 415.3476 | 257.1435 | 277.2369 | 277.3987 |  | 208.7583
[10 empty] | 0 | 257.1435 | 415.3476 | 41.094 | 41.094 | 415.3476 | 257.1435 | 277.3987 | 277.3987 |  | 208.9201
[10 empty] | 0 | 257.1435 | 415.3476 | 41.094 | 41.094 | 415.3476 | 257.1435 | 277.3987 | 277.3987 |  | 208.9201
[10 empty] | 0.099517 | 257.1933 | 415.3476 | 41.094 | 41.094 | 415.3476 | 257.243 | 277.4485 | 277.4982 |  | 208.9698
LS2 | CC02 | CAV6 | 2774 | LS2_CC02:CAV6_D2774 |  | b29 resonator | 40 |  |  | 0.32364 | 257.4048 | 415.3476 | 41.094 | 41.094 | 415.3476 | 257.5667 | 277.6601 | 277.8219 |  | 209.1814
LS2 | CC02 | BLM | 2777 | LS2_CC02:BLM_D2777 |  | outside loss monitor |  |  |  | 0 | 257.5667 | 415.3476 | 41.094 | 41.094 | 415.3476 | 257.5667 | 277.8219 | 277.8219 |  | 209.3432
[10 empty] | 0 | 257.5667 | 415.3476 | 41.094 | 41.094 | 415.3476 | 257.5667 | 277.8219 | 277.8219 |  | 209.3432
[6 empty] | bellows |  |  |  | 0.1677 | 257.6505 | 415.3476 | 41.094 | 41.094 | 415.3476 | 257.7344 | 277.9057 | 277.9896 |  | 209.4271
LS2 | CC02 | GV | 2780 | LS2_CC02:GV_D2780 |  | gate valve | 40 |  |  | 0.057404 | 257.7631 | 415.3476 | 41.094 | 41.094 | 415.3476 | 257.7918 | 278.0183 | 278.047 |  | 209.5396
[6 empty] | bellow | 40 |  |  | 0.1035 | 257.8435 | 415.3476 | 41.094 | 41.094 | 415.3476 | 257.8953 | 278.0987 | 278.1505 |  | 209.6201
LS2 | WC02 | BPM | 2782 | LS2_WC02:BPM_D2782 | BPM2 | position monitor |  |  |  | 0 | 257.8953 | 415.3476 | 41.094 | 41.094 | 415.3476 | 257.8953 | 278.1505 | 278.1505 |  | 209.6719
[6 empty] | diagnostic box |  |  |  | 0.14528 | 257.9679 | 415.3476 | 41.094 | 41.094 | 415.3476 | 258.0405 | 278.2231 | 278.2958 |  | 209.7445
LS2 | WC02 | PORT | 2783 | LS2_WC02:PORT_D2783 |  | port for future diagnostics |  |  |  | 0 | 258.0405 | 415.3476 | 41.094 | 41.094 | 415.3476 | 258.0405 | 278.2958 | 278.2958 |  | 209.8171
[6 empty] | diagnostic box |  |  |  | 0 | 258.0405 | 415.3476 | 41.094 | 41.094 | 415.3476 | 258.0405 | 278.2958 | 278.2958 |  | 209.8171
LS2 | WC02 | IP | 2783 | LS2_WC02:IP_D2783 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 258.0405 | 415.3476 | 41.094 | 41.094 | 415.3476 | 258.0405 | 278.2958 | 278.2958 |  | 209.8171
[6 empty] | bellow | 40 |  |  | 0.13122 | 258.1062 | 415.3476 | 41.094 | 41.094 | 415.3476 | 258.1718 | 278.3614 | 278.427 |  | 209.8827
LS2 | CC03 | GV | 2785 | LS2_CC03:GV_D2785 |  | gate valve | 40 |  |  | 0.057404 | 258.2005 | 415.3476 | 41.094 | 41.094 | 415.3476 | 258.2292 | 278.4557 | 278.4844 |  | 209.9771
[6 empty] | bellows |  |  |  | 0.1677 | 258.313 | 415.3476 | 41.094 | 41.094 | 415.3476 | 258.3969 | 278.5682 | 278.6521 |  | 210.0896
LS2 | CC03 | CAV1 | 2786 | LS2_CC03:CAV1_D2786 |  | b29 resonator | 40 | 0.32364 |  | 0.32364 | 258.5587 | 415.3476 | 41.094 | 41.094 | 415.3476 | 258.7205 | 278.8139 | 278.9757 |  | 210.3353
LS2 | CC03 | BLM | 2788 | LS2_CC03:BLM_D2788 |  | outside loss monitor |  |  |  | 0 | 258.7205 | 415.3476 | 41.094 | 41.094 | 415.3476 | 258.7205 | 278.9757 | 278.9757 |  | 210.4971
[10 empty] | 0 | 258.7205 | 415.3476 | 41.094 | 41.094 | 415.3476 | 258.7205 | 278.9757 | 278.9757 |  | 210.4971
[6 empty] | bellow |  |  |  | 0.099517 | 258.7703 | 415.3476 | 41.094 | 41.094 | 415.3476 | 258.82 | 279.0255 | 279.0752 |  | 210.5469
LS2 | CC03 | CAV2 | 2790 | LS2_CC03:CAV2_D2790 |  | b29 resonator | 40 |  |  | 0.32364 | 258.9819 | 415.3476 | 41.094 | 41.094 | 415.3476 | 259.1437 | 279.2371 | 279.3989 |  | 210.7584
[10 empty] | 0 | 259.1437 | 415.3476 | 41.094 | 41.094 | 415.3476 | 259.1437 | 279.3989 | 279.3989 |  | 210.9203
[10 empty] | 0 | 259.1437 | 415.3476 | 41.094 | 41.094 | 415.3476 | 259.1437 | 279.3989 | 279.3989 |  | 210.9203
[6 empty] | bellow |  |  |  | 0.099517 | 259.1934 | 415.3476 | 41.094 | 41.094 | 415.3476 | 259.2432 | 279.4487 | 279.4984 |  | 210.97
LS2 | CC03 | CAV3 | 2794 | LS2_CC03:CAV3_D2794 |  | b29 resonator | 40 |  |  | 0.32364 | 259.405 | 415.3476 | 41.094 | 41.094 | 415.3476 | 259.5668 | 279.6602 | 279.8221 |  | 211.1816
LS2 | CC03 | BLM | 2797 | LS2_CC03:BLM_D2797 |  | outside loss monitor |  |  |  | 0 | 259.5668 | 415.3476 | 41.094 | 41.094 | 415.3476 | 259.5668 | 279.8221 | 279.8221 |  | 211.3434
[10 empty] | 0 | 259.5668 | 415.3476 | 41.094 | 41.094 | 415.3476 | 259.5668 | 279.8221 | 279.8221 |  | 211.3434
[6 empty] | bellow+tube |  |  |  | 0.15525 | 259.6445 | 415.3476 | 41.094 | 41.094 | 415.3476 | 259.7221 | 279.8997 | 279.9773 |  | 211.4211
LS2 | CC03 | SOL1 | 2802 | LS2_CC03:SOL1_D2802 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 259.9721 | 415.3476 | 41.094 | 41.094 | 415.3476 | 260.2221 | 280.2273 | 280.4773 |  | 211.7487
[6 empty] | bellow+tube |  |  |  | 0.15525 | 260.2997 | 415.3476 | 41.094 | 41.094 | 415.3476 | 260.3773 | 280.5549 | 280.6326 |  | 212.0763
LS2 | CC03 | CAV4 | 2806 | LS2_CC03:CAV4_D2806 |  | b29 resonator | 40 |  |  | 0.32364 | 260.5392 | 415.3476 | 41.094 | 41.094 | 415.3476 | 260.701 | 280.7944 | 280.9562 |  | 212.3157
[10 empty] | 0 | 260.701 | 415.3476 | 41.094 | 41.094 | 415.3476 | 260.701 | 280.9562 | 280.9562 |  | 212.4776
[10 empty] | 0 | 260.701 | 415.3476 | 41.094 | 41.094 | 415.3476 | 260.701 | 280.9562 | 280.9562 |  | 212.4776
[10 empty] | 0.099517 | 260.7507 | 415.3476 | 41.094 | 41.094 | 415.3476 | 260.8005 | 281.006 | 281.0557 |  | 212.5273
LS2 | CC03 | CAV5 | 2810 | LS2_CC03:CAV5_D2810 |  | b29 resonator | 40 |  |  | 0.32364 | 260.9623 | 415.3476 | 41.094 | 41.094 | 415.3476 | 261.1241 | 281.2175 | 281.3794 |  | 212.7389
[10 empty] | 0 | 261.1241 | 415.3476 | 41.094 | 41.094 | 415.3476 | 261.1241 | 281.3794 | 281.3794 |  | 212.9007
[10 empty] | 0 | 261.1241 | 415.3476 | 41.094 | 41.094 | 415.3476 | 261.1241 | 281.3794 | 281.3794 |  | 212.9007
[10 empty] | 0.099517 | 261.1739 | 415.3476 | 41.094 | 41.094 | 415.3476 | 261.2237 | 281.4291 | 281.4789 |  | 212.9505
LS2 | CC03 | CAV6 | 2814 | LS2_CC03:CAV6_D2814 |  | b29 resonator | 40 |  |  | 0.32364 | 261.3855 | 415.3476 | 41.094 | 41.094 | 415.3476 | 261.5473 | 281.6407 | 281.8025 |  | 213.1621
LS2 | CC03 | BLM | 2816 | LS2_CC03:BLM_D2816 |  | outside loss monitor |  |  |  | 0 | 261.5473 | 415.3476 | 41.094 | 41.094 | 415.3476 | 261.5473 | 281.8025 | 281.8025 |  | 213.3239
[10 empty] | 0 | 261.5473 | 415.3476 | 41.094 | 41.094 | 415.3476 | 261.5473 | 281.8025 | 281.8025 |  | 213.3239
[6 empty] | bellows |  |  |  | 0.1677 | 261.6312 | 415.3476 | 41.094 | 41.094 | 415.3476 | 261.715 | 281.8864 | 281.9702 |  | 213.4077
LS2 | CC03 | GV | 2820 | LS2_CC03:GV_D2820 |  | gate valve | 40 |  |  | 0.057404 | 261.7437 | 415.3476 | 41.094 | 41.094 | 415.3476 | 261.7724 | 281.9989 | 282.0276 |  | 213.5203
[6 empty] | bellow | 40 |  |  | 0.1035 | 261.8242 | 415.3476 | 41.094 | 41.094 | 415.3476 | 261.8759 | 282.0794 | 282.1311 |  | 213.6008
LS2 | WC03 | BPM | 2821 | LS2_WC03:BPM_D2821 | BPM2 | position monitor |  |  |  | 0 | 261.8759 | 415.3476 | 41.094 | 41.094 | 415.3476 | 261.8759 | 282.1311 | 282.1311 |  | 213.6525
[6 empty] | diagnostic box |  |  |  | 0.14528 | 261.9486 | 415.3476 | 41.094 | 41.094 | 415.3476 | 262.0212 | 282.2038 | 282.2764 |  | 213.7251
LS2 | WC03 | PORT | 2823 | LS2_WC03:PORT_D2823 |  | port for future diagnostics |  |  |  | 0 | 262.0212 | 415.3476 | 41.094 | 41.094 | 415.3476 | 262.0212 | 282.2764 | 282.2764 |  | 213.7978
[6 empty] | diagnostic box |  |  |  | 0 | 262.0212 | 415.3476 | 41.094 | 41.094 | 415.3476 | 262.0212 | 282.2764 | 282.2764 |  | 213.7978
LS2 | WC03 | IP | 2823 | LS2_WC03:IP_D2823 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 262.0212 | 415.3476 | 41.094 | 41.094 | 415.3476 | 262.0212 | 282.2764 | 282.2764 |  | 213.7978
[6 empty] | bellow | 40 |  |  | 0.13122 | 262.0868 | 415.3476 | 41.094 | 41.094 | 415.3476 | 262.1524 | 282.342 | 282.4076 |  | 213.8634
LS2 | CC04 | GV | 2824 | LS2_CC04:GV_D2824 |  | gate valve | 40 |  |  | 0.057404 | 262.1811 | 415.3476 | 41.094 | 41.094 | 415.3476 | 262.2098 | 282.4363 | 282.465 |  | 213.9577
[6 empty] | bellows |  |  |  | 0.1677 | 262.2937 | 415.3476 | 41.094 | 41.094 | 415.3476 | 262.3775 | 282.5489 | 282.6327 |  | 214.0703
LS2 | CC04 | CAV1 | 2825 | LS2_CC04:CAV1_D2825 |  | b29 resonator | 40 | 0.32364 |  | 0.32364 | 262.5393 | 415.3476 | 41.094 | 41.094 | 415.3476 | 262.7012 | 282.7946 | 282.9564 |  | 214.3159
LS2 | CC04 | BLM | 2828 | LS2_CC04:BLM_D2828 |  | outside loss monitor |  |  |  | 0 | 262.7012 | 415.3476 | 41.094 | 41.094 | 415.3476 | 262.7012 | 282.9564 | 282.9564 |  | 214.4778
[10 empty] | 0 | 262.7012 | 415.3476 | 41.094 | 41.094 | 415.3476 | 262.7012 | 282.9564 | 282.9564 |  | 214.4778
[6 empty] | bellow |  |  |  | 0.099517 | 262.7509 | 415.3476 | 41.094 | 41.094 | 415.3476 | 262.8007 | 283.0061 | 283.0559 |  | 214.5275
LS2 | CC04 | CAV2 | 2830 | LS2_CC04:CAV2_D2830 |  | b29 resonator | 40 |  |  | 0.32364 | 262.9625 | 415.3476 | 41.094 | 41.094 | 415.3476 | 263.1243 | 283.2177 | 283.3795 |  | 214.7391
[10 empty] | 0 | 263.1243 | 415.3476 | 41.094 | 41.094 | 415.3476 | 263.1243 | 283.3795 | 283.3795 |  | 214.9009
[10 empty] | 0 | 263.1243 | 415.3476 | 41.094 | 41.094 | 415.3476 | 263.1243 | 283.3795 | 283.3795 |  | 214.9009
[6 empty] | bellow |  |  |  | 0.099517 | 263.1741 | 415.3476 | 41.094 | 41.094 | 415.3476 | 263.2238 | 283.4293 | 283.4791 |  | 214.9507
LS2 | CC04 | CAV3 | 2834 | LS2_CC04:CAV3_D2834 |  | b29 resonator | 40 |  |  | 0.32364 | 263.3857 | 415.3476 | 41.094 | 41.094 | 415.3476 | 263.5475 | 283.6409 | 283.8027 |  | 215.1623
LS2 | CC04 | BLM | 2836 | LS2_CC04:BLM_D2836 |  | outside loss monitor |  |  |  | 0 | 263.5475 | 415.3476 | 41.094 | 41.094 | 415.3476 | 263.5475 | 283.8027 | 283.8027 |  | 215.3241
[10 empty] | 0 | 263.5475 | 415.3476 | 41.094 | 41.094 | 415.3476 | 263.5475 | 283.8027 | 283.8027 |  | 215.3241
[6 empty] | bellow+tube |  |  |  | 0.15525 | 263.6251 | 415.3476 | 41.094 | 41.094 | 415.3476 | 263.7027 | 283.8803 | 283.958 |  | 215.4017
LS2 | CC04 | SOL1 | 2842 | LS2_CC04:SOL1_D2842 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 263.9527 | 415.3476 | 41.094 | 41.094 | 415.3476 | 264.2027 | 284.208 | 284.458 |  | 215.7293
[6 empty] | bellow+tube |  |  |  | 0.15525 | 264.2804 | 415.3476 | 41.094 | 41.094 | 415.3476 | 264.358 | 284.5356 | 284.6132 |  | 216.0569
LS2 | CC04 | CAV4 | 2845 | LS2_CC04:CAV4_D2845 |  | b29 resonator | 40 |  |  | 0.32364 | 264.5198 | 415.3476 | 41.094 | 41.094 | 415.3476 | 264.6816 | 284.775 | 284.9368 |  | 216.2964
[10 empty] | 0 | 264.6816 | 415.3476 | 41.094 | 41.094 | 415.3476 | 264.6816 | 284.9368 | 284.9368 |  | 216.4582
[10 empty] | 0 | 264.6816 | 415.3476 | 41.094 | 41.094 | 415.3476 | 264.6816 | 284.9368 | 284.9368 |  | 216.4582
[10 empty] | 0.099517 | 264.7314 | 415.3476 | 41.094 | 41.094 | 415.3476 | 264.7811 | 284.9866 | 285.0364 |  | 216.508
LS2 | CC04 | CAV5 | 2850 | LS2_CC04:CAV5_D2850 |  | b29 resonator | 40 |  |  | 0.32364 | 264.943 | 415.3476 | 41.094 | 41.094 | 415.3476 | 265.1048 | 285.1982 | 285.36 |  | 216.7196
[10 empty] | 0 | 265.1048 | 415.3476 | 41.094 | 41.094 | 415.3476 | 265.1048 | 285.36 | 285.36 |  | 216.8814
[10 empty] | 0 | 265.1048 | 415.3476 | 41.094 | 41.094 | 415.3476 | 265.1048 | 285.36 | 285.36 |  | 216.8814
[10 empty] | 0.099517 | 265.1546 | 415.3476 | 41.094 | 41.094 | 415.3476 | 265.2043 | 285.4098 | 285.4595 |  | 216.9311
LS2 | CC04 | CAV6 | 2854 | LS2_CC04:CAV6_D2854 |  | b29 resonator | 40 |  |  | 0.32364 | 265.3661 | 415.3476 | 41.094 | 41.094 | 415.3476 | 265.528 | 285.6213 | 285.7832 |  | 217.1427
LS2 | CC04 | BLM | 2856 | LS2_CC04:BLM_D2856 |  | outside loss monitor |  |  |  | 0 | 265.528 | 415.3476 | 41.094 | 41.094 | 415.3476 | 265.528 | 285.7832 | 285.7832 |  | 217.3045
[10 empty] | 0 | 265.528 | 415.3476 | 41.094 | 41.094 | 415.3476 | 265.528 | 285.7832 | 285.7832 |  | 217.3045
[6 empty] | bellows |  |  |  | 0.1677 | 265.6118 | 415.3476 | 41.094 | 41.094 | 415.3476 | 265.6957 | 285.867 | 285.9509 |  | 217.3884
LS2 | CC04 | GV | 2860 | LS2_CC04:GV_D2860 |  | gate valve | 40 |  |  | 0.057404 | 265.7244 | 415.3476 | 41.094 | 41.094 | 415.3476 | 265.7531 | 285.9796 | 286.0083 |  | 217.5009
[6 empty] | bellow | 40 |  |  | 0.1035 | 265.8048 | 415.3476 | 41.094 | 41.094 | 415.3476 | 265.8566 | 286.06 | 286.1118 |  | 217.5814
LS2 | WC04 | BPM | 2861 | LS2_WC04:BPM_D2861 | BPM2 | position monitor |  |  |  | 0 | 265.8566 | 415.3476 | 41.094 | 41.094 | 415.3476 | 265.8566 | 286.1118 | 286.1118 |  | 217.6331
[6 empty] | diagnostic box |  |  |  | 0.14528 | 265.9292 | 415.3476 | 41.094 | 41.094 | 415.3476 | 266.0018 | 286.1844 | 286.2571 |  | 217.7058
LS2 | WC04 | PORT | 2863 | LS2_WC04:PORT_D2863 |  | port for future diagnostics |  |  |  | 0 | 266.0018 | 415.3476 | 41.094 | 41.094 | 415.3476 | 266.0018 | 286.2571 | 286.2571 |  | 217.7784
[6 empty] | diagnostic box |  |  |  | 0 | 266.0018 | 415.3476 | 41.094 | 41.094 | 415.3476 | 266.0018 | 286.2571 | 286.2571 |  | 217.7784
LS2 | WC04 | IP | 2863 | LS2_WC04:IP_D2863 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 266.0018 | 415.3476 | 41.094 | 41.094 | 415.3476 | 266.0018 | 286.2571 | 286.2571 |  | 217.7784
[6 empty] | bellow | 40 |  |  | 0.13122 | 266.0675 | 415.3476 | 41.094 | 41.094 | 415.3476 | 266.1331 | 286.3227 | 286.3883 |  | 217.844
LS2 | CC05 | GV | 2864 | LS2_CC05:GV_D2864 |  | gate valve | 40 |  |  | 0.057404 | 266.1618 | 415.3476 | 41.094 | 41.094 | 415.3476 | 266.1905 | 286.417 | 286.4457 |  | 217.9383
[6 empty] | bellows |  |  |  | 0.1677 | 266.2743 | 415.3476 | 41.094 | 41.094 | 415.3476 | 266.3582 | 286.5295 | 286.6134 |  | 218.0509
LS2 | CC05 | CAV1 | 2865 | LS2_CC05:CAV1_D2865 |  | b29 resonator | 40 | 0.32364 |  | 0.32364 | 266.52 | 415.3476 | 41.094 | 41.094 | 415.3476 | 266.6818 | 286.7752 | 286.937 |  | 218.2966
LS2 | CC05 | BLM | 2868 | LS2_CC05:BLM_D2868 |  | outside loss monitor |  |  |  | 0 | 266.6818 | 415.3476 | 41.094 | 41.094 | 415.3476 | 266.6818 | 286.937 | 286.937 |  | 218.4584
[10 empty] | 0 | 266.6818 | 415.3476 | 41.094 | 41.094 | 415.3476 | 266.6818 | 286.937 | 286.937 |  | 218.4584
[6 empty] | bellow |  |  |  | 0.099517 | 266.7316 | 415.3476 | 41.094 | 41.094 | 415.3476 | 266.7813 | 286.9868 | 287.0365 |  | 218.5082
LS2 | CC05 | CAV2 | 2870 | LS2_CC05:CAV2_D2870 |  | b29 resonator | 40 |  |  | 0.32364 | 266.9432 | 415.3476 | 41.094 | 41.094 | 415.3476 | 267.105 | 287.1984 | 287.3602 |  | 218.7197
[10 empty] | 0 | 267.105 | 415.3476 | 41.094 | 41.094 | 415.3476 | 267.105 | 287.3602 | 287.3602 |  | 218.8816
[10 empty] | 0 | 267.105 | 415.3476 | 41.094 | 41.094 | 415.3476 | 267.105 | 287.3602 | 287.3602 |  | 218.8816
[6 empty] | bellow |  |  |  | 0.099517 | 267.1547 | 415.3476 | 41.094 | 41.094 | 415.3476 | 267.2045 | 287.4099 | 287.4597 |  | 218.9313
LS2 | CC05 | CAV3 | 2874 | LS2_CC05:CAV3_D2874 |  | b29 resonator | 40 |  |  | 0.32364 | 267.3663 | 415.3476 | 41.094 | 41.094 | 415.3476 | 267.5281 | 287.6215 | 287.7834 |  | 219.1429
LS2 | CC05 | BLM | 2876 | LS2_CC05:BLM_D2876 |  | outside loss monitor |  |  |  | 0 | 267.5281 | 415.3476 | 41.094 | 41.094 | 415.3476 | 267.5281 | 287.7834 | 287.7834 |  | 219.3047
[10 empty] | 0 | 267.5281 | 415.3476 | 41.094 | 41.094 | 415.3476 | 267.5281 | 287.7834 | 287.7834 |  | 219.3047
[6 empty] | bellow+tube |  |  |  | 0.15525 | 267.6058 | 415.3476 | 41.094 | 41.094 | 415.3476 | 267.6834 | 287.861 | 287.9386 |  | 219.3823
LS2 | CC05 | SOL1 | 2882 | LS2_CC05:SOL1_D2882 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 267.9334 | 415.3476 | 41.094 | 41.094 | 415.3476 | 268.1834 | 288.1886 | 288.4386 |  | 219.71
[6 empty] | bellow+tube |  |  |  | 0.15525 | 268.261 | 415.3476 | 41.094 | 41.094 | 415.3476 | 268.3386 | 288.5162 | 288.5939 |  | 220.0376
LS2 | CC05 | CAV4 | 2885 | LS2_CC05:CAV4_D2885 |  | b29 resonator | 40 |  |  | 0.32364 | 268.5005 | 415.3476 | 41.094 | 41.094 | 415.3476 | 268.6623 | 288.7557 | 288.9175 |  | 220.277
[10 empty] | 0 | 268.6623 | 415.3476 | 41.094 | 41.094 | 415.3476 | 268.6623 | 288.9175 | 288.9175 |  | 220.4389
[10 empty] | 0 | 268.6623 | 415.3476 | 41.094 | 41.094 | 415.3476 | 268.6623 | 288.9175 | 288.9175 |  | 220.4389
[10 empty] | 0.099517 | 268.712 | 415.3476 | 41.094 | 41.094 | 415.3476 | 268.7618 | 288.9673 | 289.017 |  | 220.4886
LS2 | CC05 | CAV5 | 2890 | LS2_CC05:CAV5_D2890 |  | b29 resonator | 40 |  |  | 0.32364 | 268.9236 | 415.3476 | 41.094 | 41.094 | 415.3476 | 269.0854 | 289.1788 | 289.3407 |  | 220.7002
[10 empty] | 0 | 269.0854 | 415.3476 | 41.094 | 41.094 | 415.3476 | 269.0854 | 289.3407 | 289.3407 |  | 220.862
[10 empty] | 0 | 269.0854 | 415.3476 | 41.094 | 41.094 | 415.3476 | 269.0854 | 289.3407 | 289.3407 |  | 220.862
[10 empty] | 0.099517 | 269.1352 | 415.3476 | 41.094 | 41.094 | 415.3476 | 269.185 | 289.3904 | 289.4402 |  | 220.9118
LS2 | CC05 | CAV6 | 2894 | LS2_CC05:CAV6_D2894 |  | b29 resonator | 40 |  |  | 0.32364 | 269.3468 | 415.3476 | 41.094 | 41.094 | 415.3476 | 269.5086 | 289.602 | 289.7638 |  | 221.1234
LS2 | CC05 | BLM | 2896 | LS2_CC05:BLM_D2896 |  | outside loss monitor |  |  |  | 0 | 269.5086 | 415.3476 | 41.094 | 41.094 | 415.3476 | 269.5086 | 289.7638 | 289.7638 |  | 221.2852
[10 empty] | 0 | 269.5086 | 415.3476 | 41.094 | 41.094 | 415.3476 | 269.5086 | 289.7638 | 289.7638 |  | 221.2852
[6 empty] | bellows |  |  |  | 0.1677 | 269.5925 | 415.3476 | 41.094 | 41.094 | 415.3476 | 269.6763 | 289.8477 | 289.9315 |  | 221.369
LS2 | CC05 | GV | 2900 | LS2_CC05:GV_D2900 |  | gate valve | 40 |  |  | 0.057404 | 269.705 | 415.3476 | 41.094 | 41.094 | 415.3476 | 269.7337 | 289.9602 | 289.9889 |  | 221.4816
[6 empty] | bellow | 40 |  |  | 0.1035 | 269.7855 | 415.3476 | 41.094 | 41.094 | 415.3476 | 269.8372 | 290.0407 | 290.0924 |  | 221.562
LS2 | WC05 | BPM | 2901 | LS2_WC05:BPM_D2901 | BPM2 | position monitor |  |  |  | 0 | 269.8372 | 415.3476 | 41.094 | 41.094 | 415.3476 | 269.8372 | 290.0924 | 290.0924 |  | 221.6138
[6 empty] | diagnostic box |  |  |  | 0.14528 | 269.9099 | 415.3476 | 41.094 | 41.094 | 415.3476 | 269.9825 | 290.1651 | 290.2377 |  | 221.6864
LS2 | WC05 | PORT | 2902 | LS2_WC05:PORT_D2902 |  | port for future diagnostics |  |  |  | 0 | 269.9825 | 415.3476 | 41.094 | 41.094 | 415.3476 | 269.9825 | 290.2377 | 290.2377 |  | 221.7591
[6 empty] | diagnostic box |  |  |  | 0 | 269.9825 | 415.3476 | 41.094 | 41.094 | 415.3476 | 269.9825 | 290.2377 | 290.2377 |  | 221.7591
LS2 | WC05 | IP | 2902 | LS2_WC05:IP_D2902 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 269.9825 | 415.3476 | 41.094 | 41.094 | 415.3476 | 269.9825 | 290.2377 | 290.2377 |  | 221.7591
[6 empty] | bellow | 40 |  |  | 0.13122 | 270.0481 | 415.3476 | 41.094 | 41.094 | 415.3476 | 270.1137 | 290.3033 | 290.3689 |  | 221.8247
LS2 | CC06 | GV | 2904 | LS2_CC06:GV_D2904 |  | gate valve | 40 |  |  | 0.057404 | 270.1424 | 415.3476 | 41.094 | 41.094 | 415.3476 | 270.1711 | 290.3976 | 290.4263 |  | 221.919
[6 empty] | bellows |  |  |  | 0.1677 | 270.255 | 415.3476 | 41.094 | 41.094 | 415.3476 | 270.3388 | 290.5102 | 290.594 |  | 222.0316
LS2 | CC06 | CAV1 | 2905 | LS2_CC06:CAV1_D2905 |  | b29 resonator | 40 | 0.32364 |  | 0.32364 | 270.5006 | 415.3476 | 41.094 | 41.094 | 415.3476 | 270.6625 | 290.7559 | 290.9177 |  | 222.2772
LS2 | CC06 | BLM | 2908 | LS2_CC06:BLM_D2908 |  | outside loss monitor |  |  |  | 0 | 270.6625 | 415.3476 | 41.094 | 41.094 | 415.3476 | 270.6625 | 290.9177 | 290.9177 |  | 222.439
[10 empty] | 0 | 270.6625 | 415.3476 | 41.094 | 41.094 | 415.3476 | 270.6625 | 290.9177 | 290.9177 |  | 222.439
[6 empty] | bellow |  |  |  | 0.099517 | 270.7122 | 415.3476 | 41.094 | 41.094 | 415.3476 | 270.762 | 290.9674 | 291.0172 |  | 222.4888
LS2 | CC06 | CAV2 | 2910 | LS2_CC06:CAV2_D2910 |  | b29 resonator | 40 |  |  | 0.32364 | 270.9238 | 415.3476 | 41.094 | 41.094 | 415.3476 | 271.0856 | 291.179 | 291.3408 |  | 222.7004
[10 empty] | 0 | 271.0856 | 415.3476 | 41.094 | 41.094 | 415.3476 | 271.0856 | 291.3408 | 291.3408 |  | 222.8622
[10 empty] | 0 | 271.0856 | 415.3476 | 41.094 | 41.094 | 415.3476 | 271.0856 | 291.3408 | 291.3408 |  | 222.8622
[6 empty] | bellow |  |  |  | 0.099517 | 271.1354 | 415.3476 | 41.094 | 41.094 | 415.3476 | 271.1851 | 291.3906 | 291.4404 |  | 222.912
LS2 | CC06 | CAV3 | 2914 | LS2_CC06:CAV3_D2914 |  | b29 resonator | 40 |  |  | 0.32364 | 271.347 | 415.3476 | 41.094 | 41.094 | 415.3476 | 271.5088 | 291.6022 | 291.764 |  | 223.1235
LS2 | CC06 | BLM | 2916 | LS2_CC06:BLM_D2916 |  | outside loss monitor |  |  |  | 0 | 271.5088 | 415.3476 | 41.094 | 41.094 | 415.3476 | 271.5088 | 291.764 | 291.764 |  | 223.2854
[10 empty] | 0 | 271.5088 | 415.3476 | 41.094 | 41.094 | 415.3476 | 271.5088 | 291.764 | 291.764 |  | 223.2854
[6 empty] | bellow+tube |  |  |  | 0.15525 | 271.5864 | 415.3476 | 41.094 | 41.094 | 415.3476 | 271.664 | 291.8416 | 291.9192 |  | 223.363
LS2 | CC06 | SOL1 | 2922 | LS2_CC06:SOL1_D2922 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 271.914 | 415.3476 | 41.094 | 41.094 | 415.3476 | 272.164 | 292.1692 | 292.4192 |  | 223.6906
[6 empty] | bellow+tube |  |  |  | 0.15525 | 272.2417 | 415.3476 | 41.094 | 41.094 | 415.3476 | 272.3193 | 292.4969 | 292.5745 |  | 224.0182
LS2 | CC06 | CAV4 | 2925 | LS2_CC06:CAV4_D2925 |  | b29 resonator | 40 |  |  | 0.32364 | 272.4811 | 415.3476 | 41.094 | 41.094 | 415.3476 | 272.6429 | 292.7363 | 292.8981 |  | 224.2577
[10 empty] | 0 | 272.6429 | 415.3476 | 41.094 | 41.094 | 415.3476 | 272.6429 | 292.8981 | 292.8981 |  | 224.4195
[10 empty] | 0 | 272.6429 | 415.3476 | 41.094 | 41.094 | 415.3476 | 272.6429 | 292.8981 | 292.8981 |  | 224.4195
[10 empty] | 0.099517 | 272.6927 | 415.3476 | 41.094 | 41.094 | 415.3476 | 272.7424 | 292.9479 | 292.9977 |  | 224.4693
LS2 | CC06 | CAV5 | 2929 | LS2_CC06:CAV5_D2929 |  | b29 resonator | 40 |  |  | 0.32364 | 272.9043 | 415.3476 | 41.094 | 41.094 | 415.3476 | 273.0661 | 293.1595 | 293.3213 |  | 224.6809
[10 empty] | 0 | 273.0661 | 415.3476 | 41.094 | 41.094 | 415.3476 | 273.0661 | 293.3213 | 293.3213 |  | 224.8427
[10 empty] | 0 | 273.0661 | 415.3476 | 41.094 | 41.094 | 415.3476 | 273.0661 | 293.3213 | 293.3213 |  | 224.8427
[10 empty] | 0.099517 | 273.1158 | 415.3476 | 41.094 | 41.094 | 415.3476 | 273.1656 | 293.3711 | 293.4208 |  | 224.8924
LS2 | CC06 | CAV6 | 2934 | LS2_CC06:CAV6_D2934 |  | b29 resonator | 40 |  |  | 0.32364 | 273.3274 | 415.3476 | 41.094 | 41.094 | 415.3476 | 273.4892 | 293.5826 | 293.7445 |  | 225.104
LS2 | CC06 | BLM | 2936 | LS2_CC06:BLM_D2936 |  | outside loss monitor |  |  |  | 0 | 273.4892 | 415.3476 | 41.094 | 41.094 | 415.3476 | 273.4892 | 293.7445 | 293.7445 |  | 225.2658
[10 empty] | 0 | 273.4892 | 415.3476 | 41.094 | 41.094 | 415.3476 | 273.4892 | 293.7445 | 293.7445 |  | 225.2658
[6 empty] | bellows |  |  |  | 0.1677 | 273.5731 | 415.3476 | 41.094 | 41.094 | 415.3476 | 273.657 | 293.8283 | 293.9122 |  | 225.3497
LS2 | CC06 | GV | 2939 | LS2_CC06:GV_D2939 |  | gate valve | 40 |  |  | 0.057404 | 273.6857 | 415.3476 | 41.094 | 41.094 | 415.3476 | 273.7144 | 293.9409 | 293.9696 |  | 225.4622
[6 empty] | bellow | 40 |  |  | 0.1035 | 273.7661 | 415.3476 | 41.094 | 41.094 | 415.3476 | 273.8179 | 294.0213 | 294.0731 |  | 225.5427
LS2 | WC06 | BPM | 2941 | LS2_WC06:BPM_D2941 | BPM2 | position monitor |  |  |  | 0 | 273.8179 | 415.3476 | 41.094 | 41.094 | 415.3476 | 273.8179 | 294.0731 | 294.0731 |  | 225.5944
[6 empty] | diagnostic box |  |  |  | 0.14528 | 273.8905 | 415.3476 | 41.094 | 41.094 | 415.3476 | 273.9631 | 294.1457 | 294.2184 |  | 225.6671
LS2 | WC06 | PORT | 2942 | LS2_WC06:PORT_D2942 |  | port for future diagnostics |  |  |  | 0 | 273.9631 | 415.3476 | 41.094 | 41.094 | 415.3476 | 273.9631 | 294.2184 | 294.2184 |  | 225.7397
[6 empty] | diagnostic box |  |  |  | 0 | 273.9631 | 415.3476 | 41.094 | 41.094 | 415.3476 | 273.9631 | 294.2184 | 294.2184 |  | 225.7397
LS2 | WC06 | IP | 2942 | LS2_WC06:IP_D2942 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 273.9631 | 415.3476 | 41.094 | 41.094 | 415.3476 | 273.9631 | 294.2184 | 294.2184 |  | 225.7397
[6 empty] | bellow | 40 |  |  | 0.13122 | 274.0287 | 415.3476 | 41.094 | 41.094 | 415.3476 | 274.0944 | 294.284 | 294.3496 |  | 225.8053
LS2 | CC07 | GV | 2944 | LS2_CC07:GV_D2944 |  | gate valve | 40 |  |  | 0.057404 | 274.1231 | 415.3476 | 41.094 | 41.094 | 415.3476 | 274.1518 | 294.3783 | 294.407 |  | 225.8996
[6 empty] | bellows |  |  |  | 0.1677 | 274.2356 | 415.3476 | 41.094 | 41.094 | 415.3476 | 274.3195 | 294.4908 | 294.5747 |  | 226.0122
LS2 | CC07 | CAV1 | 2945 | LS2_CC07:CAV1_D2945 |  | b29 resonator | 40 | 0.32364 |  | 0.32364 | 274.4813 | 415.3476 | 41.094 | 41.094 | 415.3476 | 274.6431 | 294.7365 | 294.8983 |  | 226.2579
LS2 | CC07 | BLM | 2947 | LS2_CC07:BLM_D2947 |  | outside loss monitor |  |  |  | 0 | 274.6431 | 415.3476 | 41.094 | 41.094 | 415.3476 | 274.6431 | 294.8983 | 294.8983 |  | 226.4197
[10 empty] | 0 | 274.6431 | 415.3476 | 41.094 | 41.094 | 415.3476 | 274.6431 | 294.8983 | 294.8983 |  | 226.4197
[6 empty] | bellow |  |  |  | 0.099517 | 274.6929 | 415.3476 | 41.094 | 41.094 | 415.3476 | 274.7426 | 294.9481 | 294.9978 |  | 226.4695
LS2 | CC07 | CAV2 | 2949 | LS2_CC07:CAV2_D2949 |  | b29 resonator | 40 |  |  | 0.32364 | 274.9044 | 415.3476 | 41.094 | 41.094 | 415.3476 | 275.0663 | 295.1597 | 295.3215 |  | 226.681
[10 empty] | 0 | 275.0663 | 415.3476 | 41.094 | 41.094 | 415.3476 | 275.0663 | 295.3215 | 295.3215 |  | 226.8429
[10 empty] | 0 | 275.0663 | 415.3476 | 41.094 | 41.094 | 415.3476 | 275.0663 | 295.3215 | 295.3215 |  | 226.8429
[6 empty] | bellow |  |  |  | 0.099517 | 275.116 | 415.3476 | 41.094 | 41.094 | 415.3476 | 275.1658 | 295.3712 | 295.421 |  | 226.8926
LS2 | CC07 | CAV3 | 2954 | LS2_CC07:CAV3_D2954 |  | b29 resonator | 40 |  |  | 0.32364 | 275.3276 | 415.3476 | 41.094 | 41.094 | 415.3476 | 275.4894 | 295.5828 | 295.7446 |  | 227.1042
LS2 | CC07 | BLM | 2956 | LS2_CC07:BLM_D2956 |  | outside loss monitor |  |  |  | 0 | 275.4894 | 415.3476 | 41.094 | 41.094 | 415.3476 | 275.4894 | 295.7446 | 295.7446 |  | 227.266
[10 empty] | 0 | 275.4894 | 415.3476 | 41.094 | 41.094 | 415.3476 | 275.4894 | 295.7446 | 295.7446 |  | 227.266
[6 empty] | bellow+tube |  |  |  | 0.15525 | 275.5671 | 415.3476 | 41.094 | 41.094 | 415.3476 | 275.6447 | 295.8223 | 295.8999 |  | 227.3436
LS2 | CC07 | SOL1 | 2961 | LS2_CC07:SOL1_D2961 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 275.8947 | 415.3476 | 41.094 | 41.094 | 415.3476 | 276.1447 | 296.1499 | 296.3999 |  | 227.6713
[6 empty] | bellow+tube |  |  |  | 0.15525 | 276.2223 | 415.3476 | 41.094 | 41.094 | 415.3476 | 276.2999 | 296.4775 | 296.5551 |  | 227.9989
LS2 | CC07 | CAV4 | 2965 | LS2_CC07:CAV4_D2965 |  | b29 resonator | 40 |  |  | 0.32364 | 276.4618 | 415.3476 | 41.094 | 41.094 | 415.3476 | 276.6236 | 296.717 | 296.8788 |  | 228.2383
[10 empty] | 0 | 276.6236 | 415.3476 | 41.094 | 41.094 | 415.3476 | 276.6236 | 296.8788 | 296.8788 |  | 228.4002
[10 empty] | 0 | 276.6236 | 415.3476 | 41.094 | 41.094 | 415.3476 | 276.6236 | 296.8788 | 296.8788 |  | 228.4002
[10 empty] | 0.099517 | 276.6733 | 415.3476 | 41.094 | 41.094 | 415.3476 | 276.7231 | 296.9285 | 296.9783 |  | 228.4499
LS2 | CC07 | CAV5 | 2969 | LS2_CC07:CAV5_D2969 |  | b29 resonator | 40 |  |  | 0.32364 | 276.8849 | 415.3476 | 41.094 | 41.094 | 415.3476 | 277.0467 | 297.1401 | 297.302 |  | 228.6615
[10 empty] | 0 | 277.0467 | 415.3476 | 41.094 | 41.094 | 415.3476 | 277.0467 | 297.302 | 297.302 |  | 228.8233
[10 empty] | 0 | 277.0467 | 415.3476 | 41.094 | 41.094 | 415.3476 | 277.0467 | 297.302 | 297.302 |  | 228.8233
[10 empty] | 0.099517 | 277.0965 | 415.3476 | 41.094 | 41.094 | 415.3476 | 277.1463 | 297.3517 | 297.4015 |  | 228.8731
LS2 | CC07 | CAV6 | 2974 | LS2_CC07:CAV6_D2974 |  | b29 resonator | 40 |  |  | 0.32364 | 277.3081 | 415.3476 | 41.094 | 41.094 | 415.3476 | 277.4699 | 297.5633 | 297.7251 |  | 229.0847
LS2 | CC07 | BLM | 2976 | LS2_CC07:BLM_D2976 |  | outside loss monitor |  |  |  | 0 | 277.4699 | 415.3476 | 41.094 | 41.094 | 415.3476 | 277.4699 | 297.7251 | 297.7251 |  | 229.2465
[10 empty] | 0 | 277.4699 | 415.3476 | 41.094 | 41.094 | 415.3476 | 277.4699 | 297.7251 | 297.7251 |  | 229.2465
[6 empty] | bellows |  |  |  | 0.1677 | 277.5537 | 415.3476 | 41.094 | 41.094 | 415.3476 | 277.6376 | 297.809 | 297.8928 |  | 229.3303
LS2 | CC07 | GV | 2979 | LS2_CC07:GV_D2979 |  | gate valve | 40 |  |  | 0.057404 | 277.6663 | 415.3476 | 41.094 | 41.094 | 415.3476 | 277.695 | 297.9215 | 297.9502 |  | 229.4429
[6 empty] | bellow | 40 |  |  | 0.1035 | 277.7468 | 415.3476 | 41.094 | 41.094 | 415.3476 | 277.7985 | 298.002 | 298.0537 |  | 229.5233
LS2 | WC07 | BPM | 2981 | LS2_WC07:BPM_D2981 | BPM2 | position monitor |  |  |  | 0 | 277.7985 | 415.3476 | 41.094 | 41.094 | 415.3476 | 277.7985 | 298.0537 | 298.0537 |  | 229.5751
[6 empty] | diagnostic box |  |  |  | 0.14528 | 277.8711 | 415.3476 | 41.094 | 41.094 | 415.3476 | 277.9438 | 298.1264 | 298.199 |  | 229.6477
LS2 | WC07 | PORT | 2982 | LS2_WC07:PORT_D2982 |  | port for future diagnostics |  |  |  | 0 | 277.9438 | 415.3476 | 41.094 | 41.094 | 415.3476 | 277.9438 | 298.199 | 298.199 |  | 229.7204
[6 empty] | diagnostic box |  |  |  | 0 | 277.9438 | 415.3476 | 41.094 | 41.094 | 415.3476 | 277.9438 | 298.199 | 298.199 |  | 229.7204
LS2 | WC07 | IP | 2982 | LS2_WC07:IP_D2982 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 277.9438 | 415.3476 | 41.094 | 41.094 | 415.3476 | 277.9438 | 298.199 | 298.199 |  | 229.7204
[6 empty] | bellow | 40 |  |  | 0.13122 | 278.0094 | 415.3476 | 41.094 | 41.094 | 415.3476 | 278.075 | 298.2646 | 298.3302 |  | 229.786
LS2 | CC08 | GV | 2984 | LS2_CC08:GV_D2984 |  | gate valve | 40 |  |  | 0.057404 | 278.1037 | 415.3476 | 41.094 | 41.094 | 415.3476 | 278.1324 | 298.3589 | 298.3876 |  | 229.8803
[6 empty] | bellows |  |  |  | 0.1677 | 278.2163 | 415.3476 | 41.094 | 41.094 | 415.3476 | 278.3001 | 298.4715 | 298.5553 |  | 229.9928
LS2 | CC08 | CAV1 | 2985 | LS2_CC08:CAV1_D2985 |  | b29 resonator | 40 | 0.32364 |  | 0.32364 | 278.4619 | 415.3476 | 41.094 | 41.094 | 415.3476 | 278.6238 | 298.7172 | 298.879 |  | 230.2385
LS2 | CC08 | BLM | 2987 | LS2_CC08:BLM_D2987 |  | outside loss monitor |  |  |  | 0 | 278.6238 | 415.3476 | 41.094 | 41.094 | 415.3476 | 278.6238 | 298.879 | 298.879 |  | 230.4003
[10 empty] | 0 | 278.6238 | 415.3476 | 41.094 | 41.094 | 415.3476 | 278.6238 | 298.879 | 298.879 |  | 230.4003
[6 empty] | bellow |  |  |  | 0.099517 | 278.6735 | 415.3476 | 41.094 | 41.094 | 415.3476 | 278.7233 | 298.9287 | 298.9785 |  | 230.4501
LS2 | CC08 | CAV2 | 2989 | LS2_CC08:CAV2_D2989 |  | b29 resonator | 40 |  |  | 0.32364 | 278.8851 | 415.3476 | 41.094 | 41.094 | 415.3476 | 279.0469 | 299.1403 | 299.3021 |  | 230.6617
[10 empty] | 0 | 279.0469 | 415.3476 | 41.094 | 41.094 | 415.3476 | 279.0469 | 299.3021 | 299.3021 |  | 230.8235
[10 empty] | 0 | 279.0469 | 415.3476 | 41.094 | 41.094 | 415.3476 | 279.0469 | 299.3021 | 299.3021 |  | 230.8235
[6 empty] | bellow |  |  |  | 0.099517 | 279.0967 | 415.3476 | 41.094 | 41.094 | 415.3476 | 279.1464 | 299.3519 | 299.4017 |  | 230.8733
LS2 | CC08 | CAV3 | 2994 | LS2_CC08:CAV3_D2994 |  | b29 resonator | 40 |  |  | 0.32364 | 279.3083 | 415.3476 | 41.094 | 41.094 | 415.3476 | 279.4701 | 299.5635 | 299.7253 |  | 231.0848
LS2 | CC08 | BLM | 2996 | LS2_CC08:BLM_D2996 |  | outside loss monitor |  |  |  | 0 | 279.4701 | 415.3476 | 41.094 | 41.094 | 415.3476 | 279.4701 | 299.7253 | 299.7253 |  | 231.2467
[10 empty] | 0 | 279.4701 | 415.3476 | 41.094 | 41.094 | 415.3476 | 279.4701 | 299.7253 | 299.7253 |  | 231.2467
[6 empty] | bellow+tube |  |  |  | 0.15525 | 279.5477 | 415.3476 | 41.094 | 41.094 | 415.3476 | 279.6253 | 299.8029 | 299.8805 |  | 231.3243
LS2 | CC08 | SOL1 | 3001 | LS2_CC08:SOL1_D3001 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 279.8753 | 415.3476 | 41.094 | 41.094 | 415.3476 | 280.1253 | 300.1305 | 300.3805 |  | 231.6519
[6 empty] | bellow+tube |  |  |  | 0.15525 | 280.203 | 415.3476 | 41.094 | 41.094 | 415.3476 | 280.2806 | 300.4582 | 300.5358 |  | 231.9795
LS2 | CC08 | CAV4 | 3005 | LS2_CC08:CAV4_D3005 |  | b29 resonator | 40 |  |  | 0.32364 | 280.4424 | 415.3476 | 41.094 | 41.094 | 415.3476 | 280.6042 | 300.6976 | 300.8594 |  | 232.219
[10 empty] | 0 | 280.6042 | 415.3476 | 41.094 | 41.094 | 415.3476 | 280.6042 | 300.8594 | 300.8594 |  | 232.3808
[10 empty] | 0 | 280.6042 | 415.3476 | 41.094 | 41.094 | 415.3476 | 280.6042 | 300.8594 | 300.8594 |  | 232.3808
[10 empty] | 0.099517 | 280.654 | 415.3476 | 41.094 | 41.094 | 415.3476 | 280.7037 | 300.9092 | 300.959 |  | 232.4306
LS2 | CC08 | CAV5 | 3009 | LS2_CC08:CAV5_D3009 |  | b29 resonator | 40 |  |  | 0.32364 | 280.8656 | 415.3476 | 41.094 | 41.094 | 415.3476 | 281.0274 | 301.1208 | 301.2826 |  | 232.6421
[10 empty] | 0 | 281.0274 | 415.3476 | 41.094 | 41.094 | 415.3476 | 281.0274 | 301.2826 | 301.2826 |  | 232.804
[10 empty] | 0 | 281.0274 | 415.3476 | 41.094 | 41.094 | 415.3476 | 281.0274 | 301.2826 | 301.2826 |  | 232.804
[10 empty] | 0.099517 | 281.0771 | 415.3476 | 41.094 | 41.094 | 415.3476 | 281.1269 | 301.3324 | 301.3821 |  | 232.8537
LS2 | CC08 | CAV6 | 3013 | LS2_CC08:CAV6_D3013 |  | b29 resonator | 40 |  |  | 0.32364 | 281.2887 | 415.3476 | 41.094 | 41.094 | 415.3476 | 281.4505 | 301.5439 | 301.7058 |  | 233.0653
LS2 | CC08 | BLM | 3015 | LS2_CC08:BLM_D3015 |  | outside loss monitor |  |  |  | 0 | 281.4505 | 415.3476 | 41.094 | 41.094 | 415.3476 | 281.4505 | 301.7058 | 301.7058 |  | 233.2271
[10 empty] | 0 | 281.4505 | 415.3476 | 41.094 | 41.094 | 415.3476 | 281.4505 | 301.7058 | 301.7058 |  | 233.2271
[6 empty] | bellows |  |  |  | 0.1677 | 281.5344 | 415.3476 | 41.094 | 41.094 | 415.3476 | 281.6182 | 301.7896 | 301.8735 |  | 233.311
LS2 | CC08 | GV | 3019 | LS2_CC08:GV_D3019 |  | gate valve | 40 |  |  | 0.057404 | 281.647 | 415.3476 | 41.094 | 41.094 | 415.3476 | 281.6757 | 301.9022 | 301.9309 |  | 233.4235
[6 empty] | bellow | 40 |  |  | 0.1035 | 281.7274 | 415.3476 | 41.094 | 41.094 | 415.3476 | 281.7792 | 301.9826 | 302.0344 |  | 233.504
LS2 | WC08 | BPM | 3020 | LS2_WC08:BPM_D3020 | BPM2 | position monitor |  |  |  | 0 | 281.7792 | 415.3476 | 41.094 | 41.094 | 415.3476 | 281.7792 | 302.0344 | 302.0344 |  | 233.5557
[6 empty] | diagnostic box |  |  |  | 0.14528 | 281.8518 | 415.3476 | 41.094 | 41.094 | 415.3476 | 281.9244 | 302.107 | 302.1797 |  | 233.6284
LS2 | WC08 | PORT | 3022 | LS2_WC08:PORT_D3022 |  | port for future diagnostics |  |  |  | 0 | 281.9244 | 415.3476 | 41.094 | 41.094 | 415.3476 | 281.9244 | 302.1797 | 302.1797 |  | 233.701
[6 empty] | diagnostic box |  |  |  | 0 | 281.9244 | 415.3476 | 41.094 | 41.094 | 415.3476 | 281.9244 | 302.1797 | 302.1797 |  | 233.701
LS2 | WC08 | IP | 3022 | LS2_WC08:IP_D3022 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 281.9244 | 415.3476 | 41.094 | 41.094 | 415.3476 | 281.9244 | 302.1797 | 302.1797 |  | 233.701
[6 empty] | bellow | 40 |  |  | 0.13122 | 281.99 | 415.3476 | 41.094 | 41.094 | 415.3476 | 282.0557 | 302.2453 | 302.3109 |  | 233.7666
LS2 | CC09 | GV | 3023 | LS2_CC09:GV_D3023 |  | gate valve | 40 |  |  | 0.057404 | 282.0844 | 415.3476 | 41.094 | 41.094 | 415.3476 | 282.1131 | 302.3396 | 302.3683 |  | 233.8609
[6 empty] | bellows |  |  |  | 0.1677 | 282.1969 | 415.3476 | 41.094 | 41.094 | 415.3476 | 282.2808 | 302.4521 | 302.536 |  | 233.9735
LS2 | CC09 | CAV1 | 3025 | LS2_CC09:CAV1_D3025 |  | b29 resonator | 40 | 0.32364 |  | 0.32364 | 282.4426 | 415.3476 | 41.094 | 41.094 | 415.3476 | 282.6044 | 302.6978 | 302.8596 |  | 234.2192
LS2 | CC09 | BLM | 3027 | LS2_CC09:BLM_D3027 |  | outside loss monitor |  |  |  | 0 | 282.6044 | 415.3476 | 41.094 | 41.094 | 415.3476 | 282.6044 | 302.8596 | 302.8596 |  | 234.381
[10 empty] | 0 | 282.6044 | 415.3476 | 41.094 | 41.094 | 415.3476 | 282.6044 | 302.8596 | 302.8596 |  | 234.381
[6 empty] | bellow |  |  |  | 0.099517 | 282.6542 | 415.3476 | 41.094 | 41.094 | 415.3476 | 282.7039 | 302.9094 | 302.9591 |  | 234.4307
LS2 | CC09 | CAV2 | 3029 | LS2_CC09:CAV2_D3029 |  | b29 resonator | 40 |  |  | 0.32364 | 282.8657 | 415.3476 | 41.094 | 41.094 | 415.3476 | 283.0276 | 303.121 | 303.2828 |  | 234.6423
[10 empty] | 0 | 283.0276 | 415.3476 | 41.094 | 41.094 | 415.3476 | 283.0276 | 303.2828 | 303.2828 |  | 234.8042
[10 empty] | 0 | 283.0276 | 415.3476 | 41.094 | 41.094 | 415.3476 | 283.0276 | 303.2828 | 303.2828 |  | 234.8042
[6 empty] | bellow |  |  |  | 0.099517 | 283.0773 | 415.3476 | 41.094 | 41.094 | 415.3476 | 283.1271 | 303.3325 | 303.3823 |  | 234.8539
LS2 | CC09 | CAV3 | 3033 | LS2_CC09:CAV3_D3033 |  | b29 resonator | 40 |  |  | 0.32364 | 283.2889 | 415.3476 | 41.094 | 41.094 | 415.3476 | 283.4507 | 303.5441 | 303.7059 |  | 235.0655
LS2 | CC09 | BLM | 3035 | LS2_CC09:BLM_D3035 |  | outside loss monitor |  |  |  | 0 | 283.4507 | 415.3476 | 41.094 | 41.094 | 415.3476 | 283.4507 | 303.7059 | 303.7059 |  | 235.2273
[10 empty] | 0 | 283.4507 | 415.3476 | 41.094 | 41.094 | 415.3476 | 283.4507 | 303.7059 | 303.7059 |  | 235.2273
[6 empty] | bellow+tube |  |  |  | 0.15525 | 283.5284 | 415.3476 | 41.094 | 41.094 | 415.3476 | 283.606 | 303.7836 | 303.8612 |  | 235.3049
LS2 | CC09 | SOL1 | 3041 | LS2_CC09:SOL1_D3041 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 283.856 | 415.3476 | 41.094 | 41.094 | 415.3476 | 284.106 | 304.1112 | 304.3612 |  | 235.6326
[6 empty] | bellow+tube |  |  |  | 0.15525 | 284.1836 | 415.3476 | 41.094 | 41.094 | 415.3476 | 284.2612 | 304.4388 | 304.5164 |  | 235.9602
LS2 | CC09 | CAV4 | 3044 | LS2_CC09:CAV4_D3044 |  | b29 resonator | 40 |  |  | 0.32364 | 284.423 | 415.3476 | 41.094 | 41.094 | 415.3476 | 284.5849 | 304.6783 | 304.8401 |  | 236.1996
[10 empty] | 0 | 284.5849 | 415.3476 | 41.094 | 41.094 | 415.3476 | 284.5849 | 304.8401 | 304.8401 |  | 236.3615
[10 empty] | 0 | 284.5849 | 415.3476 | 41.094 | 41.094 | 415.3476 | 284.5849 | 304.8401 | 304.8401 |  | 236.3615
[10 empty] | 0.099517 | 284.6346 | 415.3476 | 41.094 | 41.094 | 415.3476 | 284.6844 | 304.8898 | 304.9396 |  | 236.4112
LS2 | CC09 | CAV5 | 3049 | LS2_CC09:CAV5_D3049 |  | b29 resonator | 40 |  |  | 0.32364 | 284.8462 | 415.3476 | 41.094 | 41.094 | 415.3476 | 285.008 | 305.1014 | 305.2632 |  | 236.6228
[10 empty] | 0 | 285.008 | 415.3476 | 41.094 | 41.094 | 415.3476 | 285.008 | 305.2632 | 305.2632 |  | 236.7846
[10 empty] | 0 | 285.008 | 415.3476 | 41.094 | 41.094 | 415.3476 | 285.008 | 305.2632 | 305.2632 |  | 236.7846
[10 empty] | 0.099517 | 285.0578 | 415.3476 | 41.094 | 41.094 | 415.3476 | 285.1075 | 305.313 | 305.3628 |  | 236.8344
LS2 | CC09 | CAV6 | 3053 | LS2_CC09:CAV6_D3053 |  | b29 resonator | 40 |  |  | 0.32364 | 285.2694 | 415.3476 | 41.094 | 41.094 | 415.3476 | 285.4312 | 305.5246 | 305.6864 |  | 237.046
LS2 | CC09 | BLM | 3055 | LS2_CC09:BLM_D3055 |  | outside loss monitor |  |  |  | 0 | 285.4312 | 415.3476 | 41.094 | 41.094 | 415.3476 | 285.4312 | 305.6864 | 305.6864 |  | 237.2078
[10 empty] | 0 | 285.4312 | 415.3476 | 41.094 | 41.094 | 415.3476 | 285.4312 | 305.6864 | 305.6864 |  | 237.2078
[6 empty] | bellows |  |  |  | 0.1677 | 285.515 | 415.3476 | 41.094 | 41.094 | 415.3476 | 285.5989 | 305.7703 | 305.8541 |  | 237.2916
LS2 | CC09 | GV | 3059 | LS2_CC09:GV_D3059 |  | gate valve | 40 |  |  | 0.057404 | 285.6276 | 415.3476 | 41.094 | 41.094 | 415.3476 | 285.6563 | 305.8828 | 305.9115 |  | 237.4042
[6 empty] | bellow | 40 |  |  | 0.1035 | 285.7081 | 415.3476 | 41.094 | 41.094 | 415.3476 | 285.7598 | 305.9633 | 306.015 |  | 237.4846
LS2 | WC09 | BPM | 3060 | LS2_WC09:BPM_D3060 | BPM2 | position monitor |  |  |  | 0 | 285.7598 | 415.3476 | 41.094 | 41.094 | 415.3476 | 285.7598 | 306.015 | 306.015 |  | 237.5364
[6 empty] | diagnostic box |  |  |  | 0.14528 | 285.8324 | 415.3476 | 41.094 | 41.094 | 415.3476 | 285.9051 | 306.0877 | 306.1603 |  | 237.609
LS2 | WC09 | PORT | 3062 | LS2_WC09:PORT_D3062 |  | port for future diagnostics |  |  |  | 0 | 285.9051 | 415.3476 | 41.094 | 41.094 | 415.3476 | 285.9051 | 306.1603 | 306.1603 |  | 237.6817
[6 empty] | diagnostic box |  |  |  | 0 | 285.9051 | 415.3476 | 41.094 | 41.094 | 415.3476 | 285.9051 | 306.1603 | 306.1603 |  | 237.6817
LS2 | WC09 | IP | 3062 | LS2_WC09:IP_D3062 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 285.9051 | 415.3476 | 41.094 | 41.094 | 415.3476 | 285.9051 | 306.1603 | 306.1603 |  | 237.6817
[6 empty] | bellow | 40 |  |  | 0.13122 | 285.9707 | 415.3476 | 41.094 | 41.094 | 415.3476 | 286.0363 | 306.2259 | 306.2915 |  | 237.7473
LS2 | CC10 | GV | 3063 | LS2_CC10:GV_D3063 |  | gate valve | 40 |  |  | 0.057404 | 286.065 | 415.3476 | 41.094 | 41.094 | 415.3476 | 286.0937 | 306.3202 | 306.3489 |  | 237.8416
[6 empty] | bellows |  |  |  | 0.1677 | 286.1776 | 415.3476 | 41.094 | 41.094 | 415.3476 | 286.2614 | 306.4328 | 306.5166 |  | 237.9541
LS2 | CC10 | CAV1 | 3064 | LS2_CC10:CAV1_D3064 |  | b29 resonator | 40 | 0.32364 |  | 0.32364 | 286.4232 | 415.3476 | 41.094 | 41.094 | 415.3476 | 286.5851 | 306.6784 | 306.8403 |  | 238.1998
LS2 | CC10 | BLM | 3067 | LS2_CC10:BLM_D3067 |  | outside loss monitor |  |  |  | 0 | 286.5851 | 415.3476 | 41.094 | 41.094 | 415.3476 | 286.5851 | 306.8403 | 306.8403 |  | 238.3616
[10 empty] | 0 | 286.5851 | 415.3476 | 41.094 | 41.094 | 415.3476 | 286.5851 | 306.8403 | 306.8403 |  | 238.3616
[6 empty] | bellow |  |  |  | 0.099517 | 286.6348 | 415.3476 | 41.094 | 41.094 | 415.3476 | 286.6846 | 306.89 | 306.9398 |  | 238.4114
LS2 | CC10 | CAV2 | 3069 | LS2_CC10:CAV2_D3069 |  | b29 resonator | 40 |  |  | 0.32364 | 286.8464 | 415.3476 | 41.094 | 41.094 | 415.3476 | 287.0082 | 307.1016 | 307.2634 |  | 238.623
[10 empty] | 0 | 287.0082 | 415.3476 | 41.094 | 41.094 | 415.3476 | 287.0082 | 307.2634 | 307.2634 |  | 238.7848
[10 empty] | 0 | 287.0082 | 415.3476 | 41.094 | 41.094 | 415.3476 | 287.0082 | 307.2634 | 307.2634 |  | 238.7848
[6 empty] | bellow |  |  |  | 0.099517 | 287.058 | 415.3476 | 41.094 | 41.094 | 415.3476 | 287.1077 | 307.3132 | 307.3629 |  | 238.8346
LS2 | CC10 | CAV3 | 3073 | LS2_CC10:CAV3_D3073 |  | b29 resonator | 40 |  |  | 0.32364 | 287.2696 | 415.3476 | 41.094 | 41.094 | 415.3476 | 287.4314 | 307.5248 | 307.6866 |  | 239.0461
LS2 | CC10 | BLM | 3075 | LS2_CC10:BLM_D3075 |  | outside loss monitor |  |  |  | 0 | 287.4314 | 415.3476 | 41.094 | 41.094 | 415.3476 | 287.4314 | 307.6866 | 307.6866 |  | 239.208
[10 empty] | 0 | 287.4314 | 415.3476 | 41.094 | 41.094 | 415.3476 | 287.4314 | 307.6866 | 307.6866 |  | 239.208
[6 empty] | bellow+tube |  |  |  | 0.15525 | 287.509 | 415.3476 | 41.094 | 41.094 | 415.3476 | 287.5866 | 307.7642 | 307.8418 |  | 239.2856
LS2 | CC10 | SOL1 | 3081 | LS2_CC10:SOL1_D3081 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 287.8366 | 415.3476 | 41.094 | 41.094 | 415.3476 | 288.0866 | 308.0918 | 308.3418 |  | 239.6132
[6 empty] | bellow+tube |  |  |  | 0.15525 | 288.1642 | 415.3476 | 41.094 | 41.094 | 415.3476 | 288.2419 | 308.4195 | 308.4971 |  | 239.9408
LS2 | CC10 | CAV4 | 3084 | LS2_CC10:CAV4_D3084 |  | b29 resonator | 40 |  |  | 0.32364 | 288.4037 | 415.3476 | 41.094 | 41.094 | 415.3476 | 288.5655 | 308.6589 | 308.8207 |  | 240.1803
[10 empty] | 0 | 288.5655 | 415.3476 | 41.094 | 41.094 | 415.3476 | 288.5655 | 308.8207 | 308.8207 |  | 240.3421
[10 empty] | 0 | 288.5655 | 415.3476 | 41.094 | 41.094 | 415.3476 | 288.5655 | 308.8207 | 308.8207 |  | 240.3421
[10 empty] | 0.099517 | 288.6153 | 415.3476 | 41.094 | 41.094 | 415.3476 | 288.665 | 308.8705 | 308.9203 |  | 240.3919
LS2 | CC10 | CAV5 | 3089 | LS2_CC10:CAV5_D3089 |  | b29 resonator | 40 |  |  | 0.32364 | 288.8269 | 415.3476 | 41.094 | 41.094 | 415.3476 | 288.9887 | 309.0821 | 309.2439 |  | 240.6034
[10 empty] | 0 | 288.9887 | 415.3476 | 41.094 | 41.094 | 415.3476 | 288.9887 | 309.2439 | 309.2439 |  | 240.7653
[10 empty] | 0 | 288.9887 | 415.3476 | 41.094 | 41.094 | 415.3476 | 288.9887 | 309.2439 | 309.2439 |  | 240.7653
[10 empty] | 0.099517 | 289.0384 | 415.3476 | 41.094 | 41.094 | 415.3476 | 289.0882 | 309.2937 | 309.3434 |  | 240.815
LS2 | CC10 | CAV6 | 3093 | LS2_CC10:CAV6_D3093 |  | b29 resonator | 40 |  |  | 0.32364 | 289.25 | 415.3476 | 41.094 | 41.094 | 415.3476 | 289.4118 | 309.5052 | 309.6671 |  | 241.0266
LS2 | CC10 | BLM | 3095 | LS2_CC10:BLM_D3095 |  | outside loss monitor |  |  |  | 0 | 289.4118 | 415.3476 | 41.094 | 41.094 | 415.3476 | 289.4118 | 309.6671 | 309.6671 |  | 241.1884
[10 empty] | 0 | 289.4118 | 415.3476 | 41.094 | 41.094 | 415.3476 | 289.4118 | 309.6671 | 309.6671 |  | 241.1884
[6 empty] | bellows |  |  |  | 0.1677 | 289.4957 | 415.3476 | 41.094 | 41.094 | 415.3476 | 289.5795 | 309.7509 | 309.8348 |  | 241.2723
LS2 | CC10 | GV | 3099 | LS2_CC10:GV_D3099 |  | gate valve | 40 |  |  | 0.057404 | 289.6082 | 415.3476 | 41.094 | 41.094 | 415.3476 | 289.6369 | 309.8635 | 309.8922 |  | 241.3848
[6 empty] | bellow | 40 |  |  | 0.1035 | 289.6887 | 415.3476 | 41.094 | 41.094 | 415.3476 | 289.7405 | 309.9439 | 309.9957 |  | 241.4653
LS2 | WC10 | BPM | 3100 | LS2_WC10:BPM_D3100 | BPM2 | position monitor |  |  |  | 0 | 289.7405 | 415.3476 | 41.094 | 41.094 | 415.3476 | 289.7405 | 309.9957 | 309.9957 |  | 241.517
[6 empty] | diagnostic box |  |  |  | 0.14528 | 289.8131 | 415.3476 | 41.094 | 41.094 | 415.3476 | 289.8857 | 310.0683 | 310.1409 |  | 241.5897
LS2 | WC10 | PORT | 3101 | LS2_WC10:PORT_D3101 |  | port for future diagnostics |  |  |  | 0 | 289.8857 | 415.3476 | 41.094 | 41.094 | 415.3476 | 289.8857 | 310.1409 | 310.1409 |  | 241.6623
[6 empty] | diagnostic box |  |  |  | 0 | 289.8857 | 415.3476 | 41.094 | 41.094 | 415.3476 | 289.8857 | 310.1409 | 310.1409 |  | 241.6623
LS2 | WC10 | IP | 3101 | LS2_WC10:IP_D3101 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 289.8857 | 415.3476 | 41.094 | 41.094 | 415.3476 | 289.8857 | 310.1409 | 310.1409 |  | 241.6623
[6 empty] | bellow | 40 |  |  | 0.13122 | 289.9513 | 415.3476 | 41.094 | 41.094 | 415.3476 | 290.0169 | 310.2066 | 310.2722 |  | 241.7279
LS2 | CC11 | GV | 3103 | LS2_CC11:GV_D3103 |  | gate valve | 40 |  |  | 0.057404 | 290.0457 | 415.3476 | 41.094 | 41.094 | 415.3476 | 290.0744 | 310.3009 | 310.3296 |  | 241.8222
[6 empty] | bellows |  |  |  | 0.1677 | 290.1582 | 415.3476 | 41.094 | 41.094 | 415.3476 | 290.2421 | 310.4134 | 310.4973 |  | 241.9348
LS2 | CC11 | CAV1 | 3104 | LS2_CC11:CAV1_D3104 |  | b29 resonator | 40 | 0.32364 |  | 0.32364 | 290.4039 | 415.3476 | 41.094 | 41.094 | 415.3476 | 290.5657 | 310.6591 | 310.8209 |  | 242.1805
LS2 | CC11 | BLM | 3107 | LS2_CC11:BLM_D3107 |  | outside loss monitor |  |  |  | 0 | 290.5657 | 415.3476 | 41.094 | 41.094 | 415.3476 | 290.5657 | 310.8209 | 310.8209 |  | 242.3423
[10 empty] | 0 | 290.5657 | 415.3476 | 41.094 | 41.094 | 415.3476 | 290.5657 | 310.8209 | 310.8209 |  | 242.3423
[6 empty] | bellow |  |  |  | 0.099517 | 290.6155 | 415.3476 | 41.094 | 41.094 | 415.3476 | 290.6652 | 310.8707 | 310.9204 |  | 242.392
LS2 | CC11 | CAV2 | 3109 | LS2_CC11:CAV2_D3109 |  | b29 resonator | 40 |  |  | 0.32364 | 290.827 | 415.3476 | 41.094 | 41.094 | 415.3476 | 290.9889 | 311.0823 | 311.2441 |  | 242.6036
[10 empty] | 0 | 290.9889 | 415.3476 | 41.094 | 41.094 | 415.3476 | 290.9889 | 311.2441 | 311.2441 |  | 242.7654
[10 empty] | 0 | 290.9889 | 415.3476 | 41.094 | 41.094 | 415.3476 | 290.9889 | 311.2441 | 311.2441 |  | 242.7654
[6 empty] | bellow |  |  |  | 0.099517 | 291.0386 | 415.3476 | 41.094 | 41.094 | 415.3476 | 291.0884 | 311.2938 | 311.3436 |  | 242.8152
LS2 | CC11 | CAV3 | 3113 | LS2_CC11:CAV3_D3113 |  | b29 resonator | 40 |  |  | 0.32364 | 291.2502 | 415.3476 | 41.094 | 41.094 | 415.3476 | 291.412 | 311.5054 | 311.6672 |  | 243.0268
LS2 | CC11 | BLM | 3115 | LS2_CC11:BLM_D3115 |  | outside loss monitor |  |  |  | 0 | 291.412 | 415.3476 | 41.094 | 41.094 | 415.3476 | 291.412 | 311.6672 | 311.6672 |  | 243.1886
[10 empty] | 0 | 291.412 | 415.3476 | 41.094 | 41.094 | 415.3476 | 291.412 | 311.6672 | 311.6672 |  | 243.1886
[6 empty] | bellow+tube |  |  |  | 0.15525 | 291.4896 | 415.3476 | 41.094 | 41.094 | 415.3476 | 291.5673 | 311.7449 | 311.8225 |  | 243.2662
LS2 | CC11 | SOL1 | 3121 | LS2_CC11:SOL1_D3121 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 291.8173 | 415.3476 | 41.094 | 41.094 | 415.3476 | 292.0673 | 312.0725 | 312.3225 |  | 243.5939
[6 empty] | bellow+tube |  |  |  | 0.15525 | 292.1449 | 415.3476 | 41.094 | 41.094 | 415.3476 | 292.2225 | 312.4001 | 312.4777 |  | 243.9215
LS2 | CC11 | CAV4 | 3124 | LS2_CC11:CAV4_D3124 |  | b29 resonator | 40 |  |  | 0.32364 | 292.3843 | 415.3476 | 41.094 | 41.094 | 415.3476 | 292.5462 | 312.6396 | 312.8014 |  | 244.1609
[10 empty] | 0 | 292.5462 | 415.3476 | 41.094 | 41.094 | 415.3476 | 292.5462 | 312.8014 | 312.8014 |  | 244.3228
[10 empty] | 0 | 292.5462 | 415.3476 | 41.094 | 41.094 | 415.3476 | 292.5462 | 312.8014 | 312.8014 |  | 244.3228
[10 empty] | 0.099517 | 292.5959 | 415.3476 | 41.094 | 41.094 | 415.3476 | 292.6457 | 312.8511 | 312.9009 |  | 244.3725
LS2 | CC11 | CAV5 | 3129 | LS2_CC11:CAV5_D3129 |  | b29 resonator | 40 |  |  | 0.32364 | 292.8075 | 415.3476 | 41.094 | 41.094 | 415.3476 | 292.9693 | 313.0627 | 313.2245 |  | 244.5841
[10 empty] | 0 | 292.9693 | 415.3476 | 41.094 | 41.094 | 415.3476 | 292.9693 | 313.2245 | 313.2245 |  | 244.7459
[10 empty] | 0 | 292.9693 | 415.3476 | 41.094 | 41.094 | 415.3476 | 292.9693 | 313.2245 | 313.2245 |  | 244.7459
[10 empty] | 0.099517 | 293.0191 | 415.3476 | 41.094 | 41.094 | 415.3476 | 293.0688 | 313.2743 | 313.3241 |  | 244.7957
LS2 | CC11 | CAV6 | 3133 | LS2_CC11:CAV6_D3133 |  | b29 resonator | 40 |  |  | 0.32364 | 293.2307 | 415.3476 | 41.094 | 41.094 | 415.3476 | 293.3925 | 313.4859 | 313.6477 |  | 245.0073
LS2 | CC11 | BLM | 3135 | LS2_CC11:BLM_D3135 |  | outside loss monitor |  |  |  | 0 | 293.3925 | 415.3476 | 41.094 | 41.094 | 415.3476 | 293.3925 | 313.6477 | 313.6477 |  | 245.1691
[10 empty] | 0 | 293.3925 | 415.3476 | 41.094 | 41.094 | 415.3476 | 293.3925 | 313.6477 | 313.6477 |  | 245.1691
[6 empty] | bellows |  |  |  | 0.1677 | 293.4763 | 415.3476 | 41.094 | 41.094 | 415.3476 | 293.5602 | 313.7316 | 313.8154 |  | 245.2529
LS2 | CC11 | GV | 3138 | LS2_CC11:GV_D3138 |  | gate valve | 40 |  |  | 0.057404 | 293.5889 | 415.3476 | 41.094 | 41.094 | 415.3476 | 293.6176 | 313.8441 | 313.8728 |  | 245.3655
[6 empty] | bellow | 40 |  |  | 0.1035 | 293.6693 | 415.3476 | 41.094 | 41.094 | 415.3476 | 293.7211 | 313.9246 | 313.9763 |  | 245.4459
LS2 | WC11 | BPM | 3140 | LS2_WC11:BPM_D3140 | BPM2 | position monitor |  |  |  | 0 | 293.7211 | 415.3476 | 41.094 | 41.094 | 415.3476 | 293.7211 | 313.9763 | 313.9763 |  | 245.4977
[6 empty] | diagnostic box |  |  |  | 0.14528 | 293.7937 | 415.3476 | 41.094 | 41.094 | 415.3476 | 293.8664 | 314.049 | 314.1216 |  | 245.5703
LS2 | WC11 | PORT | 3141 | LS2_WC11:PORT_D3141 |  | port for future diagnostics |  |  |  | 0 | 293.8664 | 415.3476 | 41.094 | 41.094 | 415.3476 | 293.8664 | 314.1216 | 314.1216 |  | 245.643
[6 empty] | diagnostic box |  |  |  | 0 | 293.8664 | 415.3476 | 41.094 | 41.094 | 415.3476 | 293.8664 | 314.1216 | 314.1216 |  | 245.643
LS2 | WC11 | IP | 3141 | LS2_WC11:IP_D3141 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 293.8664 | 415.3476 | 41.094 | 41.094 | 415.3476 | 293.8664 | 314.1216 | 314.1216 |  | 245.643
[6 empty] | bellow | 40 |  |  | 0.13122 | 293.932 | 415.3476 | 41.094 | 41.094 | 415.3476 | 293.9976 | 314.1872 | 314.2528 |  | 245.7086
LS2 | CC12 | GV | 3143 | LS2_CC12:GV_D3143 |  | gate valve | 40 |  |  | 0.057404 | 294.0263 | 415.3476 | 41.094 | 41.094 | 415.3476 | 294.055 | 314.2815 | 314.3102 |  | 245.8029
[6 empty] | bellows |  |  |  | 0.1677 | 294.1389 | 415.3476 | 41.094 | 41.094 | 415.3476 | 294.2227 | 314.3941 | 314.4779 |  | 245.9154
LS2 | CC12 | CAV1 | 3144 | LS2_CC12:CAV1_D3144 |  | b29 resonator | 40 | 0.32364 |  | 0.32364 | 294.3845 | 415.3476 | 41.094 | 41.094 | 415.3476 | 294.5463 | 314.6397 | 314.8016 |  | 246.1611
LS2 | CC12 | BLM | 3146 | LS2_CC12:BLM_D3146 |  | outside loss monitor |  |  |  | 0 | 294.5463 | 415.3476 | 41.094 | 41.094 | 415.3476 | 294.5463 | 314.8016 | 314.8016 |  | 246.3229
[10 empty] | 0 | 294.5463 | 415.3476 | 41.094 | 41.094 | 415.3476 | 294.5463 | 314.8016 | 314.8016 |  | 246.3229
[6 empty] | bellow |  |  |  | 0.099517 | 294.5961 | 415.3476 | 41.094 | 41.094 | 415.3476 | 294.6459 | 314.8513 | 314.9011 |  | 246.3727
LS2 | CC12 | CAV2 | 3149 | LS2_CC12:CAV2_D3149 |  | b29 resonator | 40 |  |  | 0.32364 | 294.8077 | 415.3476 | 41.094 | 41.094 | 415.3476 | 294.9695 | 315.0629 | 315.2247 |  | 246.5843
[10 empty] | 0 | 294.9695 | 415.3476 | 41.094 | 41.094 | 415.3476 | 294.9695 | 315.2247 | 315.2247 |  | 246.7461
[10 empty] | 0 | 294.9695 | 415.3476 | 41.094 | 41.094 | 415.3476 | 294.9695 | 315.2247 | 315.2247 |  | 246.7461
[6 empty] | bellow |  |  |  | 0.099517 | 295.0193 | 415.3476 | 41.094 | 41.094 | 415.3476 | 295.069 | 315.2745 | 315.3242 |  | 246.7959
LS2 | CC12 | CAV3 | 3153 | LS2_CC12:CAV3_D3153 |  | b29 resonator | 40 |  |  | 0.32364 | 295.2308 | 415.3476 | 41.094 | 41.094 | 415.3476 | 295.3927 | 315.4861 | 315.6479 |  | 247.0074
LS2 | CC12 | BLM | 3155 | LS2_CC12:BLM_D3155 |  | outside loss monitor |  |  |  | 0 | 295.3927 | 415.3476 | 41.094 | 41.094 | 415.3476 | 295.3927 | 315.6479 | 315.6479 |  | 247.1693
[10 empty] | 0 | 295.3927 | 415.3476 | 41.094 | 41.094 | 415.3476 | 295.3927 | 315.6479 | 315.6479 |  | 247.1693
[6 empty] | bellow+tube |  |  |  | 0.15525 | 295.4703 | 415.3476 | 41.094 | 41.094 | 415.3476 | 295.5479 | 315.7255 | 315.8031 |  | 247.2469
LS2 | CC12 | SOL1 | 3161 | LS2_CC12:SOL1_D3161 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 295.7979 | 415.3476 | 41.094 | 41.094 | 415.3476 | 296.0479 | 316.0531 | 316.3031 |  | 247.5745
[6 empty] | bellow+tube |  |  |  | 0.15525 | 296.1255 | 415.3476 | 41.094 | 41.094 | 415.3476 | 296.2032 | 316.3808 | 316.4584 |  | 247.9021
LS2 | CC12 | CAV4 | 3164 | LS2_CC12:CAV4_D3164 |  | b29 resonator | 40 |  |  | 0.32364 | 296.365 | 415.3476 | 41.094 | 41.094 | 415.3476 | 296.5268 | 316.6202 | 316.782 |  | 248.1416
[10 empty] | 0 | 296.5268 | 415.3476 | 41.094 | 41.094 | 415.3476 | 296.5268 | 316.782 | 316.782 |  | 248.3034
[10 empty] | 0 | 296.5268 | 415.3476 | 41.094 | 41.094 | 415.3476 | 296.5268 | 316.782 | 316.782 |  | 248.3034
[10 empty] | 0.099517 | 296.5766 | 415.3476 | 41.094 | 41.094 | 415.3476 | 296.6263 | 316.8318 | 316.8815 |  | 248.3532
LS2 | CC12 | CAV5 | 3168 | LS2_CC12:CAV5_D3168 |  | b29 resonator | 40 |  |  | 0.32364 | 296.7882 | 415.3476 | 41.094 | 41.094 | 415.3476 | 296.95 | 317.0434 | 317.2052 |  | 248.5647
[10 empty] | 0 | 296.95 | 415.3476 | 41.094 | 41.094 | 415.3476 | 296.95 | 317.2052 | 317.2052 |  | 248.7266
[10 empty] | 0 | 296.95 | 415.3476 | 41.094 | 41.094 | 415.3476 | 296.95 | 317.2052 | 317.2052 |  | 248.7266
[10 empty] | 0.099517 | 296.9997 | 415.3476 | 41.094 | 41.094 | 415.3476 | 297.0495 | 317.255 | 317.3047 |  | 248.7763
LS2 | CC12 | CAV6 | 3173 | LS2_CC12:CAV6_D3173 |  | b29 resonator | 40 |  |  | 0.32364 | 297.2113 | 415.3476 | 41.094 | 41.094 | 415.3476 | 297.3731 | 317.4665 | 317.6284 |  | 248.9879
LS2 | CC12 | BLM | 3175 | LS2_CC12:BLM_D3175 |  | outside loss monitor |  |  |  | 0 | 297.3731 | 415.3476 | 41.094 | 41.094 | 415.3476 | 297.3731 | 317.6284 | 317.6284 |  | 249.1497
[10 empty] | 0 | 297.3731 | 415.3476 | 41.094 | 41.094 | 415.3476 | 297.3731 | 317.6284 | 317.6284 |  | 249.1497
[6 empty] | bellows |  |  |  | 0.1677 | 297.457 | 415.3476 | 41.094 | 41.094 | 415.3476 | 297.5408 | 317.7122 | 317.7961 |  | 249.2336
LS2 | CC12 | GV | 3178 | LS2_CC12:GV_D3178 |  | gate valve | 40 |  |  | 0.057404 | 297.5695 | 415.3476 | 41.094 | 41.094 | 415.3476 | 297.5982 | 317.8248 | 317.8535 |  | 249.3461
[6 empty] | bellow | 40 |  |  | 0.1035 | 297.65 | 415.3476 | 41.094 | 41.094 | 415.3476 | 297.7017 | 317.9052 | 317.957 |  | 249.4266
LS2 | WC12 | BPM | 3180 | LS2_WC12:BPM_D3180 | BPM2 | position monitor |  |  |  | 0 | 297.7017 | 415.3476 | 41.094 | 41.094 | 415.3476 | 297.7017 | 317.957 | 317.957 |  | 249.4783
[6 empty] | diagnostic box |  |  |  | 0.14528 | 297.7744 | 415.3476 | 41.094 | 41.094 | 415.3476 | 297.847 | 318.0296 | 318.1022 |  | 249.551
LS2 | WC12 | PM | 3181 | LS2_WC12:PM_D3181 | PMx | profile monitor (not costed) | future PM |  |  | 0 | 297.847 | 415.3476 | 41.094 | 41.094 | 415.3476 | 297.847 | 318.1022 | 318.1022 |  | 249.6236
[6 empty] | diagnostic box |  |  |  | 0 | 297.847 | 415.3476 | 41.094 | 41.094 | 415.3476 | 297.847 | 318.1022 | 318.1022 |  | 249.6236
LS2 | WC12 | IP | 3181 | LS2_WC12:IP_D3181 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 297.847 | 415.3476 | 41.094 | 41.094 | 415.3476 | 297.847 | 318.1022 | 318.1022 |  | 249.6236
[6 empty] | bellow | 40 |  |  | 0.13122 | 297.9126 | 415.3476 | 41.094 | 41.094 | 415.3476 | 297.9782 | 318.1679 | 318.2335 |  | 249.6892
LS2 | CD01 | GV | 3183 | LS2_CD01:GV_D3183 |  | gate valve | 40 |  |  | 0.057404 | 298.0069 | 415.3476 | 41.094 | 41.094 | 415.3476 | 298.0356 | 318.2622 | 318.2909 |  | 249.7835
[6 empty] | bellows |  |  |  | 0.1677 | 298.1195 | 415.3476 | 41.094 | 41.094 | 415.3476 | 298.2034 | 318.3747 | 318.4586 |  | 249.8961
LS2 | CD01 | CAV1 | 3187 | LS2_CD01:CAV1_D3187 |  | b54 resonator | 40 |  |  | 0.49566 | 298.4512 | 415.3476 | 41.094 | 41.094 | 415.3476 | 298.699 | 318.7064 | 318.9542 |  | 250.2278
LS2 | CD01 | BLM | 3190 | LS2_CD01:BLM_D3190 |  | outside loss monitor |  |  |  | 0 | 298.699 | 415.3476 | 41.094 | 41.094 | 415.3476 | 298.699 | 318.9542 | 318.9542 |  | 250.4756
[10 empty] | 0 | 298.699 | 415.3476 | 41.094 | 41.094 | 415.3476 | 298.699 | 318.9542 | 318.9542 |  | 250.4756
[6 empty] | bellow |  |  |  | 0.099517 | 298.7488 | 415.3476 | 41.094 | 41.094 | 415.3476 | 298.7985 | 319.004 | 319.0537 |  | 250.5254
LS2 | CD01 | CAV2 | 3193 | LS2_CD01:CAV2_D3193 |  | b54 resonator | 40 |  |  | 0.49566 | 299.0464 | 415.3476 | 41.094 | 41.094 | 415.3476 | 299.2942 | 319.3016 | 319.5494 |  | 250.8229
[10 empty] | 0 | 299.2942 | 415.3476 | 41.094 | 41.094 | 415.3476 | 299.2942 | 319.5494 | 319.5494 |  | 251.0708
[10 empty] | 0 | 299.2942 | 415.3476 | 41.094 | 41.094 | 415.3476 | 299.2942 | 319.5494 | 319.5494 |  | 251.0708
[6 empty] | bellow |  |  |  | 0.099517 | 299.3439 | 415.3476 | 41.094 | 41.094 | 415.3476 | 299.3937 | 319.5992 | 319.6489 |  | 251.1205
LS2 | CD01 | CAV3 | 3199 | LS2_CD01:CAV3_D3199 |  | b54 resonator | 40 |  |  | 0.49566 | 299.6415 | 415.3476 | 41.094 | 41.094 | 415.3476 | 299.8894 | 319.8967 | 320.1446 |  | 251.4181
[10 empty] | 0 | 299.8894 | 415.3476 | 41.094 | 41.094 | 415.3476 | 299.8894 | 320.1446 | 320.1446 |  | 251.6659
[10 empty] | 0 | 299.8894 | 415.3476 | 41.094 | 41.094 | 415.3476 | 299.8894 | 320.1446 | 320.1446 |  | 251.6659
[6 empty] | bellow |  |  |  | 0.099517 | 299.9391 | 415.3476 | 41.094 | 41.094 | 415.3476 | 299.9889 | 320.1943 | 320.2441 |  | 251.7157
LS2 | CD01 | CAV4 | 3205 | LS2_CD01:CAV4_D3205 |  | b54 resonator | 40 |  |  | 0.49566 | 300.2367 | 415.3476 | 41.094 | 41.094 | 415.3476 | 300.4845 | 320.4919 | 320.7397 |  | 252.0133
LS2 | CD01 | BLM | 3207 | LS2_CD01:BLM_D3207 |  | outside loss monitor |  |  |  | 0 | 300.4845 | 415.3476 | 41.094 | 41.094 | 415.3476 | 300.4845 | 320.7397 | 320.7397 |  | 252.2611
[10 empty] | 0 | 300.4845 | 415.3476 | 41.094 | 41.094 | 415.3476 | 300.4845 | 320.7397 | 320.7397 |  | 252.2611
[6 empty] | bellow |  |  |  | 0.15525 | 300.5622 | 415.3476 | 41.094 | 41.094 | 415.3476 | 300.6398 | 320.8174 | 320.895 |  | 252.3387
LS2 | CD01 | SOL1 | 3211 | LS2_CD01:SOL1_D3211 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 300.8898 | 415.3476 | 41.094 | 41.094 | 415.3476 | 301.1398 | 321.145 | 321.395 |  | 252.6664
[6 empty] | bellow |  |  |  | 0.15525 | 301.2174 | 415.3476 | 41.094 | 41.094 | 415.3476 | 301.295 | 321.4726 | 321.5502 |  | 252.994
LS2 | CD01 | CAV5 | 3218 | LS2_CD01:CAV5_D3218 |  | b54 resonator | 40 |  |  | 0.49566 | 301.5429 | 415.3476 | 41.094 | 41.094 | 415.3476 | 301.7907 | 321.7981 | 322.0459 |  | 253.3194
[10 empty] | 0 | 301.7907 | 415.3476 | 41.094 | 41.094 | 415.3476 | 301.7907 | 322.0459 | 322.0459 |  | 253.5673
[10 empty] | 0 | 301.7907 | 415.3476 | 41.094 | 41.094 | 415.3476 | 301.7907 | 322.0459 | 322.0459 |  | 253.5673
[6 empty] | bellow |  |  |  | 0.099517 | 301.8404 | 415.3476 | 41.094 | 41.094 | 415.3476 | 301.8902 | 322.0957 | 322.1454 |  | 253.617
LS2 | CD01 | CAV6 | 3224 | LS2_CD01:CAV6_D3224 |  | b54 resonator | 40 |  |  | 0.49566 | 302.138 | 415.3476 | 41.094 | 41.094 | 415.3476 | 302.3859 | 322.3932 | 322.6411 |  | 253.9146
[10 empty] | 0 | 302.3859 | 415.3476 | 41.094 | 41.094 | 415.3476 | 302.3859 | 322.6411 | 322.6411 |  | 254.1624
[10 empty] | 0 | 302.3859 | 415.3476 | 41.094 | 41.094 | 415.3476 | 302.3859 | 322.6411 | 322.6411 |  | 254.1624
[6 empty] | bellow |  |  |  | 0.099517 | 302.4356 | 415.3476 | 41.094 | 41.094 | 415.3476 | 302.4854 | 322.6908 | 322.7406 |  | 254.2122
LS2 | CD01 | CAV7 | 3230 | LS2_CD01:CAV7_D3230 |  | b54 resonator | 40 |  |  | 0.49566 | 302.7332 | 415.3476 | 41.094 | 41.094 | 415.3476 | 302.981 | 322.9884 | 323.2362 |  | 254.5098
[10 empty] | 0 | 302.981 | 415.3476 | 41.094 | 41.094 | 415.3476 | 302.981 | 323.2362 | 323.2362 |  | 254.7576
[10 empty] | 0 | 302.981 | 415.3476 | 41.094 | 41.094 | 415.3476 | 302.981 | 323.2362 | 323.2362 |  | 254.7576
[6 empty] | bellow |  |  |  | 0.099517 | 303.0308 | 415.3476 | 41.094 | 41.094 | 415.3476 | 303.0805 | 323.286 | 323.3358 |  | 254.8074
LS2 | CD01 | CAV8 | 3236 | LS2_CD01:CAV8_D3236 |  | b54 resonator | 40 |  |  | 0.49566 | 303.3284 | 415.3476 | 41.094 | 41.094 | 415.3476 | 303.5762 | 323.5836 | 323.8314 |  | 255.105
LS2 | CD01 | BLM | 3238 | LS2_CD01:BLM_D3238 |  | outside loss monitor |  |  |  | 0 | 303.5762 | 415.3476 | 41.094 | 41.094 | 415.3476 | 303.5762 | 323.8314 | 323.8314 |  | 255.3528
[10 empty] | 0 | 303.5762 | 415.3476 | 41.094 | 41.094 | 415.3476 | 303.5762 | 323.8314 | 323.8314 |  | 255.3528
[10 empty] | 0.1677 | 303.6601 | 415.3476 | 41.094 | 41.094 | 415.3476 | 303.7439 | 323.9153 | 323.9991 |  | 255.4366
LS2 | CD01 | GV | 3240 | LS2_CD01:GV_D3240 |  | gate valve | 40 |  |  | 0.057404 | 303.7726 | 415.3476 | 41.094 | 41.094 | 415.3476 | 303.8013 | 324.0278 | 324.0565 |  | 255.5492
[6 empty] | bellow | 40 |  |  | 0.1035 | 303.8531 | 415.3476 | 41.094 | 41.094 | 415.3476 | 303.9048 | 324.1083 | 324.16 |  | 255.6296
LS2 | WD01 | BPM | 3242 | LS2_WD01:BPM_D3242 | BPM2 | position monitor |  |  |  | 0 | 303.9048 | 415.3476 | 41.094 | 41.094 | 415.3476 | 303.9048 | 324.16 | 324.16 |  | 255.6814
[6 empty] | diagnostic box |  |  |  | 0.14528 | 303.9775 | 415.3476 | 41.094 | 41.094 | 415.3476 | 304.0501 | 324.2327 | 324.3053 |  | 255.754
LS2 | WD01 | PM | 3243 | LS2_WD01:PM_D3243 | PMx | profile monitor (not costed) | future PM |  |  | 0 | 304.0501 | 415.3476 | 41.094 | 41.094 | 415.3476 | 304.0501 | 324.3053 | 324.3053 |  | 255.8267
[6 empty] | diagnostic box |  |  |  | 0 | 304.0501 | 415.3476 | 41.094 | 41.094 | 415.3476 | 304.0501 | 324.3053 | 324.3053 |  | 255.8267
LS2 | WD01 | IP | 3243 | LS2_WD01:IP_D3243 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 304.0501 | 415.3476 | 41.094 | 41.094 | 415.3476 | 304.0501 | 324.3053 | 324.3053 |  | 255.8267
[6 empty] | bellow | 40 |  |  | 0.13122 | 304.1157 | 415.3476 | 41.094 | 41.094 | 415.3476 | 304.1813 | 324.3709 | 324.4365 |  | 255.8923
LS2 | CD02 | GV | 3245 | LS2_CD02:GV_D3245 |  | gate valve | 40 |  |  | 0.057404 | 304.21 | 415.3476 | 41.094 | 41.094 | 415.3476 | 304.2387 | 324.4652 | 324.4939 |  | 255.9866
[6 empty] | bellows |  |  |  | 0.1677 | 304.3226 | 415.3476 | 41.094 | 41.094 | 415.3476 | 304.4064 | 324.5778 | 324.6616 |  | 256.0992
LS2 | CD02 | CAV1 | 3249 | LS2_CD02:CAV1_D3249 |  | b54 resonator | 40 |  |  | 0.49566 | 304.6542 | 415.3476 | 41.094 | 41.094 | 415.3476 | 304.9021 | 324.9095 | 325.1573 |  | 256.4308
LS2 | CD02 | BLM | 3252 | LS2_CD02:BLM_D3252 |  | outside loss monitor |  |  |  | 0 | 304.9021 | 415.3476 | 41.094 | 41.094 | 415.3476 | 304.9021 | 325.1573 | 325.1573 |  | 256.6787
[10 empty] | 0 | 304.9021 | 415.3476 | 41.094 | 41.094 | 415.3476 | 304.9021 | 325.1573 | 325.1573 |  | 256.6787
[6 empty] | bellow |  |  |  | 0.099517 | 304.9518 | 415.3476 | 41.094 | 41.094 | 415.3476 | 305.0016 | 325.207 | 325.2568 |  | 256.7284
LS2 | CD02 | CAV2 | 3255 | LS2_CD02:CAV2_D3255 |  | b54 resonator | 40 |  |  | 0.49566 | 305.2494 | 415.3476 | 41.094 | 41.094 | 415.3476 | 305.4972 | 325.5046 | 325.7525 |  | 257.026
[10 empty] | 0 | 305.4972 | 415.3476 | 41.094 | 41.094 | 415.3476 | 305.4972 | 325.7525 | 325.7525 |  | 257.2738
[10 empty] | 0 | 305.4972 | 415.3476 | 41.094 | 41.094 | 415.3476 | 305.4972 | 325.7525 | 325.7525 |  | 257.2738
[6 empty] | bellow |  |  |  | 0.099517 | 305.547 | 415.3476 | 41.094 | 41.094 | 415.3476 | 305.5968 | 325.8022 | 325.852 |  | 257.3236
LS2 | CD02 | CAV3 | 3261 | LS2_CD02:CAV3_D3261 |  | b54 resonator | 40 |  |  | 0.49566 | 305.8446 | 415.3476 | 41.094 | 41.094 | 415.3476 | 306.0924 | 326.0998 | 326.3476 |  | 257.6212
[10 empty] | 0 | 306.0924 | 415.3476 | 41.094 | 41.094 | 415.3476 | 306.0924 | 326.3476 | 326.3476 |  | 257.869
[10 empty] | 0 | 306.0924 | 415.3476 | 41.094 | 41.094 | 415.3476 | 306.0924 | 326.3476 | 326.3476 |  | 257.869
[6 empty] | bellow |  |  |  | 0.099517 | 306.1422 | 415.3476 | 41.094 | 41.094 | 415.3476 | 306.1919 | 326.3974 | 326.4472 |  | 257.9188
LS2 | CD02 | CAV4 | 3267 | LS2_CD02:CAV4_D3267 |  | b54 resonator | 40 |  |  | 0.49566 | 306.4398 | 415.3476 | 41.094 | 41.094 | 415.3476 | 306.6876 | 326.695 | 326.9428 |  | 258.2164
LS2 | CD02 | BLM | 3269 | LS2_CD02:BLM_D3269 |  | outside loss monitor |  |  |  | 0 | 306.6876 | 415.3476 | 41.094 | 41.094 | 415.3476 | 306.6876 | 326.9428 | 326.9428 |  | 258.4642
[10 empty] | 0 | 306.6876 | 415.3476 | 41.094 | 41.094 | 415.3476 | 306.6876 | 326.9428 | 326.9428 |  | 258.4642
[6 empty] | bellow |  |  |  | 0.15525 | 306.7652 | 415.3476 | 41.094 | 41.094 | 415.3476 | 306.8428 | 327.0204 | 327.0981 |  | 258.5418
LS2 | CD02 | SOL1 | 3273 | LS2_CD02:SOL1_D3273 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 307.0928 | 415.3476 | 41.094 | 41.094 | 415.3476 | 307.3428 | 327.3481 | 327.5981 |  | 258.8694
[6 empty] | bellow |  |  |  | 0.15525 | 307.4205 | 415.3476 | 41.094 | 41.094 | 415.3476 | 307.4981 | 327.6757 | 327.7533 |  | 259.1971
LS2 | CD02 | CAV5 | 3280 | LS2_CD02:CAV5_D3280 |  | b54 resonator | 40 |  |  | 0.49566 | 307.7459 | 415.3476 | 41.094 | 41.094 | 415.3476 | 307.9937 | 328.0011 | 328.249 |  | 259.5225
[10 empty] | 0 | 307.9937 | 415.3476 | 41.094 | 41.094 | 415.3476 | 307.9937 | 328.249 | 328.249 |  | 259.7703
[10 empty] | 0 | 307.9937 | 415.3476 | 41.094 | 41.094 | 415.3476 | 307.9937 | 328.249 | 328.249 |  | 259.7703
[6 empty] | bellow |  |  |  | 0.099517 | 308.0435 | 415.3476 | 41.094 | 41.094 | 415.3476 | 308.0933 | 328.2987 | 328.3485 |  | 259.8201
LS2 | CD02 | CAV6 | 3286 | LS2_CD02:CAV6_D3286 |  | b54 resonator | 40 |  |  | 0.49566 | 308.3411 | 415.3476 | 41.094 | 41.094 | 415.3476 | 308.5889 | 328.5963 | 328.8441 |  | 260.1177
[10 empty] | 0 | 308.5889 | 415.3476 | 41.094 | 41.094 | 415.3476 | 308.5889 | 328.8441 | 328.8441 |  | 260.3655
[10 empty] | 0 | 308.5889 | 415.3476 | 41.094 | 41.094 | 415.3476 | 308.5889 | 328.8441 | 328.8441 |  | 260.3655
[6 empty] | bellow |  |  |  | 0.099517 | 308.6387 | 415.3476 | 41.094 | 41.094 | 415.3476 | 308.6884 | 328.8939 | 328.9437 |  | 260.4153
LS2 | CD02 | CAV7 | 3292 | LS2_CD02:CAV7_D3292 |  | b54 resonator | 40 |  |  | 0.49566 | 308.9363 | 415.3476 | 41.094 | 41.094 | 415.3476 | 309.1841 | 329.1915 | 329.4393 |  | 260.7129
[10 empty] | 0 | 309.1841 | 415.3476 | 41.094 | 41.094 | 415.3476 | 309.1841 | 329.4393 | 329.4393 |  | 260.9607
[10 empty] | 0 | 309.1841 | 415.3476 | 41.094 | 41.094 | 415.3476 | 309.1841 | 329.4393 | 329.4393 |  | 260.9607
[6 empty] | bellow |  |  |  | 0.099517 | 309.2339 | 415.3476 | 41.094 | 41.094 | 415.3476 | 309.2836 | 329.4891 | 329.5388 |  | 261.0104
LS2 | CD02 | CAV8 | 3298 | LS2_CD02:CAV8_D3298 |  | b54 resonator | 40 |  |  | 0.49566 | 309.5314 | 415.3476 | 41.094 | 41.094 | 415.3476 | 309.7793 | 329.7867 | 330.0345 |  | 261.308
LS2 | CD02 | BLM | 3300 | LS2_CD02:BLM_D3300 |  | outside loss monitor |  |  |  | 0 | 309.7793 | 415.3476 | 41.094 | 41.094 | 415.3476 | 309.7793 | 330.0345 | 330.0345 |  | 261.5559
[10 empty] | 0 | 309.7793 | 415.3476 | 41.094 | 41.094 | 415.3476 | 309.7793 | 330.0345 | 330.0345 |  | 261.5559
[10 empty] | 0.1677 | 309.8631 | 415.3476 | 41.094 | 41.094 | 415.3476 | 309.947 | 330.1183 | 330.2022 |  | 261.6397
LS2 | CD02 | GV | 3302 | LS2_CD02:GV_D3302 |  | gate valve | 40 |  |  | 0.057404 | 309.9757 | 415.3476 | 41.094 | 41.094 | 415.3476 | 310.0044 | 330.2309 | 330.2596 |  | 261.7523
[6 empty] | bellow | 40 |  |  | 0.1035 | 310.0561 | 415.3476 | 41.094 | 41.094 | 415.3476 | 310.1079 | 330.3113 | 330.3631 |  | 261.8327
LS2 | WD02 | BPM | 3304 | LS2_WD02:BPM_D3304 | BPM2 | position monitor |  |  |  | 0 | 310.1079 | 415.3476 | 41.094 | 41.094 | 415.3476 | 310.1079 | 330.3631 | 330.3631 |  | 261.8845
[6 empty] | diagnostic box |  |  |  | 0.14528 | 310.1805 | 415.3476 | 41.094 | 41.094 | 415.3476 | 310.2532 | 330.4357 | 330.5084 |  | 261.9571
LS2 | WD02 | PM | 3305 | LS2_WD02:PM_D3305 | PMx | profile monitor (not costed) | future PM |  |  | 0 | 310.2532 | 415.3476 | 41.094 | 41.094 | 415.3476 | 310.2532 | 330.5084 | 330.5084 |  | 262.0297
[6 empty] | diagnostic box |  |  |  | 0 | 310.2532 | 415.3476 | 41.094 | 41.094 | 415.3476 | 310.2532 | 330.5084 | 330.5084 |  | 262.0297
LS2 | WD02 | IP | 3305 | LS2_WD02:IP_D3305 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 310.2532 | 415.3476 | 41.094 | 41.094 | 415.3476 | 310.2532 | 330.5084 | 330.5084 |  | 262.0297
[6 empty] | bellow | 40 |  |  | 0.13122 | 310.3188 | 415.3476 | 41.094 | 41.094 | 415.3476 | 310.3844 | 330.574 | 330.6396 |  | 262.0954
LS2 | CD03 | GV | 3307 | LS2_CD03:GV_D3307 |  | gate valve | 40 |  |  | 0.057404 | 310.4131 | 415.3476 | 41.094 | 41.094 | 415.3476 | 310.4418 | 330.6683 | 330.697 |  | 262.1897
[6 empty] | bellows |  |  |  | 0.1677 | 310.5256 | 415.3476 | 41.094 | 41.094 | 415.3476 | 310.6095 | 330.7808 | 330.8647 |  | 262.3022
LS2 | CD03 | CAV1 | 3311 | LS2_CD03:CAV1_D3311 |  | b54 resonator | 40 |  |  | 0.49566 | 310.8573 | 415.3476 | 41.094 | 41.094 | 415.3476 | 311.1051 | 331.1125 | 331.3604 |  | 262.6339
LS2 | CD03 | BLM | 3314 | LS2_CD03:BLM_D3314 |  | outside loss monitor |  |  |  | 0 | 311.1051 | 415.3476 | 41.094 | 41.094 | 415.3476 | 311.1051 | 331.3604 | 331.3604 |  | 262.8817
[10 empty] | 0 | 311.1051 | 415.3476 | 41.094 | 41.094 | 415.3476 | 311.1051 | 331.3604 | 331.3604 |  | 262.8817
[6 empty] | bellow |  |  |  | 0.099517 | 311.1549 | 415.3476 | 41.094 | 41.094 | 415.3476 | 311.2047 | 331.4101 | 331.4599 |  | 262.9315
LS2 | CD03 | CAV2 | 3317 | LS2_CD03:CAV2_D3317 |  | b54 resonator | 40 |  |  | 0.49566 | 311.4525 | 415.3476 | 41.094 | 41.094 | 415.3476 | 311.7003 | 331.7077 | 331.9555 |  | 263.2291
[10 empty] | 0 | 311.7003 | 415.3476 | 41.094 | 41.094 | 415.3476 | 311.7003 | 331.9555 | 331.9555 |  | 263.4769
[10 empty] | 0 | 311.7003 | 415.3476 | 41.094 | 41.094 | 415.3476 | 311.7003 | 331.9555 | 331.9555 |  | 263.4769
[6 empty] | bellow |  |  |  | 0.099517 | 311.7501 | 415.3476 | 41.094 | 41.094 | 415.3476 | 311.7998 | 332.0053 | 332.055 |  | 263.5267
LS2 | CD03 | CAV3 | 3323 | LS2_CD03:CAV3_D3323 |  | b54 resonator | 40 |  |  | 0.49566 | 312.0477 | 415.3476 | 41.094 | 41.094 | 415.3476 | 312.2955 | 332.3029 | 332.5507 |  | 263.8242
[10 empty] | 0 | 312.2955 | 415.3476 | 41.094 | 41.094 | 415.3476 | 312.2955 | 332.5507 | 332.5507 |  | 264.0721
[10 empty] | 0 | 312.2955 | 415.3476 | 41.094 | 41.094 | 415.3476 | 312.2955 | 332.5507 | 332.5507 |  | 264.0721
[6 empty] | bellow |  |  |  | 0.099517 | 312.3452 | 415.3476 | 41.094 | 41.094 | 415.3476 | 312.395 | 332.6005 | 332.6502 |  | 264.1218
LS2 | CD03 | CAV4 | 3329 | LS2_CD03:CAV4_D3329 |  | b54 resonator | 40 |  |  | 0.49566 | 312.6428 | 415.3476 | 41.094 | 41.094 | 415.3476 | 312.8907 | 332.898 | 333.1459 |  | 264.4194
LS2 | CD03 | BLM | 3331 | LS2_CD03:BLM_D3331 |  | outside loss monitor |  |  |  | 0 | 312.8907 | 415.3476 | 41.094 | 41.094 | 415.3476 | 312.8907 | 333.1459 | 333.1459 |  | 264.6672
[10 empty] | 0 | 312.8907 | 415.3476 | 41.094 | 41.094 | 415.3476 | 312.8907 | 333.1459 | 333.1459 |  | 264.6672
[6 empty] | bellow |  |  |  | 0.15525 | 312.9683 | 415.3476 | 41.094 | 41.094 | 415.3476 | 313.0459 | 333.2235 | 333.3011 |  | 264.7449
LS2 | CD03 | SOL1 | 3336 | LS2_CD03:SOL1_D3336 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 313.2959 | 415.3476 | 41.094 | 41.094 | 415.3476 | 313.5459 | 333.5511 | 333.8011 |  | 265.0725
[6 empty] | bellow |  |  |  | 0.15525 | 313.6235 | 415.3476 | 41.094 | 41.094 | 415.3476 | 313.7012 | 333.8788 | 333.9564 |  | 265.4001
LS2 | CD03 | CAV5 | 3342 | LS2_CD03:CAV5_D3342 |  | b54 resonator | 40 |  |  | 0.49566 | 313.949 | 415.3476 | 41.094 | 41.094 | 415.3476 | 314.1968 | 334.2042 | 334.452 |  | 265.7256
[10 empty] | 0 | 314.1968 | 415.3476 | 41.094 | 41.094 | 415.3476 | 314.1968 | 334.452 | 334.452 |  | 265.9734
[10 empty] | 0 | 314.1968 | 415.3476 | 41.094 | 41.094 | 415.3476 | 314.1968 | 334.452 | 334.452 |  | 265.9734
[6 empty] | bellow |  |  |  | 0.099517 | 314.2466 | 415.3476 | 41.094 | 41.094 | 415.3476 | 314.2963 | 334.5018 | 334.5515 |  | 266.0232
LS2 | CD03 | CAV6 | 3348 | LS2_CD03:CAV6_D3348 |  | b54 resonator | 40 |  |  | 0.49566 | 314.5442 | 415.3476 | 41.094 | 41.094 | 415.3476 | 314.792 | 334.7994 | 335.0472 |  | 266.3207
[10 empty] | 0 | 314.792 | 415.3476 | 41.094 | 41.094 | 415.3476 | 314.792 | 335.0472 | 335.0472 |  | 266.5686
[10 empty] | 0 | 314.792 | 415.3476 | 41.094 | 41.094 | 415.3476 | 314.792 | 335.0472 | 335.0472 |  | 266.5686
[6 empty] | bellow |  |  |  | 0.099517 | 314.8417 | 415.3476 | 41.094 | 41.094 | 415.3476 | 314.8915 | 335.097 | 335.1467 |  | 266.6183
LS2 | CD03 | CAV7 | 3354 | LS2_CD03:CAV7_D3354 |  | b54 resonator | 40 |  |  | 0.49566 | 315.1393 | 415.3476 | 41.094 | 41.094 | 415.3476 | 315.3872 | 335.3945 | 335.6424 |  | 266.9159
[10 empty] | 0 | 315.3872 | 415.3476 | 41.094 | 41.094 | 415.3476 | 315.3872 | 335.6424 | 335.6424 |  | 267.1637
[10 empty] | 0 | 315.3872 | 415.3476 | 41.094 | 41.094 | 415.3476 | 315.3872 | 335.6424 | 335.6424 |  | 267.1637
[6 empty] | bellow |  |  |  | 0.099517 | 315.4369 | 415.3476 | 41.094 | 41.094 | 415.3476 | 315.4867 | 335.6921 | 335.7419 |  | 267.2135
LS2 | CD03 | CAV8 | 3360 | LS2_CD03:CAV8_D3360 |  | b54 resonator | 40 |  |  | 0.49566 | 315.7345 | 415.3476 | 41.094 | 41.094 | 415.3476 | 315.9823 | 335.9897 | 336.2376 |  | 267.5111
LS2 | CD03 | BLM | 3362 | LS2_CD03:BLM_D3362 |  | outside loss monitor |  |  |  | 0 | 315.9823 | 415.3476 | 41.094 | 41.094 | 415.3476 | 315.9823 | 336.2376 | 336.2376 |  | 267.7589
[10 empty] | 0 | 315.9823 | 415.3476 | 41.094 | 41.094 | 415.3476 | 315.9823 | 336.2376 | 336.2376 |  | 267.7589
[10 empty] | 0.1677 | 316.0662 | 415.3476 | 41.094 | 41.094 | 415.3476 | 316.15 | 336.3214 | 336.4053 |  | 267.8428
LS2 | CD03 | GV | 3364 | LS2_CD03:GV_D3364 |  | gate valve | 40 |  |  | 0.057404 | 316.1787 | 415.3476 | 41.094 | 41.094 | 415.3476 | 316.2074 | 336.434 | 336.4627 |  | 267.9553
[6 empty] | bellow | 40 |  |  | 0.1035 | 316.2592 | 415.3476 | 41.094 | 41.094 | 415.3476 | 316.3109 | 336.5144 | 336.5662 |  | 268.0358
LS2 | WD03 | BPM | 3366 | LS2_WD03:BPM_D3366 | BPM2 | position monitor |  |  |  | 0 | 316.3109 | 415.3476 | 41.094 | 41.094 | 415.3476 | 316.3109 | 336.5662 | 336.5662 |  | 268.0875
[6 empty] | diagnostic box |  |  |  | 0.14528 | 316.3836 | 415.3476 | 41.094 | 41.094 | 415.3476 | 316.4562 | 336.6388 | 336.7114 |  | 268.1602
LS2 | WD03 | PM | 3367 | LS2_WD03:PM_D3367 | PMx | profile monitor (not costed) | future PM |  |  | 0 | 316.4562 | 415.3476 | 41.094 | 41.094 | 415.3476 | 316.4562 | 336.7114 | 336.7114 |  | 268.2328
[6 empty] | diagnostic box |  |  |  | 0 | 316.4562 | 415.3476 | 41.094 | 41.094 | 415.3476 | 316.4562 | 336.7114 | 336.7114 |  | 268.2328
LS2 | WD03 | IP | 3367 | LS2_WD03:IP_D3367 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 316.4562 | 415.3476 | 41.094 | 41.094 | 415.3476 | 316.4562 | 336.7114 | 336.7114 |  | 268.2328
[6 empty] | bellow | 40 |  |  | 0.13122 | 316.5218 | 415.3476 | 41.094 | 41.094 | 415.3476 | 316.5874 | 336.7771 | 336.8427 |  | 268.2984
LS2 | CD04 | GV | 3369 | LS2_CD04:GV_D3369 |  | gate valve | 40 |  |  | 0.057404 | 316.6161 | 415.3476 | 41.094 | 41.094 | 415.3476 | 316.6448 | 336.8714 | 336.9001 |  | 268.3927
[6 empty] | bellows |  |  |  | 0.1677 | 316.7287 | 415.3476 | 41.094 | 41.094 | 415.3476 | 316.8126 | 336.9839 | 337.0678 |  | 268.5053
LS2 | CD04 | CAV1 | 3373 | LS2_CD04:CAV1_D3373 |  | b54 resonator | 40 |  |  | 0.49566 | 317.0604 | 415.3476 | 41.094 | 41.094 | 415.3476 | 317.3082 | 337.3156 | 337.5634 |  | 268.837
LS2 | CD04 | BLM | 3376 | LS2_CD04:BLM_D3376 |  | outside loss monitor |  |  |  | 0 | 317.3082 | 415.3476 | 41.094 | 41.094 | 415.3476 | 317.3082 | 337.5634 | 337.5634 |  | 269.0848
[10 empty] | 0 | 317.3082 | 415.3476 | 41.094 | 41.094 | 415.3476 | 317.3082 | 337.5634 | 337.5634 |  | 269.0848
[6 empty] | bellow |  |  |  | 0.099517 | 317.358 | 415.3476 | 41.094 | 41.094 | 415.3476 | 317.4077 | 337.6132 | 337.6629 |  | 269.1346
LS2 | CD04 | CAV2 | 3379 | LS2_CD04:CAV2_D3379 |  | b54 resonator | 40 |  |  | 0.49566 | 317.6556 | 415.3476 | 41.094 | 41.094 | 415.3476 | 317.9034 | 337.9108 | 338.1586 |  | 269.4321
[10 empty] | 0 | 317.9034 | 415.3476 | 41.094 | 41.094 | 415.3476 | 317.9034 | 338.1586 | 338.1586 |  | 269.68
[10 empty] | 0 | 317.9034 | 415.3476 | 41.094 | 41.094 | 415.3476 | 317.9034 | 338.1586 | 338.1586 |  | 269.68
[6 empty] | bellow |  |  |  | 0.099517 | 317.9531 | 415.3476 | 41.094 | 41.094 | 415.3476 | 318.0029 | 338.2084 | 338.2581 |  | 269.7297
LS2 | CD04 | CAV3 | 3385 | LS2_CD04:CAV3_D3385 |  | b54 resonator | 40 |  |  | 0.49566 | 318.2507 | 415.3476 | 41.094 | 41.094 | 415.3476 | 318.4986 | 338.5059 | 338.7538 |  | 270.0273
[10 empty] | 0 | 318.4986 | 415.3476 | 41.094 | 41.094 | 415.3476 | 318.4986 | 338.7538 | 338.7538 |  | 270.2751
[10 empty] | 0 | 318.4986 | 415.3476 | 41.094 | 41.094 | 415.3476 | 318.4986 | 338.7538 | 338.7538 |  | 270.2751
[6 empty] | bellow |  |  |  | 0.099517 | 318.5483 | 415.3476 | 41.094 | 41.094 | 415.3476 | 318.5981 | 338.8035 | 338.8533 |  | 270.3249
LS2 | CD04 | CAV4 | 3391 | LS2_CD04:CAV4_D3391 |  | b54 resonator | 40 |  |  | 0.49566 | 318.8459 | 415.3476 | 41.094 | 41.094 | 415.3476 | 319.0937 | 339.1011 | 339.3489 |  | 270.6225
LS2 | CD04 | BLM | 3393 | LS2_CD04:BLM_D3393 |  | outside loss monitor |  |  |  | 0 | 319.0937 | 415.3476 | 41.094 | 41.094 | 415.3476 | 319.0937 | 339.3489 | 339.3489 |  | 270.8703
[10 empty] | 0 | 319.0937 | 415.3476 | 41.094 | 41.094 | 415.3476 | 319.0937 | 339.3489 | 339.3489 |  | 270.8703
[6 empty] | bellow |  |  |  | 0.15525 | 319.1714 | 415.3476 | 41.094 | 41.094 | 415.3476 | 319.249 | 339.4266 | 339.5042 |  | 270.9479
LS2 | CD04 | SOL1 | 3398 | LS2_CD04:SOL1_D3398 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 319.499 | 415.3476 | 41.094 | 41.094 | 415.3476 | 319.749 | 339.7542 | 340.0042 |  | 271.2756
[6 empty] | bellow |  |  |  | 0.15525 | 319.8266 | 415.3476 | 41.094 | 41.094 | 415.3476 | 319.9042 | 340.0818 | 340.1594 |  | 271.6032
LS2 | CD04 | CAV5 | 3404 | LS2_CD04:CAV5_D3404 |  | b54 resonator | 40 |  |  | 0.49566 | 320.1521 | 415.3476 | 41.094 | 41.094 | 415.3476 | 320.3999 | 340.4073 | 340.6551 |  | 271.9286
[10 empty] | 0 | 320.3999 | 415.3476 | 41.094 | 41.094 | 415.3476 | 320.3999 | 340.6551 | 340.6551 |  | 272.1765
[10 empty] | 0 | 320.3999 | 415.3476 | 41.094 | 41.094 | 415.3476 | 320.3999 | 340.6551 | 340.6551 |  | 272.1765
[6 empty] | bellow |  |  |  | 0.099517 | 320.4496 | 415.3476 | 41.094 | 41.094 | 415.3476 | 320.4994 | 340.7049 | 340.7546 |  | 272.2262
LS2 | CD04 | CAV6 | 3410 | LS2_CD04:CAV6_D3410 |  | b54 resonator | 40 |  |  | 0.49566 | 320.7472 | 415.3476 | 41.094 | 41.094 | 415.3476 | 320.9951 | 341.0024 | 341.2503 |  | 272.5238
[10 empty] | 0 | 320.9951 | 415.3476 | 41.094 | 41.094 | 415.3476 | 320.9951 | 341.2503 | 341.2503 |  | 272.7716
[10 empty] | 0 | 320.9951 | 415.3476 | 41.094 | 41.094 | 415.3476 | 320.9951 | 341.2503 | 341.2503 |  | 272.7716
[6 empty] | bellow |  |  |  | 0.099517 | 321.0448 | 415.3476 | 41.094 | 41.094 | 415.3476 | 321.0946 | 341.3 | 341.3498 |  | 272.8214
LS2 | CD04 | CAV7 | 3416 | LS2_CD04:CAV7_D3416 |  | b54 resonator | 40 |  |  | 0.49566 | 321.3424 | 415.3476 | 41.094 | 41.094 | 415.3476 | 321.5902 | 341.5976 | 341.8454 |  | 273.119
[10 empty] | 0 | 321.5902 | 415.3476 | 41.094 | 41.094 | 415.3476 | 321.5902 | 341.8454 | 341.8454 |  | 273.3668
[10 empty] | 0 | 321.5902 | 415.3476 | 41.094 | 41.094 | 415.3476 | 321.5902 | 341.8454 | 341.8454 |  | 273.3668
[6 empty] | bellow |  |  |  | 0.099517 | 321.64 | 415.3476 | 41.094 | 41.094 | 415.3476 | 321.6897 | 341.8952 | 341.945 |  | 273.4166
LS2 | CD04 | CAV8 | 3422 | LS2_CD04:CAV8_D3422 |  | b54 resonator | 40 |  |  | 0.49566 | 321.9376 | 415.3476 | 41.094 | 41.094 | 415.3476 | 322.1854 | 342.1928 | 342.4406 |  | 273.7142
LS2 | CD04 | BLM | 3424 | LS2_CD04:BLM_D3424 |  | outside loss monitor |  |  |  | 0 | 322.1854 | 415.3476 | 41.094 | 41.094 | 415.3476 | 322.1854 | 342.4406 | 342.4406 |  | 273.962
[10 empty] | 0 | 322.1854 | 415.3476 | 41.094 | 41.094 | 415.3476 | 322.1854 | 342.4406 | 342.4406 |  | 273.962
[10 empty] | 0.1677 | 322.2693 | 415.3476 | 41.094 | 41.094 | 415.3476 | 322.3531 | 342.5245 | 342.6083 |  | 274.0458
LS2 | CD04 | GV | 3426 | LS2_CD04:GV_D3426 |  | gate valve | 40 |  |  | 0.057404 | 322.3818 | 415.3476 | 41.094 | 41.094 | 415.3476 | 322.4105 | 342.637 | 342.6657 |  | 274.1584
[6 empty] | bellow | 40 |  |  | 0.1035 | 322.4623 | 415.3476 | 41.094 | 41.094 | 415.3476 | 322.514 | 342.7175 | 342.7692 |  | 274.2388
LS2 | WD04 | BPM | 3428 | LS2_WD04:BPM_D3428 | BPM2 | position monitor |  |  |  | 0 | 322.514 | 415.3476 | 41.094 | 41.094 | 415.3476 | 322.514 | 342.7692 | 342.7692 |  | 274.2906
[6 empty] | diagnostic box |  |  |  | 0.14528 | 322.5867 | 415.3476 | 41.094 | 41.094 | 415.3476 | 322.6593 | 342.8419 | 342.9145 |  | 274.3632
LS2 | WD04 | PORT | 3429 | LS2_WD04:PORT_D3429 |  | port for future diagnostics |  |  |  | 0 | 322.6593 | 415.3476 | 41.094 | 41.094 | 415.3476 | 322.6593 | 342.9145 | 342.9145 |  | 274.4359
[6 empty] | diagnostic box |  |  |  | 0 | 322.6593 | 415.3476 | 41.094 | 41.094 | 415.3476 | 322.6593 | 342.9145 | 342.9145 |  | 274.4359
LS2 | WD04 | IP | 3429 | LS2_WD04:IP_D3429 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 322.6593 | 415.3476 | 41.094 | 41.094 | 415.3476 | 322.6593 | 342.9145 | 342.9145 |  | 274.4359
[6 empty] | bellow | 40 |  |  | 0.13122 | 322.7249 | 415.3476 | 41.094 | 41.094 | 415.3476 | 322.7905 | 342.9801 | 343.0457 |  | 274.5015
LS2 | CD05 | GV | 3431 | LS2_CD05:GV_D3431 |  | gate valve | 40 |  |  | 0.057404 | 322.8192 | 415.3476 | 41.094 | 41.094 | 415.3476 | 322.8479 | 343.0744 | 343.1031 |  | 274.5958
[6 empty] | bellows |  |  |  | 0.1677 | 322.9318 | 415.3476 | 41.094 | 41.094 | 415.3476 | 323.0156 | 343.187 | 343.2708 |  | 274.7084
LS2 | CD05 | CAV1 | 3435 | LS2_CD05:CAV1_D3435 |  | b54 resonator | 40 |  |  | 0.49566 | 323.2634 | 415.3476 | 41.094 | 41.094 | 415.3476 | 323.5113 | 343.5187 | 343.7665 |  | 275.04
LS2 | CD05 | BLM | 3438 | LS2_CD05:BLM_D3438 |  | outside loss monitor |  |  |  | 0 | 323.5113 | 415.3476 | 41.094 | 41.094 | 415.3476 | 323.5113 | 343.7665 | 343.7665 |  | 275.2879
[10 empty] | 0 | 323.5113 | 415.3476 | 41.094 | 41.094 | 415.3476 | 323.5113 | 343.7665 | 343.7665 |  | 275.2879
[6 empty] | bellow |  |  |  | 0.099517 | 323.561 | 415.3476 | 41.094 | 41.094 | 415.3476 | 323.6108 | 343.8162 | 343.866 |  | 275.3376
LS2 | CD05 | CAV2 | 3441 | LS2_CD05:CAV2_D3441 |  | b54 resonator | 40 |  |  | 0.49566 | 323.8586 | 415.3476 | 41.094 | 41.094 | 415.3476 | 324.1064 | 344.1138 | 344.3617 |  | 275.6352
[10 empty] | 0 | 324.1064 | 415.3476 | 41.094 | 41.094 | 415.3476 | 324.1064 | 344.3617 | 344.3617 |  | 275.883
[10 empty] | 0 | 324.1064 | 415.3476 | 41.094 | 41.094 | 415.3476 | 324.1064 | 344.3617 | 344.3617 |  | 275.883
[6 empty] | bellow |  |  |  | 0.099517 | 324.1562 | 415.3476 | 41.094 | 41.094 | 415.3476 | 324.206 | 344.4114 | 344.4612 |  | 275.9328
LS2 | CD05 | CAV3 | 3447 | LS2_CD05:CAV3_D3447 |  | b54 resonator | 40 |  |  | 0.49566 | 324.4538 | 415.3476 | 41.094 | 41.094 | 415.3476 | 324.7016 | 344.709 | 344.9568 |  | 276.2304
[10 empty] | 0 | 324.7016 | 415.3476 | 41.094 | 41.094 | 415.3476 | 324.7016 | 344.9568 | 344.9568 |  | 276.4782
[10 empty] | 0 | 324.7016 | 415.3476 | 41.094 | 41.094 | 415.3476 | 324.7016 | 344.9568 | 344.9568 |  | 276.4782
[6 empty] | bellow |  |  |  | 0.099517 | 324.7514 | 415.3476 | 41.094 | 41.094 | 415.3476 | 324.8011 | 345.0066 | 345.0564 |  | 276.528
LS2 | CD05 | CAV4 | 3453 | LS2_CD05:CAV4_D3453 |  | b54 resonator | 40 |  |  | 0.49566 | 325.049 | 415.3476 | 41.094 | 41.094 | 415.3476 | 325.2968 | 345.3042 | 345.552 |  | 276.8256
LS2 | CD05 | BLM | 3456 | LS2_CD05:BLM_D3456 |  | outside loss monitor |  |  |  | 0 | 325.2968 | 415.3476 | 41.094 | 41.094 | 415.3476 | 325.2968 | 345.552 | 345.552 |  | 277.0734
[10 empty] | 0 | 325.2968 | 415.3476 | 41.094 | 41.094 | 415.3476 | 325.2968 | 345.552 | 345.552 |  | 277.0734
[6 empty] | bellow |  |  |  | 0.15525 | 325.3744 | 415.3476 | 41.094 | 41.094 | 415.3476 | 325.452 | 345.6296 | 345.7073 |  | 277.151
LS2 | CD05 | SOL1 | 3460 | LS2_CD05:SOL1_D3460 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 325.702 | 415.3476 | 41.094 | 41.094 | 415.3476 | 325.952 | 345.9573 | 346.2073 |  | 277.4786
[6 empty] | bellow |  |  |  | 0.15525 | 326.0297 | 415.3476 | 41.094 | 41.094 | 415.3476 | 326.1073 | 346.2849 | 346.3625 |  | 277.8063
LS2 | CD05 | CAV5 | 3466 | LS2_CD05:CAV5_D3466 |  | b54 resonator | 40 |  |  | 0.49566 | 326.3551 | 415.3476 | 41.094 | 41.094 | 415.3476 | 326.6029 | 346.6103 | 346.8582 |  | 278.1317
[10 empty] | 0 | 326.6029 | 415.3476 | 41.094 | 41.094 | 415.3476 | 326.6029 | 346.8582 | 346.8582 |  | 278.3795
[10 empty] | 0 | 326.6029 | 415.3476 | 41.094 | 41.094 | 415.3476 | 326.6029 | 346.8582 | 346.8582 |  | 278.3795
[6 empty] | bellow |  |  |  | 0.099517 | 326.6527 | 415.3476 | 41.094 | 41.094 | 415.3476 | 326.7025 | 346.9079 | 346.9577 |  | 278.4293
LS2 | CD05 | CAV6 | 3472 | LS2_CD05:CAV6_D3472 |  | b54 resonator | 40 |  |  | 0.49566 | 326.9503 | 415.3476 | 41.094 | 41.094 | 415.3476 | 327.1981 | 347.2055 | 347.4533 |  | 278.7269
[10 empty] | 0 | 327.1981 | 415.3476 | 41.094 | 41.094 | 415.3476 | 327.1981 | 347.4533 | 347.4533 |  | 278.9747
[10 empty] | 0 | 327.1981 | 415.3476 | 41.094 | 41.094 | 415.3476 | 327.1981 | 347.4533 | 347.4533 |  | 278.9747
[6 empty] | bellow |  |  |  | 0.099517 | 327.2479 | 415.3476 | 41.094 | 41.094 | 415.3476 | 327.2976 | 347.5031 | 347.5529 |  | 279.0245
LS2 | CD05 | CAV7 | 3478 | LS2_CD05:CAV7_D3478 |  | b54 resonator | 40 |  |  | 0.49566 | 327.5455 | 415.3476 | 41.094 | 41.094 | 415.3476 | 327.7933 | 347.8007 | 348.0485 |  | 279.3221
[10 empty] | 0 | 327.7933 | 415.3476 | 41.094 | 41.094 | 415.3476 | 327.7933 | 348.0485 | 348.0485 |  | 279.5699
[10 empty] | 0 | 327.7933 | 415.3476 | 41.094 | 41.094 | 415.3476 | 327.7933 | 348.0485 | 348.0485 |  | 279.5699
[6 empty] | bellow |  |  |  | 0.099517 | 327.8431 | 415.3476 | 41.094 | 41.094 | 415.3476 | 327.8928 | 348.0983 | 348.148 |  | 279.6196
LS2 | CD05 | CAV8 | 3484 | LS2_CD05:CAV8_D3484 |  | b54 resonator | 40 |  |  | 0.49566 | 328.1406 | 415.3476 | 41.094 | 41.094 | 415.3476 | 328.3885 | 348.3959 | 348.6437 |  | 279.9172
LS2 | CD05 | BLM | 3486 | LS2_CD05:BLM_D3486 |  | outside loss monitor |  |  |  | 0 | 328.3885 | 415.3476 | 41.094 | 41.094 | 415.3476 | 328.3885 | 348.6437 | 348.6437 |  | 280.1651
[10 empty] | 0 | 328.3885 | 415.3476 | 41.094 | 41.094 | 415.3476 | 328.3885 | 348.6437 | 348.6437 |  | 280.1651
[10 empty] | 0.1677 | 328.4723 | 415.3476 | 41.094 | 41.094 | 415.3476 | 328.5562 | 348.7275 | 348.8114 |  | 280.2489
LS2 | CD05 | GV | 3488 | LS2_CD05:GV_D3488 |  | gate valve | 40 |  |  | 0.057404 | 328.5849 | 415.3476 | 41.094 | 41.094 | 415.3476 | 328.6136 | 348.8401 | 348.8688 |  | 280.3615
[6 empty] | bellow | 40 |  |  | 0.1035 | 328.6653 | 415.3476 | 41.094 | 41.094 | 415.3476 | 328.7171 | 348.9205 | 348.9723 |  | 280.4419
LS2 | WD05 | BPM | 3490 | LS2_WD05:BPM_D3490 | BPM2 | position monitor |  |  |  | 0 | 328.7171 | 415.3476 | 41.094 | 41.094 | 415.3476 | 328.7171 | 348.9723 | 348.9723 |  | 280.4937
[6 empty] | diagnostic box |  |  |  | 0.14528 | 328.7897 | 415.3476 | 41.094 | 41.094 | 415.3476 | 328.8624 | 349.0449 | 349.1176 |  | 280.5663
LS2 | WD05 | PORT | 3491 | LS2_WD05:PORT_D3491 |  | port for future diagnostics |  |  |  | 0 | 328.8624 | 415.3476 | 41.094 | 41.094 | 415.3476 | 328.8624 | 349.1176 | 349.1176 |  | 280.6389
[6 empty] | diagnostic box |  |  |  | 0 | 328.8624 | 415.3476 | 41.094 | 41.094 | 415.3476 | 328.8624 | 349.1176 | 349.1176 |  | 280.6389
LS2 | WD05 | IP | 3491 | LS2_WD05:IP_D3491 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 328.8624 | 415.3476 | 41.094 | 41.094 | 415.3476 | 328.8624 | 349.1176 | 349.1176 |  | 280.6389
[6 empty] | bellow | 40 |  |  | 0.13122 | 328.928 | 415.3476 | 41.094 | 41.094 | 415.3476 | 328.9936 | 349.1832 | 349.2488 |  | 280.7046
LS2 | CD06 | GV | 3493 | LS2_CD06:GV_D3493 |  | gate valve | 40 |  |  | 0.057404 | 329.0223 | 415.3476 | 41.094 | 41.094 | 415.3476 | 329.051 | 349.2775 | 349.3062 |  | 280.7989
[6 empty] | bellows |  |  |  | 0.1677 | 329.1348 | 415.3476 | 41.094 | 41.094 | 415.3476 | 329.2187 | 349.39 | 349.4739 |  | 280.9114
LS2 | CD06 | CAV1 | 3497 | LS2_CD06:CAV1_D3497 |  | b54 resonator | 40 |  |  | 0.49566 | 329.4665 | 415.3476 | 41.094 | 41.094 | 415.3476 | 329.7143 | 349.7217 | 349.9696 |  | 281.2431
LS2 | CD06 | BLM | 3500 | LS2_CD06:BLM_D3500 |  | outside loss monitor |  |  |  | 0 | 329.7143 | 415.3476 | 41.094 | 41.094 | 415.3476 | 329.7143 | 349.9696 | 349.9696 |  | 281.4909
[10 empty] | 0 | 329.7143 | 415.3476 | 41.094 | 41.094 | 415.3476 | 329.7143 | 349.9696 | 349.9696 |  | 281.4909
[6 empty] | bellow |  |  |  | 0.099517 | 329.7641 | 415.3476 | 41.094 | 41.094 | 415.3476 | 329.8139 | 350.0193 | 350.0691 |  | 281.5407
LS2 | CD06 | CAV2 | 3503 | LS2_CD06:CAV2_D3503 |  | b54 resonator | 40 |  |  | 0.49566 | 330.0617 | 415.3476 | 41.094 | 41.094 | 415.3476 | 330.3095 | 350.3169 | 350.5647 |  | 281.8383
[10 empty] | 0 | 330.3095 | 415.3476 | 41.094 | 41.094 | 415.3476 | 330.3095 | 350.5647 | 350.5647 |  | 282.0861
[10 empty] | 0 | 330.3095 | 415.3476 | 41.094 | 41.094 | 415.3476 | 330.3095 | 350.5647 | 350.5647 |  | 282.0861
[6 empty] | bellow |  |  |  | 0.099517 | 330.3593 | 415.3476 | 41.094 | 41.094 | 415.3476 | 330.409 | 350.6145 | 350.6642 |  | 282.1359
LS2 | CD06 | CAV3 | 3509 | LS2_CD06:CAV3_D3509 |  | b54 resonator | 40 |  |  | 0.49566 | 330.6569 | 415.3476 | 41.094 | 41.094 | 415.3476 | 330.9047 | 350.9121 | 351.1599 |  | 282.4334
[10 empty] | 0 | 330.9047 | 415.3476 | 41.094 | 41.094 | 415.3476 | 330.9047 | 351.1599 | 351.1599 |  | 282.6813
[10 empty] | 0 | 330.9047 | 415.3476 | 41.094 | 41.094 | 415.3476 | 330.9047 | 351.1599 | 351.1599 |  | 282.6813
[6 empty] | bellow |  |  |  | 0.099517 | 330.9544 | 415.3476 | 41.094 | 41.094 | 415.3476 | 331.0042 | 351.2097 | 351.2594 |  | 282.731
LS2 | CD06 | CAV4 | 3515 | LS2_CD06:CAV4_D3515 |  | b54 resonator | 40 |  |  | 0.49566 | 331.252 | 415.3476 | 41.094 | 41.094 | 415.3476 | 331.4999 | 351.5072 | 351.7551 |  | 283.0286
LS2 | CD06 | BLM | 3518 | LS2_CD06:BLM_D3518 |  | outside loss monitor |  |  |  | 0 | 331.4999 | 415.3476 | 41.094 | 41.094 | 415.3476 | 331.4999 | 351.7551 | 351.7551 |  | 283.2764
[10 empty] | 0 | 331.4999 | 415.3476 | 41.094 | 41.094 | 415.3476 | 331.4999 | 351.7551 | 351.7551 |  | 283.2764
[6 empty] | bellow |  |  |  | 0.15525 | 331.5775 | 415.3476 | 41.094 | 41.094 | 415.3476 | 331.6551 | 351.8327 | 351.9103 |  | 283.3541
LS2 | CD06 | SOL1 | 3522 | LS2_CD06:SOL1_D3522 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 331.9051 | 415.3476 | 41.094 | 41.094 | 415.3476 | 332.1551 | 352.1603 | 352.4103 |  | 283.6817
[6 empty] | bellow |  |  |  | 0.15525 | 332.2327 | 415.3476 | 41.094 | 41.094 | 415.3476 | 332.3104 | 352.4879 | 352.5656 |  | 284.0093
LS2 | CD06 | CAV5 | 3528 | LS2_CD06:CAV5_D3528 |  | b54 resonator | 40 |  |  | 0.49566 | 332.5582 | 415.3476 | 41.094 | 41.094 | 415.3476 | 332.806 | 352.8134 | 353.0612 |  | 284.3348
[10 empty] | 0 | 332.806 | 415.3476 | 41.094 | 41.094 | 415.3476 | 332.806 | 353.0612 | 353.0612 |  | 284.5826
[10 empty] | 0 | 332.806 | 415.3476 | 41.094 | 41.094 | 415.3476 | 332.806 | 353.0612 | 353.0612 |  | 284.5826
[6 empty] | bellow |  |  |  | 0.099517 | 332.8558 | 415.3476 | 41.094 | 41.094 | 415.3476 | 332.9055 | 353.111 | 353.1607 |  | 284.6324
LS2 | CD06 | CAV6 | 3534 | LS2_CD06:CAV6_D3534 |  | b54 resonator | 40 |  |  | 0.49566 | 333.1534 | 415.3476 | 41.094 | 41.094 | 415.3476 | 333.4012 | 353.4086 | 353.6564 |  | 284.9299
[10 empty] | 0 | 333.4012 | 415.3476 | 41.094 | 41.094 | 415.3476 | 333.4012 | 353.6564 | 353.6564 |  | 285.1778
[10 empty] | 0 | 333.4012 | 415.3476 | 41.094 | 41.094 | 415.3476 | 333.4012 | 353.6564 | 353.6564 |  | 285.1778
[6 empty] | bellow |  |  |  | 0.099517 | 333.4509 | 415.3476 | 41.094 | 41.094 | 415.3476 | 333.5007 | 353.7062 | 353.7559 |  | 285.2275
LS2 | CD06 | CAV7 | 3540 | LS2_CD06:CAV7_D3540 |  | b54 resonator | 40 |  |  | 0.49566 | 333.7485 | 415.3476 | 41.094 | 41.094 | 415.3476 | 333.9964 | 354.0037 | 354.2516 |  | 285.5251
[10 empty] | 0 | 333.9964 | 415.3476 | 41.094 | 41.094 | 415.3476 | 333.9964 | 354.2516 | 354.2516 |  | 285.7729
[10 empty] | 0 | 333.9964 | 415.3476 | 41.094 | 41.094 | 415.3476 | 333.9964 | 354.2516 | 354.2516 |  | 285.7729
[6 empty] | bellow |  |  |  | 0.099517 | 334.0461 | 415.3476 | 41.094 | 41.094 | 415.3476 | 334.0959 | 354.3013 | 354.3511 |  | 285.8227
LS2 | CD06 | CAV8 | 3546 | LS2_CD06:CAV8_D3546 |  | b54 resonator | 40 |  |  | 0.49566 | 334.3437 | 415.3476 | 41.094 | 41.094 | 415.3476 | 334.5915 | 354.5989 | 354.8467 |  | 286.1203
LS2 | CD06 | BLM | 3548 | LS2_CD06:BLM_D3548 |  | outside loss monitor |  |  |  | 0 | 334.5915 | 415.3476 | 41.094 | 41.094 | 415.3476 | 334.5915 | 354.8467 | 354.8467 |  | 286.3681
[10 empty] | 0 | 334.5915 | 415.3476 | 41.094 | 41.094 | 415.3476 | 334.5915 | 354.8467 | 354.8467 |  | 286.3681
[10 empty] | 0.1677 | 334.6754 | 415.3476 | 41.094 | 41.094 | 415.3476 | 334.7592 | 354.9306 | 355.0145 |  | 286.452
LS2 | CD06 | GV | 3550 | LS2_CD06:GV_D3550 |  | gate valve | 40 |  |  | 0.057404 | 334.7879 | 415.3476 | 41.094 | 41.094 | 415.3476 | 334.8166 | 355.0432 | 355.0719 |  | 286.5645
[6 empty] | bellow | 40 |  |  | 0.1035 | 334.8684 | 415.3476 | 41.094 | 41.094 | 415.3476 | 334.9201 | 355.1236 | 355.1754 |  | 286.645
LS2 | WD06 | BPM | 3552 | LS2_WD06:BPM_D3552 | BPM2 | position monitor |  |  |  | 0 | 334.9201 | 415.3476 | 41.094 | 41.094 | 415.3476 | 334.9201 | 355.1754 | 355.1754 |  | 286.6967
[6 empty] | diagnostic box |  |  |  | 0.14528 | 334.9928 | 415.3476 | 41.094 | 41.094 | 415.3476 | 335.0654 | 355.248 | 355.3206 |  | 286.7694
LS2 | WD06 | PORT | 3553 | LS2_WD06:PORT_D3553 |  | port for future diagnostics |  |  |  | 0 | 335.0654 | 415.3476 | 41.094 | 41.094 | 415.3476 | 335.0654 | 355.3206 | 355.3206 |  | 286.842
[6 empty] | diagnostic box |  |  |  | 0 | 335.0654 | 415.3476 | 41.094 | 41.094 | 415.3476 | 335.0654 | 355.3206 | 355.3206 |  | 286.842
LS2 | WD06 | IP | 3553 | LS2_WD06:IP_D3553 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 335.0654 | 415.3476 | 41.094 | 41.094 | 415.3476 | 335.0654 | 355.3206 | 355.3206 |  | 286.842
[6 empty] | bellow | 40 |  |  | 0.13122 | 335.131 | 415.3476 | 41.094 | 41.094 | 415.3476 | 335.1966 | 355.3862 | 355.4519 |  | 286.9076
LS2 | CD07 | GV | 3555 | LS2_CD07:GV_D3555 |  | gate valve | 40 |  |  | 0.057404 | 335.2253 | 415.3476 | 41.094 | 41.094 | 415.3476 | 335.254 | 355.4806 | 355.5093 |  | 287.0019
[6 empty] | bellows |  |  |  | 0.1677 | 335.3379 | 415.3476 | 41.094 | 41.094 | 415.3476 | 335.4217 | 355.5931 | 355.677 |  | 287.1145
LS2 | CD07 | CAV1 | 3559 | LS2_CD07:CAV1_D3559 |  | b54 resonator | 40 |  |  | 0.49566 | 335.6696 | 415.3476 | 41.094 | 41.094 | 415.3476 | 335.9174 | 355.9248 | 356.1726 |  | 287.4462
LS2 | CD07 | BLM | 3562 | LS2_CD07:BLM_D3562 |  | outside loss monitor |  |  |  | 0 | 335.9174 | 415.3476 | 41.094 | 41.094 | 415.3476 | 335.9174 | 356.1726 | 356.1726 |  | 287.694
[10 empty] | 0 | 335.9174 | 415.3476 | 41.094 | 41.094 | 415.3476 | 335.9174 | 356.1726 | 356.1726 |  | 287.694
[6 empty] | bellow |  |  |  | 0.099517 | 335.9672 | 415.3476 | 41.094 | 41.094 | 415.3476 | 336.0169 | 356.2224 | 356.2721 |  | 287.7437
LS2 | CD07 | CAV2 | 3565 | LS2_CD07:CAV2_D3565 |  | b54 resonator | 40 |  |  | 0.49566 | 336.2647 | 415.3476 | 41.094 | 41.094 | 415.3476 | 336.5126 | 356.52 | 356.7678 |  | 288.0413
[10 empty] | 0 | 336.5126 | 415.3476 | 41.094 | 41.094 | 415.3476 | 336.5126 | 356.7678 | 356.7678 |  | 288.2892
[10 empty] | 0 | 336.5126 | 415.3476 | 41.094 | 41.094 | 415.3476 | 336.5126 | 356.7678 | 356.7678 |  | 288.2892
[6 empty] | bellow |  |  |  | 0.099517 | 336.5623 | 415.3476 | 41.094 | 41.094 | 415.3476 | 336.6121 | 356.8176 | 356.8673 |  | 288.3389
LS2 | CD07 | CAV3 | 3571 | LS2_CD07:CAV3_D3571 |  | b54 resonator | 40 |  |  | 0.49566 | 336.8599 | 415.3476 | 41.094 | 41.094 | 415.3476 | 337.1078 | 357.1151 | 357.363 |  | 288.6365
[10 empty] | 0 | 337.1078 | 415.3476 | 41.094 | 41.094 | 415.3476 | 337.1078 | 357.363 | 357.363 |  | 288.8843
[10 empty] | 0 | 337.1078 | 415.3476 | 41.094 | 41.094 | 415.3476 | 337.1078 | 357.363 | 357.363 |  | 288.8843
[6 empty] | bellow |  |  |  | 0.099517 | 337.1575 | 415.3476 | 41.094 | 41.094 | 415.3476 | 337.2073 | 357.4127 | 357.4625 |  | 288.9341
LS2 | CD07 | CAV4 | 3577 | LS2_CD07:CAV4_D3577 |  | b54 resonator | 40 |  |  | 0.49566 | 337.4551 | 415.3476 | 41.094 | 41.094 | 415.3476 | 337.7029 | 357.7103 | 357.9581 |  | 289.2317
LS2 | CD07 | BLM | 3580 | LS2_CD07:BLM_D3580 |  | outside loss monitor |  |  |  | 0 | 337.7029 | 415.3476 | 41.094 | 41.094 | 415.3476 | 337.7029 | 357.9581 | 357.9581 |  | 289.4795
[10 empty] | 0 | 337.7029 | 415.3476 | 41.094 | 41.094 | 415.3476 | 337.7029 | 357.9581 | 357.9581 |  | 289.4795
[6 empty] | bellow |  |  |  | 0.15525 | 337.7805 | 415.3476 | 41.094 | 41.094 | 415.3476 | 337.8582 | 358.0358 | 358.1134 |  | 289.5571
LS2 | CD07 | SOL1 | 3584 | LS2_CD07:SOL1_D3584 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 338.1082 | 415.3476 | 41.094 | 41.094 | 415.3476 | 338.3582 | 358.3634 | 358.6134 |  | 289.8848
[6 empty] | bellow |  |  |  | 0.15525 | 338.4358 | 415.3476 | 41.094 | 41.094 | 415.3476 | 338.5134 | 358.691 | 358.7686 |  | 290.2124
LS2 | CD07 | CAV5 | 3590 | LS2_CD07:CAV5_D3590 |  | b54 resonator | 40 |  |  | 0.49566 | 338.7613 | 415.3476 | 41.094 | 41.094 | 415.3476 | 339.0091 | 359.0165 | 359.2643 |  | 290.5378
[10 empty] | 0 | 339.0091 | 415.3476 | 41.094 | 41.094 | 415.3476 | 339.0091 | 359.2643 | 359.2643 |  | 290.7857
[10 empty] | 0 | 339.0091 | 415.3476 | 41.094 | 41.094 | 415.3476 | 339.0091 | 359.2643 | 359.2643 |  | 290.7857
[6 empty] | bellow |  |  |  | 0.099517 | 339.0588 | 415.3476 | 41.094 | 41.094 | 415.3476 | 339.1086 | 359.3141 | 359.3638 |  | 290.8354
LS2 | CD07 | CAV6 | 3596 | LS2_CD07:CAV6_D3596 |  | b54 resonator | 40 |  |  | 0.49566 | 339.3564 | 415.3476 | 41.094 | 41.094 | 415.3476 | 339.6043 | 359.6116 | 359.8595 |  | 291.133
[10 empty] | 0 | 339.6043 | 415.3476 | 41.094 | 41.094 | 415.3476 | 339.6043 | 359.8595 | 359.8595 |  | 291.3808
[10 empty] | 0 | 339.6043 | 415.3476 | 41.094 | 41.094 | 415.3476 | 339.6043 | 359.8595 | 359.8595 |  | 291.3808
[6 empty] | bellow |  |  |  | 0.099517 | 339.654 | 415.3476 | 41.094 | 41.094 | 415.3476 | 339.7038 | 359.9092 | 359.959 |  | 291.4306
LS2 | CD07 | CAV7 | 3602 | LS2_CD07:CAV7_D3602 |  | b54 resonator | 40 |  |  | 0.49566 | 339.9516 | 415.3476 | 41.094 | 41.094 | 415.3476 | 340.1994 | 360.2068 | 360.4546 |  | 291.7282
[10 empty] | 0 | 340.1994 | 415.3476 | 41.094 | 41.094 | 415.3476 | 340.1994 | 360.4546 | 360.4546 |  | 291.976
[10 empty] | 0 | 340.1994 | 415.3476 | 41.094 | 41.094 | 415.3476 | 340.1994 | 360.4546 | 360.4546 |  | 291.976
[6 empty] | bellow |  |  |  | 0.099517 | 340.2492 | 415.3476 | 41.094 | 41.094 | 415.3476 | 340.2989 | 360.5044 | 360.5542 |  | 292.0258
LS2 | CD07 | CAV8 | 3608 | LS2_CD07:CAV8_D3608 |  | b54 resonator | 40 |  |  | 0.49566 | 340.5468 | 415.3476 | 41.094 | 41.094 | 415.3476 | 340.7946 | 360.802 | 361.0498 |  | 292.3234
LS2 | CD07 | BLM | 3610 | LS2_CD07:BLM_D3610 |  | outside loss monitor |  |  |  | 0 | 340.7946 | 415.3476 | 41.094 | 41.094 | 415.3476 | 340.7946 | 361.0498 | 361.0498 |  | 292.5712
[10 empty] | 0 | 340.7946 | 415.3476 | 41.094 | 41.094 | 415.3476 | 340.7946 | 361.0498 | 361.0498 |  | 292.5712
[10 empty] | 0.1677 | 340.8785 | 415.3476 | 41.094 | 41.094 | 415.3476 | 340.9623 | 361.1337 | 361.2175 |  | 292.655
LS2 | CD07 | GV | 3612 | LS2_CD07:GV_D3612 |  | gate valve | 40 |  |  | 0.057404 | 340.991 | 415.3476 | 41.094 | 41.094 | 415.3476 | 341.0197 | 361.2462 | 361.2749 |  | 292.7676
[6 empty] | bellow | 40 |  |  | 0.1035 | 341.0715 | 415.3476 | 41.094 | 41.094 | 415.3476 | 341.1232 | 361.3267 | 361.3784 |  | 292.848
LS2 | WD07 | BPM | 3614 | LS2_WD07:BPM_D3614 | BPM2 | position monitor |  |  |  | 0 | 341.1232 | 415.3476 | 41.094 | 41.094 | 415.3476 | 341.1232 | 361.3784 | 361.3784 |  | 292.8998
[6 empty] | diagnostic box |  |  |  | 0.14528 | 341.1958 | 415.3476 | 41.094 | 41.094 | 415.3476 | 341.2685 | 361.4511 | 361.5237 |  | 292.9724
LS2 | WD07 | PORT | 3615 | LS2_WD07:PORT_D3615 |  | port for future diagnostics |  |  |  | 0 | 341.2685 | 415.3476 | 41.094 | 41.094 | 415.3476 | 341.2685 | 361.5237 | 361.5237 |  | 293.0451
[6 empty] | diagnostic box |  |  |  | 0 | 341.2685 | 415.3476 | 41.094 | 41.094 | 415.3476 | 341.2685 | 361.5237 | 361.5237 |  | 293.0451
LS2 | WD07 | IP | 3615 | LS2_WD07:IP_D3615 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 341.2685 | 415.3476 | 41.094 | 41.094 | 415.3476 | 341.2685 | 361.5237 | 361.5237 |  | 293.0451
[6 empty] | bellow | 40 |  |  | 0.13122 | 341.3341 | 415.3476 | 41.094 | 41.094 | 415.3476 | 341.3997 | 361.5893 | 361.6549 |  | 293.1107
LS2 | CD08 | GV | 3617 | LS2_CD08:GV_D3617 |  | gate valve | 40 |  |  | 0.057404 | 341.4284 | 415.3476 | 41.094 | 41.094 | 415.3476 | 341.4571 | 361.6836 | 361.7123 |  | 293.205
[6 empty] | bellows |  |  |  | 0.1677 | 341.541 | 415.3476 | 41.094 | 41.094 | 415.3476 | 341.6248 | 361.7962 | 361.88 |  | 293.3175
LS2 | CD08 | CAV1 | 3621 | LS2_CD08:CAV1_D3621 |  | b54 resonator | 40 |  |  | 0.49566 | 341.8726 | 415.3476 | 41.094 | 41.094 | 415.3476 | 342.1205 | 362.1279 | 362.3757 |  | 293.6492
LS2 | CD08 | BLM | 3624 | LS2_CD08:BLM_D3624 |  | outside loss monitor |  |  |  | 0 | 342.1205 | 415.3476 | 41.094 | 41.094 | 415.3476 | 342.1205 | 362.3757 | 362.3757 |  | 293.8971
[10 empty] | 0 | 342.1205 | 415.3476 | 41.094 | 41.094 | 415.3476 | 342.1205 | 362.3757 | 362.3757 |  | 293.8971
[6 empty] | bellow |  |  |  | 0.099517 | 342.1702 | 415.3476 | 41.094 | 41.094 | 415.3476 | 342.22 | 362.4254 | 362.4752 |  | 293.9468
LS2 | CD08 | CAV2 | 3627 | LS2_CD08:CAV2_D3627 |  | b54 resonator | 40 |  |  | 0.49566 | 342.4678 | 415.3476 | 41.094 | 41.094 | 415.3476 | 342.7156 | 362.723 | 362.9709 |  | 294.2444
[10 empty] | 0 | 342.7156 | 415.3476 | 41.094 | 41.094 | 415.3476 | 342.7156 | 362.9709 | 362.9709 |  | 294.4922
[10 empty] | 0 | 342.7156 | 415.3476 | 41.094 | 41.094 | 415.3476 | 342.7156 | 362.9709 | 362.9709 |  | 294.4922
[6 empty] | bellow |  |  |  | 0.099517 | 342.7654 | 415.3476 | 41.094 | 41.094 | 415.3476 | 342.8152 | 363.0206 | 363.0704 |  | 294.542
LS2 | CD08 | CAV3 | 3633 | LS2_CD08:CAV3_D3633 |  | b54 resonator | 40 |  |  | 0.49566 | 343.063 | 415.3476 | 41.094 | 41.094 | 415.3476 | 343.3108 | 363.3182 | 363.566 |  | 294.8396
[10 empty] | 0 | 343.3108 | 415.3476 | 41.094 | 41.094 | 415.3476 | 343.3108 | 363.566 | 363.566 |  | 295.0874
[10 empty] | 0 | 343.3108 | 415.3476 | 41.094 | 41.094 | 415.3476 | 343.3108 | 363.566 | 363.566 |  | 295.0874
[6 empty] | bellow |  |  |  | 0.099517 | 343.3606 | 415.3476 | 41.094 | 41.094 | 415.3476 | 343.4103 | 363.6158 | 363.6655 |  | 295.1372
LS2 | CD08 | CAV4 | 3639 | LS2_CD08:CAV4_D3639 |  | b54 resonator | 40 |  |  | 0.49566 | 343.6582 | 415.3476 | 41.094 | 41.094 | 415.3476 | 343.906 | 363.9134 | 364.1612 |  | 295.4347
LS2 | CD08 | BLM | 3642 | LS2_CD08:BLM_D3642 |  | outside loss monitor |  |  |  | 0 | 343.906 | 415.3476 | 41.094 | 41.094 | 415.3476 | 343.906 | 364.1612 | 364.1612 |  | 295.6826
[10 empty] | 0 | 343.906 | 415.3476 | 41.094 | 41.094 | 415.3476 | 343.906 | 364.1612 | 364.1612 |  | 295.6826
[6 empty] | bellow |  |  |  | 0.15525 | 343.9836 | 415.3476 | 41.094 | 41.094 | 415.3476 | 344.0612 | 364.2388 | 364.3165 |  | 295.7602
LS2 | CD08 | SOL1 | 3646 | LS2_CD08:SOL1_D3646 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 344.3112 | 415.3476 | 41.094 | 41.094 | 415.3476 | 344.5612 | 364.5665 | 364.8165 |  | 296.0878
[6 empty] | bellow |  |  |  | 0.15525 | 344.6389 | 415.3476 | 41.094 | 41.094 | 415.3476 | 344.7165 | 364.8941 | 364.9717 |  | 296.4155
LS2 | CD08 | CAV5 | 3652 | LS2_CD08:CAV5_D3652 |  | b54 resonator | 40 |  |  | 0.49566 | 344.9643 | 415.3476 | 41.094 | 41.094 | 415.3476 | 345.2121 | 365.2195 | 365.4674 |  | 296.7409
[10 empty] | 0 | 345.2121 | 415.3476 | 41.094 | 41.094 | 415.3476 | 345.2121 | 365.4674 | 365.4674 |  | 296.9887
[10 empty] | 0 | 345.2121 | 415.3476 | 41.094 | 41.094 | 415.3476 | 345.2121 | 365.4674 | 365.4674 |  | 296.9887
[6 empty] | bellow |  |  |  | 0.099517 | 345.2619 | 415.3476 | 41.094 | 41.094 | 415.3476 | 345.3117 | 365.5171 | 365.5669 |  | 297.0385
LS2 | CD08 | CAV6 | 3658 | LS2_CD08:CAV6_D3658 |  | b54 resonator | 40 |  |  | 0.49566 | 345.5595 | 415.3476 | 41.094 | 41.094 | 415.3476 | 345.8073 | 365.8147 | 366.0625 |  | 297.3361
[10 empty] | 0 | 345.8073 | 415.3476 | 41.094 | 41.094 | 415.3476 | 345.8073 | 366.0625 | 366.0625 |  | 297.5839
[10 empty] | 0 | 345.8073 | 415.3476 | 41.094 | 41.094 | 415.3476 | 345.8073 | 366.0625 | 366.0625 |  | 297.5839
[6 empty] | bellow |  |  |  | 0.099517 | 345.8571 | 415.3476 | 41.094 | 41.094 | 415.3476 | 345.9068 | 366.1123 | 366.1621 |  | 297.6337
LS2 | CD08 | CAV7 | 3664 | LS2_CD08:CAV7_D3664 |  | b54 resonator | 40 |  |  | 0.49566 | 346.1547 | 415.3476 | 41.094 | 41.094 | 415.3476 | 346.4025 | 366.4099 | 366.6577 |  | 297.9313
[10 empty] | 0 | 346.4025 | 415.3476 | 41.094 | 41.094 | 415.3476 | 346.4025 | 366.6577 | 366.6577 |  | 298.1791
[10 empty] | 0 | 346.4025 | 415.3476 | 41.094 | 41.094 | 415.3476 | 346.4025 | 366.6577 | 366.6577 |  | 298.1791
[6 empty] | bellow |  |  |  | 0.099517 | 346.4523 | 415.3476 | 41.094 | 41.094 | 415.3476 | 346.502 | 366.7075 | 366.7572 |  | 298.2288
LS2 | CD08 | CAV8 | 3670 | LS2_CD08:CAV8_D3670 |  | b54 resonator | 40 |  |  | 0.49566 | 346.7498 | 415.3476 | 41.094 | 41.094 | 415.3476 | 346.9977 | 367.0051 | 367.2529 |  | 298.5264
LS2 | CD08 | BLM | 3673 | LS2_CD08:BLM_D3673 |  | outside loss monitor |  |  |  | 0 | 346.9977 | 415.3476 | 41.094 | 41.094 | 415.3476 | 346.9977 | 367.2529 | 367.2529 |  | 298.7743
[10 empty] | 0 | 346.9977 | 415.3476 | 41.094 | 41.094 | 415.3476 | 346.9977 | 367.2529 | 367.2529 |  | 298.7743
[10 empty] | 0.1677 | 347.0815 | 415.3476 | 41.094 | 41.094 | 415.3476 | 347.1654 | 367.3367 | 367.4206 |  | 298.8581
LS2 | CD08 | GV | 3674 | LS2_CD08:GV_D3674 |  | gate valve | 40 |  |  | 0.057404 | 347.1941 | 415.3476 | 41.094 | 41.094 | 415.3476 | 347.2228 | 367.4493 | 367.478 |  | 298.9707
[6 empty] | bellow | 40 |  |  | 0.1035 | 347.2745 | 415.3476 | 41.094 | 41.094 | 415.3476 | 347.3263 | 367.5297 | 367.5815 |  | 299.0511
LS2 | WD08 | BPM | 3676 | LS2_WD08:BPM_D3676 | BPM2 | position monitor |  |  |  | 0 | 347.3263 | 415.3476 | 41.094 | 41.094 | 415.3476 | 347.3263 | 367.5815 | 367.5815 |  | 299.1029
[6 empty] | diagnostic box |  |  |  | 0.14528 | 347.3989 | 415.3476 | 41.094 | 41.094 | 415.3476 | 347.4716 | 367.6541 | 367.7268 |  | 299.1755
LS2 | WD08 | PORT | 3677 | LS2_WD08:PORT_D3677 |  | port for future diagnostics |  |  |  | 0 | 347.4716 | 415.3476 | 41.094 | 41.094 | 415.3476 | 347.4716 | 367.7268 | 367.7268 |  | 299.2481
[6 empty] | diagnostic box |  |  |  | 0 | 347.4716 | 415.3476 | 41.094 | 41.094 | 415.3476 | 347.4716 | 367.7268 | 367.7268 |  | 299.2481
LS2 | WD08 | IP | 3677 | LS2_WD08:IP_D3677 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 347.4716 | 415.3476 | 41.094 | 41.094 | 415.3476 | 347.4716 | 367.7268 | 367.7268 |  | 299.2481
[6 empty] | bellow | 40 |  |  | 0.13122 | 347.5372 | 415.3476 | 41.094 | 41.094 | 415.3476 | 347.6028 | 367.7924 | 367.858 |  | 299.3138
LS2 | CD09 | GV | 3679 | LS2_CD09:GV_D3679 |  | gate valve | 40 |  |  | 0.057404 | 347.6315 | 415.3476 | 41.094 | 41.094 | 415.3476 | 347.6602 | 367.8867 | 367.9154 |  | 299.4081
[6 empty] | bellows |  |  |  | 0.1677 | 347.744 | 415.3476 | 41.094 | 41.094 | 415.3476 | 347.8279 | 367.9992 | 368.0831 |  | 299.5206
LS2 | CD09 | CAV1 | 3683 | LS2_CD09:CAV1_D3683 |  | b54 resonator | 40 |  |  | 0.49566 | 348.0757 | 415.3476 | 41.094 | 41.094 | 415.3476 | 348.3235 | 368.3309 | 368.5788 |  | 299.8523
LS2 | CD09 | BLM | 3686 | LS2_CD09:BLM_D3686 |  | outside loss monitor |  |  |  | 0 | 348.3235 | 415.3476 | 41.094 | 41.094 | 415.3476 | 348.3235 | 368.5788 | 368.5788 |  | 300.1001
[10 empty] | 0 | 348.3235 | 415.3476 | 41.094 | 41.094 | 415.3476 | 348.3235 | 368.5788 | 368.5788 |  | 300.1001
[6 empty] | bellow |  |  |  | 0.099517 | 348.3733 | 415.3476 | 41.094 | 41.094 | 415.3476 | 348.4231 | 368.6285 | 368.6783 |  | 300.1499
LS2 | CD09 | CAV2 | 3689 | LS2_CD09:CAV2_D3689 |  | b54 resonator | 40 |  |  | 0.49566 | 348.6709 | 415.3476 | 41.094 | 41.094 | 415.3476 | 348.9187 | 368.9261 | 369.1739 |  | 300.4475
[10 empty] | 0 | 348.9187 | 415.3476 | 41.094 | 41.094 | 415.3476 | 348.9187 | 369.1739 | 369.1739 |  | 300.6953
[10 empty] | 0 | 348.9187 | 415.3476 | 41.094 | 41.094 | 415.3476 | 348.9187 | 369.1739 | 369.1739 |  | 300.6953
[6 empty] | bellow |  |  |  | 0.099517 | 348.9685 | 415.3476 | 41.094 | 41.094 | 415.3476 | 349.0182 | 369.2237 | 369.2734 |  | 300.7451
LS2 | CD09 | CAV3 | 3695 | LS2_CD09:CAV3_D3695 |  | b54 resonator | 40 |  |  | 0.49566 | 349.2661 | 415.3476 | 41.094 | 41.094 | 415.3476 | 349.5139 | 369.5213 | 369.7691 |  | 301.0426
[10 empty] | 0 | 349.5139 | 415.3476 | 41.094 | 41.094 | 415.3476 | 349.5139 | 369.7691 | 369.7691 |  | 301.2905
[10 empty] | 0 | 349.5139 | 415.3476 | 41.094 | 41.094 | 415.3476 | 349.5139 | 369.7691 | 369.7691 |  | 301.2905
[6 empty] | bellow |  |  |  | 0.099517 | 349.5636 | 415.3476 | 41.094 | 41.094 | 415.3476 | 349.6134 | 369.8189 | 369.8686 |  | 301.3402
LS2 | CD09 | CAV4 | 3701 | LS2_CD09:CAV4_D3701 |  | b54 resonator | 40 |  |  | 0.49566 | 349.8612 | 415.3476 | 41.094 | 41.094 | 415.3476 | 350.1091 | 370.1164 | 370.3643 |  | 301.6378
LS2 | CD09 | BLM | 3704 | LS2_CD09:BLM_D3704 |  | outside loss monitor |  |  |  | 0 | 350.1091 | 415.3476 | 41.094 | 41.094 | 415.3476 | 350.1091 | 370.3643 | 370.3643 |  | 301.8856
[10 empty] | 0 | 350.1091 | 415.3476 | 41.094 | 41.094 | 415.3476 | 350.1091 | 370.3643 | 370.3643 |  | 301.8856
[6 empty] | bellow |  |  |  | 0.15525 | 350.1867 | 415.3476 | 41.094 | 41.094 | 415.3476 | 350.2643 | 370.4419 | 370.5195 |  | 301.9633
LS2 | CD09 | SOL1 | 3708 | LS2_CD09:SOL1_D3708 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 350.5143 | 415.3476 | 41.094 | 41.094 | 415.3476 | 350.7643 | 370.7695 | 371.0195 |  | 302.2909
[6 empty] | bellow |  |  |  | 0.15525 | 350.8419 | 415.3476 | 41.094 | 41.094 | 415.3476 | 350.9196 | 371.0971 | 371.1748 |  | 302.6185
LS2 | CD09 | CAV5 | 3714 | LS2_CD09:CAV5_D3714 |  | b54 resonator | 40 |  |  | 0.49566 | 351.1674 | 415.3476 | 41.094 | 41.094 | 415.3476 | 351.4152 | 371.4226 | 371.6704 |  | 302.944
[10 empty] | 0 | 351.4152 | 415.3476 | 41.094 | 41.094 | 415.3476 | 351.4152 | 371.6704 | 371.6704 |  | 303.1918
[10 empty] | 0 | 351.4152 | 415.3476 | 41.094 | 41.094 | 415.3476 | 351.4152 | 371.6704 | 371.6704 |  | 303.1918
[6 empty] | bellow |  |  |  | 0.099517 | 351.465 | 415.3476 | 41.094 | 41.094 | 415.3476 | 351.5147 | 371.7202 | 371.7699 |  | 303.2416
LS2 | CD09 | CAV6 | 3720 | LS2_CD09:CAV6_D3720 |  | b54 resonator | 40 |  |  | 0.49566 | 351.7626 | 415.3476 | 41.094 | 41.094 | 415.3476 | 352.0104 | 372.0178 | 372.2656 |  | 303.5391
[10 empty] | 0 | 352.0104 | 415.3476 | 41.094 | 41.094 | 415.3476 | 352.0104 | 372.2656 | 372.2656 |  | 303.787
[10 empty] | 0 | 352.0104 | 415.3476 | 41.094 | 41.094 | 415.3476 | 352.0104 | 372.2656 | 372.2656 |  | 303.787
[6 empty] | bellow |  |  |  | 0.099517 | 352.0601 | 415.3476 | 41.094 | 41.094 | 415.3476 | 352.1099 | 372.3154 | 372.3651 |  | 303.8367
LS2 | CD09 | CAV7 | 3726 | LS2_CD09:CAV7_D3726 |  | b54 resonator | 40 |  |  | 0.49566 | 352.3577 | 415.3476 | 41.094 | 41.094 | 415.3476 | 352.6056 | 372.6129 | 372.8608 |  | 304.1343
[10 empty] | 0 | 352.6056 | 415.3476 | 41.094 | 41.094 | 415.3476 | 352.6056 | 372.8608 | 372.8608 |  | 304.3821
[10 empty] | 0 | 352.6056 | 415.3476 | 41.094 | 41.094 | 415.3476 | 352.6056 | 372.8608 | 372.8608 |  | 304.3821
[6 empty] | bellow |  |  |  | 0.099517 | 352.6553 | 415.3476 | 41.094 | 41.094 | 415.3476 | 352.7051 | 372.9105 | 372.9603 |  | 304.4319
LS2 | CD09 | CAV8 | 3732 | LS2_CD09:CAV8_D3732 |  | b54 resonator | 40 |  |  | 0.49566 | 352.9529 | 415.3476 | 41.094 | 41.094 | 415.3476 | 353.2007 | 373.2081 | 373.4559 |  | 304.7295
LS2 | CD09 | BLM | 3735 | LS2_CD09:BLM_D3735 |  | outside loss monitor |  |  |  | 0 | 353.2007 | 415.3476 | 41.094 | 41.094 | 415.3476 | 353.2007 | 373.4559 | 373.4559 |  | 304.9773
[10 empty] | 0 | 353.2007 | 415.3476 | 41.094 | 41.094 | 415.3476 | 353.2007 | 373.4559 | 373.4559 |  | 304.9773
[10 empty] | 0.1677 | 353.2846 | 415.3476 | 41.094 | 41.094 | 415.3476 | 353.3684 | 373.5398 | 373.6237 |  | 305.0612
LS2 | CD09 | GV | 3737 | LS2_CD09:GV_D3737 |  | gate valve | 40 |  |  | 0.057404 | 353.3971 | 415.3476 | 41.094 | 41.094 | 415.3476 | 353.4258 | 373.6524 | 373.6811 |  | 305.1737
[6 empty] | bellow | 40 |  |  | 0.1035 | 353.4776 | 415.3476 | 41.094 | 41.094 | 415.3476 | 353.5293 | 373.7328 | 373.7846 |  | 305.2542
LS2 | WD09 | BPM | 3738 | LS2_WD09:BPM_D3738 | BPM2 | position monitor |  |  |  | 0 | 353.5293 | 415.3476 | 41.094 | 41.094 | 415.3476 | 353.5293 | 373.7846 | 373.7846 |  | 305.3059
[6 empty] | diagnostic box |  |  |  | 0.14528 | 353.602 | 415.3476 | 41.094 | 41.094 | 415.3476 | 353.6746 | 373.8572 | 373.9298 |  | 305.3786
LS2 | WD09 | PORT | 3739 | LS2_WD09:PORT_D3739 |  | port for future diagnostics |  |  |  | 0 | 353.6746 | 415.3476 | 41.094 | 41.094 | 415.3476 | 353.6746 | 373.9298 | 373.9298 |  | 305.4512
[6 empty] | diagnostic box |  |  |  | 0 | 353.6746 | 415.3476 | 41.094 | 41.094 | 415.3476 | 353.6746 | 373.9298 | 373.9298 |  | 305.4512
LS2 | WD09 | IP | 3739 | LS2_WD09:IP_D3739 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 353.6746 | 415.3476 | 41.094 | 41.094 | 415.3476 | 353.6746 | 373.9298 | 373.9298 |  | 305.4512
[6 empty] | bellow | 40 |  |  | 0.13122 | 353.7402 | 415.3476 | 41.094 | 41.094 | 415.3476 | 353.8058 | 373.9954 | 374.0611 |  | 305.5168
LS2 | CD10 | GV | 3741 | LS2_CD10:GV_D3741 |  | gate valve | 40 |  |  | 0.057404 | 353.8345 | 415.3476 | 41.094 | 41.094 | 415.3476 | 353.8632 | 374.0898 | 374.1185 |  | 305.6111
[6 empty] | bellows |  |  |  | 0.1677 | 353.9471 | 415.3476 | 41.094 | 41.094 | 415.3476 | 354.0309 | 374.2023 | 374.2862 |  | 305.7237
LS2 | CD10 | CAV1 | 3745 | LS2_CD10:CAV1_D3745 |  | b54 resonator | 40 |  |  | 0.49566 | 354.2788 | 415.3476 | 41.094 | 41.094 | 415.3476 | 354.5266 | 374.534 | 374.7818 |  | 306.0554
LS2 | CD10 | BLM | 3748 | LS2_CD10:BLM_D3748 |  | outside loss monitor |  |  |  | 0 | 354.5266 | 415.3476 | 41.094 | 41.094 | 415.3476 | 354.5266 | 374.7818 | 374.7818 |  | 306.3032
[10 empty] | 0 | 354.5266 | 415.3476 | 41.094 | 41.094 | 415.3476 | 354.5266 | 374.7818 | 374.7818 |  | 306.3032
[6 empty] | bellow |  |  |  | 0.099517 | 354.5764 | 415.3476 | 41.094 | 41.094 | 415.3476 | 354.6261 | 374.8316 | 374.8813 |  | 306.3529
LS2 | CD10 | CAV2 | 3751 | LS2_CD10:CAV2_D3751 |  | b54 resonator | 40 |  |  | 0.49566 | 354.8739 | 415.3476 | 41.094 | 41.094 | 415.3476 | 355.1218 | 375.1292 | 375.377 |  | 306.6505
[10 empty] | 0 | 355.1218 | 415.3476 | 41.094 | 41.094 | 415.3476 | 355.1218 | 375.377 | 375.377 |  | 306.8984
[10 empty] | 0 | 355.1218 | 415.3476 | 41.094 | 41.094 | 415.3476 | 355.1218 | 375.377 | 375.377 |  | 306.8984
[6 empty] | bellow |  |  |  | 0.099517 | 355.1715 | 415.3476 | 41.094 | 41.094 | 415.3476 | 355.2213 | 375.4268 | 375.4765 |  | 306.9481
LS2 | CD10 | CAV3 | 3757 | LS2_CD10:CAV3_D3757 |  | b54 resonator | 40 |  |  | 0.49566 | 355.4691 | 415.3476 | 41.094 | 41.094 | 415.3476 | 355.7169 | 375.7243 | 375.9722 |  | 307.2457
[10 empty] | 0 | 355.7169 | 415.3476 | 41.094 | 41.094 | 415.3476 | 355.7169 | 375.9722 | 375.9722 |  | 307.4935
[10 empty] | 0 | 355.7169 | 415.3476 | 41.094 | 41.094 | 415.3476 | 355.7169 | 375.9722 | 375.9722 |  | 307.4935
[6 empty] | bellow |  |  |  | 0.099517 | 355.7667 | 415.3476 | 41.094 | 41.094 | 415.3476 | 355.8165 | 376.0219 | 376.0717 |  | 307.5433
LS2 | CD10 | CAV4 | 3763 | LS2_CD10:CAV4_D3763 |  | b54 resonator | 40 |  |  | 0.49566 | 356.0643 | 415.3476 | 41.094 | 41.094 | 415.3476 | 356.3121 | 376.3195 | 376.5673 |  | 307.8409
LS2 | CD10 | BLM | 3766 | LS2_CD10:BLM_D3766 |  | outside loss monitor |  |  |  | 0 | 356.3121 | 415.3476 | 41.094 | 41.094 | 415.3476 | 356.3121 | 376.5673 | 376.5673 |  | 308.0887
[10 empty] | 0 | 356.3121 | 415.3476 | 41.094 | 41.094 | 415.3476 | 356.3121 | 376.5673 | 376.5673 |  | 308.0887
[6 empty] | bellow |  |  |  | 0.15525 | 356.3897 | 415.3476 | 41.094 | 41.094 | 415.3476 | 356.4674 | 376.645 | 376.7226 |  | 308.1663
LS2 | CD10 | SOL1 | 3770 | LS2_CD10:SOL1_D3770 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 356.7174 | 415.3476 | 41.094 | 41.094 | 415.3476 | 356.9674 | 376.9726 | 377.2226 |  | 308.494
[6 empty] | bellow |  |  |  | 0.15525 | 357.045 | 415.3476 | 41.094 | 41.094 | 415.3476 | 357.1226 | 377.3002 | 377.3778 |  | 308.8216
LS2 | CD10 | CAV5 | 3776 | LS2_CD10:CAV5_D3776 |  | b54 resonator | 40 |  |  | 0.49566 | 357.3704 | 415.3476 | 41.094 | 41.094 | 415.3476 | 357.6183 | 377.6257 | 377.8735 |  | 309.147
[10 empty] | 0 | 357.6183 | 415.3476 | 41.094 | 41.094 | 415.3476 | 357.6183 | 377.8735 | 377.8735 |  | 309.3949
[10 empty] | 0 | 357.6183 | 415.3476 | 41.094 | 41.094 | 415.3476 | 357.6183 | 377.8735 | 377.8735 |  | 309.3949
[6 empty] | bellow |  |  |  | 0.099517 | 357.668 | 415.3476 | 41.094 | 41.094 | 415.3476 | 357.7178 | 377.9233 | 377.973 |  | 309.4446
LS2 | CD10 | CAV6 | 3782 | LS2_CD10:CAV6_D3782 |  | b54 resonator | 40 |  |  | 0.49566 | 357.9656 | 415.3476 | 41.094 | 41.094 | 415.3476 | 358.2135 | 378.2208 | 378.4687 |  | 309.7422
[10 empty] | 0 | 358.2135 | 415.3476 | 41.094 | 41.094 | 415.3476 | 358.2135 | 378.4687 | 378.4687 |  | 309.99
[10 empty] | 0 | 358.2135 | 415.3476 | 41.094 | 41.094 | 415.3476 | 358.2135 | 378.4687 | 378.4687 |  | 309.99
[6 empty] | bellow |  |  |  | 0.099517 | 358.2632 | 415.3476 | 41.094 | 41.094 | 415.3476 | 358.313 | 378.5184 | 378.5682 |  | 310.0398
LS2 | CD10 | CAV7 | 3788 | LS2_CD10:CAV7_D3788 |  | b54 resonator | 40 |  |  | 0.49566 | 358.5608 | 415.3476 | 41.094 | 41.094 | 415.3476 | 358.8086 | 378.816 | 379.0638 |  | 310.3374
[10 empty] | 0 | 358.8086 | 415.3476 | 41.094 | 41.094 | 415.3476 | 358.8086 | 379.0638 | 379.0638 |  | 310.5852
[10 empty] | 0 | 358.8086 | 415.3476 | 41.094 | 41.094 | 415.3476 | 358.8086 | 379.0638 | 379.0638 |  | 310.5852
[6 empty] | bellow |  |  |  | 0.099517 | 358.8584 | 415.3476 | 41.094 | 41.094 | 415.3476 | 358.9081 | 379.1136 | 379.1634 |  | 310.635
LS2 | CD10 | CAV8 | 3794 | LS2_CD10:CAV8_D3794 |  | b54 resonator | 40 |  |  | 0.49566 | 359.156 | 415.3476 | 41.094 | 41.094 | 415.3476 | 359.4038 | 379.4112 | 379.659 |  | 310.9326
LS2 | CD10 | BLM | 3797 | LS2_CD10:BLM_D3797 |  | outside loss monitor |  |  |  | 0 | 359.4038 | 415.3476 | 41.094 | 41.094 | 415.3476 | 359.4038 | 379.659 | 379.659 |  | 311.1804
[10 empty] | 0 | 359.4038 | 415.3476 | 41.094 | 41.094 | 415.3476 | 359.4038 | 379.659 | 379.659 |  | 311.1804
[10 empty] | 0.1677 | 359.4876 | 415.3476 | 41.094 | 41.094 | 415.3476 | 359.5715 | 379.7429 | 379.8267 |  | 311.2642
LS2 | CD10 | GV | 3799 | LS2_CD10:GV_D3799 |  | gate valve | 40 |  |  | 0.057404 | 359.6002 | 415.3476 | 41.094 | 41.094 | 415.3476 | 359.6289 | 379.8554 | 379.8841 |  | 311.3768
[6 empty] | bellow | 40 |  |  | 0.1035 | 359.6807 | 415.3476 | 41.094 | 41.094 | 415.3476 | 359.7324 | 379.9359 | 379.9876 |  | 311.4572
LS2 | WD10 | BPM | 3800 | LS2_WD10:BPM_D3800 | BPM2 | position monitor |  |  |  | 0 | 359.7324 | 415.3476 | 41.094 | 41.094 | 415.3476 | 359.7324 | 379.9876 | 379.9876 |  | 311.509
[6 empty] | diagnostic box |  |  |  | 0.14528 | 359.805 | 415.3476 | 41.094 | 41.094 | 415.3476 | 359.8777 | 380.0603 | 380.1329 |  | 311.5816
LS2 | WD10 | PORT | 3801 | LS2_WD10:PORT_D3801 |  | port for future diagnostics |  |  |  | 0 | 359.8777 | 415.3476 | 41.094 | 41.094 | 415.3476 | 359.8777 | 380.1329 | 380.1329 |  | 311.6543
[6 empty] | diagnostic box |  |  |  | 0 | 359.8777 | 415.3476 | 41.094 | 41.094 | 415.3476 | 359.8777 | 380.1329 | 380.1329 |  | 311.6543
LS2 | WD10 | IP | 3801 | LS2_WD10:IP_D3801 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 359.8777 | 415.3476 | 41.094 | 41.094 | 415.3476 | 359.8777 | 380.1329 | 380.1329 |  | 311.6543
[6 empty] | bellow | 40 |  |  | 0.13122 | 359.9433 | 415.3476 | 41.094 | 41.094 | 415.3476 | 360.0089 | 380.1985 | 380.2641 |  | 311.7199
LS2 | CD11 | GV | 3803 | LS2_CD11:GV_D3803 |  | gate valve | 40 |  |  | 0.057404 | 360.0376 | 415.3476 | 41.094 | 41.094 | 415.3476 | 360.0663 | 380.2928 | 380.3215 |  | 311.8142
[6 empty] | bellows |  |  |  | 0.1677 | 360.1502 | 415.3476 | 41.094 | 41.094 | 415.3476 | 360.234 | 380.4054 | 380.4892 |  | 311.9267
LS2 | CD11 | CAV1 | 3807 | LS2_CD11:CAV1_D3807 |  | b54 resonator | 40 |  |  | 0.49566 | 360.4818 | 415.3476 | 41.094 | 41.094 | 415.3476 | 360.7297 | 380.7371 | 380.9849 |  | 312.2584
LS2 | CD11 | BLM | 3810 | LS2_CD11:BLM_D3810 |  | outside loss monitor |  |  |  | 0 | 360.7297 | 415.3476 | 41.094 | 41.094 | 415.3476 | 360.7297 | 380.9849 | 380.9849 |  | 312.5063
[10 empty] | 0 | 360.7297 | 415.3476 | 41.094 | 41.094 | 415.3476 | 360.7297 | 380.9849 | 380.9849 |  | 312.5063
[6 empty] | bellow |  |  |  | 0.099517 | 360.7794 | 415.3476 | 41.094 | 41.094 | 415.3476 | 360.8292 | 381.0346 | 381.0844 |  | 312.556
LS2 | CD11 | CAV2 | 3813 | LS2_CD11:CAV2_D3813 |  | b54 resonator | 40 |  |  | 0.49566 | 361.077 | 415.3476 | 41.094 | 41.094 | 415.3476 | 361.3248 | 381.3322 | 381.5801 |  | 312.8536
[10 empty] | 0 | 361.3248 | 415.3476 | 41.094 | 41.094 | 415.3476 | 361.3248 | 381.5801 | 381.5801 |  | 313.1014
[10 empty] | 0 | 361.3248 | 415.3476 | 41.094 | 41.094 | 415.3476 | 361.3248 | 381.5801 | 381.5801 |  | 313.1014
[6 empty] | bellow |  |  |  | 0.099517 | 361.3746 | 415.3476 | 41.094 | 41.094 | 415.3476 | 361.4244 | 381.6298 | 381.6796 |  | 313.1512
LS2 | CD11 | CAV3 | 3819 | LS2_CD11:CAV3_D3819 |  | b54 resonator | 40 |  |  | 0.49566 | 361.6722 | 415.3476 | 41.094 | 41.094 | 415.3476 | 361.92 | 381.9274 | 382.1752 |  | 313.4488
[10 empty] | 0 | 361.92 | 415.3476 | 41.094 | 41.094 | 415.3476 | 361.92 | 382.1752 | 382.1752 |  | 313.6966
[10 empty] | 0 | 361.92 | 415.3476 | 41.094 | 41.094 | 415.3476 | 361.92 | 382.1752 | 382.1752 |  | 313.6966
[6 empty] | bellow |  |  |  | 0.099517 | 361.9698 | 415.3476 | 41.094 | 41.094 | 415.3476 | 362.0195 | 382.225 | 382.2747 |  | 313.7464
LS2 | CD11 | CAV4 | 3825 | LS2_CD11:CAV4_D3825 |  | b54 resonator | 40 |  |  | 0.49566 | 362.2674 | 415.3476 | 41.094 | 41.094 | 415.3476 | 362.5152 | 382.5226 | 382.7704 |  | 314.0439
LS2 | CD11 | BLM | 3828 | LS2_CD11:BLM_D3828 |  | outside loss monitor |  |  |  | 0 | 362.5152 | 415.3476 | 41.094 | 41.094 | 415.3476 | 362.5152 | 382.7704 | 382.7704 |  | 314.2918
[10 empty] | 0 | 362.5152 | 415.3476 | 41.094 | 41.094 | 415.3476 | 362.5152 | 382.7704 | 382.7704 |  | 314.2918
[6 empty] | bellow |  |  |  | 0.15525 | 362.5928 | 415.3476 | 41.094 | 41.094 | 415.3476 | 362.6704 | 382.848 | 382.9257 |  | 314.3694
LS2 | CD11 | SOL1 | 3832 | LS2_CD11:SOL1_D3832 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 362.9204 | 415.3476 | 41.094 | 41.094 | 415.3476 | 363.1704 | 383.1757 | 383.4257 |  | 314.697
[6 empty] | bellow |  |  |  | 0.15525 | 363.2481 | 415.3476 | 41.094 | 41.094 | 415.3476 | 363.3257 | 383.5033 | 383.5809 |  | 315.0246
LS2 | CD11 | CAV5 | 3838 | LS2_CD11:CAV5_D3838 |  | b54 resonator | 40 |  |  | 0.49566 | 363.5735 | 415.3476 | 41.094 | 41.094 | 415.3476 | 363.8213 | 383.8287 | 384.0766 |  | 315.3501
[10 empty] | 0 | 363.8213 | 415.3476 | 41.094 | 41.094 | 415.3476 | 363.8213 | 384.0766 | 384.0766 |  | 315.5979
[10 empty] | 0 | 363.8213 | 415.3476 | 41.094 | 41.094 | 415.3476 | 363.8213 | 384.0766 | 384.0766 |  | 315.5979
[6 empty] | bellow |  |  |  | 0.099517 | 363.8711 | 415.3476 | 41.094 | 41.094 | 415.3476 | 363.9209 | 384.1263 | 384.1761 |  | 315.6477
LS2 | CD11 | CAV6 | 3844 | LS2_CD11:CAV6_D3844 |  | b54 resonator | 40 |  |  | 0.49566 | 364.1687 | 415.3476 | 41.094 | 41.094 | 415.3476 | 364.4165 | 384.4239 | 384.6717 |  | 315.9453
[10 empty] | 0 | 364.4165 | 415.3476 | 41.094 | 41.094 | 415.3476 | 364.4165 | 384.6717 | 384.6717 |  | 316.1931
[10 empty] | 0 | 364.4165 | 415.3476 | 41.094 | 41.094 | 415.3476 | 364.4165 | 384.6717 | 384.6717 |  | 316.1931
[6 empty] | bellow |  |  |  | 0.099517 | 364.4663 | 415.3476 | 41.094 | 41.094 | 415.3476 | 364.516 | 384.7215 | 384.7712 |  | 316.2429
LS2 | CD11 | CAV7 | 3850 | LS2_CD11:CAV7_D3850 |  | b54 resonator | 40 |  |  | 0.49566 | 364.7639 | 415.3476 | 41.094 | 41.094 | 415.3476 | 365.0117 | 385.0191 | 385.2669 |  | 316.5404
[10 empty] | 0 | 365.0117 | 415.3476 | 41.094 | 41.094 | 415.3476 | 365.0117 | 385.2669 | 385.2669 |  | 316.7883
[10 empty] | 0 | 365.0117 | 415.3476 | 41.094 | 41.094 | 415.3476 | 365.0117 | 385.2669 | 385.2669 |  | 316.7883
[6 empty] | bellow |  |  |  | 0.099517 | 365.0614 | 415.3476 | 41.094 | 41.094 | 415.3476 | 365.1112 | 385.3167 | 385.3664 |  | 316.838
LS2 | CD11 | CAV8 | 3856 | LS2_CD11:CAV8_D3856 |  | b54 resonator | 40 |  |  | 0.49566 | 365.359 | 415.3476 | 41.094 | 41.094 | 415.3476 | 365.6069 | 385.6143 | 385.8621 |  | 317.1356
LS2 | CD11 | BLM | 3859 | LS2_CD11:BLM_D3859 |  | outside loss monitor |  |  |  | 0 | 365.6069 | 415.3476 | 41.094 | 41.094 | 415.3476 | 365.6069 | 385.8621 | 385.8621 |  | 317.3834
[10 empty] | 0 | 365.6069 | 415.3476 | 41.094 | 41.094 | 415.3476 | 365.6069 | 385.8621 | 385.8621 |  | 317.3834
[10 empty] | 0.1677 | 365.6907 | 415.3476 | 41.094 | 41.094 | 415.3476 | 365.7746 | 385.9459 | 386.0298 |  | 317.4673
LS2 | CD11 | GV | 3861 | LS2_CD11:GV_D3861 |  | gate valve | 40 |  |  | 0.057404 | 365.8033 | 415.3476 | 41.094 | 41.094 | 415.3476 | 365.832 | 386.0585 | 386.0872 |  | 317.5799
[6 empty] | bellow | 40 |  |  | 0.1035 | 365.8837 | 415.3476 | 41.094 | 41.094 | 415.3476 | 365.9355 | 386.1389 | 386.1907 |  | 317.6603
LS2 | WD11 | BPM | 3862 | LS2_WD11:BPM_D3862 | BPM2 | position monitor |  |  |  | 0 | 365.9355 | 415.3476 | 41.094 | 41.094 | 415.3476 | 365.9355 | 386.1907 | 386.1907 |  | 317.7121
[6 empty] | diagnostic box |  |  |  | 0.14528 | 366.0081 | 415.3476 | 41.094 | 41.094 | 415.3476 | 366.0808 | 386.2633 | 386.336 |  | 317.7847
LS2 | WD11 | PORT | 3863 | LS2_WD11:PORT_D3863 |  | port for future diagnostics |  |  |  | 0 | 366.0808 | 415.3476 | 41.094 | 41.094 | 415.3476 | 366.0808 | 386.336 | 386.336 |  | 317.8573
[6 empty] | diagnostic box |  |  |  | 0 | 366.0808 | 415.3476 | 41.094 | 41.094 | 415.3476 | 366.0808 | 386.336 | 386.336 |  | 317.8573
LS2 | WD11 | IP | 3863 | LS2_WD11:IP_D3863 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 366.0808 | 415.3476 | 41.094 | 41.094 | 415.3476 | 366.0808 | 386.336 | 386.336 |  | 317.8573
[6 empty] | bellow | 40 |  |  | 0.13122 | 366.1464 | 415.3476 | 41.094 | 41.094 | 415.3476 | 366.212 | 386.4016 | 386.4672 |  | 317.9229
LS2 | CD12 | GV | 3865 | LS2_CD12:GV_D3865 |  | gate valve | 40 |  |  | 0.057404 | 366.2407 | 415.3476 | 41.094 | 41.094 | 415.3476 | 366.2694 | 386.4959 | 386.5246 |  | 318.0173
[6 empty] | bellows |  |  |  | 0.1677 | 366.3532 | 415.3476 | 41.094 | 41.094 | 415.3476 | 366.4371 | 386.6084 | 386.6923 |  | 318.1298
LS2 | CD12 | CAV1 | 3869 | LS2_CD12:CAV1_D3869 |  | b54 resonator | 40 |  |  | 0.49566 | 366.6849 | 415.3476 | 41.094 | 41.094 | 415.3476 | 366.9327 | 386.9401 | 387.188 |  | 318.4615
LS2 | CD12 | BLM | 3872 | LS2_CD12:BLM_D3872 |  | outside loss monitor |  |  |  | 0 | 366.9327 | 415.3476 | 41.094 | 41.094 | 415.3476 | 366.9327 | 387.188 | 387.188 |  | 318.7093
[10 empty] | 0 | 366.9327 | 415.3476 | 41.094 | 41.094 | 415.3476 | 366.9327 | 387.188 | 387.188 |  | 318.7093
[6 empty] | bellow |  |  |  | 0.099517 | 366.9825 | 415.3476 | 41.094 | 41.094 | 415.3476 | 367.0323 | 387.2377 | 387.2875 |  | 318.7591
LS2 | CD12 | CAV2 | 3875 | LS2_CD12:CAV2_D3875 |  | b54 resonator | 40 |  |  | 0.49566 | 367.2801 | 415.3476 | 41.094 | 41.094 | 415.3476 | 367.5279 | 387.5353 | 387.7831 |  | 319.0567
[10 empty] | 0 | 367.5279 | 415.3476 | 41.094 | 41.094 | 415.3476 | 367.5279 | 387.7831 | 387.7831 |  | 319.3045
[10 empty] | 0 | 367.5279 | 415.3476 | 41.094 | 41.094 | 415.3476 | 367.5279 | 387.7831 | 387.7831 |  | 319.3045
[6 empty] | bellow |  |  |  | 0.099517 | 367.5777 | 415.3476 | 41.094 | 41.094 | 415.3476 | 367.6274 | 387.8329 | 387.8826 |  | 319.3543
LS2 | CD12 | CAV3 | 3881 | LS2_CD12:CAV3_D3881 |  | b54 resonator | 40 |  |  | 0.49566 | 367.8753 | 415.3476 | 41.094 | 41.094 | 415.3476 | 368.1231 | 388.1305 | 388.3783 |  | 319.6518
[10 empty] | 0 | 368.1231 | 415.3476 | 41.094 | 41.094 | 415.3476 | 368.1231 | 388.3783 | 388.3783 |  | 319.8997
[10 empty] | 0 | 368.1231 | 415.3476 | 41.094 | 41.094 | 415.3476 | 368.1231 | 388.3783 | 388.3783 |  | 319.8997
[6 empty] | bellow |  |  |  | 0.099517 | 368.1728 | 415.3476 | 41.094 | 41.094 | 415.3476 | 368.2226 | 388.4281 | 388.4778 |  | 319.9494
LS2 | CD12 | CAV4 | 3887 | LS2_CD12:CAV4_D3887 |  | b54 resonator | 40 |  |  | 0.49566 | 368.4704 | 415.3476 | 41.094 | 41.094 | 415.3476 | 368.7183 | 388.7256 | 388.9735 |  | 320.247
LS2 | CD12 | BLM | 3890 | LS2_CD12:BLM_D3890 |  | outside loss monitor |  |  |  | 0 | 368.7183 | 415.3476 | 41.094 | 41.094 | 415.3476 | 368.7183 | 388.9735 | 388.9735 |  | 320.4948
[10 empty] | 0 | 368.7183 | 415.3476 | 41.094 | 41.094 | 415.3476 | 368.7183 | 388.9735 | 388.9735 |  | 320.4948
[6 empty] | bellow |  |  |  | 0.15525 | 368.7959 | 415.3476 | 41.094 | 41.094 | 415.3476 | 368.8735 | 389.0511 | 389.1287 |  | 320.5725
LS2 | CD12 | SOL1 | 3894 | LS2_CD12:SOL1_D3894 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 369.1235 | 415.3476 | 41.094 | 41.094 | 415.3476 | 369.3735 | 389.3787 | 389.6287 |  | 320.9001
[6 empty] | bellow |  |  |  | 0.15525 | 369.4511 | 415.3476 | 41.094 | 41.094 | 415.3476 | 369.5288 | 389.7063 | 389.784 |  | 321.2277
LS2 | CD12 | CAV5 | 3900 | LS2_CD12:CAV5_D3900 |  | b54 resonator | 40 |  |  | 0.49566 | 369.7766 | 415.3476 | 41.094 | 41.094 | 415.3476 | 370.0244 | 390.0318 | 390.2796 |  | 321.5532
[10 empty] | 0 | 370.0244 | 415.3476 | 41.094 | 41.094 | 415.3476 | 370.0244 | 390.2796 | 390.2796 |  | 321.801
[10 empty] | 0 | 370.0244 | 415.3476 | 41.094 | 41.094 | 415.3476 | 370.0244 | 390.2796 | 390.2796 |  | 321.801
[6 empty] | bellow |  |  |  | 0.099517 | 370.0742 | 415.3476 | 41.094 | 41.094 | 415.3476 | 370.1239 | 390.3294 | 390.3791 |  | 321.8508
LS2 | CD12 | CAV6 | 3906 | LS2_CD12:CAV6_D3906 |  | b54 resonator | 40 |  |  | 0.49566 | 370.3718 | 415.3476 | 41.094 | 41.094 | 415.3476 | 370.6196 | 390.627 | 390.8748 |  | 322.1483
[10 empty] | 0 | 370.6196 | 415.3476 | 41.094 | 41.094 | 415.3476 | 370.6196 | 390.8748 | 390.8748 |  | 322.3962
[10 empty] | 0 | 370.6196 | 415.3476 | 41.094 | 41.094 | 415.3476 | 370.6196 | 390.8748 | 390.8748 |  | 322.3962
[6 empty] | bellow |  |  |  | 0.099517 | 370.6693 | 415.3476 | 41.094 | 41.094 | 415.3476 | 370.7191 | 390.9246 | 390.9743 |  | 322.4459
LS2 | CD12 | CAV7 | 3912 | LS2_CD12:CAV7_D3912 |  | b54 resonator | 40 |  |  | 0.49566 | 370.9669 | 415.3476 | 41.094 | 41.094 | 415.3476 | 371.2148 | 391.2221 | 391.47 |  | 322.7435
[10 empty] | 0 | 371.2148 | 415.3476 | 41.094 | 41.094 | 415.3476 | 371.2148 | 391.47 | 391.47 |  | 322.9913
[10 empty] | 0 | 371.2148 | 415.3476 | 41.094 | 41.094 | 415.3476 | 371.2148 | 391.47 | 391.47 |  | 322.9913
[6 empty] | bellow |  |  |  | 0.099517 | 371.2645 | 415.3476 | 41.094 | 41.094 | 415.3476 | 371.3143 | 391.5197 | 391.5695 |  | 323.0411
LS2 | CD12 | CAV8 | 3918 | LS2_CD12:CAV8_D3918 |  | b54 resonator | 40 |  |  | 0.49566 | 371.5621 | 415.3476 | 41.094 | 41.094 | 415.3476 | 371.8099 | 391.8173 | 392.0651 |  | 323.3387
LS2 | CD12 | BLM | 3921 | LS2_CD12:BLM_D3921 |  | outside loss monitor |  |  |  | 0 | 371.8099 | 415.3476 | 41.094 | 41.094 | 415.3476 | 371.8099 | 392.0651 | 392.0651 |  | 323.5865
[10 empty] | 0 | 371.8099 | 415.3476 | 41.094 | 41.094 | 415.3476 | 371.8099 | 392.0651 | 392.0651 |  | 323.5865
[10 empty] | 0.1677 | 371.8938 | 415.3476 | 41.094 | 41.094 | 415.3476 | 371.9776 | 392.149 | 392.2328 |  | 323.6704
LS2 | CD12 | GV | 3923 | LS2_CD12:GV_D3923 |  | gate valve | 40 |  |  | 0.057404 | 372.0063 | 415.3476 | 41.094 | 41.094 | 415.3476 | 372.035 | 392.2616 | 392.2903 |  | 323.7829
[6 empty] | bellow | 40 |  |  | 0.1035 | 372.0868 | 415.3476 | 41.094 | 41.094 | 415.3476 | 372.1385 | 392.342 | 392.3938 |  | 323.8634
LS2 | WD12 | BPM | 3924 | LS2_WD12:BPM_D3924 | BPM2 | position monitor |  |  |  | 0 | 372.1385 | 415.3476 | 41.094 | 41.094 | 415.3476 | 372.1385 | 392.3938 | 392.3938 |  | 323.9151
[6 empty] | diagnostic box |  |  |  | 0.14528 | 372.2112 | 415.3476 | 41.094 | 41.094 | 415.3476 | 372.2838 | 392.4664 | 392.539 |  | 323.9878
LS2 | WD12 | PORT | 3925 | LS2_WD12:PORT_D3925 |  | port for future diagnostics |  |  |  | 0 | 372.2838 | 415.3476 | 41.094 | 41.094 | 415.3476 | 372.2838 | 392.539 | 392.539 |  | 324.0604
[6 empty] | diagnostic box |  |  |  | 0 | 372.2838 | 415.3476 | 41.094 | 41.094 | 415.3476 | 372.2838 | 392.539 | 392.539 |  | 324.0604
LS2 | WD12 | IP | 3925 | LS2_WD12:IP_D3925 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 372.2838 | 415.3476 | 41.094 | 41.094 | 415.3476 | 372.2838 | 392.539 | 392.539 |  | 324.0604
[6 empty] | bellow | 40 |  |  | 0.13122 | 372.3494 | 415.3476 | 41.094 | 41.094 | 415.3476 | 372.415 | 392.6046 | 392.6703 |  | 324.126
SEGMENT 2 END (12 b29 CMs +12 b54 CMs) | [9 empty] | 0 | 372.415 | 415.3476 | 41.094 | 41.094 | 415.3476 | 372.415 | 392.6703 | 392.6703 |  | 324.1916
[8 empty] | 0.003903 |  | 0.037404 | 372.4337 | 415.3476 | 41.094 | 41.094 | 415.3476 | 372.4524 | 392.689 | 392.7077 |  | 324.2103
FS2 | BTS | FAV | 3927 | FS2_BTS:FAV_D3927 | FV | fast valve | 40 | 2.176e-06 | 372.4524 | 0 | 372.4524 | 415.3476 | 41.094 | 41.094 | 415.3476 | 372.4524 | 392.7077 | 392.7077 | 0.067196 | 324.229
[10 empty] | 0.1626 | 372.5337 | 415.3476 | 41.094 | 41.094 | 415.3476 | 372.615 | 392.789 | 392.8703 |  | 324.3103
FS2 | BTS | DCHV | 3930 | FS2_BTS:DCHV_D3930 | C2 | correctors-X,Y | 50 |  | 372.715 | 0.2 | 372.715 | 415.3476 | 41.094 | 41.094 | 415.3476 | 372.815 | 392.9703 | 393.0703 |  | 324.4916
[10 empty] | 0.2 | 372.915 | 415.3476 | 41.094 | 41.094 | 415.3476 | 373.015 | 393.1703 | 393.2703 |  | 324.6916
FS2 | BTS | QH | 3934 | FS2_BTS:QH_D3934 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 373.14 | 415.3476 | 41.094 | 41.094 | 415.3476 | 373.265 | 393.3953 | 393.5203 |  | 324.9166
[6 empty] | coil out + leads |  |  |  | 0.098 | 373.314 | 415.3476 | 41.094 | 41.094 | 415.3476 | 373.363 | 393.5693 | 393.6183 |  | 325.0906
[8 empty] | 3.176e-06 |  | 0.001485 | 373.3638 | 415.3476 | 41.094 | 41.094 | 415.3476 | 373.3645 | 393.619 | 393.6197 |  | 325.1404
FS2 | BTS | BCM | 3936 | FS2_BTS:BCM_D3936 | BCM3 | current monitor |  | 373.3645 | 373.3645 | 0 | 373.3645 | 415.3476 | 41.094 | 41.094 | 415.3476 | 373.3645 | 393.6197 | 393.6197 | 0.1 | 325.1411
[10 empty] | 0.17852 | 373.4538 | 415.3476 | 41.094 | 41.094 | 415.3476 | 373.543 | 393.709 | 393.7983 |  | 325.2304
[6 empty] | coil out |  |  |  | 0.047 | 373.5665 | 415.3476 | 41.094 | 41.094 | 415.3476 | 373.59 | 393.8218 | 393.8453 |  | 325.3431
FS2 | BTS | QV | 3940 | FS2_BTS:QV_D3940 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 373.715 | 415.3476 | 41.094 | 41.094 | 415.3476 | 373.84 | 393.9703 | 394.0953 |  | 325.4916
[8 empty] | 3.176e-06 |  | 0.21871 | 373.9494 | 415.3476 | 41.094 | 41.094 | 415.3476 | 374.0587 | 394.2046 | 394.314 |  | 325.726
FS2 | BTS | BPM | 3943 | FS2_BTS:BPM_D3943 | BPM3 | position monitor | 47.5 | 374.0587 | 374.0587 | 0 | 374.0587 | 415.3476 | 41.094 | 41.094 | 415.3476 | 374.0587 | 394.314 | 394.314 | 0.137 | 325.8353
[10 empty] | 0.081294 | 374.0994 | 415.3476 | 41.094 | 41.094 | 415.3476 | 374.14 | 394.3546 | 394.3953 |  | 325.876
FS2 | BTS | DCHV | 3945 | FS2_BTS:DCHV_D3945 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 374.24 | 415.3476 | 41.094 | 41.094 | 415.3476 | 374.34 | 394.4953 | 394.5953 |  | 326.0166
[10 empty] | 0 | 374.34 | 415.3476 | 41.094 | 41.094 | 415.3476 | 374.34 | 394.5953 | 394.5953 |  | 326.1166
FS2 | BTS | BPM | 3946 | [6 empty] | 0 | 374.34 | 415.3476 | 41.094 | 41.094 | 415.3476 | 374.34 | 394.5953 | 394.5953 | 20170214 BPM moved upstream | 326.1166
[10 empty] | 0.25 | 374.465 | 415.3476 | 41.094 | 41.094 | 415.3476 | 374.59 | 394.7203 | 394.8453 |  | 326.2416
FS2 | BTS | QV | 3950 | FS2_BTS:QV_D3950 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 374.715 | 415.3476 | 41.094 | 41.094 | 415.3476 | 374.84 | 394.9703 | 395.0953 |  | 326.4916
[10 empty] | 0 | 374.84 | 415.3476 | 41.094 | 41.094 | 415.3476 | 374.84 | 395.0953 | 395.0953 |  | 326.6166
[10 empty] | 0.25 | 374.965 | 415.3476 | 41.094 | 41.094 | 415.3476 | 375.09 | 395.2203 | 395.3453 |  | 326.7416
[10 empty] | 0 | 375.09 | 415.3476 | 41.094 | 41.094 | 415.3476 | 375.09 | 395.3453 | 395.3453 |  | 326.8666
FS2 | BTS | QH | 3955 | FS2_BTS:QH_D3955 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 375.215 | 415.3476 | 41.094 | 41.094 | 415.3476 | 375.34 | 395.4703 | 395.5953 |  | 326.9916
[10 empty] | 0.062353 | 375.3712 | 415.3476 | 41.094 | 41.094 | 415.3476 | 375.4024 | 395.6264 | 395.6576 |  | 327.1478
[6 empty] | bellow |  | -0.0016469 |  | 0.1035 | 375.4541 | 415.3476 | 41.094 | 41.094 | 415.3476 | 375.5059 | 395.7094 | 395.7611 | 20170427 adjusted based on ME | 327.2307
FS2 | BTS | BPM | 3958 | FS2_BTS:BPM_D3958 | BPM3 | position monitor |  | 2.03e-06 | 375.5059 | 0 | 375.5059 | 415.3476 | 41.094 | 41.094 | 415.3476 | 375.5059 | 395.7611 | 395.7611 |  | 327.2825
[6 empty] | diagnostic box |  |  |  | 0.1718 | 375.5918 | 415.3476 | 41.094 | 41.094 | 415.3476 | 375.6777 | 395.847 | 395.9329 |  | 327.3684
FS2 | BTS | PM | 3959 | FS2_BTS:PM_D3959 | PM3 | profile monitor |  |  |  | 0 | 375.6777 | 415.3476 | 41.094 | 41.094 | 415.3476 | 375.6777 | 395.9329 | 395.9329 |  | 327.4543
[6 empty] | diagnostic box |  | 0.026517 |  | 0 | 375.6777 | 415.3476 | 41.094 | 41.094 | 415.3476 | 375.6777 | 395.9329 | 395.9329 |  | 327.4543
FS2 | BTS | FAVS | 3959 | FS2_BTS:FAVS_D3959 | FVS | fast valve sensor |  | 2.03e-06 | 375.6777 | 0 | 375.6777 | 415.3476 | 41.094 | 41.094 | 415.3476 | 375.6777 | 395.9329 | 395.9329 |  | 327.4543
FS2 | BTS | IP | 3959 | FS2_BTS:IP_D3959 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 375.6777 | 415.3476 | 41.094 | 41.094 | 415.3476 | 375.6777 | 395.9329 | 395.9329 |  | 327.4543
[6 empty] | bellow |  |  |  | 0.13122 | 375.7433 | 415.3476 | 41.094 | 41.094 | 415.3476 | 375.8089 | 395.9985 | 396.0641 |  | 327.5199
[8 empty] | 3.176e-06 |  | 0.069773 | 375.8438 | 415.3476 | 41.094 | 41.094 | 415.3476 | 375.8787 | 396.099 | 396.1339 | 20170427 adjusted based on ME | 327.6204
FS2 | BTS | DCHV | 3962 | FS2_BTS:DCHV_D3962 | C2 | correctors-X,Y | 50 | 375.9787 | 375.9787 | 0.2 | 375.9787 | 415.3476 | 41.094 | 41.094 | 415.3476 | 376.0787 | 396.2339 | 396.3339 |  | 327.7553
[8 empty] | 0.039879 |  | 0.11309 | 376.1352 | 415.3476 | 41.094 | 41.094 | 415.3476 | 376.1918 | 396.3904 | 396.447 |  | 327.9118
FS2 | BTS | GV | 3965 | FS2_BTS:GV_D3965 | GV_M | gate valve | 63 | 2.176e-06 | 376.2293 | 0.075 | 376.2293 | 415.3476 | 41.094 | 41.094 | 415.3476 | 376.2668 | 396.4845 | 396.522 | 0.075 | 328.0059 | [19 empty] | straight line
Third beam dump line starts in "V" column | [9 empty] | 0.20827 | 376.3709 | 415.3476 | 41.094 | 41.094 | 415.3476 | 376.475 | 396.6261 | 396.7303 |  | 328.1475 | dump3 line start (FS2 straight line dump) | [13 empty] | 41.094 | 415.3476 | 376.475 |  |  | 2013/06/20 length updated from 5.283 to 5.0992 Nathan Bultman
FS2 | BBS | DH | 3979 | FS2_BBS:DH_D3979 | D3 | dipole-45deg, D3-1 | 35 x 100 |  |  | 2.3375 | 377.5273 | 415.7834 | 41.094 | 41.094 | 416.2193 | 378.5795 | 397.899 | 399.0678 |  | 329.4204 | FS2 | BTS | DH | 3979 | FS2_BTS:DH_D3979 | D3 | dipole-45deg, D3-1 | 40 x 100 | 2.03e-06 |  | 2.2419 | 378.717 | 415.3476 | 41.094 | 41.094 | 415.3476 | 378.717 | 397.8512 | 398.9722
[10 empty] | 0.4 | 378.7209 | 416.3607 | 41.094 | 41.094 | 416.5021 | 378.8624 | 399.2678 | 399.4678 |  | 330.7891 | FS2 | BTS | GV | 3990 | FS2_BTS:GV_D3990 | GV_M | gate valve | 63 | 378.7545 | 378.7545 | 0.075 | 378.7545 | 415.3476 | 41.094 | 41.094 | 415.3476 | 378.792 | 399.0097 | 399.0472
FS2 | BBS | QH | 3996 | FS2_BBS:QH_D3996 | Q4 | quadrupole-X,33cm,B2 | 100 |  |  | 0.33 | 378.979 | 416.6188 | 41.094 | 41.094 | 416.7354 | 379.0957 | 399.6328 | 399.7978 |  | 331.1541 | [8 empty] | 2.03e-06 |  | 0.16876 | 378.8764 | 415.3476 | 41.094 | 41.094 | 415.3476 | 378.9607 | 399.1316 | 399.2159
[10 empty] | 0.12 | 379.1381 | 416.7779 | 41.094 | 41.094 | 416.8203 | 379.1806 | 399.8578 | 399.9178 |  | 331.3791 | FS2 | BTS | Q | 3994 | FS2_BTS:Q_D3994 | Q3 | quadrupole-40cm |  | 379.1607 | 379.1607 | 0.4 | 379.1607 | 415.3476 | 41.094 | 41.094 | 415.3476 | 379.3607 | 399.4159 | 399.6159
FS2 | BBS | S | 4000 | FS2_BBS:S_D4000 | H1 | hexapole-X+ | 100 |  |  | 0.1 | 379.2159 | 416.8557 | 41.094 | 41.094 | 416.891 | 379.2513 | 399.9678 | 400.0178 |  | 331.4891 | [10 empty] | 0 | 379.3607 | 415.3476 | 41.094 | 41.094 | 415.3476 | 379.3607 | 399.6159 | 399.6159
[10 empty] | 0.15 | 379.3043 | 416.944 | 41.094 | 41.094 | 416.9971 | 379.3573 | 400.0928 | 400.1678 |  | 331.6141 | [10 empty] | 0 | 379.3607 | 415.3476 | 41.094 | 41.094 | 415.3476 | 379.3607 | 399.6159 | 399.6159
FS2 | BBS | QV | 4004 | FS2_BBS:QV_D4004 | Q5 | quadrupole-Y,40cm,B2 | 100 | 0 | 0.2 | 0.4 | 379.4988 | 417.1385 | 41.094 | 41.094 | 417.2799 | 379.6402 | 400.3678 | 400.5678 |  | 331.8891 | [10 empty] | 0.22731 | 379.4744 | 415.3476 | 41.094 | 41.094 | 415.3476 | 379.588 | 399.7296 | 399.8433
[8 empty] | 0.36323 | 0.11323 | 0.11323 | 379.6802 | 417.32 | 41.094 | 41.094 | 417.36 | 379.7203 | 400.6244 | 400.681 | 20170216 .13 -> .11323 | 332.1457 | FS2 | BTS | Q | 4000 | FS2_BTS:Q_D4000 | Q2 | quadrupole-26cm |  |  |  | 0.26 | 379.718 | 415.3476 | 41.094 | 41.094 | 415.3476 | 379.848 | 399.9733 | 400.1033
FS2 | BBS | S | 4007 | FS2_BBS:S_D4007 | H1 | hexapole-X- | 100 |  | 0.05 | 0.1 | 379.7556 | 417.3953 | 41.094 | 41.094 | 417.4307 | 379.791 | 400.731 | 400.781 |  | 332.2524 | [10 empty] | 0 | 379.848 | 415.3476 | 41.094 | 41.094 | 415.3476 | 379.848 | 400.1033 | 400.1033
[8 empty] | 1e-06 |  | 0.095715 | 379.8248 | 417.4645 | 41.094 | 41.094 | 417.4984 | 379.8586 | 400.8288 | 400.8767 | 20170216 .07 -> .09572 | 332.3502 | [10 empty] | 0.38824 | 380.0422 | 415.3476 | 41.094 | 41.094 | 415.3476 | 380.2363 | 400.2974 | 400.4915
[8 empty] | 0.24572 | 0.095716 | 0 | 379.8586 | 417.4984 | 41.094 | 41.094 | 417.4984 | 379.8586 | 400.8767 | 400.8767 |  | 332.3981 | [10 empty] | 0 | 380.2363 | 415.3476 | 41.094 | 41.094 | 415.3476 | 380.2363 | 400.4915 | 400.4915
FS2 | BBS | DCHV | 4010 | FS2_BBS:DCHV_D4010 | C3 | correctors-X,Y | 100 | 0 | 0.1 | 0.2 | 379.9294 | 417.5691 | 41.094 | 41.094 | 417.6398 | 380.0001 | 400.9767 | 401.0767 |  | 332.4981 | [6 empty] | bellow |  | 2.03e-06 |  | 0.1035 | 380.288 | 415.3476 | 41.094 | 41.094 | 415.3476 | 380.3398 | 400.5432 | 400.595
[8 empty] | 0.37605 | 0.11105 | 0.11105 | 380.0393 | 417.6791 | 41.094 | 41.094 | 417.7183 | 380.0786 | 401.1322 | 401.1878 | 20170216 .12 -> .111055 | 332.6536 | FS2 | BTS | BPM | 4006 | FS2_BTS:BPM_D4006 | BPM4 | position monitor | 100 | 380.3398 | 380.3398 | 0 | 380.3398 | 415.3476 | 41.094 | 41.094 | 415.3476 | 380.3398 | 400.595 | 400.595 | 20170427 adjusted based on ME, but .1657 is different from .1513 other areas
FS2 | BBS | QH | 4014 | FS2_BBS:QH_D4014 | Q4 | quadrupole-X,33cm,B2 | 100 | 0 | 0.165 | 0.33 | 380.1953 | 417.835 | 41.094 | 41.094 | 417.9517 | 380.3119 | 401.3528 | 401.5178 |  | 332.8741 | [6 empty] | diagnostic box |  |  |  | 0.31096 | 380.4953 | 415.3476 | 41.094 | 41.094 | 415.3476 | 380.6507 | 400.7505 | 400.906
[8 empty] | 0.40344 | 0.23844 | 0.23844 | 380.3962 | 418.036 | 41.094 | 41.094 | 418.1203 | 380.4805 | 401.637 | 401.7562 | 20170216 .157931 -> .23844 | 333.1583 | FS2 | BTS | PM | 4009 | FS2_BTS:PM_D4009 | PM5 | profile monitor | 100 |  |  | 0 | 380.6507 | 415.3476 | 41.094 | 41.094 | 415.3476 | 380.6507 | 400.906 | 400.906
FS2 | BBS | IP | 4018 | FS2_BBS:IP_D4018 | IP_L | ion pump 300 l/s |  | 2.5e-07 | 0 | 0 | 380.4805 | 418.1203 | 41.094 | 41.094 | 418.1203 | 380.4805 | 401.7562 | 401.7562 |  | 333.2776 | [6 empty] | diagnostic box |  |  |  | 0 | 380.6507 | 415.3476 | 41.094 | 41.094 | 415.3476 | 380.6507 | 400.906 | 400.906
[8 empty] | 0.19262 | 0.15055 | 0.15055 | 380.5338 | 418.1735 | 41.094 | 41.094 | 418.2267 | 380.587 | 401.8315 | 401.9067 | 20170216 .0 -> .150549 | 333.3528 | [8 empty] | 2.03e-06 | 0.16567 | 0 | 380.6507 | 415.3476 | 41.094 | 41.094 | 415.3476 | 380.6507 | 400.906 | 400.906
FS2 | BBS | BPM | 4019 | FS2_BBS:BPM_D4019 | BPM4 | position monitor R=50mm | 100 |  | 0.042069 | 0.084138 | 380.6167 | 418.2565 | 41.094 | 41.094 | 418.2862 | 380.6465 | 401.9488 | 401.9909 | 20170216 BPM moved after box | 333.4702 | FS2 | BTS | IP | 4009 | FS2_BTS:IP_D4009 | IP_L | ion pump 300 l/s |  | 380.6507 | 380.6507 | 0 | 380.6507 | 415.3476 | 41.094 | 41.094 | 415.3476 | 380.6507 | 400.906 | 400.906
[10 empty] | 0.2548 | 380.7366 | 418.3763 | 41.094 | 41.094 | 418.4664 | 380.8267 | 402.1183 | 402.2457 | 20170216 .485861 -> .254803 | 333.6397 | [6 empty] | bellow |  |  |  | 0.13122 | 380.7163 | 415.3476 | 41.094 | 41.094 | 415.3476 | 380.782 | 400.9716 | 401.0372 | 20170427 adjusted based on ME
FS2 | BBS | DH | 4034 | FS2_BBS:DH_D4034 | D3 | dipole-45deg, D3-2 | 35 x 100 |  |  | 2.3375 | 381.2625 | 419.5186 | 41.094 | 41.094 | 420.5709 | 381.6984 | 403.4144 | 404.5832 |  | 334.9358 | [10 empty] | 0.79228 | 381.1781 | 415.3476 | 41.094 | 41.094 | 415.3476 | 381.5742 | 401.4333 | 401.8295
[10 empty] | 0.50759 | 381.6984 | 420.8247 | 41.094 | 41.094 | 421.0785 | 381.6984 | 404.837 | 405.0908 | 20170216 space adjusted | 336.3583 | FS2 | BTS | BD | 4018 | FS2_BTS:BD_D4018 |  | dump3 |  |  |  | 0 | 381.5742 | 415.3476 | 41.094 | 41.094 | 415.3476 | 381.5742 | 401.8295 | 401.8295
FS2 | BBS | PM | 4051 | FS2_BBS:PM_D4051 | PM5 | profile monitor | 100 |  |  | 0 | 381.6984 | 421.0785 | 41.094 | 41.094 | 421.0785 | 381.6984 | 405.0908 | 405.0908 |  | 336.6121 | dump3 lind end (FS2 straight line dump)
FS2 | BBS | IP | 4051 | FS2_BBS:IP_D4051 | IP_L | ion pump 300 l/s |  |  |  | 0 | 381.6984 | 421.0785 | 41.094 | 41.094 | 421.0785 | 381.6984 | 405.0908 | 405.0908 |  | 336.6121
FS2 | BBS | FAVS | 4051 | FS2_BBS:FAVS_D4051 | FVS | fast valve sensor |  |  |  | 0 | 381.6984 | 421.0785 | 41.094 | 41.094 | 421.0785 | 381.6984 | 405.0908 | 405.0908 |  | 336.6121
[10 empty] | 0.21854 | 381.6984 | 421.1877 | 41.094 | 41.094 | 421.297 | 381.6984 | 405.2 | 405.3093 | 20170216 space adjusted | 336.7214
FS2 | BBS | BPM | 4054 | FS2_BBS:BPM_D4054 | BPM4 | position monitor R=50mm | 100 |  |  | 0.084138 | 381.6984 | 421.3391 | 41.094 | 41.094 | 421.3812 | 381.6984 | 405.3514 | 405.3934 | 20170216 BPM moved after box | 336.8728
[10 empty] | 0.021733 | 381.6984 | 421.392 | 41.094 | 41.094 | 421.4029 | 381.6984 | 405.4043 | 405.4152 | 20170216 space adjusted | 336.9257
FS2 | BBS | DCHV | 4055 | FS2_BBS:DCHV_D4055 | C3 | correctors-X,Y | 100 |  |  | 0.2 | 381.6984 | 421.5029 | 41.094 | 41.094 | 421.6029 | 381.6984 | 405.5152 | 405.6152 |  | 337.0366
[10 empty] | 0.368 | 381.6984 | 421.7869 | 41.094 | 41.094 | 421.9709 | 381.6984 | 405.7992 | 405.9832 |  | 337.3206
FS2 | BBS | DH | 4072 | FS2_BBS:DH_D4072 | D3 | dipole-45deg, D3-3 | 35 x 100 |  |  | 2.3375 | 381.2625 | 423.0231 | 41.094 | 41.094 | 424.0754 | 380.8267 | 407.1519 | 408.3207 |  | 338.6733
[10 empty] | 0.32554 | 380.7116 | 424.1905 | 41.094 | 41.094 | 424.3056 | 380.5965 | 408.4834 | 408.6462 | 20170216 this space NOT symmetric | 340.0048
FS2 | BBS | BPM | 4087 | FS2_BBS:BPM_D4087 | BPM4 | position monitor R=50mm | 100 |  |  | 0.084138 | 380.5667 | 424.3353 | 41.094 | 41.094 | 424.3651 | 380.537 | 408.6883 | 408.7304 |  | 340.2097
[10 empty] | 0.15055 | 380.4838 | 424.4183 | 41.094 | 41.094 | 424.4715 | 380.4305 | 408.8056 | 408.8809 |  | 340.327
FS2 | BBS | IP | 4089 | FS2_BBS:IP_D4089 | IP_L | ion pump 300 l/s |  | 0 | 0 | 0 | 380.4305 | 424.4715 | 41.094 | 41.094 | 424.4715 | 380.4305 | 408.8809 | 408.8809 |  | 340.4023
[8 empty] | 0.33271 | 0.16771 | 0.16771 | 380.3712 | 424.5308 | 41.094 | 41.094 | 424.5901 | 380.3119 | 408.9648 | 409.0486 | 20170216 this space NOT symmetric | 340.4861
FS2 | BBS | QH | 4092 | FS2_BBS:QH_D4092 | Q4 | quadrupole-X,33cm,B2 | 100 |  | 0.165 | 0.33 | 380.1953 | 424.7068 | 41.094 | 41.094 | 424.8234 | 380.0786 | 409.2136 | 409.3786 |  | 340.735
[10 empty] | 0.11105 | 380.0393 | 424.8627 | 41.094 | 41.094 | 424.902 | 380.0001 | 409.4341 | 409.4897 |  | 340.9555
FS2 | BBS | DCHV | 4096 | FS2_BBS:DCHV_D4096 | C3 | correctors-X,Y | 100 |  |  | 0.2 | 379.9294 | 424.9727 | 41.094 | 41.094 | 425.0434 | 379.8586 | 409.5897 | 409.6897 |  | 341.111
[10 empty] | 0 | 379.8586 | 425.0434 | 41.094 | 41.094 | 425.0434 | 379.8586 | 409.6897 | 409.6897 |  | 341.211
[10 empty] | 0.095715 | 379.8248 | 425.0772 | 41.094 | 41.094 | 425.1111 | 379.791 | 409.7375 | 409.7854 |  | 341.2589
FS2 | BBS | S | 4098 | FS2_BBS:S_D4098 | H1 | hexapole-X- | 100 |  |  | 0.1 | 379.7556 | 425.1464 | 41.094 | 41.094 | 425.1818 | 379.7203 | 409.8354 | 409.8854 |  | 341.3568
[10 empty] | 0.11323 | 379.6802 | 425.2218 | 41.094 | 41.094 | 425.2619 | 379.6402 | 409.942 | 409.9986 |  | 341.4634
FS2 | BBS | QV | 4102 | FS2_BBS:QV_D4102 | Q5 | quadrupole-Y,40cm,B2 | 100 |  |  | 0.4 | 379.4988 | 425.4033 | 41.094 | 41.094 | 425.5447 | 379.3573 | 410.1986 | 410.3986 |  | 341.72
[10 empty] | 0.15 | 379.3043 | 425.5977 | 41.094 | 41.094 | 425.6508 | 379.2513 | 410.4736 | 410.5486 |  | 341.995
FS2 | BBS | S | 4106 | FS2_BBS:S_D4106 | H1 | hexapole-X+ | 100 |  |  | 0.1 | 379.2159 | 425.6861 | 41.094 | 41.094 | 425.7215 | 379.1806 | 410.5986 | 410.6486 |  | 342.12
[10 empty] | 0.12 | 379.1381 | 425.7639 | 41.094 | 41.094 | 425.8063 | 379.0957 | 410.7086 | 410.7686 |  | 342.23
FS2 | BBS | QH | 4109 | FS2_BBS:QH_D4109 | Q4 | quadrupole-X,33cm,B2 | 100 |  |  | 0.33 | 378.979 | 425.923 | 41.094 | 41.094 | 426.0397 | 378.8624 | 410.9336 | 411.0986 |  | 342.455
[10 empty] | 0.4 | 378.7209 | 426.1811 | 41.094 | 41.094 | 426.3225 | 378.5795 | 411.2986 | 411.4986 |  | 342.82
FS2 | BBS | DH | 4127 | FS2_BBS:DH_D4127 | D3 | dipole-45deg, D3-4 | 35 x 100 |  |  | 2.3375 | 377.5273 | 426.7584 | 41.094 | 41.094 | 427.1942 | 376.475 | 412.6674 | 413.8361 |  | 344.1887
[7 empty] | -0.17121 | 2.176e-06 |  | 0.20827 | 376.3709 | 427.1942 | 41.094 | 41.094 | 427.1942 | 376.2668 | 413.9402 | 414.0444 | 20170214 more space needed | 345.4616
FS2 | BMS | GV | 4141 | FS2_BMS:GV_D4141 | GV_M | gate valve | 63 | 376.2293 | 376.2293 | 0.075 | 376.2293 | 427.1942 | 41.094 | 41.094 | 427.1942 | 376.1918 | 414.0819 | 414.1194 |  | 345.6032
[6 empty] | bellow | 40 | 2.176e-06 |  | 0.094701 | 376.1444 | 427.1942 | 41.094 | 41.094 | 427.1942 | 376.0971 | 414.1667 | 414.2141 |  | 345.6881
FS2 | BMS | BPM | 4142 | FS2_BMS:BPM_D4142 | BPM3 | position monitor |  | 376.0971 | 376.0971 | 0 | 376.0971 | 427.1942 | 41.094 | 41.094 | 427.1942 | 376.0971 | 414.2141 | 414.2141 |  | 345.7354
[6 empty] | diagnostic box |  |  |  | 0.14528 | 376.0244 | 427.1942 | 41.094 | 41.094 | 427.1942 | 375.9518 | 414.2867 | 414.3594 |  | 345.8081
FS2 | BMS | PM | 4144 | FS2_BMS:PM_D4144 | PM3 | profile monitor |  | 2.176e-06 |  | 0 | 375.9518 | 427.1942 | 41.094 | 41.094 | 427.1942 | 375.9518 | 414.3594 | 414.3594 |  | 345.8807
[6 empty] | diagnostic box |  | 375.9518 | 375.9518 | 0 | 375.9518 | 427.1942 | 41.094 | 41.094 | 427.1942 | 375.9518 | 414.3594 | 414.3594 |  | 345.8807
FS2 | BMS | FAVS | 4144 | FS2_BMS:FAVS_D4144 | FVS | fast valve sensor |  |  |  | 0 | 375.9518 | 427.1942 | 41.094 | 41.094 | 427.1942 | 375.9518 | 414.3594 | 414.3594 |  | 345.8807
FS2 | BMS | IP | 4144 | FS2_BMS:IP_D4144 | IP_S | ion pump, 75 l/s |  | -8.8818e-15 |  | 0 | 375.9518 | 427.1942 | 41.094 | 41.094 | 427.1942 | 375.9518 | 414.3594 | 414.3594 |  | 345.8807
[6 empty] | bellow | 40 | 1.135 |  | 0.099 | 375.9023 | 427.1942 | 41.094 | 41.094 | 427.1942 | 375.8528 | 414.4089 | 414.4584 |  | 345.9302
[8 empty] | 1.135 |  | 0.087751 | 375.8089 | 427.1942 | 41.094 | 41.094 | 427.1942 | 375.765 | 414.5022 | 414.5461 |  | 346.0236
FS2 | BMS | DCHV | 4146 | FS2_BMS:DCHV_D4146 | C2 | correctors-X,Y | 50 |  | 375.665 | 0.2 | 375.665 | 427.1942 | 41.094 | 41.094 | 427.1942 | 375.565 | 414.6461 | 414.7461 |  | 346.1675
[10 empty] | 0.225 | 375.4525 | 427.1942 | 41.094 | 41.094 | 427.1942 | 375.34 | 414.8586 | 414.9711 | 20170214 0.2 --> 0.225 per ME | 346.38
FS2 | BMS | QH | 4151 | FS2_BMS:QH_D4151 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 375.215 | 427.1942 | 41.094 | 41.094 | 427.1942 | 375.09 | 415.0961 | 415.2211 |  | 346.6175
[10 empty] | 0.25 | 374.965 | 427.1942 | 41.094 | 41.094 | 427.1942 | 374.84 | 415.3461 | 415.4711 |  | 346.8675
FS2 | BMS | QV | 4156 | FS2_BMS:QV_D4156 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 374.715 | 427.1942 | 41.094 | 41.094 | 427.1942 | 374.59 | 415.5961 | 415.7211 |  | 347.1175
[10 empty] | 0.3 | 374.44 | 427.1942 | 41.094 | 41.094 | 427.1942 | 374.29 | 415.8711 | 416.0211 |  | 347.3925
FS2 | BMS | DCHV | 4161 | FS2_BMS:DCHV_D4161 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 374.19 | 427.1942 | 41.094 | 41.094 | 427.1942 | 374.09 | 416.1211 | 416.2211 |  | 347.6425
[8 empty] | 3.176e-06 |  | 0.19316 | 373.9935 | 427.1942 | 41.094 | 41.094 | 427.1942 | 373.8969 | 416.3177 | 416.4143 |  | 347.8391
FS2 | BMS | BPM | 4164 | FS2_BMS:BPM_D4164 | BPM3 | position monitor | 47.5 | 373.8969 | 373.8969 | 0 | 373.8969 | 427.1942 | 41.094 | 41.094 | 427.1942 | 373.8969 | 416.4143 | 416.4143 | 0.137 | 347.9356
[10 empty] | 0.056842 | 373.8685 | 427.1942 | 41.094 | 41.094 | 427.1942 | 373.84 | 416.4427 | 416.4711 |  | 347.9641
FS2 | BMS | QV | 4166 | FS2_BMS:QV_D4166 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 373.715 | 427.1942 | 41.094 | 41.094 | 427.1942 | 373.59 | 416.5961 | 416.7211 |  | 348.1175
[8 empty] | 3.176e-06 |  | 0.18331 | 373.4984 | 427.1942 | 41.094 | 41.094 | 427.1942 | 373.4067 | 416.8128 | 416.9044 |  | 348.3341
FS2 | BBS | BCM | 4169 | FS2_BBS:BCM_D4169 | BCM3 | current monitor |  | 373.4067 | 373.4067 | 0 | 373.4067 | 427.1942 | 41.094 | 41.094 | 427.1942 | 373.4067 | 416.9044 | 416.9044 | 0.1 | 348.4258
[10 empty] | 0.14169 | 373.3359 | 427.1942 | 41.094 | 41.094 | 427.1942 | 373.265 | 416.9753 | 417.0461 |  | 348.4966
FS2 | BMS | QH | 4172 | FS2_BMS:QH_D4172 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 373.14 | 427.1942 | 41.094 | 41.094 | 427.1942 | 373.015 | 417.1711 | 417.2961 |  | 348.6925
[10 empty] | 0.13938 | 372.9453 | 427.1942 | 41.094 | 41.094 | 427.1942 | 372.8757 | 417.3658 | 417.4355 |  | 348.8872
[10 empty] | 0.1 | 372.8257 | 427.1942 | 41.094 | 41.094 | 427.1942 | 372.7757 | 417.4855 | 417.5355 |  | 349.0069
[8 empty] | 3.176e-06 |  | 0 | 372.7757 | 427.1942 | 41.094 | 41.094 | 427.1942 | 372.7757 | 417.5355 | 417.5355 |  | 349.0569
FS2 | BMS | FAV | 4175 | FS2_BMS:FAV_D4175 | FV | fast valve | 40 | 372.7757 | 372.7757 | 0 | 372.7757 | 427.1942 | 41.094 | 41.094 | 427.1942 | 372.7757 | 417.5355 | 417.5355 | 0.067196 | 349.0569
[10 empty] | 0.037495 | 372.7569 | 427.1942 | 41.094 | 41.094 | 427.1942 | 372.7382 | 417.5542 | 417.573 |  | 349.0756
[6 empty] | bellow | 40 | -0.0044369 |  | 0.1035 | 372.6864 | 427.1942 | 41.094 | 41.094 | 427.1942 | 372.6347 | 417.6247 | 417.6765 |  | 349.1461
FS2 | BMS | BPM | 4177 | FS2_BMS:BPM_D4177 | BPM2 | position monitor |  | 2.03e-06 | 372.6347 | 0 | 372.6347 | 427.1942 | 41.094 | 41.094 | 427.1942 | 372.6347 | 417.6765 | 417.6765 |  | 349.1979
[6 empty] | diagnostic box | 0.11862 |  |  | 0.14528 | 372.562 | 427.1942 | 41.094 | 41.094 | 427.1942 | 372.4894 | 417.7491 | 417.8218 |  | 349.2705
FS2 | BMS | PM | 4178 | FS2_BMS:PM_D4178 | PMx | profile monitor (not costed) | future PM |  |  | 0 | 372.4894 | 427.1942 | 41.094 | 41.094 | 427.1942 | 372.4894 | 417.8218 | 417.8218 |  | 349.3431
[6 empty] | diagnostic box |  | 2.03e-06 | 372.4894 | 0 | 372.4894 | 427.1942 | 41.094 | 41.094 | 427.1942 | 372.4894 | 417.8218 | 417.8218 |  | 349.3431
FS2 | BMS | IP | 4178 | FS2_BMS:IP_D4178 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 372.4894 | 427.1942 | 41.094 | 41.094 | 427.1942 | 372.4894 | 417.8218 | 417.8218 |  | 349.3431
[10 empty] | 0 | 372.4894 | 427.1942 | 41.094 | 41.094 | 427.1942 | 372.4894 | 417.8218 | 417.8218 |  | 349.3431
[6 empty] | bellow | 40 |  |  | 0.13122 | 372.4238 | 427.1942 | 41.094 | 41.094 | 427.1942 | 372.3582 | 417.8874 | 417.953 |  | 349.4087
[9 empty] | ??? | 0.0086001 | 372.3539 | 427.1942 | 41.094 | 41.094 | 427.1942 | 372.3496 | 417.9573 | 417.9616 | 20120312: why this odd number? | 349.4787
FS2 | CG01 | GV | 4180 | FS2_CG01:GV_D4180 |  | gate valve | 40 |  |  | 0.057404 | 372.3209 | 427.1942 | 41.094 | 41.094 | 427.1942 | 372.2922 | 417.9903 | 418.019 |  | 349.5117
[6 empty] | bellows |  |  |  | 0.17589 | 372.2042 | 427.1942 | 41.094 | 41.094 | 427.1942 | 372.1163 | 418.1069 | 418.1949 |  | 349.6283
FS2 | CG01 | CAV1 | 4184 | FS2_CG01:CAV1_D4184 |  | b54 resonator | 40 |  |  | 0.49566 | 371.8684 | 427.1942 | 41.094 | 41.094 | 427.1942 | 371.6206 | 418.4427 | 418.6905 |  | 349.9641
[6 empty] | bellow |  |  |  | 0.099517 | 371.5709 | 427.1942 | 41.094 | 41.094 | 427.1942 | 371.5211 | 418.7403 | 418.7901 |  | 350.2617
FS2 | CG01 | CAV2 | 4190 | FS2_CG01:CAV2_D4190 |  | b54 resonator | 40 |  |  | 0.49566 | 371.2733 | 427.1942 | 41.094 | 41.094 | 427.1942 | 371.0254 | 419.0379 | 419.2857 |  | 350.5593
[6 empty] | bellow |  |  |  | 0.099517 | 370.9757 | 427.1942 | 41.094 | 41.094 | 427.1942 | 370.9259 | 419.3355 | 419.3852 |  | 350.8568
[10 empty] | 0.20048 | 370.8257 | 427.1942 | 41.094 | 41.094 | 427.1942 | 370.7254 | 419.4855 | 419.5857 |  | 351.0068
[6 empty] | bellow |  |  |  | 0.099517 | 370.6757 | 427.1942 | 41.094 | 41.094 | 427.1942 | 370.6259 | 419.6355 | 419.6852 |  | 351.1568
FS2 | CG01 | CAV3 | 4199 | FS2_CG01:CAV3_D4199 |  | b54 resonator | 40 |  |  | 0.49566 | 370.3781 | 427.1942 | 41.094 | 41.094 | 427.1942 | 370.1303 | 419.9331 | 420.1809 |  | 351.4544
[6 empty] | bellow |  |  |  | 0.099517 | 370.0805 | 427.1942 | 41.094 | 41.094 | 427.1942 | 370.0307 | 420.2306 | 420.2804 |  | 351.752
FS2 | CG01 | CAV4 | 4205 | FS2_CG01:CAV4_D4205 |  | b54 resonator | 40 |  |  | 0.49566 | 369.7829 | 427.1942 | 41.094 | 41.094 | 427.1942 | 369.5351 | 420.5282 | 420.7761 |  | 352.0496
[6 empty] | bellows |  | 3.176e-06 |  | 0.17589 | 369.4471 | 427.1942 | 41.094 | 41.094 | 427.1942 | 369.3592 | 420.864 | 420.952 |  | 352.3854
FS2 | CG01 | GV | 4210 | FS2_CG01:GV_D4210 |  | gate valve | 40 | 369.3018 | 369.3018 | 0.057404 | 369.3305 | 427.1942 | 41.094 | 41.094 | 427.1942 | 369.3018 | 420.9807 | 421.0094 |  | 352.502
[7 empty] | 40 |  |  | 0 | 369.3018 | 427.1942 | 41.094 | 41.094 | 427.1942 | 369.3018 | 421.0094 | 421.0094 |  | 352.5307
[10 empty] | 0 | 369.3018 | 427.1942 | 41.094 | 41.094 | 427.1942 | 369.3018 | 421.0094 | 421.0094 |  | 352.5307
[10 empty] | 0 | 369.3018 | 427.1942 | 41.094 | 41.094 | 427.1942 | 369.3018 | 421.0094 | 421.0094 |  | 352.5307
[10 empty] | 0 | 369.3018 | 427.1942 | 41.094 | 41.094 | 427.1942 | 369.3018 | 421.0094 | 421.0094 |  | 352.5307
[10 empty] | 0 | 369.3018 | 427.1942 | 41.094 | 41.094 | 427.1942 | 369.3018 | 421.0094 | 421.0094 |  | 352.5307
[8 empty] | 0 |  | 0.18837 | 369.2076 | 427.1942 | 41.094 | 41.094 | 427.1942 | 369.1134 | 421.1035 | 421.1977 | 20170214 box moved downstream | 352.6249
[8 empty] | 3.9382 | 3.9382 | 0.036554 | 369.0951 | 427.1942 | 41.094 | 41.094 | 427.1942 | 369.0769 | 421.216 | 421.2343 |  | 352.7374
FS2 | BMS | DCHV | 4213 | FS2_BMS:DCHV_D4213 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 368.9769 | 427.1942 | 41.094 | 41.094 | 427.1942 | 368.8769 | 421.3343 | 421.4343 |  | 352.8557
[8 empty] | 3.176e-06 |  | 0.19472 | 368.7795 | 427.1942 | 41.094 | 41.094 | 427.1942 | 368.6821 | 421.5316 | 421.629 |  | 353.053
FS2 | BMS | BPM | 4216 | FS2_BMS:BPM_D4216 | BPM3 | position monitor | 47.5 | 368.6821 | 368.6821 | 0 | 368.6821 | 427.1942 | 41.094 | 41.094 | 427.1942 | 368.6821 | 421.629 | 421.629 |  | 353.1504
[10 empty] | 0.05528 | 368.6545 | 427.1942 | 41.094 | 41.094 | 427.1942 | 368.6269 | 421.6566 | 421.6843 |  | 353.178
FS2 | BMS | QH | 4218 | FS2_BMS:QH_D4218 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 368.5019 | 427.1942 | 41.094 | 41.094 | 427.1942 | 368.3769 | 421.8093 | 421.9343 |  | 353.3307
[10 empty] | 0 | 368.3769 | 427.1942 | 41.094 | 41.094 | 427.1942 | 368.3769 | 421.9343 | 421.9343 | 20170214 split between the bellow and nipple | 353.4557
[6 empty] | bellow | 40 |  |  | 0.28777 | 368.233 | 427.1942 | 41.094 | 41.094 | 427.1942 | 368.0891 | 422.0782 | 422.2221 |  | 353.5995
[6 empty] | vacuum box |  | 3.176e-06 |  | 0 | 368.0891 | 427.1942 | 41.094 | 41.094 | 427.1942 | 368.0891 | 422.2221 | 422.2221 |  | 353.7434
FS2 | BMS | FAVS | 3537 | FS2_BMS:FAVS_D3537 | FVS | fast valve sensor |  | 368.0891 | 368.0891 | 0 | 368.0891 | 427.1942 | 41.094 | 41.094 | 427.1942 | 368.0891 | 422.2221 | 422.2221 |  | 353.7434
[6 empty] | vacuum box |  |  |  | 0 | 368.0891 | 427.1942 | 41.094 | 41.094 | 427.1942 | 368.0891 | 422.2221 | 422.2221 |  | 353.7434
FS2 | BMS | IP | 3537 | FS2_BMS:IP_D3537 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 368.0891 | 427.1942 | 41.094 | 41.094 | 427.1942 | 368.0891 | 422.2221 | 422.2221 |  | 353.7434
[6 empty] | bellow | 40 |  |  | 0.21223 | 367.983 | 427.1942 | 41.094 | 41.094 | 427.1942 | 367.8769 | 422.3282 | 422.4343 |  | 353.8495
FS2 | BMS | QV | 4226 | FS2_BMS:QV_D4226 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 367.7519 | 427.1942 | 41.094 | 41.094 | 427.1942 | 367.6269 | 422.5593 | 422.6843 |  | 354.0807
[10 empty] | 2.5923 | 366.3307 | 427.1942 | 41.094 | 41.094 | 427.1942 | 365.0346 | 423.9804 | 425.2766 |  | 355.5018
FS2 | BMS | BD | 4253 | FS2_BMS:BD_D4253 | dump | dump 25W no current reading |  | 2.176e-06 | 365.0346 | 0 | 365.0346 | 427.1942 | 41.094 | 41.094 | 427.1942 | 365.0346 | 425.2766 | 425.2766 | 20170214 check position with ME | 356.7979
[10 empty] | 2.3777 | 363.8457 | 427.1942 | 41.094 | 41.094 | 427.1942 | 362.6569 | 426.4654 | 427.6543 |  | 357.9868
FS2 | BMS | LPM | 4277 | FS2_BMS:LPM_D4277 | LPMx | bunch length monitor (not costed) | future LPM2 |  |  | 0 | 362.6569 | 427.1942 | 41.094 | 41.094 | 427.1942 | 362.6569 | 427.6543 | 427.6543 |  | 359.1757
FS2 | BMS | IP | 4277 | FS2_BMS:IP_D4277 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 362.6569 | 427.1942 | 41.094 | 41.094 | 427.1942 | 362.6569 | 427.6543 | 427.6543 |  | 359.1757
[10 empty] | 0.32183 | 362.496 | 427.1942 | 41.094 | 41.094 | 427.1942 | 362.335 | 427.8152 | 427.9761 |  | 359.3366
FS2 | BMS | DCHV | 4281 | FS2_BMS:DCHV_D4281 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 362.235 | 427.1942 | 41.094 | 41.094 | 427.1942 | 362.135 | 428.0761 | 428.1761 |  | 359.5975
[8 empty] | 3.176e-06 |  | 0.12715 | 362.0715 | 427.1942 | 41.094 | 41.094 | 427.1942 | 362.0079 | 428.2397 | 428.3033 |  | 359.7611
FS2 | BMS | BPM | 4283 | FS2_BMS:BPM_D4283 | BPM3 | position monitor | 47.5 | 362.0079 | 362.0079 | 0 | 362.0079 | 427.1942 | 41.094 | 41.094 | 427.1942 | 362.0079 | 428.3033 | 428.3033 | 0.137 | 359.8246
[10 empty] | 0.12285 | 361.9465 | 427.1942 | 41.094 | 41.094 | 427.1942 | 361.885 | 428.3647 | 428.4261 |  | 359.8861
FS2 | BMS | QH | 4286 | FS2_BMS:QH_D4286 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 361.76 | 427.1942 | 41.094 | 41.094 | 427.1942 | 361.635 | 428.5511 | 428.6761 |  | 360.0725
[10 empty] | 0.75 | 361.26 | 427.1942 | 41.094 | 41.094 | 427.1942 | 360.885 | 429.0511 | 429.4261 |  | 360.5725
FS2 | BMS | QV | 4296 | FS2_BMS:QV_D4296 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 360.76 | 427.1942 | 41.094 | 41.094 | 427.1942 | 360.635 | 429.5511 | 429.6761 |  | 361.0725
[10 empty] | 0.87 | 360.2 | 427.1942 | 41.094 | 41.094 | 427.1942 | 359.765 | 430.1111 | 430.5461 |  | 361.6325
[10 empty] | 0 | 359.765 | 427.1942 | 41.094 | 41.094 | 427.1942 | 359.765 | 430.5461 | 430.5461 | dump moved upstream | 362.0675
[10 empty] | 2 | 358.765 | 427.1942 | 41.094 | 41.094 | 427.1942 | 357.765 | 431.5461 | 432.5461 |  | 363.0675
[6 empty] | bellow | 40 | -3.3824e-05 |  | 0.1035 | 357.7133 | 427.1942 | 41.094 | 41.094 | 427.1942 | 357.6615 | 432.5979 | 432.6496 |  | 364.1192
FS2 | BMS | BPM | 4326 | FS2_BMS:BPM_D4326 | BPM3 | position monitor |  | 357.6615 | 357.6615 | 0 | 357.6615 | 427.1942 | 41.094 | 41.094 | 427.1942 | 357.6615 | 432.6496 | 432.6496 |  | 364.171
[6 empty] | diagnostic box |  |  |  | 0.14528 | 357.5889 | 427.1942 | 41.094 | 41.094 | 427.1942 | 357.5163 | 432.7223 | 432.7949 |  | 364.2436
FS2 | BMS | PM | 4328 | FS2_BMS:PM_D4328 | PMx | profile monitor (not costed) | future PM |  |  | 0 | 357.5163 | 427.1942 | 41.094 | 41.094 | 427.1942 | 357.5163 | 432.7949 | 432.7949 |  | 364.3163
[6 empty] | diagnostic box |  | -3.3824e-05 |  | 0 | 357.5163 | 427.1942 | 41.094 | 41.094 | 427.1942 | 357.5163 | 432.7949 | 432.7949 |  | 364.3163
FS2 | BMS | IP | 4328 | FS2_BMS:IP_D4328 | IP_S | ion pump, 75 l/s |  | 357.5162 | 357.5163 | 0 | 357.5163 | 427.1942 | 41.094 | 41.094 | 427.1942 | 357.5163 | 432.7949 | 432.7949 |  | 364.3163
[6 empty] | bellow | 40 |  |  | 0.13122 | 357.4506 | 427.1942 | 41.094 | 41.094 | 427.1942 | 357.385 | 432.8605 | 432.9261 |  | 364.3819
[10 empty] | 0 | 357.385 | 427.1942 | 41.094 | 41.094 | 427.1942 | 357.385 | 432.9261 | 432.9261 |  | 364.4475
LS3 | CD01 | GV | 4330 | LS3_CD01:GV_D4330 |  | gate valve | 40 |  |  | 0.057404 | 357.3563 | 427.1942 | 41.094 | 41.094 | 427.1942 | 357.3276 | 432.9548 | 432.9835 |  | 364.4762
[6 empty] | bellows |  |  |  | 0.1677 | 357.2438 | 427.1942 | 41.094 | 41.094 | 427.1942 | 357.1599 | 433.0674 | 433.1512 |  | 364.5887
LS3 | CD01 | CAV1 | 4334 | LS3_CD01:CAV1_D4334 |  | b54 resonator | 40 |  |  | 0.49566 | 356.9121 | 427.1942 | 41.094 | 41.094 | 427.1942 | 356.6643 | 433.399 | 433.6469 |  | 364.9204
LS3 | CD01 | BLM | 4336 | LS3_CD01:BLM_D4336 |  | outside loss monitor |  |  |  | 0 | 356.6643 | 427.1942 | 41.094 | 41.094 | 427.1942 | 356.6643 | 433.6469 | 433.6469 |  | 365.1682
[10 empty] | 0 | 356.6643 | 427.1942 | 41.094 | 41.094 | 427.1942 | 356.6643 | 433.6469 | 433.6469 |  | 365.1682
[6 empty] | bellow |  |  |  | 0.099517 | 356.6145 | 427.1942 | 41.094 | 41.094 | 427.1942 | 356.5648 | 433.6966 | 433.7464 |  | 365.218
LS3 | CD01 | CAV2 | 4340 | LS3_CD01:CAV2_D4340 |  | b54 resonator | 40 |  |  | 0.49566 | 356.3169 | 427.1942 | 41.094 | 41.094 | 427.1942 | 356.0691 | 433.9942 | 434.242 |  | 365.5156
[10 empty] | 0 | 356.0691 | 427.1942 | 41.094 | 41.094 | 427.1942 | 356.0691 | 434.242 | 434.242 |  | 365.7634
[10 empty] | 0 | 356.0691 | 427.1942 | 41.094 | 41.094 | 427.1942 | 356.0691 | 434.242 | 434.242 |  | 365.7634
[6 empty] | bellow |  |  |  | 0.099517 | 356.0193 | 427.1942 | 41.094 | 41.094 | 427.1942 | 355.9696 | 434.2918 | 434.3416 |  | 365.8132
LS3 | CD01 | CAV3 | 4346 | LS3_CD01:CAV3_D4346 |  | b54 resonator | 40 |  |  | 0.49566 | 355.7218 | 427.1942 | 41.094 | 41.094 | 427.1942 | 355.4739 | 434.5894 | 434.8372 |  | 366.1108
[10 empty] | 0 | 355.4739 | 427.1942 | 41.094 | 41.094 | 427.1942 | 355.4739 | 434.8372 | 434.8372 |  | 366.3586
[10 empty] | 0 | 355.4739 | 427.1942 | 41.094 | 41.094 | 427.1942 | 355.4739 | 434.8372 | 434.8372 |  | 366.3586
[6 empty] | bellow |  |  |  | 0.099517 | 355.4242 | 427.1942 | 41.094 | 41.094 | 427.1942 | 355.3744 | 434.887 | 434.9367 |  | 366.4084
LS3 | CD01 | CAV4 | 4352 | LS3_CD01:CAV4_D4352 |  | b54 resonator | 40 |  |  | 0.49566 | 355.1266 | 427.1942 | 41.094 | 41.094 | 427.1942 | 354.8788 | 435.1846 | 435.4324 |  | 366.7059
LS3 | CD01 | BLM | 4354 | LS3_CD01:BLM_D4354 |  | outside loss monitor |  |  |  | 0 | 354.8788 | 427.1942 | 41.094 | 41.094 | 427.1942 | 354.8788 | 435.4324 | 435.4324 |  | 366.9538
[10 empty] | 0 | 354.8788 | 427.1942 | 41.094 | 41.094 | 427.1942 | 354.8788 | 435.4324 | 435.4324 |  | 366.9538
[6 empty] | bellow |  |  |  | 0.15525 | 354.8011 | 427.1942 | 41.094 | 41.094 | 427.1942 | 354.7235 | 435.51 | 435.5876 |  | 367.0314
LS3 | CD01 | SOL1 | 4358 | LS3_CD01:SOL1_D4358 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 354.4735 | 427.1942 | 41.094 | 41.094 | 427.1942 | 354.2235 | 435.8376 | 436.0876 |  | 367.359
[6 empty] | bellow |  |  |  | 0.15525 | 354.1459 | 427.1942 | 41.094 | 41.094 | 427.1942 | 354.0683 | 436.1653 | 436.2429 |  | 367.6866
LS3 | CD01 | CAV5 | 4365 | LS3_CD01:CAV5_D4365 |  | b54 resonator | 40 |  |  | 0.49566 | 353.8204 | 427.1942 | 41.094 | 41.094 | 427.1942 | 353.5726 | 436.4907 | 436.7386 |  | 368.0121
[10 empty] | 0 | 353.5726 | 427.1942 | 41.094 | 41.094 | 427.1942 | 353.5726 | 436.7386 | 436.7386 |  | 368.2599
[10 empty] | 0 | 353.5726 | 427.1942 | 41.094 | 41.094 | 427.1942 | 353.5726 | 436.7386 | 436.7386 |  | 368.2599
[6 empty] | bellow |  |  |  | 0.099517 | 353.5228 | 427.1942 | 41.094 | 41.094 | 427.1942 | 353.4731 | 436.7883 | 436.8381 |  | 368.3097
LS3 | CD01 | CAV6 | 4371 | LS3_CD01:CAV6_D4371 |  | b54 resonator | 40 |  |  | 0.49566 | 353.2253 | 427.1942 | 41.094 | 41.094 | 427.1942 | 352.9774 | 437.0859 | 437.3337 |  | 368.6073
[10 empty] | 0 | 352.9774 | 427.1942 | 41.094 | 41.094 | 427.1942 | 352.9774 | 437.3337 | 437.3337 |  | 368.8551
[10 empty] | 0 | 352.9774 | 427.1942 | 41.094 | 41.094 | 427.1942 | 352.9774 | 437.3337 | 437.3337 |  | 368.8551
[6 empty] | bellow |  |  |  | 0.099517 | 352.9277 | 427.1942 | 41.094 | 41.094 | 427.1942 | 352.8779 | 437.3835 | 437.4332 |  | 368.9049
LS3 | CD01 | CAV7 | 4377 | LS3_CD01:CAV7_D4377 |  | b54 resonator | 40 |  |  | 0.49566 | 352.6301 | 427.1942 | 41.094 | 41.094 | 427.1942 | 352.3823 | 437.6811 | 437.9289 |  | 369.2024
[10 empty] | 0 | 352.3823 | 427.1942 | 41.094 | 41.094 | 427.1942 | 352.3823 | 437.9289 | 437.9289 |  | 369.4503
[10 empty] | 0 | 352.3823 | 427.1942 | 41.094 | 41.094 | 427.1942 | 352.3823 | 437.9289 | 437.9289 |  | 369.4503
[6 empty] | bellow |  |  |  | 0.099517 | 352.3325 | 427.1942 | 41.094 | 41.094 | 427.1942 | 352.2827 | 437.9787 | 438.0284 |  | 369.5
LS3 | CD01 | CAV8 | 4383 | LS3_CD01:CAV8_D4383 |  | b54 resonator | 40 |  |  | 0.49566 | 352.0349 | 427.1942 | 41.094 | 41.094 | 427.1942 | 351.7871 | 438.2762 | 438.5241 |  | 369.7976
LS3 | CD01 | BLM | 4385 | LS3_CD01:BLM_D4385 |  | outside loss monitor |  |  |  | 0 | 351.7871 | 427.1942 | 41.094 | 41.094 | 427.1942 | 351.7871 | 438.5241 | 438.5241 |  | 370.0454
[10 empty] | 0 | 351.7871 | 427.1942 | 41.094 | 41.094 | 427.1942 | 351.7871 | 438.5241 | 438.5241 |  | 370.0454
[10 empty] | 0.1677 | 351.7032 | 427.1942 | 41.094 | 41.094 | 427.1942 | 351.6194 | 438.6079 | 438.6918 |  | 370.1293
LS3 | CD01 | GV | 4387 | LS3_CD01:GV_D4387 |  | gate valve | 40 |  |  | 0.057404 | 351.5907 | 427.1942 | 41.094 | 41.094 | 427.1942 | 351.562 | 438.7205 | 438.7492 |  | 370.2418
[6 empty] | bellow | 40 |  |  | 0.1035 | 351.5102 | 427.1942 | 41.094 | 41.094 | 427.1942 | 351.4585 | 438.8009 | 438.8527 |  | 370.3223
LS3 | WD01 | BPM | 4389 | LS3_WD01:BPM_D4389 | BPM2 | position monitor |  |  |  | 0 | 351.4585 | 427.1942 | 41.094 | 41.094 | 427.1942 | 351.4585 | 438.8527 | 438.8527 |  | 370.374
[6 empty] | diagnostic box |  |  |  | 0.14528 | 351.3858 | 427.1942 | 41.094 | 41.094 | 427.1942 | 351.3132 | 438.9253 | 438.998 |  | 370.4467
LS3 | WD01 | PORT | 4390 | LS3_WD01:PORT_D4390 |  | port for future diagnostics |  |  |  | 0 | 351.3132 | 427.1942 | 41.094 | 41.094 | 427.1942 | 351.3132 | 438.998 | 438.998 |  | 370.5193
[6 empty] | diagnostic box |  |  |  | 0 | 351.3132 | 427.1942 | 41.094 | 41.094 | 427.1942 | 351.3132 | 438.998 | 438.998 |  | 370.5193
LS3 | WD01 | IP | 4390 | LS3_WD01:IP_D4390 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 351.3132 | 427.1942 | 41.094 | 41.094 | 427.1942 | 351.3132 | 438.998 | 438.998 |  | 370.5193
[6 empty] | bellow | 40 |  |  | 0.13122 | 351.2476 | 427.1942 | 41.094 | 41.094 | 427.1942 | 351.182 | 439.0636 | 439.1292 |  | 370.5849
LS3 | CD02 | GV | 4392 | LS3_CD02:GV_D4392 |  | gate valve | 40 |  |  | 0.057404 | 351.1533 | 427.1942 | 41.094 | 41.094 | 427.1942 | 351.1246 | 439.1579 | 439.1866 |  | 370.6792
[6 empty] | bellows |  |  |  | 0.1677 | 351.0407 | 427.1942 | 41.094 | 41.094 | 427.1942 | 350.9569 | 439.2704 | 439.3543 |  | 370.7918
LS3 | CD02 | CAV1 | 4396 | LS3_CD02:CAV1_D4396 |  | b54 resonator | 40 |  |  | 0.49566 | 350.709 | 427.1942 | 41.094 | 41.094 | 427.1942 | 350.4612 | 439.6021 | 439.8499 |  | 371.1235
LS3 | CD02 | BLM | 4398 | LS3_CD02:BLM_D4398 |  | outside loss monitor |  |  |  | 0 | 350.4612 | 427.1942 | 41.094 | 41.094 | 427.1942 | 350.4612 | 439.8499 | 439.8499 |  | 371.3713
[10 empty] | 0 | 350.4612 | 427.1942 | 41.094 | 41.094 | 427.1942 | 350.4612 | 439.8499 | 439.8499 |  | 371.3713
[6 empty] | bellow |  |  |  | 0.099517 | 350.4114 | 427.1942 | 41.094 | 41.094 | 427.1942 | 350.3617 | 439.8997 | 439.9495 |  | 371.4211
LS3 | CD02 | CAV2 | 4402 | LS3_CD02:CAV2_D4402 |  | b54 resonator | 40 |  |  | 0.49566 | 350.1139 | 427.1942 | 41.094 | 41.094 | 427.1942 | 349.866 | 440.1973 | 440.4451 |  | 371.7187
[10 empty] | 0 | 349.866 | 427.1942 | 41.094 | 41.094 | 427.1942 | 349.866 | 440.4451 | 440.4451 |  | 371.9665
[10 empty] | 0 | 349.866 | 427.1942 | 41.094 | 41.094 | 427.1942 | 349.866 | 440.4451 | 440.4451 |  | 371.9665
[6 empty] | bellow |  |  |  | 0.099517 | 349.8163 | 427.1942 | 41.094 | 41.094 | 427.1942 | 349.7665 | 440.4949 | 440.5446 |  | 372.0162
LS3 | CD02 | CAV3 | 4408 | LS3_CD02:CAV3_D4408 |  | b54 resonator | 40 |  |  | 0.49566 | 349.5187 | 427.1942 | 41.094 | 41.094 | 427.1942 | 349.2709 | 440.7925 | 441.0403 |  | 372.3138
[10 empty] | 0 | 349.2709 | 427.1942 | 41.094 | 41.094 | 427.1942 | 349.2709 | 441.0403 | 441.0403 |  | 372.5617
[10 empty] | 0 | 349.2709 | 427.1942 | 41.094 | 41.094 | 427.1942 | 349.2709 | 441.0403 | 441.0403 |  | 372.5617
[6 empty] | bellow |  |  |  | 0.099517 | 349.2211 | 427.1942 | 41.094 | 41.094 | 427.1942 | 349.1713 | 441.09 | 441.1398 |  | 372.6114
LS3 | CD02 | CAV4 | 4414 | LS3_CD02:CAV4_D4414 |  | b54 resonator | 40 |  |  | 0.49566 | 348.9235 | 427.1942 | 41.094 | 41.094 | 427.1942 | 348.6757 | 441.3876 | 441.6355 |  | 372.909
LS3 | CD02 | BLM | 4416 | LS3_CD02:BLM_D4416 |  | outside loss monitor |  |  |  | 0 | 348.6757 | 427.1942 | 41.094 | 41.094 | 427.1942 | 348.6757 | 441.6355 | 441.6355 |  | 373.1568
[10 empty] | 0 | 348.6757 | 427.1942 | 41.094 | 41.094 | 427.1942 | 348.6757 | 441.6355 | 441.6355 |  | 373.1568
[6 empty] | bellow |  |  |  | 0.15525 | 348.5981 | 427.1942 | 41.094 | 41.094 | 427.1942 | 348.5204 | 441.7131 | 441.7907 |  | 373.2345
LS3 | CD02 | SOL1 | 4420 | LS3_CD02:SOL1_D4420 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 348.2704 | 427.1942 | 41.094 | 41.094 | 427.1942 | 348.0204 | 442.0407 | 442.2907 |  | 373.5621
[6 empty] | bellow |  |  |  | 0.15525 | 347.9428 | 427.1942 | 41.094 | 41.094 | 427.1942 | 347.8652 | 442.3683 | 442.446 |  | 373.8897
LS3 | CD02 | CAV5 | 4427 | LS3_CD02:CAV5_D4427 |  | b54 resonator | 40 |  |  | 0.49566 | 347.6174 | 427.1942 | 41.094 | 41.094 | 427.1942 | 347.3695 | 442.6938 | 442.9416 |  | 374.2152
[10 empty] | 0 | 347.3695 | 427.1942 | 41.094 | 41.094 | 427.1942 | 347.3695 | 442.9416 | 442.9416 |  | 374.463
[10 empty] | 0 | 347.3695 | 427.1942 | 41.094 | 41.094 | 427.1942 | 347.3695 | 442.9416 | 442.9416 |  | 374.463
[6 empty] | bellow |  |  |  | 0.099517 | 347.3198 | 427.1942 | 41.094 | 41.094 | 427.1942 | 347.27 | 442.9914 | 443.0411 |  | 374.5127
LS3 | CD02 | CAV6 | 4433 | LS3_CD02:CAV6_D4433 |  | b54 resonator | 40 |  |  | 0.49566 | 347.0222 | 427.1942 | 41.094 | 41.094 | 427.1942 | 346.7744 | 443.289 | 443.5368 |  | 374.8103
[10 empty] | 0 | 346.7744 | 427.1942 | 41.094 | 41.094 | 427.1942 | 346.7744 | 443.5368 | 443.5368 |  | 375.0582
[10 empty] | 0 | 346.7744 | 427.1942 | 41.094 | 41.094 | 427.1942 | 346.7744 | 443.5368 | 443.5368 |  | 375.0582
[6 empty] | bellow |  |  |  | 0.099517 | 346.7246 | 427.1942 | 41.094 | 41.094 | 427.1942 | 346.6748 | 443.5865 | 443.6363 |  | 375.1079
LS3 | CD02 | CAV7 | 4439 | LS3_CD02:CAV7_D4439 |  | b54 resonator | 40 |  |  | 0.49566 | 346.427 | 427.1942 | 41.094 | 41.094 | 427.1942 | 346.1792 | 443.8841 | 444.132 |  | 375.4055
[10 empty] | 0 | 346.1792 | 427.1942 | 41.094 | 41.094 | 427.1942 | 346.1792 | 444.132 | 444.132 |  | 375.6533
[10 empty] | 0 | 346.1792 | 427.1942 | 41.094 | 41.094 | 427.1942 | 346.1792 | 444.132 | 444.132 |  | 375.6533
[6 empty] | bellow |  |  |  | 0.099517 | 346.1294 | 427.1942 | 41.094 | 41.094 | 427.1942 | 346.0797 | 444.1817 | 444.2315 |  | 375.7031
LS3 | CD02 | CAV8 | 4445 | LS3_CD02:CAV8_D4445 |  | b54 resonator | 40 |  |  | 0.49566 | 345.8318 | 427.1942 | 41.094 | 41.094 | 427.1942 | 345.584 | 444.4793 | 444.7271 |  | 376.0007
LS3 | CD02 | BLM | 4447 | LS3_CD02:BLM_D4447 |  | outside loss monitor |  |  |  | 0 | 345.584 | 427.1942 | 41.094 | 41.094 | 427.1942 | 345.584 | 444.7271 | 444.7271 |  | 376.2485
[10 empty] | 0 | 345.584 | 427.1942 | 41.094 | 41.094 | 427.1942 | 345.584 | 444.7271 | 444.7271 |  | 376.2485
[10 empty] | 0.1677 | 345.5002 | 427.1942 | 41.094 | 41.094 | 427.1942 | 345.4163 | 444.811 | 444.8948 |  | 376.3324
LS3 | CD02 | GV | 4449 | LS3_CD02:GV_D4449 |  | gate valve | 40 |  |  | 0.057404 | 345.3876 | 427.1942 | 41.094 | 41.094 | 427.1942 | 345.3589 | 444.9235 | 444.9522 |  | 376.4449
[6 empty] | bellow | 40 |  |  | 0.1035 | 345.3072 | 427.1942 | 41.094 | 41.094 | 427.1942 | 345.2554 | 445.004 | 445.0557 |  | 376.5254
LS3 | WD02 | BPM | 4451 | LS3_WD02:BPM_D4451 | BPM2 | position monitor |  |  |  | 0 | 345.2554 | 427.1942 | 41.094 | 41.094 | 427.1942 | 345.2554 | 445.0557 | 445.0557 |  | 376.5771
[6 empty] | diagnostic box |  |  |  | 0.14528 | 345.1828 | 427.1942 | 41.094 | 41.094 | 427.1942 | 345.1101 | 445.1284 | 445.201 |  | 376.6498
LS3 | WD02 | PORT | 4452 | LS3_WD02:PORT_D4452 |  | port for future diagnostics |  |  |  | 0 | 345.1101 | 427.1942 | 41.094 | 41.094 | 427.1942 | 345.1101 | 445.201 | 445.201 |  | 376.7224
[6 empty] | diagnostic box |  |  |  | 0 | 345.1101 | 427.1942 | 41.094 | 41.094 | 427.1942 | 345.1101 | 445.201 | 445.201 |  | 376.7224
LS3 | WD02 | IP | 4452 | LS3_WD02:IP_D4452 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 345.1101 | 427.1942 | 41.094 | 41.094 | 427.1942 | 345.1101 | 445.201 | 445.201 |  | 376.7224
[6 empty] | bellow | 40 |  |  | 0.13122 | 345.0445 | 427.1942 | 41.094 | 41.094 | 427.1942 | 344.9789 | 445.2666 | 445.3322 |  | 376.788
LS3 | CD03 | GV | 4454 | LS3_CD03:GV_D4454 |  | gate valve | 40 |  |  | 0.057404 | 344.9502 | 427.1942 | 41.094 | 41.094 | 427.1942 | 344.9215 | 445.3609 | 445.3896 |  | 376.8823
[6 empty] | bellows |  |  |  | 0.1677 | 344.8376 | 427.1942 | 41.094 | 41.094 | 427.1942 | 344.7538 | 445.4735 | 445.5574 |  | 376.9949
LS3 | CD03 | CAV1 | 4458 | LS3_CD03:CAV1_D4458 |  | b54 resonator | 40 |  |  | 0.49566 | 344.506 | 427.1942 | 41.094 | 41.094 | 427.1942 | 344.2581 | 445.8052 | 446.053 |  | 377.3265
LS3 | CD03 | BLM | 4461 | LS3_CD03:BLM_D4461 |  | outside loss monitor |  |  |  | 0 | 344.2581 | 427.1942 | 41.094 | 41.094 | 427.1942 | 344.2581 | 446.053 | 446.053 |  | 377.5744
[10 empty] | 0 | 344.2581 | 427.1942 | 41.094 | 41.094 | 427.1942 | 344.2581 | 446.053 | 446.053 |  | 377.5744
[6 empty] | bellow |  |  |  | 0.099517 | 344.2084 | 427.1942 | 41.094 | 41.094 | 427.1942 | 344.1586 | 446.1028 | 446.1525 |  | 377.6241
LS3 | CD03 | CAV2 | 4464 | LS3_CD03:CAV2_D4464 |  | b54 resonator | 40 |  |  | 0.49566 | 343.9108 | 427.1942 | 41.094 | 41.094 | 427.1942 | 343.663 | 446.4004 | 446.6482 |  | 377.9217
[10 empty] | 0 | 343.663 | 427.1942 | 41.094 | 41.094 | 427.1942 | 343.663 | 446.6482 | 446.6482 |  | 378.1696
[10 empty] | 0 | 343.663 | 427.1942 | 41.094 | 41.094 | 427.1942 | 343.663 | 446.6482 | 446.6482 |  | 378.1696
[6 empty] | bellow |  |  |  | 0.099517 | 343.6132 | 427.1942 | 41.094 | 41.094 | 427.1942 | 343.5635 | 446.6979 | 446.7477 |  | 378.2193
LS3 | CD03 | CAV3 | 4470 | LS3_CD03:CAV3_D4470 |  | b54 resonator | 40 |  |  | 0.49566 | 343.3156 | 427.1942 | 41.094 | 41.094 | 427.1942 | 343.0678 | 446.9955 | 447.2434 |  | 378.5169
[10 empty] | 0 | 343.0678 | 427.1942 | 41.094 | 41.094 | 427.1942 | 343.0678 | 447.2434 | 447.2434 |  | 378.7647
[10 empty] | 0 | 343.0678 | 427.1942 | 41.094 | 41.094 | 427.1942 | 343.0678 | 447.2434 | 447.2434 |  | 378.7647
[6 empty] | bellow |  |  |  | 0.099517 | 343.018 | 427.1942 | 41.094 | 41.094 | 427.1942 | 342.9683 | 447.2931 | 447.3429 |  | 378.8145
LS3 | CD03 | CAV4 | 4476 | LS3_CD03:CAV4_D4476 |  | b54 resonator | 40 |  |  | 0.49566 | 342.7204 | 427.1942 | 41.094 | 41.094 | 427.1942 | 342.4726 | 447.5907 | 447.8385 |  | 379.1121
LS3 | CD03 | BLM | 4478 | LS3_CD03:BLM_D4478 |  | outside loss monitor |  |  |  | 0 | 342.4726 | 427.1942 | 41.094 | 41.094 | 427.1942 | 342.4726 | 447.8385 | 447.8385 |  | 379.3599
[10 empty] | 0 | 342.4726 | 427.1942 | 41.094 | 41.094 | 427.1942 | 342.4726 | 447.8385 | 447.8385 |  | 379.3599
[6 empty] | bellow |  |  |  | 0.15525 | 342.395 | 427.1942 | 41.094 | 41.094 | 427.1942 | 342.3174 | 447.9162 | 447.9938 |  | 379.4375
LS3 | CD03 | SOL1 | 4482 | LS3_CD03:SOL1_D4482 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 342.0674 | 427.1942 | 41.094 | 41.094 | 427.1942 | 341.8174 | 448.2438 | 448.4938 |  | 379.7651
[6 empty] | bellow |  |  |  | 0.15525 | 341.7397 | 427.1942 | 41.094 | 41.094 | 427.1942 | 341.6621 | 448.5714 | 448.649 |  | 380.0928
LS3 | CD03 | CAV5 | 4489 | LS3_CD03:CAV5_D4489 |  | b54 resonator | 40 |  |  | 0.49566 | 341.4143 | 427.1942 | 41.094 | 41.094 | 427.1942 | 341.1665 | 448.8969 | 449.1447 |  | 380.4182
[10 empty] | 0 | 341.1665 | 427.1942 | 41.094 | 41.094 | 427.1942 | 341.1665 | 449.1447 | 449.1447 |  | 380.6661
[10 empty] | 0 | 341.1665 | 427.1942 | 41.094 | 41.094 | 427.1942 | 341.1665 | 449.1447 | 449.1447 |  | 380.6661
[6 empty] | bellow |  |  |  | 0.099517 | 341.1167 | 427.1942 | 41.094 | 41.094 | 427.1942 | 341.0669 | 449.1944 | 449.2442 |  | 380.7158
LS3 | CD03 | CAV6 | 4495 | LS3_CD03:CAV6_D4495 |  | b54 resonator | 40 |  |  | 0.49566 | 340.8191 | 427.1942 | 41.094 | 41.094 | 427.1942 | 340.5713 | 449.492 | 449.7399 |  | 381.0134
[10 empty] | 0 | 340.5713 | 427.1942 | 41.094 | 41.094 | 427.1942 | 340.5713 | 449.7399 | 449.7399 |  | 381.2612
[10 empty] | 0 | 340.5713 | 427.1942 | 41.094 | 41.094 | 427.1942 | 340.5713 | 449.7399 | 449.7399 |  | 381.2612
[6 empty] | bellow |  |  |  | 0.099517 | 340.5215 | 427.1942 | 41.094 | 41.094 | 427.1942 | 340.4718 | 449.7896 | 449.8394 |  | 381.311
LS3 | CD03 | CAV7 | 4501 | LS3_CD03:CAV7_D4501 |  | b54 resonator | 40 |  |  | 0.49566 | 340.2239 | 427.1942 | 41.094 | 41.094 | 427.1942 | 339.9761 | 450.0872 | 450.335 |  | 381.6086
[10 empty] | 0 | 339.9761 | 427.1942 | 41.094 | 41.094 | 427.1942 | 339.9761 | 450.335 | 450.335 |  | 381.8564
[10 empty] | 0 | 339.9761 | 427.1942 | 41.094 | 41.094 | 427.1942 | 339.9761 | 450.335 | 450.335 |  | 381.8564
[6 empty] | bellow |  |  |  | 0.099517 | 339.9264 | 427.1942 | 41.094 | 41.094 | 427.1942 | 339.8766 | 450.3848 | 450.4345 |  | 381.9062
LS3 | CD03 | CAV8 | 4507 | LS3_CD03:CAV8_D4507 |  | b54 resonator | 40 |  |  | 0.49566 | 339.6288 | 427.1942 | 41.094 | 41.094 | 427.1942 | 339.3809 | 450.6824 | 450.9302 |  | 382.2037
LS3 | CD03 | BLM | 4509 | LS3_CD03:BLM_D4509 |  | outside loss monitor |  |  |  | 0 | 339.3809 | 427.1942 | 41.094 | 41.094 | 427.1942 | 339.3809 | 450.9302 | 450.9302 |  | 382.4516
[10 empty] | 0 | 339.3809 | 427.1942 | 41.094 | 41.094 | 427.1942 | 339.3809 | 450.9302 | 450.9302 |  | 382.4516
[10 empty] | 0.1677 | 339.2971 | 427.1942 | 41.094 | 41.094 | 427.1942 | 339.2132 | 451.0141 | 451.0979 |  | 382.5354
LS3 | CD03 | GV | 4511 | LS3_CD03:GV_D4511 |  | gate valve | 40 |  |  | 0.057404 | 339.1845 | 427.1942 | 41.094 | 41.094 | 427.1942 | 339.1558 | 451.1266 | 451.1553 |  | 382.648
[6 empty] | bellow | 40 |  |  | 0.1035 | 339.1041 | 427.1942 | 41.094 | 41.094 | 427.1942 | 339.0523 | 451.2071 | 451.2588 |  | 382.7284
LS3 | WD03 | BPM | 4513 | LS3_WD03:BPM_D4513 | BPM2 | position monitor |  |  |  | 0 | 339.0523 | 427.1942 | 41.094 | 41.094 | 427.1942 | 339.0523 | 451.2588 | 451.2588 |  | 382.7802
[6 empty] | diagnostic box |  |  |  | 0.14528 | 338.9797 | 427.1942 | 41.094 | 41.094 | 427.1942 | 338.9071 | 451.3315 | 451.4041 |  | 382.8528
LS3 | WD03 | PORT | 4514 | LS3_WD03:PORT_D4514 |  | port for future diagnostics |  |  |  | 0 | 338.9071 | 427.1942 | 41.094 | 41.094 | 427.1942 | 338.9071 | 451.4041 | 451.4041 |  | 382.9255
[6 empty] | diagnostic box |  |  |  | 0 | 338.9071 | 427.1942 | 41.094 | 41.094 | 427.1942 | 338.9071 | 451.4041 | 451.4041 |  | 382.9255
LS3 | WD03 | IP | 4514 | LS3_WD03:IP_D4514 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 338.9071 | 427.1942 | 41.094 | 41.094 | 427.1942 | 338.9071 | 451.4041 | 451.4041 |  | 382.9255
[6 empty] | bellow | 40 |  |  | 0.13122 | 338.8414 | 427.1942 | 41.094 | 41.094 | 427.1942 | 338.7758 | 451.4697 | 451.5353 |  | 382.9911
LS3 | CD04 | GV | 4516 | LS3_CD04:GV_D4516 |  | gate valve | 40 |  |  | 0.057404 | 338.7471 | 427.1942 | 41.094 | 41.094 | 427.1942 | 338.7184 | 451.564 | 451.5927 |  | 383.0854
[6 empty] | bellows |  |  |  | 0.1677 | 338.6346 | 427.1942 | 41.094 | 41.094 | 427.1942 | 338.5507 | 451.6766 | 451.7604 |  | 383.1979
LS3 | CD04 | CAV1 | 4520 | LS3_CD04:CAV1_D4520 |  | b54 resonator | 40 |  |  | 0.49566 | 338.3029 | 427.1942 | 41.094 | 41.094 | 427.1942 | 338.0551 | 452.0082 | 452.2561 |  | 383.5296
LS3 | CD04 | BLM | 4523 | LS3_CD04:BLM_D4523 |  | outside loss monitor |  |  |  | 0 | 338.0551 | 427.1942 | 41.094 | 41.094 | 427.1942 | 338.0551 | 452.2561 | 452.2561 |  | 383.7774
[10 empty] | 0 | 338.0551 | 427.1942 | 41.094 | 41.094 | 427.1942 | 338.0551 | 452.2561 | 452.2561 |  | 383.7774
[6 empty] | bellow |  |  |  | 0.099517 | 338.0053 | 427.1942 | 41.094 | 41.094 | 427.1942 | 337.9556 | 452.3058 | 452.3556 |  | 383.8272
LS3 | CD04 | CAV2 | 4526 | LS3_CD04:CAV2_D4526 |  | b54 resonator | 40 |  |  | 0.49566 | 337.7077 | 427.1942 | 41.094 | 41.094 | 427.1942 | 337.4599 | 452.6034 | 452.8512 |  | 384.1248
[10 empty] | 0 | 337.4599 | 427.1942 | 41.094 | 41.094 | 427.1942 | 337.4599 | 452.8512 | 452.8512 |  | 384.3726
[10 empty] | 0 | 337.4599 | 427.1942 | 41.094 | 41.094 | 427.1942 | 337.4599 | 452.8512 | 452.8512 |  | 384.3726
[6 empty] | bellow |  |  |  | 0.099517 | 337.4101 | 427.1942 | 41.094 | 41.094 | 427.1942 | 337.3604 | 452.901 | 452.9508 |  | 384.4224
LS3 | CD04 | CAV3 | 4532 | LS3_CD04:CAV3_D4532 |  | b54 resonator | 40 |  |  | 0.49566 | 337.1126 | 427.1942 | 41.094 | 41.094 | 427.1942 | 336.8647 | 453.1986 | 453.4464 |  | 384.72
[10 empty] | 0 | 336.8647 | 427.1942 | 41.094 | 41.094 | 427.1942 | 336.8647 | 453.4464 | 453.4464 |  | 384.9678
[10 empty] | 0 | 336.8647 | 427.1942 | 41.094 | 41.094 | 427.1942 | 336.8647 | 453.4464 | 453.4464 |  | 384.9678
[6 empty] | bellow |  |  |  | 0.099517 | 336.815 | 427.1942 | 41.094 | 41.094 | 427.1942 | 336.7652 | 453.4962 | 453.5459 |  | 385.0175
LS3 | CD04 | CAV4 | 4538 | LS3_CD04:CAV4_D4538 |  | b54 resonator | 40 |  |  | 0.49566 | 336.5174 | 427.1942 | 41.094 | 41.094 | 427.1942 | 336.2696 | 453.7938 | 454.0416 |  | 385.3151
LS3 | CD04 | BLM | 4540 | LS3_CD04:BLM_D4540 |  | outside loss monitor |  |  |  | 0 | 336.2696 | 427.1942 | 41.094 | 41.094 | 427.1942 | 336.2696 | 454.0416 | 454.0416 |  | 385.563
[10 empty] | 0 | 336.2696 | 427.1942 | 41.094 | 41.094 | 427.1942 | 336.2696 | 454.0416 | 454.0416 |  | 385.563
[6 empty] | bellow |  |  |  | 0.15525 | 336.1919 | 427.1942 | 41.094 | 41.094 | 427.1942 | 336.1143 | 454.1192 | 454.1968 |  | 385.6406
LS3 | CD04 | SOL1 | 4544 | LS3_CD04:SOL1_D4544 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 335.8643 | 427.1942 | 41.094 | 41.094 | 427.1942 | 335.6143 | 454.4468 | 454.6968 |  | 385.9682
[6 empty] | bellow |  |  |  | 0.15525 | 335.5367 | 427.1942 | 41.094 | 41.094 | 427.1942 | 335.4591 | 454.7745 | 454.8521 |  | 386.2958
LS3 | CD04 | CAV5 | 4551 | LS3_CD04:CAV5_D4551 |  | b54 resonator | 40 |  |  | 0.49566 | 335.2112 | 427.1942 | 41.094 | 41.094 | 427.1942 | 334.9634 | 455.0999 | 455.3477 |  | 386.6213
[10 empty] | 0 | 334.9634 | 427.1942 | 41.094 | 41.094 | 427.1942 | 334.9634 | 455.3477 | 455.3477 |  | 386.8691
[10 empty] | 0 | 334.9634 | 427.1942 | 41.094 | 41.094 | 427.1942 | 334.9634 | 455.3477 | 455.3477 |  | 386.8691
[6 empty] | bellow |  |  |  | 0.099517 | 334.9136 | 427.1942 | 41.094 | 41.094 | 427.1942 | 334.8639 | 455.3975 | 455.4473 |  | 386.9189
LS3 | CD04 | CAV6 | 4557 | LS3_CD04:CAV6_D4557 |  | b54 resonator | 40 |  |  | 0.49566 | 334.6161 | 427.1942 | 41.094 | 41.094 | 427.1942 | 334.3682 | 455.6951 | 455.9429 |  | 387.2165
[10 empty] | 0 | 334.3682 | 427.1942 | 41.094 | 41.094 | 427.1942 | 334.3682 | 455.9429 | 455.9429 |  | 387.4643
[10 empty] | 0 | 334.3682 | 427.1942 | 41.094 | 41.094 | 427.1942 | 334.3682 | 455.9429 | 455.9429 |  | 387.4643
[6 empty] | bellow |  |  |  | 0.099517 | 334.3185 | 427.1942 | 41.094 | 41.094 | 427.1942 | 334.2687 | 455.9927 | 456.0424 |  | 387.5141
LS3 | CD04 | CAV7 | 4563 | LS3_CD04:CAV7_D4563 |  | b54 resonator | 40 |  |  | 0.49566 | 334.0209 | 427.1942 | 41.094 | 41.094 | 427.1942 | 333.7731 | 456.2903 | 456.5381 |  | 387.8116
[10 empty] | 0 | 333.7731 | 427.1942 | 41.094 | 41.094 | 427.1942 | 333.7731 | 456.5381 | 456.5381 |  | 388.0595
[10 empty] | 0 | 333.7731 | 427.1942 | 41.094 | 41.094 | 427.1942 | 333.7731 | 456.5381 | 456.5381 |  | 388.0595
[6 empty] | bellow |  |  |  | 0.099517 | 333.7233 | 427.1942 | 41.094 | 41.094 | 427.1942 | 333.6735 | 456.5879 | 456.6376 |  | 388.1092
LS3 | CD04 | CAV8 | 4569 | LS3_CD04:CAV8_D4569 |  | b54 resonator | 40 |  |  | 0.49566 | 333.4257 | 427.1942 | 41.094 | 41.094 | 427.1942 | 333.1779 | 456.8854 | 457.1333 |  | 388.4068
LS3 | CD04 | BLM | 4571 | LS3_CD04:BLM_D4571 |  | outside loss monitor |  |  |  | 0 | 333.1779 | 427.1942 | 41.094 | 41.094 | 427.1942 | 333.1779 | 457.1333 | 457.1333 |  | 388.6546
[10 empty] | 0 | 333.1779 | 427.1942 | 41.094 | 41.094 | 427.1942 | 333.1779 | 457.1333 | 457.1333 |  | 388.6546
[10 empty] | 0.1677 | 333.094 | 427.1942 | 41.094 | 41.094 | 427.1942 | 333.0102 | 457.2171 | 457.301 |  | 388.7385
LS3 | CD04 | GV | 4573 | LS3_CD04:GV_D4573 |  | gate valve | 40 |  |  | 0.057404 | 332.9815 | 427.1942 | 41.094 | 41.094 | 427.1942 | 332.9528 | 457.3297 | 457.3584 |  | 388.851
[6 empty] | bellow | 40 |  |  | 0.1035 | 332.901 | 427.1942 | 41.094 | 41.094 | 427.1942 | 332.8493 | 457.4101 | 457.4619 |  | 388.9315
LS3 | WD04 | BPM | 4575 | LS3_WD04:BPM_D4575 | BPM2 | position monitor |  |  |  | 0 | 332.8493 | 427.1942 | 41.094 | 41.094 | 427.1942 | 332.8493 | 457.4619 | 457.4619 |  | 388.9832
[6 empty] | diagnostic box |  |  |  | 0.14528 | 332.7766 | 427.1942 | 41.094 | 41.094 | 427.1942 | 332.704 | 457.5345 | 457.6072 |  | 389.0559
LS3 | WD04 | PORT | 4576 | LS3_WD04:PORT_D4576 |  | port for future diagnostics |  |  |  | 0 | 332.704 | 427.1942 | 41.094 | 41.094 | 427.1942 | 332.704 | 457.6072 | 457.6072 |  | 389.1285
[6 empty] | diagnostic box |  |  |  | 0 | 332.704 | 427.1942 | 41.094 | 41.094 | 427.1942 | 332.704 | 457.6072 | 457.6072 |  | 389.1285
LS3 | WD04 | IP | 4576 | LS3_WD04:IP_D4576 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 332.704 | 427.1942 | 41.094 | 41.094 | 427.1942 | 332.704 | 457.6072 | 457.6072 |  | 389.1285
[6 empty] | bellow | 40 |  |  | 0.13122 | 332.6384 | 427.1942 | 41.094 | 41.094 | 427.1942 | 332.5728 | 457.6728 | 457.7384 |  | 389.1941
LS3 | CD05 | GV | 4578 | LS3_CD05:GV_D4578 |  | gate valve | 40 |  |  | 0.057404 | 332.5441 | 427.1942 | 41.094 | 41.094 | 427.1942 | 332.5154 | 457.7671 | 457.7958 |  | 389.2884
[6 empty] | bellows |  |  |  | 0.1677 | 332.4315 | 427.1942 | 41.094 | 41.094 | 427.1942 | 332.3477 | 457.8796 | 457.9635 |  | 389.401
LS3 | CD05 | CAV1 | 4582 | LS3_CD05:CAV1_D4582 |  | b54 resonator | 40 |  |  | 0.49566 | 332.0998 | 427.1942 | 41.094 | 41.094 | 427.1942 | 331.852 | 458.2113 | 458.4591 |  | 389.7327
LS3 | CD05 | BLM | 4585 | LS3_CD05:BLM_D4585 |  | outside loss monitor |  |  |  | 0 | 331.852 | 427.1942 | 41.094 | 41.094 | 427.1942 | 331.852 | 458.4591 | 458.4591 |  | 389.9805
[10 empty] | 0 | 331.852 | 427.1942 | 41.094 | 41.094 | 427.1942 | 331.852 | 458.4591 | 458.4591 |  | 389.9805
[6 empty] | bellow |  |  |  | 0.099517 | 331.8023 | 427.1942 | 41.094 | 41.094 | 427.1942 | 331.7525 | 458.5089 | 458.5587 |  | 390.0303
LS3 | CD05 | CAV2 | 4588 | LS3_CD05:CAV2_D4588 |  | b54 resonator | 40 |  |  | 0.49566 | 331.5047 | 427.1942 | 41.094 | 41.094 | 427.1942 | 331.2568 | 458.8065 | 459.0543 |  | 390.3279
[10 empty] | 0 | 331.2568 | 427.1942 | 41.094 | 41.094 | 427.1942 | 331.2568 | 459.0543 | 459.0543 |  | 390.5757
[10 empty] | 0 | 331.2568 | 427.1942 | 41.094 | 41.094 | 427.1942 | 331.2568 | 459.0543 | 459.0543 |  | 390.5757
[6 empty] | bellow |  |  |  | 0.099517 | 331.2071 | 427.1942 | 41.094 | 41.094 | 427.1942 | 331.1573 | 459.1041 | 459.1538 |  | 390.6254
LS3 | CD05 | CAV3 | 4594 | LS3_CD05:CAV3_D4594 |  | b54 resonator | 40 |  |  | 0.49566 | 330.9095 | 427.1942 | 41.094 | 41.094 | 427.1942 | 330.6617 | 459.4017 | 459.6495 |  | 390.923
[10 empty] | 0 | 330.6617 | 427.1942 | 41.094 | 41.094 | 427.1942 | 330.6617 | 459.6495 | 459.6495 |  | 391.1709
[10 empty] | 0 | 330.6617 | 427.1942 | 41.094 | 41.094 | 427.1942 | 330.6617 | 459.6495 | 459.6495 |  | 391.1709
[6 empty] | bellow |  |  |  | 0.099517 | 330.6119 | 427.1942 | 41.094 | 41.094 | 427.1942 | 330.5621 | 459.6992 | 459.749 |  | 391.2206
LS3 | CD05 | CAV4 | 4600 | LS3_CD05:CAV4_D4600 |  | b54 resonator | 40 |  |  | 0.49566 | 330.3143 | 427.1942 | 41.094 | 41.094 | 427.1942 | 330.0665 | 459.9968 | 460.2447 |  | 391.5182
LS3 | CD05 | BLM | 4602 | LS3_CD05:BLM_D4602 |  | outside loss monitor |  |  |  | 0 | 330.0665 | 427.1942 | 41.094 | 41.094 | 427.1942 | 330.0665 | 460.2447 | 460.2447 |  | 391.766
[10 empty] | 0 | 330.0665 | 427.1942 | 41.094 | 41.094 | 427.1942 | 330.0665 | 460.2447 | 460.2447 |  | 391.766
[6 empty] | bellow |  |  |  | 0.15525 | 329.9889 | 427.1942 | 41.094 | 41.094 | 427.1942 | 329.9112 | 460.3223 | 460.3999 |  | 391.8437
LS3 | CD05 | SOL1 | 4606 | LS3_CD05:SOL1_D4606 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 329.6612 | 427.1942 | 41.094 | 41.094 | 427.1942 | 329.4112 | 460.6499 | 460.8999 |  | 392.1713
[6 empty] | bellow |  |  |  | 0.15525 | 329.3336 | 427.1942 | 41.094 | 41.094 | 427.1942 | 329.256 | 460.9775 | 461.0552 |  | 392.4989
LS3 | CD05 | CAV5 | 4613 | LS3_CD05:CAV5_D4613 |  | b54 resonator | 40 |  |  | 0.49566 | 329.0082 | 427.1942 | 41.094 | 41.094 | 427.1942 | 328.7603 | 461.303 | 461.5508 |  | 392.8244
[10 empty] | 0 | 328.7603 | 427.1942 | 41.094 | 41.094 | 427.1942 | 328.7603 | 461.5508 | 461.5508 |  | 393.0722
[10 empty] | 0 | 328.7603 | 427.1942 | 41.094 | 41.094 | 427.1942 | 328.7603 | 461.5508 | 461.5508 |  | 393.0722
[6 empty] | bellow |  |  |  | 0.099517 | 328.7106 | 427.1942 | 41.094 | 41.094 | 427.1942 | 328.6608 | 461.6006 | 461.6503 |  | 393.1219
LS3 | CD05 | CAV6 | 4619 | LS3_CD05:CAV6_D4619 |  | b54 resonator | 40 |  |  | 0.49566 | 328.413 | 427.1942 | 41.094 | 41.094 | 427.1942 | 328.1652 | 461.8982 | 462.146 |  | 393.4195
[10 empty] | 0 | 328.1652 | 427.1942 | 41.094 | 41.094 | 427.1942 | 328.1652 | 462.146 | 462.146 |  | 393.6674
[10 empty] | 0 | 328.1652 | 427.1942 | 41.094 | 41.094 | 427.1942 | 328.1652 | 462.146 | 462.146 |  | 393.6674
[6 empty] | bellow |  |  |  | 0.099517 | 328.1154 | 427.1942 | 41.094 | 41.094 | 427.1942 | 328.0656 | 462.1957 | 462.2455 |  | 393.7171
LS3 | CD05 | CAV7 | 4625 | LS3_CD05:CAV7_D4625 |  | b54 resonator | 40 |  |  | 0.49566 | 327.8178 | 427.1942 | 41.094 | 41.094 | 427.1942 | 327.57 | 462.4933 | 462.7412 |  | 394.0147
[10 empty] | 0 | 327.57 | 427.1942 | 41.094 | 41.094 | 427.1942 | 327.57 | 462.7412 | 462.7412 |  | 394.2625
[10 empty] | 0 | 327.57 | 427.1942 | 41.094 | 41.094 | 427.1942 | 327.57 | 462.7412 | 462.7412 |  | 394.2625
[6 empty] | bellow |  |  |  | 0.099517 | 327.5202 | 427.1942 | 41.094 | 41.094 | 427.1942 | 327.4705 | 462.7909 | 462.8407 |  | 394.3123
LS3 | CD05 | CAV8 | 4631 | LS3_CD05:CAV8_D4631 |  | b54 resonator | 40 |  |  | 0.49566 | 327.2226 | 427.1942 | 41.094 | 41.094 | 427.1942 | 326.9748 | 463.0885 | 463.3363 |  | 394.6099
LS3 | CD05 | BLM | 4633 | LS3_CD05:BLM_D4633 |  | outside loss monitor |  |  |  | 0 | 326.9748 | 427.1942 | 41.094 | 41.094 | 427.1942 | 326.9748 | 463.3363 | 463.3363 |  | 394.8577
[10 empty] | 0 | 326.9748 | 427.1942 | 41.094 | 41.094 | 427.1942 | 326.9748 | 463.3363 | 463.3363 |  | 394.8577
[10 empty] | 0.1677 | 326.891 | 427.1942 | 41.094 | 41.094 | 427.1942 | 326.8071 | 463.4202 | 463.504 |  | 394.9416
LS3 | CD05 | GV | 4635 | LS3_CD05:GV_D4635 |  | gate valve | 40 |  |  | 0.057404 | 326.7784 | 427.1942 | 41.094 | 41.094 | 427.1942 | 326.7497 | 463.5327 | 463.5614 |  | 395.0541
[6 empty] | bellow | 40 |  |  | 0.1035 | 326.698 | 427.1942 | 41.094 | 41.094 | 427.1942 | 326.6462 | 463.6132 | 463.6649 |  | 395.1346
LS3 | WD05 | BPM | 4637 | LS3_WD05:BPM_D4637 | BPM2 | position monitor |  |  |  | 0 | 326.6462 | 427.1942 | 41.094 | 41.094 | 427.1942 | 326.6462 | 463.6649 | 463.6649 |  | 395.1863
[6 empty] | diagnostic box |  |  |  | 0.14528 | 326.5736 | 427.1942 | 41.094 | 41.094 | 427.1942 | 326.5009 | 463.7376 | 463.8102 |  | 395.259
LS3 | WD05 | PORT | 4638 | LS3_WD05:PORT_D4638 |  | port for future diagnostics |  |  |  | 0 | 326.5009 | 427.1942 | 41.094 | 41.094 | 427.1942 | 326.5009 | 463.8102 | 463.8102 |  | 395.3316
[6 empty] | diagnostic box |  |  |  | 0 | 326.5009 | 427.1942 | 41.094 | 41.094 | 427.1942 | 326.5009 | 463.8102 | 463.8102 |  | 395.3316
LS3 | WD05 | IP | 4638 | LS3_WD05:IP_D4638 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 326.5009 | 427.1942 | 41.094 | 41.094 | 427.1942 | 326.5009 | 463.8102 | 463.8102 |  | 395.3316
[6 empty] | bellow | 40 |  |  | 0.13122 | 326.4353 | 427.1942 | 41.094 | 41.094 | 427.1942 | 326.3697 | 463.8758 | 463.9414 |  | 395.3972
LS3 | CD06 | GV | 4640 | LS3_CD06:GV_D4640 |  | gate valve | 40 |  |  | 0.057404 | 326.341 | 427.1942 | 41.094 | 41.094 | 427.1942 | 326.3123 | 463.9701 | 463.9988 |  | 395.4915
[6 empty] | bellows |  |  |  | 0.1677 | 326.2285 | 427.1942 | 41.094 | 41.094 | 427.1942 | 326.1446 | 464.0827 | 464.1665 |  | 395.6041
LS3 | CD06 | CAV1 | 4644 | LS3_CD06:CAV1_D4644 |  | b54 resonator | 40 |  |  | 0.49566 | 325.8968 | 427.1942 | 41.094 | 41.094 | 427.1942 | 325.6489 | 464.4144 | 464.6622 |  | 395.9357
LS3 | CD06 | BLM | 4647 | LS3_CD06:BLM_D4647 |  | outside loss monitor |  |  |  | 0 | 325.6489 | 427.1942 | 41.094 | 41.094 | 427.1942 | 325.6489 | 464.6622 | 464.6622 |  | 396.1836
[10 empty] | 0 | 325.6489 | 427.1942 | 41.094 | 41.094 | 427.1942 | 325.6489 | 464.6622 | 464.6622 |  | 396.1836
[6 empty] | bellow |  |  |  | 0.099517 | 325.5992 | 427.1942 | 41.094 | 41.094 | 427.1942 | 325.5494 | 464.712 | 464.7617 |  | 396.2333
LS3 | CD06 | CAV2 | 4650 | LS3_CD06:CAV2_D4650 |  | b54 resonator | 40 |  |  | 0.49566 | 325.3016 | 427.1942 | 41.094 | 41.094 | 427.1942 | 325.0538 | 465.0096 | 465.2574 |  | 396.5309
[10 empty] | 0 | 325.0538 | 427.1942 | 41.094 | 41.094 | 427.1942 | 325.0538 | 465.2574 | 465.2574 |  | 396.7787
[10 empty] | 0 | 325.0538 | 427.1942 | 41.094 | 41.094 | 427.1942 | 325.0538 | 465.2574 | 465.2574 |  | 396.7787
[6 empty] | bellow |  |  |  | 0.099517 | 325.004 | 427.1942 | 41.094 | 41.094 | 427.1942 | 324.9543 | 465.3071 | 465.3569 |  | 396.8285
LS3 | CD06 | CAV3 | 4656 | LS3_CD06:CAV3_D4656 |  | b54 resonator | 40 |  |  | 0.49566 | 324.7064 | 427.1942 | 41.094 | 41.094 | 427.1942 | 324.4586 | 465.6047 | 465.8526 |  | 397.1261
[10 empty] | 0 | 324.4586 | 427.1942 | 41.094 | 41.094 | 427.1942 | 324.4586 | 465.8526 | 465.8526 |  | 397.3739
[10 empty] | 0 | 324.4586 | 427.1942 | 41.094 | 41.094 | 427.1942 | 324.4586 | 465.8526 | 465.8526 |  | 397.3739
[6 empty] | bellow |  |  |  | 0.099517 | 324.4088 | 427.1942 | 41.094 | 41.094 | 427.1942 | 324.3591 | 465.9023 | 465.9521 |  | 397.4237
LS3 | CD06 | CAV4 | 4662 | LS3_CD06:CAV4_D4662 |  | b54 resonator | 40 |  |  | 0.49566 | 324.1113 | 427.1942 | 41.094 | 41.094 | 427.1942 | 323.8634 | 466.1999 | 466.4477 |  | 397.7213
LS3 | CD06 | BLM | 4664 | LS3_CD06:BLM_D4664 |  | outside loss monitor |  |  |  | 0 | 323.8634 | 427.1942 | 41.094 | 41.094 | 427.1942 | 323.8634 | 466.4477 | 466.4477 |  | 397.9691
[10 empty] | 0 | 323.8634 | 427.1942 | 41.094 | 41.094 | 427.1942 | 323.8634 | 466.4477 | 466.4477 |  | 397.9691
[6 empty] | bellow |  |  |  | 0.15525 | 323.7858 | 427.1942 | 41.094 | 41.094 | 427.1942 | 323.7082 | 466.5253 | 466.603 |  | 398.0467
LS3 | CD06 | SOL1 | 4669 | LS3_CD06:SOL1_D4669 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 323.4582 | 427.1942 | 41.094 | 41.094 | 427.1942 | 323.2082 | 466.853 | 467.103 |  | 398.3743
[6 empty] | bellow |  |  |  | 0.15525 | 323.1305 | 427.1942 | 41.094 | 41.094 | 427.1942 | 323.0529 | 467.1806 | 467.2582 |  | 398.702
LS3 | CD06 | CAV5 | 4675 | LS3_CD06:CAV5_D4675 |  | b54 resonator | 40 |  |  | 0.49566 | 322.8051 | 427.1942 | 41.094 | 41.094 | 427.1942 | 322.5573 | 467.5061 | 467.7539 |  | 399.0274
[10 empty] | 0 | 322.5573 | 427.1942 | 41.094 | 41.094 | 427.1942 | 322.5573 | 467.7539 | 467.7539 |  | 399.2753
[10 empty] | 0 | 322.5573 | 427.1942 | 41.094 | 41.094 | 427.1942 | 322.5573 | 467.7539 | 467.7539 |  | 399.2753
[6 empty] | bellow |  |  |  | 0.099517 | 322.5075 | 427.1942 | 41.094 | 41.094 | 427.1942 | 322.4578 | 467.8036 | 467.8534 |  | 399.325
LS3 | CD06 | CAV6 | 4681 | LS3_CD06:CAV6_D4681 |  | b54 resonator | 40 |  |  | 0.49566 | 322.2099 | 427.1942 | 41.094 | 41.094 | 427.1942 | 321.9621 | 468.1012 | 468.3491 |  | 399.6226
[10 empty] | 0 | 321.9621 | 427.1942 | 41.094 | 41.094 | 427.1942 | 321.9621 | 468.3491 | 468.3491 |  | 399.8704
[10 empty] | 0 | 321.9621 | 427.1942 | 41.094 | 41.094 | 427.1942 | 321.9621 | 468.3491 | 468.3491 |  | 399.8704
[6 empty] | bellow |  |  |  | 0.099517 | 321.9123 | 427.1942 | 41.094 | 41.094 | 427.1942 | 321.8626 | 468.3988 | 468.4486 |  | 399.9202
LS3 | CD06 | CAV7 | 4687 | LS3_CD06:CAV7_D4687 |  | b54 resonator | 40 |  |  | 0.49566 | 321.6147 | 427.1942 | 41.094 | 41.094 | 427.1942 | 321.3669 | 468.6964 | 468.9442 |  | 400.2178
[10 empty] | 0 | 321.3669 | 427.1942 | 41.094 | 41.094 | 427.1942 | 321.3669 | 468.9442 | 468.9442 |  | 400.4656
[10 empty] | 0 | 321.3669 | 427.1942 | 41.094 | 41.094 | 427.1942 | 321.3669 | 468.9442 | 468.9442 |  | 400.4656
[6 empty] | bellow |  |  |  | 0.099517 | 321.3172 | 427.1942 | 41.094 | 41.094 | 427.1942 | 321.2674 | 468.994 | 469.0437 |  | 400.5154
LS3 | CD06 | CAV8 | 4693 | LS3_CD06:CAV8_D4693 |  | b54 resonator | 40 |  |  | 0.49566 | 321.0196 | 427.1942 | 41.094 | 41.094 | 427.1942 | 320.7717 | 469.2916 | 469.5394 |  | 400.8129
LS3 | CD06 | BLM | 4695 | LS3_CD06:BLM_D4695 |  | outside loss monitor |  |  |  | 0 | 320.7717 | 427.1942 | 41.094 | 41.094 | 427.1942 | 320.7717 | 469.5394 | 469.5394 |  | 401.0608
[10 empty] | 0 | 320.7717 | 427.1942 | 41.094 | 41.094 | 427.1942 | 320.7717 | 469.5394 | 469.5394 |  | 401.0608
[10 empty] | 0.1677 | 320.6879 | 427.1942 | 41.094 | 41.094 | 427.1942 | 320.604 | 469.6233 | 469.7071 |  | 401.1446
LS3 | CD06 | GV | 4697 | LS3_CD06:GV_D4697 |  | gate valve | 40 |  |  | 0.057404 | 320.5753 | 427.1942 | 41.094 | 41.094 | 427.1942 | 320.5466 | 469.7358 | 469.7645 |  | 401.2572
[6 empty] | bellow | 40 |  |  | 0.1035 | 320.4949 | 427.1942 | 41.094 | 41.094 | 427.1942 | 320.4431 | 469.8163 | 469.868 |  | 401.3376
LS3 | WD06 | BPM | 4699 | LS3_WD06:BPM_D4699 | BPM2 | position monitor |  |  |  | 0 | 320.4431 | 427.1942 | 41.094 | 41.094 | 427.1942 | 320.4431 | 469.868 | 469.868 |  | 401.3894
[6 empty] | diagnostic box |  |  |  | 0.14528 | 320.3705 | 427.1942 | 41.094 | 41.094 | 427.1942 | 320.2979 | 469.9407 | 470.0133 |  | 401.462
LS3 | WD06 | PORT | 4700 | LS3_WD06:PORT_D4700 |  | port for future diagnostics |  |  |  | 0 | 320.2979 | 427.1942 | 41.094 | 41.094 | 427.1942 | 320.2979 | 470.0133 | 470.0133 |  | 401.5347
[6 empty] | diagnostic box |  |  |  | 0 | 320.2979 | 427.1942 | 41.094 | 41.094 | 427.1942 | 320.2979 | 470.0133 | 470.0133 |  | 401.5347
LS3 | WD06 | IP | 4700 | LS3_WD06:IP_D4700 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 320.2979 | 427.1942 | 41.094 | 41.094 | 427.1942 | 320.2979 | 470.0133 | 470.0133 |  | 401.5347
[6 empty] | bellow | 40 |  |  | 0.13122 | 320.2322 | 427.1942 | 41.094 | 41.094 | 427.1942 | 320.1666 | 470.0789 | 470.1445 |  | 401.6003
LINAC SEGMENT 3 CM END | [9 empty] | 0 | 320.1666 | 427.1942 | 41.094 | 41.094 | 427.1942 | 320.1666 | 470.1445 | 470.1445 |  | 401.6659
[8 empty] | -0.00046882 |  | 0 | 320.1666 | 427.1942 | 41.094 | 41.094 | 427.1942 | 320.1666 | 470.1445 | 470.1445 |  | 401.6659
LS3 | BTS | FAV | 4702 | LS3_BTS:FAV_D4702 | FV | fast valve | 40 | 320.1326 | 320.133 | 0.067196 | 320.133 | 427.1942 | 41.094 | 41.094 | 427.1942 | 320.0994 | 470.1781 | 470.2117 |  | 401.6995
[10 empty] | 0.5558 | 319.8215 | 427.1942 | 41.094 | 41.094 | 427.1942 | 319.5436 | 470.4896 | 470.7675 |  | 402.011
LS3 | BTS | DCHV | 4709 | LS3_BTS:DCHV_D4709 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 319.4436 | 427.1942 | 41.094 | 41.094 | 427.1942 | 319.3436 | 470.8675 | 470.9675 |  | 402.3889
[10 empty] | 0.2 | 319.2436 | 427.1942 | 41.094 | 41.094 | 427.1942 | 319.1436 | 471.0675 | 471.1675 |  | 402.5889
LS3 | BTS | QV | 4713 | LS3_BTS:QV_D4713 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 319.0186 | 427.1942 | 41.094 | 41.094 | 427.1942 | 318.8936 | 471.2925 | 471.4175 |  | 402.8139
[10 empty] | 0.25 | 318.7686 | 427.1942 | 41.094 | 41.094 | 427.1942 | 318.6436 | 471.5425 | 471.6675 |  | 403.0639
LS3 | BTS | QH | 4718 | LS3_BTS:QH_D4718 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 318.5186 | 427.1942 | 41.094 | 41.094 | 427.1942 | 318.3936 | 471.7925 | 471.9175 |  | 403.3139
[10 empty] | 0.66963 | 318.0588 | 427.1942 | 41.094 | 41.094 | 427.1942 | 317.724 | 472.2523 | 472.5871 |  | 403.7737
LS3 | BTS | GV | 4726 | LS3_BTS:GV_D4726 | GV_M | gate valve | 63 | 2.176e-06 | 317.6865 | 0.075 | 317.6865 | 427.1942 | 41.094 | 41.094 | 427.1942 | 317.649 | 472.6246 | 472.6621 |  | 404.146
[10 empty] | 2.2843 | 316.5068 | 427.1942 | 41.094 | 41.094 | 427.1942 | 315.3647 | 473.8043 | 474.9465 | 20170217 space adjusted | 405.3257
LS3 | BTS | DCHV | 4750 | LS3_BTS:DCHV_D4750 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 315.2647 | 427.1942 | 41.094 | 41.094 | 427.1942 | 315.1647 | 475.0465 | 475.1465 |  | 406.5678
[10 empty] | 0.13321 | 315.0981 | 427.1942 | 41.094 | 41.094 | 427.1942 | 315.0315 | 475.2131 | 475.2797 | 20170217 space adjusted | 406.7345
LS3 | BTS | BPM | 4753 | LS3_BTS:BPM_D4753 | BPM3 | position monitor | 47.5 |  |  | 0.02 | 315.0215 | 427.1942 | 41.094 | 41.094 | 427.1942 | 315.0115 | 475.2897 | 475.2997 | 20170217 changed from .137 to .02 m (button diameter) | 406.8111
[10 empty] | 0.04782 | 314.9876 | 427.1942 | 41.094 | 41.094 | 427.1942 | 314.9636 | 475.3236 | 475.3475 | 20170217 space adjusted | 406.845
LS3 | BTS | QV | 4755 | LS3_BTS:QV_D4755 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 314.8386 | 427.1942 | 41.094 | 41.094 | 427.1942 | 314.7136 | 475.4725 | 475.5975 |  | 406.9939
[10 empty] | 0.25 | 314.5886 | 427.1942 | 41.094 | 41.094 | 427.1942 | 314.4636 | 475.7225 | 475.8475 |  | 407.2439
LS3 | BTS | QH | 4760 | LS3_BTS:QH_D4760 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 314.3386 | 427.1942 | 41.094 | 41.094 | 427.1942 | 314.2136 | 475.9725 | 476.0975 |  | 407.4939
[10 empty] | 0.72812 | 313.8496 | 427.1942 | 41.094 | 41.094 | 427.1942 | 313.4855 | 476.4616 | 476.8256 | 20170217 space adjusted | 407.9829
[6 empty] | bellow |  | 3.176e-06 | -0.067115 | 0.1035 | 313.4338 | 427.1942 | 41.094 | 41.094 | 427.1942 | 313.382 | 476.8774 | 476.9291 |  | 408.3987
LS3 | BTS | BPM | 4769 | LS3_BTS:BPM_D4769 | BPM3 | position monitor |  | 313.382 | 313.382 | 0 | 313.382 | 427.1942 | 41.094 | 41.094 | 427.1942 | 313.382 | 476.9291 | 476.9291 |  | 408.4505
[6 empty] | diagnostic box |  |  |  | 0.14528 | 313.3094 | 427.1942 | 41.094 | 41.094 | 427.1942 | 313.2367 | 477.0018 | 477.0744 |  | 408.5231
LS3 | BTS | PM | 4771 | LS3_BTS:PM_D4771 | PM3 | profile monitor |  | 4.176e-06 |  | 0 | 313.2367 | 427.1942 | 41.094 | 41.094 | 427.1942 | 313.2367 | 477.0744 | 477.0744 |  | 408.5958
[6 empty] | diagnostic box |  | 313.2367 | 313.2367 | 0 | 313.2367 | 427.1942 | 41.094 | 41.094 | 427.1942 | 313.2367 | 477.0744 | 477.0744 |  | 408.5958
LS3 | BTS | IP | 4771 | LS3_BTS:IP_D4771 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 313.2367 | 427.1942 | 41.094 | 41.094 | 427.1942 | 313.2367 | 477.0744 | 477.0744 |  | 408.5958
LS3 | BTS | FAVS | 4771 | LS3_BTS:FAVS_D4771 | FVS | fast valve sensor |  |  |  | 0 | 313.2367 | 427.1942 | 41.094 | 41.094 | 427.1942 | 313.2367 | 477.0744 | 477.0744 |  | 408.5958
[6 empty] | bellow |  |  |  | 0.13122 | 313.1711 | 427.1942 | 41.094 | 41.094 | 427.1942 | 313.1055 | 477.14 | 477.2056 |  | 408.6614
[10 empty] | 0.59388 | 312.8086 | 427.1942 | 41.094 | 41.094 | 427.1942 | 312.5116 | 477.5026 | 477.7995 | 20170217 space adjusted | 409.0239
LS3 | BTS | DCHV | 4779 | LS3_BTS:DCHV_D4779 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 312.4116 | 427.1942 | 41.094 | 41.094 | 427.1942 | 312.3116 | 477.8995 | 477.9995 |  | 409.4209
[10 empty] | 0.2 | 312.2116 | 427.1942 | 41.094 | 41.094 | 427.1942 | 312.1116 | 478.0995 | 478.1995 |  | 409.6209
LS3 | BTS | QH | 4783 | LS3_BTS:QH_D4783 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 311.9866 | 427.1942 | 41.094 | 41.094 | 427.1942 | 311.8616 | 478.3245 | 478.4495 |  | 409.8459
[8 empty] | 3.176e-06 | -0.20458 | 1.2046 | 311.2594 | 427.1942 | 41.094 | 41.094 | 427.1942 | 310.6571 | 479.0518 | 479.6541 | 20170217 space adjusted | 410.5732
LS3 | BTS | PM | 4797 | LS3_BTS:PM_D4797 | PM3 | profile monitor |  | 310.6571 | 310.6571 | 0 | 310.6571 | 427.1942 | 41.094 | 41.094 | 427.1942 | 310.6571 | 479.6541 | 479.6541 |  | 411.1755
LS3 | BTS | IP | 4797 | LS3_BTS:IP_D4797 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 310.6571 | 427.1942 | 41.094 | 41.094 | 427.1942 | 310.6571 | 479.6541 | 479.6541 |  | 411.1755
[10 empty] | 0.79543 | 310.2594 | 427.1942 | 41.094 | 41.094 | 427.1942 | 309.8616 | 480.0518 | 480.4495 | 20170217 space adjusted | 411.5732
LS3 | BTS | QV | 4806 | LS3_BTS:QV_D4806 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 309.7366 | 427.1942 | 41.094 | 41.094 | 427.1942 | 309.6116 | 480.5745 | 480.6995 |  | 412.0959
[8 empty] | 3.176e-06 | -0.35958 | 2.0116 | 308.6058 | 427.1942 | 41.094 | 41.094 | 427.1942 | 307.6001 | 481.7053 | 482.7111 |  | 413.2267
LS3 | BTS | PM | 4827 | LS3_BTS:PM_D4827 | PM3 | profile monitor |  | 307.6001 | 307.6001 | 0 | 307.6001 | 427.1942 | 41.094 | 41.094 | 427.1942 | 307.6001 | 482.7111 | 482.7111 |  | 414.2325
LS3 | BTS | IP | 4827 | LS3_BTS:IP_D4827 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 307.6001 | 427.1942 | 41.094 | 41.094 | 427.1942 | 307.6001 | 482.7111 | 482.7111 |  | 414.2325
[8 empty] | 3.176e-06 | 0.001019 | 1.2914 | 306.9544 | 427.1942 | 41.094 | 41.094 | 427.1942 | 306.3087 | 483.3568 | 484.0025 | 20170217 space adjusted | 414.8782
LS3 | BTS | DCHV | 4841 | LS3_BTS:DCHV_D4841 | C2 | correctors-X,Y | 50 | 306.2087 | 306.2087 | 0.2 | 306.2087 | 427.1942 | 41.094 | 41.094 | 427.1942 | 306.1087 | 484.1025 | 484.2025 |  | 415.6239
[8 empty] | 3.176e-06 | -0.05468 | 0.1332 | 306.0421 | 427.1942 | 41.094 | 41.094 | 427.1942 | 305.9755 | 484.2691 | 484.3357 | 20170217 space adjusted | 415.7905
LS3 | BTS | BPM | 4843 | LS3_BTS:BPM_D4843 | BPM3 | position monitor | 47.5 | 305.9655 | 305.9655 | 0.02 | 305.9655 | 427.1942 | 41.094 | 41.094 | 427.1942 | 305.9555 | 484.3457 | 484.3557 | 20170217 changed from .137 to .02 m (button diameter) | 415.8671
[8 empty] | 0.0585 |  | 0.04782 | 305.9316 | 427.1942 | 41.094 | 41.094 | 427.1942 | 305.9076 | 484.3796 | 484.4035 | 20170217 space adjusted | 415.901
LS3 | BTS | QH | 4845 | LS3_BTS:QH_D4845 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 305.7826 | 427.1942 | 41.094 | 41.094 | 427.1942 | 305.6576 | 484.5285 | 484.6535 |  | 416.0499
[10 empty] | 1.5 | 304.9076 | 427.1942 | 41.094 | 41.094 | 427.1942 | 304.1576 | 485.4035 | 486.1535 |  | 416.9249
LS3 | BTS | PM | 4862 | LS3_BTS:PM_D4862 | PM3 | profile monitor |  |  |  | 0 | 304.1576 | 427.1942 | 41.094 | 41.094 | 427.1942 | 304.1576 | 486.1535 | 486.1535 |  | 417.6749
LS3 | BTS | IP | 4862 | LS3_BTS:IP_D4862 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 304.1576 | 427.1942 | 41.094 | 41.094 | 427.1942 | 304.1576 | 486.1535 | 486.1535 |  | 417.6749
[10 empty] | 0.5 | 303.9076 | 427.1942 | 41.094 | 41.094 | 427.1942 | 303.6576 | 486.4035 | 486.6535 |  | 417.9249
LS3 | BTS | QV | 4868 | LS3_BTS:QV_D4868 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 303.5326 | 427.1942 | 41.094 | 41.094 | 427.1942 | 303.4076 | 486.7785 | 486.9035 |  | 418.2999
[8 empty] | 3.176e-06 | 1.3481 | 0.12643 | 303.3444 | 427.1942 | 41.094 | 41.094 | 427.1942 | 303.2812 | 486.9667 | 487.0299 | 20170217 space adjusted | 418.4881
LS3 | BTS | GV | 4871 | LS3_BTS:GV_D4871 | GV_M | gate valve | 63 | 303.2437 | 303.2437 | 0.075 | 303.2437 | 427.1942 | 41.094 | 41.094 | 427.1942 | 303.2062 | 487.0674 | 487.1049 |  | 418.5888
[10 empty] | 1.3481 | 302.5322 | 427.1942 | 41.094 | 41.094 | 427.1942 | 301.8581 | 487.779 | 488.453 | 20170217 space adjusted | 419.3003
[6 empty] | bellow |  | 3.176e-06 | -0.075651 | 0.17915 | 301.7686 | 427.1942 | 41.094 | 41.094 | 427.1942 | 301.679 | 488.5426 | 488.6322 | 20170217 space adjusted | 420.064
LS3 | BTS | BPM | 4886 | LS3_BTS:BPM_D4886 | BPM3 | position monitor |  | 301.679 | 301.679 | 0 | 301.679 | 427.1942 | 41.094 | 41.094 | 427.1942 | 301.679 | 488.6322 | 488.6322 |  | 420.1535
[6 empty] | diagnostic box |  | 3.176e-06 |  | 0.14528 | 301.6063 | 427.1942 | 41.094 | 41.094 | 427.1942 | 301.5337 | 488.7048 | 488.7774 |  | 420.2262
LS3 | BTS | PORT | 4888 | LS3_BTS:PORT_D4888 |  | port for future diagnostics |  | 301.5337 | 301.5337 | 0 | 301.5337 | 427.1942 | 41.094 | 41.094 | 427.1942 | 301.5337 | 488.7774 | 488.7774 |  | 420.2988
[6 empty] | diagnostic box |  |  |  | 0 | 301.5337 | 427.1942 | 41.094 | 41.094 | 427.1942 | 301.5337 | 488.7774 | 488.7774 |  | 420.2988
LS3 | BTS | IP | 4888 | LS3_BTS:IP_D4888 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 301.5337 | 427.1942 | 41.094 | 41.094 | 427.1942 | 301.5337 | 488.7774 | 488.7774 |  | 420.2988
[6 empty] | bellow |  |  |  | 0.13122 | 301.4681 | 427.1942 | 41.094 | 41.094 | 427.1942 | 301.4025 | 488.8431 | 488.9087 |  | 420.3644
[10 empty] | 1.2988 | 300.7531 | 427.1942 | 41.094 | 41.094 | 427.1942 | 300.1036 | 489.5581 | 490.2075 | 20170217 space adjusted | 421.0795
LS3 | BTS | DCHV | 4903 | LS3_BTS:DCHV_D4903 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 300.0036 | 427.1942 | 41.094 | 41.094 | 427.1942 | 299.9036 | 490.3075 | 490.4075 |  | 421.8289
[10 empty] | 0.2 | 299.8036 | 427.1942 | 41.094 | 41.094 | 427.1942 | 299.7036 | 490.5075 | 490.6075 |  | 422.0289
LS3 | BTS | QH | 4907 | LS3_BTS:QH_D4907 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 299.5786 | 427.1942 | 41.094 | 41.094 | 427.1942 | 299.4536 | 490.7325 | 490.8575 |  | 422.2539
[10 empty] | 2 | 298.4536 | 427.1942 | 41.094 | 41.094 | 427.1942 | 297.4536 | 491.8575 | 492.8575 |  | 423.3789
LS3 | BTS | QV | 4930 | LS3_BTS:QV_D4930 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 297.3286 | 427.1942 | 41.094 | 41.094 | 427.1942 | 297.2036 | 492.9825 | 493.1075 |  | 424.5039
[10 empty] | 1.5495 | 296.4289 | 427.1942 | 41.094 | 41.094 | 427.1942 | 295.6541 | 493.8823 | 494.657 |  | 425.4036
[10 empty] | 0.096 | 295.6061 | 427.1942 | 41.094 | 41.094 | 427.1942 | 295.5581 | 494.705 | 494.753 |  | 426.2264
[10 empty] | 0 | 295.5581 | 427.1942 | 41.094 | 41.094 | 427.1942 | 295.5581 | 494.753 | 494.753 |  | 426.2744
[6 empty] | bellow |  |  |  | 0.094 | 295.5111 | 427.1942 | 41.094 | 41.094 | 427.1942 | 295.4641 | 494.8 | 494.847 |  | 426.3214
LS3 | BTS | PORT | 4948 | LS3_BTS:PORT_D4948 |  | port for future diagnostics |  |  |  | 0 | 295.4641 | 427.1942 | 41.094 | 41.094 | 427.1942 | 295.4641 | 494.847 | 494.847 |  | 426.3684
[6 empty] | vacuum box |  | 2.53e-06 | 0.090994 | 6e-06 | 295.4641 | 427.1942 | 41.094 | 41.094 | 427.1942 | 295.4641 | 494.847 | 494.847 | 20170217 space adjusted | 426.3684
LS3 | BTS | IP | 4948 | LS3_BTS:IP_D4948 | IP_S | ion pump, 75 l/s |  | 295.4641 | 295.4641 | 0 | 295.4641 | 427.1942 | 41.094 | 41.094 | 427.1942 | 295.4641 | 494.847 | 494.847 |  | 426.3684
[6 empty] | bellow |  |  |  | 0.18999 | 295.3691 | 427.1942 | 41.094 | 41.094 | 427.1942 | 295.2741 | 494.942 | 495.037 | 20170217 space adjusted | 426.4634
[8 empty] | 3.176e-06 | 0.001019 | 1.3735 | 294.5874 | 427.1942 | 41.094 | 41.094 | 427.1942 | 293.9007 | 495.7237 | 496.4105 | 20170217 space adjusted | 427.2451
LS3 | BTS | DCHV | 4965 | LS3_BTS:DCHV_D4965 | C2 | correctors-X,Y | 50 | 293.8007 | 293.8007 | 0.2 | 293.8007 | 427.1942 | 41.094 | 41.094 | 427.1942 | 293.7007 | 496.5105 | 496.6105 |  | 428.0319
[8 empty] | 3.176e-06 | -0.05468 | 0.1332 | 293.6341 | 427.1942 | 41.094 | 41.094 | 427.1942 | 293.5675 | 496.6771 | 496.7437 | 20170217 space adjusted | 428.1985
LS3 | BTS | BPM | 4968 | LS3_BTS:BPM_D4968 | BPM3 | position monitor | 47.5 | 293.5575 | 293.5575 | 0.02 | 293.5575 | 427.1942 | 41.094 | 41.094 | 427.1942 | 293.5475 | 496.7537 | 496.7637 | 20170217 changed from .137 to .02 m (button diameter) | 428.2751
[8 empty] | 0.0585 |  | 0.04782 | 293.5236 | 427.1942 | 41.094 | 41.094 | 427.1942 | 293.4996 | 496.7876 | 496.8115 | 20170217 space adjusted | 428.309
LS3 | BTS | QH | 4969 | LS3_BTS:QH_D4969 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 293.3746 | 427.1942 | 41.094 | 41.094 | 427.1942 | 293.2496 | 496.9365 | 497.0615 |  | 428.4579
[10 empty] | 2 | 292.2496 | 427.1942 | 41.094 | 41.094 | 427.1942 | 291.2496 | 498.0615 | 499.0615 |  | 429.5829
LS3 | BTS | QV | 4992 | LS3_BTS:QV_D4992 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 291.1246 | 427.1942 | 41.094 | 41.094 | 427.1942 | 290.9996 | 499.1865 | 499.3115 |  | 430.7079
[10 empty] | 1.6251 | 290.1871 | 427.1942 | 41.094 | 41.094 | 427.1942 | 289.3745 | 500.1241 | 500.9367 | 20170217 space adjusted | 431.6455
[6 empty] | bellow |  | 3.176e-06 | -0.075649 | 0.1035 | 289.3227 | 427.1942 | 41.094 | 41.094 | 427.1942 | 289.271 | 500.9884 | 501.0402 |  | 432.5098
LS3 | BTS | BPM | 5010 | LS3_BTS:BPM_D5010 | BPM3 | position monitor |  | 289.271 | 289.271 | 0 | 289.271 | 427.1942 | 41.094 | 41.094 | 427.1942 | 289.271 | 501.0402 | 501.0402 |  | 432.5615
[6 empty] | diagnostic box |  | 3.176e-06 |  | 0.14528 | 289.1983 | 427.1942 | 41.094 | 41.094 | 427.1942 | 289.1257 | 501.1128 | 501.1854 |  | 432.6342
LS3 | BTS | PORT | 5012 | LS3_BTS:PORT_D5012 |  | port for future diagnostics |  | 289.1257 | 289.1257 | 0 | 289.1257 | 427.1942 | 41.094 | 41.094 | 427.1942 | 289.1257 | 501.1854 | 501.1854 |  | 432.7068
[6 empty] | diagnostic box |  |  |  | 0 | 289.1257 | 427.1942 | 41.094 | 41.094 | 427.1942 | 289.1257 | 501.1854 | 501.1854 |  | 432.7068
LS3 | BTS | IP | 5012 | LS3_BTS:IP_D5012 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 289.1257 | 427.1942 | 41.094 | 41.094 | 427.1942 | 289.1257 | 501.1854 | 501.1854 |  | 432.7068
[6 empty] | bellow |  |  |  | 0.13122 | 289.0601 | 427.1942 | 41.094 | 41.094 | 427.1942 | 288.9945 | 501.251 | 501.3167 |  | 432.7724
[10 empty] | 1.2989 | 288.3451 | 427.1942 | 41.094 | 41.094 | 427.1942 | 287.6956 | 501.9661 | 502.6155 | 20170217 space adjusted | 433.4875
LS3 | BTS | DCHV | 5027 | LS3_BTS:DCHV_D5027 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 287.5956 | 427.1942 | 41.094 | 41.094 | 427.1942 | 287.4956 | 502.7155 | 502.8155 |  | 434.2369
[10 empty] | 0.2 | 287.3956 | 427.1942 | 41.094 | 41.094 | 427.1942 | 287.2956 | 502.9155 | 503.0155 |  | 434.4369
LS3 | BTS | QH | 5031 | LS3_BTS:QH_D5031 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 287.1706 | 427.1942 | 41.094 | 41.094 | 427.1942 | 287.0456 | 503.1405 | 503.2655 |  | 434.6619
[10 empty] | 2 | 286.0456 | 427.1942 | 41.094 | 41.094 | 427.1942 | 285.0456 | 504.2655 | 505.2655 |  | 435.7869
LS3 | BTS | QV | 5054 | LS3_BTS:QV_D5054 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 284.9206 | 427.1942 | 41.094 | 41.094 | 427.1942 | 284.7956 | 505.3905 | 505.5155 |  | 436.9119
[10 empty] | 1.5495 | 284.0209 | 427.1942 | 41.094 | 41.094 | 427.1942 | 283.2461 | 506.2903 | 507.065 |  | 437.8116
[10 empty] | 0.096 | 283.1981 | 427.1942 | 41.094 | 41.094 | 427.1942 | 283.1501 | 507.113 | 507.161 |  | 438.6344
[10 empty] | 0 | 283.1501 | 427.1942 | 41.094 | 41.094 | 427.1942 | 283.1501 | 507.161 | 507.161 |  | 438.6824
[6 empty] | bellow |  |  |  | 0.094 | 283.1031 | 427.1942 | 41.094 | 41.094 | 427.1942 | 283.0561 | 507.208 | 507.255 |  | 438.7294
LS3 | BTS | PORT | 5073 | LS3_BTS:PORT_D5073 |  | port for future diagnostics |  |  |  | 0 | 283.0561 | 427.1942 | 41.094 | 41.094 | 427.1942 | 283.0561 | 507.255 | 507.255 |  | 438.7764
[6 empty] | vacuum box |  | 2.53e-06 | 0.090994 | 6e-06 | 283.0561 | 427.1942 | 41.094 | 41.094 | 427.1942 | 283.0561 | 507.255 | 507.255 | 20170217 space adjusted | 438.7764
LS3 | BTS | IP | 5073 | LS3_BTS:IP_D5073 | IP_S | ion pump, 75 l/s |  | 283.0561 | 283.0561 | 0 | 283.0561 | 427.1942 | 41.094 | 41.094 | 427.1942 | 283.0561 | 507.255 | 507.255 |  | 438.7764
[6 empty] | bellow |  |  |  | 0.18999 | 282.9611 | 427.1942 | 41.094 | 41.094 | 427.1942 | 282.8661 | 507.35 | 507.445 | 20170217 space adjusted | 438.8714
[10 empty] | 1.3735 | 282.1794 | 427.1942 | 41.094 | 41.094 | 427.1942 | 281.4927 | 508.1317 | 508.8185 |  | 439.6531
[8 empty] | 3.176e-06 | 0.001019 | 0 | 281.4927 | 427.1942 | 41.094 | 41.094 | 427.1942 | 281.4927 | 508.8185 | 508.8185 |  | 440.3399
LS3 | BTS | DCHV | 5089 | LS3_BTS:DCHV_D5089 | C2 | correctors-X,Y | 50 | 281.3927 | 281.3927 | 0.2 | 281.3927 | 427.1942 | 41.094 | 41.094 | 427.1942 | 281.2927 | 508.9185 | 509.0185 |  | 440.4399
[8 empty] | 3.176e-06 | -0.05468 | 0.1332 | 281.2261 | 427.1942 | 41.094 | 41.094 | 427.1942 | 281.1595 | 509.0851 | 509.1517 | 20170217 space adjusted | 440.6065
LS3 | BTS | BPM | 5092 | LS3_BTS:BPM_D5092 | BPM3 | position monitor | 47.5 | 281.1495 | 281.1495 | 0.02 | 281.1495 | 427.1942 | 41.094 | 41.094 | 427.1942 | 281.1395 | 509.1617 | 509.1717 | 20170217 changed from .137 to .02 m (button diameter) | 440.6831
[8 empty] | 0.0585 |  | 0.04782 | 281.1156 | 427.1942 | 41.094 | 41.094 | 427.1942 | 281.0916 | 509.1956 | 509.2195 | 20170217 space adjusted | 440.717
LS3 | BTS | QH | 5093 | LS3_BTS:QH_D5093 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 280.9666 | 427.1942 | 41.094 | 41.094 | 427.1942 | 280.8416 | 509.3445 | 509.4695 |  | 440.8659
[10 empty] | 2 | 279.8416 | 427.1942 | 41.094 | 41.094 | 427.1942 | 278.8416 | 510.4695 | 511.4695 |  | 441.9909
LS3 | BTS | QV | 5116 | LS3_BTS:QV_D5116 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 278.7166 | 427.1942 | 41.094 | 41.094 | 427.1942 | 278.5916 | 511.5945 | 511.7195 |  | 443.1159
[10 empty] | 1.6251 | 277.7791 | 427.1942 | 41.094 | 41.094 | 427.1942 | 276.9665 | 512.5321 | 513.3447 |  | 444.0535
[6 empty] | bellow |  | -0.075645 |  | 0.1035 | 276.9147 | 427.1942 | 41.094 | 41.094 | 427.1942 | 276.863 | 513.3964 | 513.4482 |  | 444.9178
LS3 | BTS | BPM | 5134 | LS3_BTS:BPM_D5134 | BPM3 | position monitor |  | 2.03e-06 | 276.863 | 0 | 276.863 | 427.1942 | 41.094 | 41.094 | 427.1942 | 276.863 | 513.4482 | 513.4482 |  | 444.9695
[6 empty] | diagnostic box |  |  |  | 0.14528 | 276.7904 | 427.1942 | 41.094 | 41.094 | 427.1942 | 276.7177 | 513.5208 | 513.5934 |  | 445.0422
LS3 | BTS | PORT | 5136 | LS3_BTS:PORT_D5136 |  | port for future diagnostics |  |  |  | 0 | 276.7177 | 427.1942 | 41.094 | 41.094 | 427.1942 | 276.7177 | 513.5934 | 513.5934 |  | 445.1148
[6 empty] | diagnostic box |  |  |  | 0 | 276.7177 | 427.1942 | 41.094 | 41.094 | 427.1942 | 276.7177 | 513.5934 | 513.5934 |  | 445.1148
LS3 | BTS | IP | 5136 | LS3_BTS:IP_D5136 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 276.7177 | 427.1942 | 41.094 | 41.094 | 427.1942 | 276.7177 | 513.5934 | 513.5934 |  | 445.1148
[6 empty] | bellow |  |  |  | 0.13122 | 276.6521 | 427.1942 | 41.094 | 41.094 | 427.1942 | 276.5865 | 513.659 | 513.7247 |  | 445.1804
[10 empty] | 1.2989 | 275.9371 | 427.1942 | 41.094 | 41.094 | 427.1942 | 275.2876 | 514.3741 | 515.0235 |  | 445.8955
LS3 | BTS | DCHV | 5151 | LS3_BTS:DCHV_D5151 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 275.1876 | 427.1942 | 41.094 | 41.094 | 427.1942 | 275.0876 | 515.1235 | 515.2235 |  | 446.6449
[10 empty] | 0.2 | 274.9876 | 427.1942 | 41.094 | 41.094 | 427.1942 | 274.8876 | 515.3235 | 515.4235 |  | 446.8449
LS3 | BTS | QH | 5155 | LS3_BTS:QH_D5155 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 274.7626 | 427.1942 | 41.094 | 41.094 | 427.1942 | 274.6376 | 515.5485 | 515.6735 |  | 447.0699
[10 empty] | 2 | 273.6376 | 427.1942 | 41.094 | 41.094 | 427.1942 | 272.6376 | 516.6735 | 517.6735 |  | 448.1949
LS3 | BTS | QV | 5178 | LS3_BTS:QV_D5178 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 272.5126 | 427.1942 | 41.094 | 41.094 | 427.1942 | 272.3876 | 517.7985 | 517.9235 |  | 449.3199
[10 empty] | 1.5495 | 271.6129 | 427.1942 | 41.094 | 41.094 | 427.1942 | 270.8381 | 518.6983 | 519.473 |  | 450.2196
[10 empty] | 0.096 | 270.7901 | 427.1942 | 41.094 | 41.094 | 427.1942 | 270.7421 | 519.521 | 519.569 |  | 451.0424
[10 empty] | 0 | 270.7421 | 427.1942 | 41.094 | 41.094 | 427.1942 | 270.7421 | 519.569 | 519.569 |  | 451.0904
[6 empty] | bellow |  |  |  | 0.094 | 270.6951 | 427.1942 | 41.094 | 41.094 | 427.1942 | 270.6481 | 519.616 | 519.663 |  | 451.1374
LS3 | BTS | PORT | 5197 | LS3_BTS:PORT_D5197 |  | port for future diagnostics |  |  |  | 0 | 270.6481 | 427.1942 | 41.094 | 41.094 | 427.1942 | 270.6481 | 519.663 | 519.663 |  | 451.1844
[6 empty] | vacuum box |  |  |  | 0.091 | 270.6026 | 427.1942 | 41.094 | 41.094 | 427.1942 | 270.5571 | 519.7085 | 519.754 |  | 451.2299
LS3 | BTS | IP | 5198 | LS3_BTS:IP_D5198 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 270.5571 | 427.1942 | 41.094 | 41.094 | 427.1942 | 270.5571 | 519.754 | 519.754 |  | 451.2754
[6 empty] | bellow |  |  |  | 0.099 | 270.5076 | 427.1942 | 41.094 | 41.094 | 427.1942 | 270.4581 | 519.8035 | 519.853 |  | 451.3249
[10 empty] | 1.3745 | 269.7709 | 427.1942 | 41.094 | 41.094 | 427.1942 | 269.0836 | 520.5403 | 521.2275 |  | 452.0616
[10 empty] | 0 | 269.0836 | 427.1942 | 41.094 | 41.094 | 427.1942 | 269.0836 | 521.2275 | 521.2275 |  | 452.7489
LS3 | BTS | DCHV | 5213 | LS3_BTS:DCHV_D5213 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 268.9836 | 427.1942 | 41.094 | 41.094 | 427.1942 | 268.8836 | 521.3275 | 521.4275 |  | 452.8489
[10 empty] | 0.14218 | 268.8126 | 427.1942 | 41.094 | 41.094 | 427.1942 | 268.7415 | 521.4986 | 521.5697 |  | 453.02
LS3 | BTS | BPM | 5216 | LS3_BTS:BPM_D5216 | BPM3 | position monitor | 47.5 | -2.176e-06 | 268.7415 | 0 | 268.7415 | 427.1942 | 41.094 | 41.094 | 427.1942 | 268.7415 | 521.5697 | 521.5697 |  | 453.0911
[10 empty] | 0.057821 | 268.7126 | 427.1942 | 41.094 | 41.094 | 427.1942 | 268.6836 | 521.5986 | 521.6275 |  | 453.12
LS3 | BTS | QH | 5218 | LS3_BTS:QH_D5218 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 268.5586 | 427.1942 | 41.094 | 41.094 | 427.1942 | 268.4336 | 521.7525 | 521.8775 |  | 453.2739
[10 empty] | 2 | 267.4336 | 427.1942 | 41.094 | 41.094 | 427.1942 | 266.4336 | 522.8775 | 523.8775 |  | 454.3989
LS3 | BTS | QV | 5240 | LS3_BTS:QV_D5240 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 266.3086 | 427.1942 | 41.094 | 41.094 | 427.1942 | 266.1836 | 524.0025 | 524.1275 |  | 455.5239
[10 empty] | 1.6252 | 265.3711 | 427.1942 | 41.094 | 41.094 | 427.1942 | 264.5585 | 524.9401 | 525.7527 |  | 456.4615
[6 empty] | bellow |  | -0.075663 |  | 0.1035 | 264.5067 | 427.1942 | 41.094 | 41.094 | 427.1942 | 264.455 | 525.8044 | 525.8562 |  | 457.3258
LS3 | BTS | BPM | 5259 | LS3_BTS:BPM_D5259 | BPM3 | position monitor |  | 2.03e-06 | 264.455 | 0 | 264.455 | 427.1942 | 41.094 | 41.094 | 427.1942 | 264.455 | 525.8562 | 525.8562 |  | 457.3775
[6 empty] | diagnostic box |  |  |  | 0.14528 | 264.3823 | 427.1942 | 41.094 | 41.094 | 427.1942 | 264.3097 | 525.9288 | 526.0015 |  | 457.4502
LS3 | BTS | PORT | 5260 | LS3_BTS:PORT_D5260 |  | port for future diagnostics |  |  |  | 0 | 264.3097 | 427.1942 | 41.094 | 41.094 | 427.1942 | 264.3097 | 526.0015 | 526.0015 |  | 457.5228
[6 empty] | diagnostic box |  | 2.03e-06 | 264.3097 | 0 | 264.3097 | 427.1942 | 41.094 | 41.094 | 427.1942 | 264.3097 | 526.0015 | 526.0015 |  | 457.5228
LS3 | BTS | IP | 5260 | LS3_BTS:IP_D5260 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 264.3097 | 427.1942 | 41.094 | 41.094 | 427.1942 | 264.3097 | 526.0015 | 526.0015 |  | 457.5228
[6 empty] | bellow |  |  |  | 0.13122 | 264.2441 | 427.1942 | 41.094 | 41.094 | 427.1942 | 264.1785 | 526.0671 | 526.1327 |  | 457.5884
[10 empty] | 1.2988 | 263.5291 | 427.1942 | 41.094 | 41.094 | 427.1942 | 262.8796 | 526.7821 | 527.4315 |  | 458.3035
LS3 | BTS | DCHV | 5275 | LS3_BTS:DCHV_D5275 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 262.7796 | 427.1942 | 41.094 | 41.094 | 427.1942 | 262.6796 | 527.5315 | 527.6315 |  | 459.0529
[10 empty] | 0.2 | 262.5796 | 427.1942 | 41.094 | 41.094 | 427.1942 | 262.4796 | 527.7315 | 527.8315 |  | 459.2529
LS3 | BTS | QH | 5280 | LS3_BTS:QH_D5280 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 262.3546 | 427.1942 | 41.094 | 41.094 | 427.1942 | 262.2296 | 527.9565 | 528.0815 |  | 459.4779
[10 empty] | 2 | 261.2296 | 427.1942 | 41.094 | 41.094 | 427.1942 | 260.2296 | 529.0815 | 530.0815 |  | 460.6029
LS3 | BTS | QV | 5302 | LS3_BTS:QV_D5302 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 260.1046 | 427.1942 | 41.094 | 41.094 | 427.1942 | 259.9796 | 530.2065 | 530.3315 |  | 461.7279
[10 empty] | 1.5495 | 259.2049 | 427.1942 | 41.094 | 41.094 | 427.1942 | 258.4301 | 531.1063 | 531.881 |  | 462.6276
[10 empty] | 0.096 | 258.3821 | 427.1942 | 41.094 | 41.094 | 427.1942 | 258.3341 | 531.929 | 531.977 |  | 463.4504
[10 empty] | 0 | 258.3341 | 427.1942 | 41.094 | 41.094 | 427.1942 | 258.3341 | 531.977 | 531.977 |  | 463.4984
[6 empty] | bellow |  |  |  | 0.094 | 258.2871 | 427.1942 | 41.094 | 41.094 | 427.1942 | 258.2401 | 532.024 | 532.071 |  | 463.5454
LS3 | BTS | PORT | 5321 | LS3_BTS:PORT_D5321 |  | port for future diagnostics |  |  |  | 0 | 258.2401 | 427.1942 | 41.094 | 41.094 | 427.1942 | 258.2401 | 532.071 | 532.071 |  | 463.5924
[6 empty] | vacuum box |  |  |  | 0 | 258.2401 | 427.1942 | 41.094 | 41.094 | 427.1942 | 258.2401 | 532.071 | 532.071 |  | 463.5924
LS3 | BTS | IP | 5321 | LS3_BTS:IP_D5321 | IP_S | ion pump, 75 l/s |  | 2.176e-06 | 258.2401 | 0 | 258.2401 | 427.1942 | 41.094 | 41.094 | 427.1942 | 258.2401 | 532.071 | 532.071 |  | 463.5924
[6 empty] | bellow |  |  |  | 0.19 | 258.1451 | 427.1942 | 41.094 | 41.094 | 427.1942 | 258.0501 | 532.166 | 532.261 |  | 463.6874
[10 empty] | 1.3745 | 257.3629 | 427.1942 | 41.094 | 41.094 | 427.1942 | 256.6756 | 532.9483 | 533.6355 |  | 464.4696
[10 empty] | 0 | 256.6756 | 427.1942 | 41.094 | 41.094 | 427.1942 | 256.6756 | 533.6355 | 533.6355 |  | 465.1569
LS3 | BTS | DCHV | 5337 | LS3_BTS:DCHV_D5337 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 256.5756 | 427.1942 | 41.094 | 41.094 | 427.1942 | 256.4756 | 533.7355 | 533.8355 |  | 465.2569
[10 empty] | 0.14219 | 256.4045 | 427.1942 | 41.094 | 41.094 | 427.1942 | 256.3334 | 533.9066 | 533.9777 |  | 465.428
LS3 | BTS | BPM | 5340 | LS3_BTS:BPM_D5340 | BPM3 | position monitor | 47.5 | 2.176e-06 | 256.3335 | 0 | 256.3334 | 427.1942 | 41.094 | 41.094 | 427.1942 | 256.3334 | 533.9777 | 533.9777 |  | 465.4991
[10 empty] | 0.057809 | 256.3045 | 427.1942 | 41.094 | 41.094 | 427.1942 | 256.2756 | 534.0066 | 534.0355 |  | 465.528
LS3 | BTS | QH | 5342 | LS3_BTS:QH_D5342 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 256.1506 | 427.1942 | 41.094 | 41.094 | 427.1942 | 256.0256 | 534.1605 | 534.2855 |  | 465.6819
[10 empty] | 2 | 255.0256 | 427.1942 | 41.094 | 41.094 | 427.1942 | 254.0256 | 535.2855 | 536.2855 |  | 466.8069
LS3 | BTS | QV | 5364 | LS3_BTS:QV_D5364 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 253.9006 | 427.1942 | 41.094 | 41.094 | 427.1942 | 253.7756 | 536.4105 | 536.5355 |  | 467.9319
[10 empty] | 1.4852 | 253.033 | 427.1942 | 41.094 | 41.094 | 427.1942 | 252.2904 | 537.2781 | 538.0207 |  | 468.7995
[6 empty] | bellow |  |  |  | 0.1035 | 252.2387 | 427.1942 | 41.094 | 41.094 | 427.1942 | 252.1869 | 538.0725 | 538.1242 |  | 469.5938
LS3 | BTS | BPM | 5381 | LS3_BTS:BPM_D5381 | BPM3 | position monitor |  | 2.176e-06 | 252.1869 | 0 | 252.1869 | 427.1942 | 41.094 | 41.094 | 427.1942 | 252.1869 | 538.1242 | 538.1242 |  | 469.6456
[6 empty] | diagnostic box |  |  |  | 0.14528 | 252.1143 | 427.1942 | 41.094 | 41.094 | 427.1942 | 252.0416 | 538.1969 | 538.2695 |  | 469.7182
LS3 | BTS | PORT | 5383 | LS3_BTS:PORT_D5383 |  | port for future diagnostics |  |  |  | 0 | 252.0416 | 427.1942 | 41.094 | 41.094 | 427.1942 | 252.0416 | 538.2695 | 538.2695 |  | 469.7909
[6 empty] | diagnostic box |  |  |  | 0 | 252.0416 | 427.1942 | 41.094 | 41.094 | 427.1942 | 252.0416 | 538.2695 | 538.2695 |  | 469.7909
LS3 | BTS | IP | 5383 | LS3_BTS:IP_D5383 | IP_S | ion pump, 75 l/s |  | 2.176e-06 | 252.0416 | 0 | 252.0416 | 427.1942 | 41.094 | 41.094 | 427.1942 | 252.0416 | 538.2695 | 538.2695 |  | 469.7909
[6 empty] | bellow |  |  |  | 0.13122 | 251.976 | 427.1942 | 41.094 | 41.094 | 427.1942 | 251.9104 | 538.3351 | 538.4007 |  | 469.8565
[10 empty] | 0.35978 | 251.7305 | 427.1942 | 41.094 | 41.094 | 427.1942 | 251.5506 | 538.5806 | 538.7605 |  | 470.102
LS3 | BTS | DCHV | 5389 | LS3_BTS:DCHV_D5389 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 251.4506 | 427.1942 | 41.094 | 41.094 | 427.1942 | 251.3506 | 538.8605 | 538.9605 |  | 470.3819
[10 empty] | 0.2 | 251.2506 | 427.1942 | 41.094 | 41.094 | 427.1942 | 251.1506 | 539.0605 | 539.1605 |  | 470.5819
LS3 | BTS | QH | 5393 | LS3_BTS:QH_D5393 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 251.0256 | 427.1942 | 41.094 | 41.094 | 427.1942 | 250.9006 | 539.2855 | 539.4105 |  | 470.8069
[10 empty] | 1.25 | 250.2756 | 427.1942 | 41.094 | 41.094 | 427.1942 | 249.6506 | 540.0355 | 540.6605 |  | 471.5569
LS3 | BTS | QV | 5408 | LS3_BTS:QV_D5408 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 249.5256 | 427.1942 | 41.094 | 41.094 | 427.1942 | 249.4006 | 540.7855 | 540.9105 |  | 472.3069
[8 empty] | 3.176e-06 | 0.001019 | 1.779 | 248.5112 | 427.1942 | 41.094 | 41.094 | 427.1942 | 247.6217 | 541.8 | 542.6895 | 20170217 space adjusted | 473.3214
LS3 | BTS | DCHV | 5428 | LS3_BTS:DCHV_D5428 | C2 | correctors-X,Y | 50 | 247.5217 | 247.5217 | 0.2 | 247.5217 | 427.1942 | 41.094 | 41.094 | 427.1942 | 247.4217 | 542.7895 | 542.8895 |  | 474.3109
[8 empty] | 3.176e-06 | -0.05468 | 0.1332 | 247.3551 | 427.1942 | 41.094 | 41.094 | 427.1942 | 247.2885 | 542.9561 | 543.0227 | 20170217 space adjusted | 474.4775
LS3 | BTS | BPM | 5430 | LS3_BTS:BPM_D5430 | BPM3 | position monitor | 47.5 | 247.2785 | 247.2785 | 0.02 | 247.2785 | 427.1942 | 41.094 | 41.094 | 427.1942 | 247.2685 | 543.0327 | 543.0427 | 20170217 changed from .137 to .02 m (button diameter) | 474.5541
[8 empty] | 0.0585 |  | 0.04782 | 247.2446 | 427.1942 | 41.094 | 41.094 | 427.1942 | 247.2206 | 543.0666 | 543.0905 | 20170217 space adjusted | 474.588
LS3 | BTS | QH | 5432 | LS3_BTS:QH_D5432 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 247.0956 | 427.1942 | 41.094 | 41.094 | 427.1942 | 246.9706 | 543.2155 | 543.3405 |  | 474.7369
[10 empty] | 0.5 | 246.7206 | 427.1942 | 41.094 | 41.094 | 427.1942 | 246.4706 | 543.5905 | 543.8405 |  | 475.1119
LS3 | BTS | QV | 5440 | LS3_BTS:QV_D5440 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 246.3456 | 427.1942 | 41.094 | 41.094 | 427.1942 | 246.2206 | 543.9655 | 544.0905 |  | 475.4869
[10 empty] | 0.27322 | 246.084 | 427.1942 | 41.094 | 41.094 | 427.1942 | 245.9474 | 544.2271 | 544.3637 | 20170217 space adjusted | 475.7485
[6 empty] | bellow |  | 2.99e-06 | -0.053068 | 0.1035 | 245.8957 | 427.1942 | 41.094 | 41.094 | 427.1942 | 245.8439 | 544.4155 | 544.4672 |  | 475.9369
LS3 | BTS | BPM | 5445 | LS3_BTS:BPM_D5445 | BPM3 | position monitor |  | 245.8439 | 245.8439 | 0 | 245.8439 | 427.1942 | 41.094 | 41.094 | 427.1942 | 245.8439 | 544.4672 | 544.4672 |  | 475.9886
[6 empty] | diagnostic box |  | 2.99e-06 |  | 0.14528 | 245.7713 | 427.1942 | 41.094 | 41.094 | 427.1942 | 245.6986 | 544.5399 | 544.6125 |  | 476.0612
LS3 | BTS | PORT | 5446 | LS3_BTS:PORT_D5446 |  | port for future diagnostics |  | 245.6986 | 245.6986 | 0 | 245.6986 | 427.1942 | 41.094 | 41.094 | 427.1942 | 245.6986 | 544.6125 | 544.6125 |  | 476.1339
[6 empty] | diagnostic box |  |  |  | 0 | 245.6986 | 427.1942 | 41.094 | 41.094 | 427.1942 | 245.6986 | 544.6125 | 544.6125 |  | 476.1339
LS3 | BTS | IP | 5446 | LS3_BTS:IP_D5446 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 245.6986 | 427.1942 | 41.094 | 41.094 | 427.1942 | 245.6986 | 544.6125 | 544.6125 |  | 476.1339
[6 empty] | bellow |  |  |  | 0.13122 | 245.633 | 427.1942 | 41.094 | 41.094 | 427.1942 | 245.5674 | 544.6781 | 544.7437 |  | 476.1995
[10 empty] | 0.24678 | 245.444 | 427.1942 | 41.094 | 41.094 | 427.1942 | 245.3206 | 544.8671 | 544.9905 | 20170217 space adjusted | 476.3885
[10 empty] | 1.6 | 244.5206 | 427.1942 | 41.094 | 41.094 | 427.1942 | 243.7206 | 545.7905 | 546.5905 |  | 477.3119
BDS | BTS | DCHV | 5467 | BDS_BTS:DCHV_D5467 | C3 | correctors-X,Y | 50 |  |  | 0.2 | 243.6206 | 427.1942 | 41.094 | 41.094 | 427.1942 | 243.5206 | 546.6905 | 546.7905 | 20170217 C3 used to replace C2 | 478.2119
[10 empty] | 0.2 | 243.4206 | 427.1942 | 41.094 | 41.094 | 427.1942 | 243.3206 | 546.8905 | 546.9905 |  | 478.4119
BDS | BTS | QH | 5471 | BDS_BTS:QH_D5471 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 243.1956 | 427.1942 | 41.094 | 41.094 | 427.1942 | 243.0706 | 547.1155 | 547.2405 |  | 478.6369
[10 empty] | 0.5 | 242.8206 | 427.1942 | 41.094 | 41.094 | 427.1942 | 242.5706 | 547.4905 | 547.7405 |  | 479.0119
BDS | BTS | QV | 5479 | BDS_BTS:QV_D5479 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 242.4456 | 427.1942 | 41.094 | 41.094 | 427.1942 | 242.3206 | 547.8655 | 547.9905 |  | 479.3869
[10 empty] | 1.55 | 241.5456 | 427.1942 | 41.094 | 41.094 | 427.1942 | 240.7706 | 548.7655 | 549.5405 |  | 480.2869
BDS | BTS | DCHV | 5496 | BDS_BTS:DCHV_D5496 | C3 | correctors-X,Y | 50 |  |  | 0.2 | 240.6706 | 427.1942 | 41.094 | 41.094 | 427.1942 | 240.5706 | 549.6405 | 549.7405 | 20170217 C3 used to replace C2 | 481.1619
[8 empty] | 2.976e-06 | -0.054676 | 0.18218 | 240.4796 | 427.1942 | 41.094 | 41.094 | 427.1942 | 240.3885 | 549.8316 | 549.9227 | 20170217 space adjusted | 481.353
BDS | BTS | BPM | 5499 | BDS_BTS:BPM_D5499 | BPM3 | position monitor | 47.5 | 240.3785 | 240.3785 | 0.02 | 240.3785 | 427.1942 | 41.094 | 41.094 | 427.1942 | 240.3685 | 549.9327 | 549.9427 | 20170217 changed from .137 to .02 m (button diameter) | 481.4541
[8 empty] | 0.0585 |  | 0.047824 | 240.3446 | 427.1942 | 41.094 | 41.094 | 427.1942 | 240.3206 | 549.9666 | 549.9905 | 20170217 space adjusted | 481.488
BDS | BTS | QV | 5501 | BDS_BTS:QV_D5501 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 240.1956 | 427.1942 | 41.094 | 41.094 | 427.1942 | 240.0706 | 550.1155 | 550.2405 |  | 481.6369
[10 empty] | 0.5 | 239.8206 | 427.1942 | 41.094 | 41.094 | 427.1942 | 239.5706 | 550.4905 | 550.7405 |  | 482.0119
BDS | BTS | QH | 5509 | BDS_BTS:QH_D5509 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 239.4456 | 427.1942 | 41.094 | 41.094 | 427.1942 | 239.3206 | 550.8655 | 550.9905 |  | 482.3869
[10 empty] | 0.17008 | 239.2356 | 427.1942 | 41.094 | 41.094 | 427.1942 | 239.1506 | 551.0755 | 551.1606 | 20170217 space adjusted | 482.5969
[6 empty] | bellow |  | 2.99e-06 | -0.085079 | 0.1035 | 239.0988 | 427.1942 | 41.094 | 41.094 | 427.1942 | 239.0471 | 551.2123 | 551.2641 |  | 482.7337
BDS | BTS | BPM | 5513 | BDS_BTS:BPM_D5513 | BPM3 | position monitor |  | 239.0471 | 239.0471 | 0 | 239.0471 | 427.1942 | 41.094 | 41.094 | 427.1942 | 239.0471 | 551.2641 | 551.2641 |  | 482.7855
[6 empty] | diagnostic box |  |  |  | 0.14528 | 238.9744 | 427.1942 | 41.094 | 41.094 | 427.1942 | 238.9018 | 551.3367 | 551.4094 |  | 482.8581
BDS | BTS | PM | 5514 | BDS_BTS:PM_D5514 | PM3 | profile monitor |  |  |  | 0 | 238.9018 | 427.1942 | 41.094 | 41.094 | 427.1942 | 238.9018 | 551.4094 | 551.4094 |  | 482.9307
[6 empty] | diagnostic box |  |  |  | 0 | 238.9018 | 427.1942 | 41.094 | 41.094 | 427.1942 | 238.9018 | 551.4094 | 551.4094 |  | 482.9307
BDS | BTS | IP | 5514 | BDS_BTS:IP_D5514 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 238.9018 | 427.1942 | 41.094 | 41.094 | 427.1942 | 238.9018 | 551.4094 | 551.4094 |  | 482.9307
[6 empty] | bellow |  |  |  | 0.13122 | 238.8362 | 427.1942 | 41.094 | 41.094 | 427.1942 | 238.7706 | 551.475 | 551.5406 |  | 482.9963
[10 empty] | 0 | 238.7706 | 427.1942 | 41.094 | 41.094 | 427.1942 | 238.7706 | 551.5406 | 551.5406 |  | 483.062
[8 empty] | 3.176e-06 | -0.19144 | 0.52464 | 238.5082 | 427.1942 | 41.094 | 41.094 | 427.1942 | 238.2459 | 551.8029 | 552.0652 | 20170217 space adjusted | 483.3243
BDS | BTS | BCM | 5521 | BDS_BTS:BCM_D5521 | BCM3 | current monitor |  | 238.1959 | 238.1959 | 0.1 | 238.1959 | 427.1942 | 41.094 | 41.094 | 427.1942 | 238.1459 | 552.1152 | 552.1652 |  | 483.6366
[8 empty] | 3.176e-06 | 0.32498 | 2.4836 | 236.9041 | 427.1942 | 41.094 | 41.094 | 427.1942 | 235.6623 | 553.407 | 554.6488 | 20170217 space adjusted | 484.9284
BDS | BTS | GV | 5547 | BDS_BTS:GV_D5547 | GV_M | gate valve | 63 | 235.6248 | 235.6248 | 0.075 | 235.6248 | 427.1942 | 41.094 | 41.094 | 427.1942 | 235.5873 | 554.6863 | 554.7238 |  | 486.2077
[10 empty] | 0.32498 | 235.4249 | 427.1942 | 41.094 | 41.094 | 427.1942 | 235.2624 | 554.8863 | 555.0488 | 20170217 space adjusted | 486.4077
BDS | BTS | QV | 5552 | BDS_BTS:QV_D5552 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 235.1374 | 427.1942 | 41.094 | 41.094 | 427.1942 | 235.0124 | 555.1738 | 555.2988 |  | 486.6952
[10 empty] | 0.5 | 234.7624 | 427.1942 | 41.094 | 41.094 | 427.1942 | 234.5124 | 555.5488 | 555.7988 |  | 487.0702
BDS | BTS | QH | 5559 | BDS_BTS:QH_D5559 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 234.3874 | 427.1942 | 41.094 | 41.094 | 427.1942 | 234.2624 | 555.9238 | 556.0488 |  | 487.4452
[10 empty] | 0.125 | 234.1999 | 427.1942 | 41.094 | 41.094 | 427.1942 | 234.1374 | 556.1113 | 556.1738 |  | 487.6327
BDS | BTS | DCHV | 5563 | BDS_BTS:DCHV_D5563 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 234.0374 | 427.1942 | 41.094 | 41.094 | 427.1942 | 233.9374 | 556.2738 | 556.3738 |  | 487.7952
[10 empty] | 0.03057 | 233.9221 | 427.1942 | 41.094 | 41.094 | 427.1942 | 233.9068 | 556.3891 | 556.4044 | 20170217 space adjusted | 487.9104
[6 empty] | bellow |  |  | 0.058777 | 0.1035 | 233.855 | 427.1942 | 41.094 | 41.094 | 427.1942 | 233.8033 | 556.4561 | 556.5079 |  | 487.9775
BDS | BTS | BPM | 5565 | BDS_BTS:BPM_D5565 | BPM3 | position monitor |  | 2.176e-06 | 233.8033 | 0 | 233.8033 | 427.1942 | 41.094 | 41.094 | 427.1942 | 233.8033 | 556.5079 | 556.5079 |  | 488.0292
[6 empty] | diagnostic box |  |  |  | 0.14528 | 233.7307 | 427.1942 | 41.094 | 41.094 | 427.1942 | 233.658 | 556.5805 | 556.6531 |  | 488.1019
BDS | BTS | PM | 5567 | BDS_BTS:PM_D5567 | PM3 | profile monitor |  |  |  | 0 | 233.658 | 427.1942 | 41.094 | 41.094 | 427.1942 | 233.658 | 556.6531 | 556.6531 |  | 488.1745
[6 empty] | diagnostic box |  |  |  | 0 | 233.658 | 427.1942 | 41.094 | 41.094 | 427.1942 | 233.658 | 556.6531 | 556.6531 |  | 488.1745
BDS | BTS | IP | 5567 | BDS_BTS:IP_D5567 | IP_S | ion pump, 75 l/s |  | 2.176e-06 | 233.658 | 0 | 233.658 | 427.1942 | 41.094 | 41.094 | 427.1942 | 233.658 | 556.6531 | 556.6531 |  | 488.1745
[10 empty] | 0 | 233.658 | 427.1942 | 41.094 | 41.094 | 427.1942 | 233.658 | 556.6531 | 556.6531 |  | 488.1745 | [14 empty] | 41.094 | 427.1942 | 233.3624
[6 empty] | bellow |  |  |  | 0.13122 | 233.5924 | 427.1942 | 41.094 | 41.094 | 427.1942 | 233.5268 | 556.7187 | 556.7844 |  | 488.2401 | [14 empty] | 41.094 | 427.1942 | 231.5562 |  |  | straight line
Fourth beam dump line starts in "V" column | [9 empty] | 0.16443 | 233.4446 | 427.1942 | 41.094 | 41.094 | 427.1942 | 233.3624 | 556.8666 | 556.9488 | 20170217 space adjusted | 488.3879 | dump4 line start (BDS straight line dump) | [13 empty] | 41.094 | 427.1942 | 233.3624 |  |  | 2013/06/20 length updated from 10.47 to 9.1406 Nathan Bultman
BDS | BBS | DH | 5578 | BDS_BBS:DH_D5578 | D4 | dipole-17.5deg, D4-1 | 35 x 50 |  |  | 1.6503 | 232.55 | 427.3193 | 41.094 | 41.094 | 427.4443 | 231.7377 | 557.7739 | 558.599 |  | 489.2953 | BDS | BTS | DH | 5579 | BDS_BTS:DH_D5579 | D4 | dipole-17.5deg, D4-1 | 35 x 50 | 2.03e-06 | 1.8061 | 1.8061 | 231.5562 | 427.1942 | 41.094 | 41.094 | 427.1942 | 231.5562 | 557.8519 | 558.7549
[10 empty] | 0.78617 | 231.3628 | 427.5625 | 41.094 | 41.094 | 427.6807 | 230.9879 | 558.9921 | 559.3852 |  | 490.5135 | BDS | BTS | GV | 5588 | BDS_BTS:GV_D5588 | GV_M | gate valve | 63 | 231.5187 | 231.5187 | 0.075 | 231.5187 | 427.1942 | 41.094 | 41.094 | 427.1942 | 231.4812 | 558.7924 | 558.8299
[10 empty] | 0.3114 | 230.8394 | 427.7275 | 41.094 | 41.094 | 427.7743 | 230.6909 | 559.5409 | 559.6966 |  | 491.0623 | [9 empty] | 0.98885 | 0 | 231.4812 | 427.1942 | 41.094 | 41.094 | 427.1942 | 231.4812 | 558.8299 | 558.8299
[10 empty] | 0.8 | 230.3094 | 427.8946 | 41.094 | 41.094 | 428.0149 | 229.9279 | 560.0966 | 560.4966 |  | 491.618 | [10 empty] | 0 | 231.4812 | 427.1942 | 41.094 | 41.094 | 427.1942 | 231.4812 | 558.8299 | 558.8299
BDS | BBS | S | 5606 | BDS_BBS:S_D5606 | H2 | hexapole-X+ | 100 |  |  | 0.2 | 229.8325 | 428.045 | 41.094 | 41.094 | 428.075 | 229.7372 | 560.5966 | 560.6966 |  | 492.118 | [10 empty] | 0 | 231.4812 | 427.1942 | 41.094 | 41.094 | 427.1942 | 231.4812 | 558.8299 | 558.8299
[10 empty] | 0.2 | 229.6418 | 428.1051 | 41.094 | 41.094 | 428.1352 | 229.5464 | 560.7966 | 560.8966 |  | 492.318 | [10 empty] | 0 | 231.4812 | 427.1942 | 41.094 | 41.094 | 427.1942 | 231.4812 | 558.8299 | 558.8299
BDS | BBS | QH | 5611 | BDS_BBS:QH_D5611 | Q5 | quadrupole-X,40cm,B3 | 100 |  |  | 0.4 | 229.3557 | 428.1953 | 41.094 | 41.094 | 428.2555 | 229.1649 | 561.0966 | 561.2966 |  | 492.618 | [10 empty] | 0.98885 | 230.9868 | 427.1942 | 41.094 | 41.094 | 427.1942 | 230.4924 | 559.3244 | 559.8188
[10 empty] | 0.2 | 229.0696 | 428.2855 | 41.094 | 41.094 | 428.3156 | 228.9742 | 561.3966 | 561.4966 |  | 492.918 | BDS | BTS | Q | 5599 | BDS_BTS:Q_D5599 | Q1 | quadrupole-25cm |  |  |  | 0.25 | 230.3674 | 427.1942 | 41.094 | 41.094 | 427.1942 | 230.2424 | 559.9438 | 560.0688
BDS | BBS | DCHV | 5616 | BDS_BBS:DCHV_D5616 | C3 | correctors-X,Y | 100 |  |  | 0.2 | 228.8788 | 428.3457 | 41.094 | 41.094 | 428.3757 | 228.7834 | 561.5966 | 561.6966 |  | 493.118 | [10 empty] | 0.75 | 229.8674 | 427.1942 | 41.094 | 41.094 | 427.1942 | 229.4924 | 560.4438 | 560.8188
[10 empty] | 0.2 | 228.6881 | 428.4058 | 41.094 | 41.094 | 428.4359 | 228.5927 | 561.7966 | 561.8966 |  | 493.318 | BDS | BTS | Q | 5609 | BDS_BTS:Q_D5609 | Q2 | quadrupole-26cm |  |  |  | 0.26 | 229.3624 | 427.1942 | 41.094 | 41.094 | 427.1942 | 229.2324 | 560.9488 | 561.0788
[10 empty] | 0 | 228.5927 | 428.4359 | 41.094 | 41.094 | 428.4359 | 228.5927 | 561.8966 | 561.8966 |  | 493.418 | [10 empty] | 3.7274 | 227.3687 | 427.1942 | 41.094 | 41.094 | 427.1942 | 225.505 | 562.9425 | 564.8062
BDS | BBS | QV | 5621 | BDS_BBS:QV_D5621 | Q5 | quadrupole-Y,40cm,B3 | 100 | 0 | 0.2 | 0.4 | 228.402 | 428.496 | 41.094 | 41.094 | 428.5562 | 228.2112 | 562.0966 | 562.2966 |  | 493.618 | [6 empty] | bellow |  | 2.03e-06 |  | 0.1035 | 225.4532 | 427.1942 | 41.094 | 41.094 | 427.1942 | 225.4015 | 564.8579 | 564.9097
[8 empty] | 0.43015 | 0.19308 | 0.19308 | 228.1191 | 428.5852 | 41.094 | 41.094 | 428.6142 | 228.0271 | 562.3931 | 562.4897 | 20170216 space reduced by 0.79852 | 493.9145 | BDS | BTS | BPM | 5649 | BDS_BTS:BPM_D5649 | BPM4 | position monitor | 100 | 225.4015 | 225.4015 | 0 | 225.4015 | 427.1942 | 41.094 | 41.094 | 427.1942 | 225.4015 | 564.9097 | 564.9097 | 20170427 adjusted based on ME
BDS | BBS | BPM | 5625 | BDS_BBS:BPM_D5625 | BPM4 | position monitor | 100 | 0 | 0.037069 | 0.074138 | 227.9917 | 428.6254 | 41.094 | 41.094 | 428.6365 | 227.9564 | 562.5268 | 562.5638 |  | 494.0481 | [6 empty] | diagnostic box |  |  |  | 0.2966 | 225.2532 | 427.1942 | 41.094 | 41.094 | 427.1942 | 225.1049 | 565.058 | 565.2063 | PM5 NOT COSTED
[8 empty] | 0.27004 | 0.23297 | 0.23297 | 227.8453 | 428.6715 | 41.094 | 41.094 | 428.7066 | 227.7342 | 562.6803 | 562.7968 | 20170216 0.049913 m added | 494.2017 | BDS | BTS | PM | 5652 | BDS_BTS:PM_D5652 | PMx | profile monitor (not costed) | 100 |  |  | 0 | 225.1049 | 427.1942 | 41.094 | 41.094 | 427.1942 | 225.1049 | 565.2063 | 565.2063
BDS | BBS | PM | 5628 | BDS_BBS:PM_D5628 | PM5 | profile monitor | 100 |  | 0 | 0 | 227.7342 | 428.7066 | 41.094 | 41.094 | 428.7066 | 227.7342 | 562.7968 | 562.7968 |  | 494.3182 | [6 empty] | diagnostic box |  | -0.15132 |  | 0 | 225.1049 | 427.1942 | 41.094 | 41.094 | 427.1942 | 225.1049 | 565.2063 | 565.2063
BDS | BBS | IP | 5628 | BDS_BBS:IP_D5628 | IP_L | ion pump, 300 l/s |  |  |  | 0 | 227.7342 | 428.7066 | 41.094 | 41.094 | 428.7066 | 227.7342 | 562.7968 | 562.7968 |  | 494.3182 | BDS | BTS | IP | 5652 | BDS_BTS:IP_D5652 | IP_L | ion pump 300 l/s |  | 2.03e-06 | 225.1049 | 0 | 225.1049 | 427.1942 | 41.094 | 41.094 | 427.1942 | 225.1049 | 565.2063 | 565.2063
[10 empty] | 0.44982 | 227.5197 | 428.7742 | 41.094 | 41.094 | 428.8418 | 227.3052 | 563.0217 | 563.2466 | 20170216 0.029939 m added | 494.5431 | [10 empty] | 0 | 225.1049 | 427.1942 | 41.094 | 41.094 | 427.1942 | 225.1049 | 565.2063 | 565.2063
BDS | BBS | DH | 5641 | BDS_BBS:DH_D5641 | D4 | dipole-17.5deg, D4-2 | 35 x 50 |  |  | 1.6503 | 226.568 | 429.2054 | 41.094 | 41.094 | 429.5689 | 225.8309 | 564.0717 | 564.8969 |  | 495.5931 | [6 empty] | bellow |  |  |  | 0.13122 | 225.0393 | 427.1942 | 41.094 | 41.094 | 427.1942 | 224.9737 | 565.2719 | 565.3375 | 20170427 adjusted based on ME
[10 empty] | 0.33331 | 225.6943 | 429.6645 | 41.094 | 41.094 | 429.7601 | 225.5578 | 565.0635 | 565.2302 | 20170217 space adjusted | 496.5849 | [10 empty] | 0.75191 | 224.5977 | 427.1942 | 41.094 | 41.094 | 427.1942 | 224.2218 | 565.7134 | 566.0894
BDS | BBS | BPM | 5653 | BDS_BBS:BPM_D5653 | BPM4 | position monitor | 100 |  | 0.037069 | 0.074138 | 225.5275 | 429.7813 | 41.094 | 41.094 | 429.8026 | 225.4971 | 565.2672 | 565.3043 |  | 496.7886 | BDS | BTS | BD | 5661 | BDS_BTS:BD_D5661 |  | dump4 |  |  |  | 0 | 224.2218 | 427.1942 | 41.094 | 41.094 | 427.1942 | 224.2218 | 566.0894 | 566.0894
[10 empty] | 0.00706 | 225.4942 | 429.8046 | 41.094 | 41.094 | 429.8066 | 225.4913 | 565.3078 | 565.3114 | 20170217 space adjusted | 496.8292 | dump4 line end (BDS straight line dump)
BDS | BBS | PM | 5653 | BDS_BBS:PM_D5653 | PM5 | profile monitor | 100 | 0 |  | 0 | 225.4913 | 429.8066 | 41.094 | 41.094 | 429.8066 | 225.4913 | 565.3114 | 565.3114 |  | 496.8327
BDS | BBS | IP | 5653 | BDS_BBS:IP_D5653 | IP_L | ion pump, 300 l/s |  | 0.34736 | 0.34736 | 0 | 225.4913 | 429.8066 | 41.094 | 41.094 | 429.8066 | 225.4913 | 565.3114 | 565.3114 |  | 496.8327
[8 empty] | 2.9523e-06 |  | 0.24736 | 225.39 | 429.8776 | 41.094 | 41.094 | 429.9485 | 225.2887 | 565.4351 | 565.5587 | 20170217 space adjusted | 496.9564
BDS | BBS | DCHV | 5657 | BDS_BBS:DCHV_D5657 | C3 | correctors-X,Y | 100 | 225.2068 | 225.2068 | 0.2 | 225.2068 | 430.0059 | 41.094 | 41.094 | 430.0632 | 225.1249 | 565.6587 | 565.7587 |  | 497.1801
[7 empty] | C3 Y not match !!! | 430.0037 | 430.0059 | 0.23813 | 225.0273 | 430.1315 | 41.094 | 41.094 | 430.1998 | 224.9298 | 565.8778 | 565.9969 | 20170217 space adjusted | 497.3992
BDS | BBS | DH | 5668 | BDS_BBS:DH_D5668 | D4 | dipole-17.5deg, D4-3 | 35 x 50 |  |  | 1.6503 | 224.3361 | 430.7682 | 41.094 | 41.094 | 431.3366 | 223.7423 | 566.822 | 567.6471 |  | 498.3434
[8 empty] | 0 |  | 0.3069 | 223.6489 | 431.4583 | 41.094 | 41.094 | 431.58 | 223.5555 | 567.8006 | 567.954 | 20170217 space adjusted | 499.3219
BDS | BBS | BPM | 5680 | BDS_BBS:BPM_D5680 | BPM4 | position monitor |  | 0.27004 | 0.27004 | 0.074138 | 223.5329 | 431.6095 | 41.094 | 41.094 | 431.6389 | 223.5104 | 567.9911 | 568.0281 |  | 499.5124
[10 empty] | 0.23297 | 223.4395 | 431.7313 | 41.094 | 41.094 | 431.8237 | 223.3686 | 568.1446 | 568.2611 |  | 499.666
BDS | BBS | PM | 5683 | BDS_BBS:PM_D5683 | PMx | profile monitor (not costed) | future PM5 |  |  | 0 | 223.3686 | 431.8237 | 41.094 | 41.094 | 431.8237 | 223.3686 | 568.2611 | 568.2611 |  | 499.7825
BDS | BBS | IP | 5683 | BDS_BBS:IP_D5683 | IP_L | ion pump, 300 l/s |  | 0 | 0.093103 | 0 | 223.3686 | 431.8237 | 41.094 | 41.094 | 431.8237 | 223.3686 | 568.2611 | 568.2611 |  | 499.7825
[8 empty] | 0.536 | 0.536 | 0.336 | 223.2663 | 431.957 | 41.094 | 41.094 | 432.0903 | 223.164 | 568.4291 | 568.5971 | 20170217 space adjusted | 499.9505
BDS | BBS | QV | 5688 | BDS_BBS:QV_D5688 | Q5 | quadrupole-Y,40cm,B3 | 100 |  |  | 0.4 | 223.0423 | 432.2489 | 41.094 | 41.094 | 432.4076 | 222.9205 | 568.7971 | 568.9971 |  | 500.3185
[10 empty] | 0.2 | 222.8596 | 432.4869 | 41.094 | 41.094 | 432.5663 | 222.7988 | 569.0971 | 569.1971 |  | 500.6185
BDS | BBS | DCHV | 5693 | BDS_BBS:DCHV_D5693 | C3 | correctors-X,Y | 100 |  |  | 0.2 | 222.7379 | 432.6456 | 41.094 | 41.094 | 432.7249 | 222.677 | 569.2971 | 569.3971 |  | 500.8185
[10 empty] | 0.2 | 222.6161 | 432.8043 | 41.094 | 41.094 | 432.8836 | 222.5553 | 569.4971 | 569.5971 |  | 501.0185
BDS | BBS | QH | 5698 | BDS_BBS:QH_D5698 | Q5 | quadrupole-X,40cm,B3 | 100 |  |  | 0.4 | 222.4335 | 433.0423 | 41.094 | 41.094 | 433.2009 | 222.3118 | 569.7971 | 569.9971 |  | 501.3185
[10 empty] | 0.2 | 222.2509 | 433.2803 | 41.094 | 41.094 | 433.3596 | 222.19 | 570.0971 | 570.1971 |  | 501.6185
BDS | BBS | S | 5703 | BDS_BBS:S_D5703 | H2 | hexapole-X+ | 100 |  |  | 0.2 | 222.1291 | 433.439 | 41.094 | 41.094 | 433.5183 | 222.0682 | 570.2971 | 570.3971 |  | 501.8185
[10 empty] | 1.8976 | 221.4907 | 434.271 | 41.094 | 41.094 | 435.0237 | 220.9131 | 571.3459 | 572.2947 |  | 502.8673
BDS | BBS | DH | 5731 | BDS_BBS:DH_D5731 | D4 | dipole-17.5deg, D4-4 | 35 x 50 |  |  | 1.6503 | 220.5177 | 435.7443 | 41.094 | 41.094 | 436.4649 | 220.1224 | 573.1198 | 573.9449 |  | 504.6412
[10 empty] | 0.15235 | 220.0964 | 436.5365 | 41.094 | 41.094 | 436.6081 | 220.0703 | 574.0211 | 574.0973 | 20170217 space adjusted | 505.5425
[6 empty] | bellow |  | -3.5722e-06 |  | 0.1035 | 220.0526 | 436.6567 | 41.094 | 41.094 | 436.7054 | 220.0349 | 574.149 | 574.2008 |  | 505.6704
BDS | FFS | BPM | 5742 | BDS_FFS:BPM_D5742 | BPM3 | position monitor | 50 | 220.0349 | 220.0349 | 0 | 220.0349 | 436.7054 | 41.094 | 41.094 | 436.7054 | 220.0349 | 574.2008 | 574.2008 |  | 505.7222
[6 empty] | diagnostic box |  |  |  | 0.14528 | 220.0101 | 436.7736 | 41.094 | 41.094 | 436.8419 | 219.9852 | 574.2734 | 574.3461 |  | 505.7948
BDS | FFS | PM | 5743 | BDS_FFS:PM_D5743 | PM3 | profile monitor | 50 | 0 | -0.059133 | 0 | 219.9852 | 436.8419 | 41.094 | 41.094 | 436.8419 | 219.9852 | 574.3461 | 574.3461 |  | 505.8674
[6 empty] | diagnostic box |  | 0.29927 | 0.29927 | 0 | 219.9852 | 436.8419 | 41.094 | 41.094 | 436.8419 | 219.9852 | 574.3461 | 574.3461 |  | 505.8674
BDS | FFS | IP | 5743 | BDS_FFS:IP_D5743 | IP_S | ion pump, 75 l/s |  | -3.5435e-06 |  | 0 | 219.9852 | 436.8419 | 41.094 | 41.094 | 436.8419 | 219.9852 | 574.3461 | 574.3461 |  | 505.8674
[6 empty] | bellow |  | 219.8828 | 219.8829 | 0.13122 | 219.9628 | 436.9035 | 41.094 | 41.094 | 436.9652 | 219.9403 | 574.4117 | 574.4773 |  | 505.9331
[7 empty] | C2 Y matched! | 437.1231 | 437.1231 | 0.068049 | 219.9287 | 436.9972 | 41.094 | 41.094 | 437.0291 | 219.9171 | 574.5113 | 574.5453 | 20170217 space adjusted | 506.0327
BDS | FFS | DCHV | 5746 | BDS_FFS:DCHV_D5746 | C2 | correctors-X,Y | 50 | 0 |  | 0.2 | 219.8829 | 437.1231 | 41.094 | 41.094 | 437.2171 | 219.8487 | 574.6453 | 574.7453 | moved 1" upstream | 506.1667
[8 empty] | 0.3246 | 0.3246 | 0.0996 | 219.8316 | 437.2639 | 41.094 | 41.094 | 437.3107 | 219.8146 | 574.7951 | 574.8449 | 20170217 space adjusted | 506.3165
BDS | FFS | QH | 5750 | BDS_FFS:QH_D5750 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 219.7718 | 437.4281 | 41.094 | 41.094 | 437.5456 | 219.7291 | 574.9699 | 575.0949 |  | 506.4913
[10 empty] | 0.5 | 219.6436 | 437.7805 | 41.094 | 41.094 | 438.0154 | 219.5581 | 575.3449 | 575.5949 |  | 506.8663
BDS | FFS | QV | 5757 | BDS_FFS:QV_D5757 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 219.5153 | 438.1329 | 41.094 | 41.094 | 438.2504 | 219.4726 | 575.7199 | 575.8449 |  | 507.2413
[10 empty] | 1.2675 | 219.2558 | 438.8459 | 41.094 | 41.094 | 439.4414 | 219.0391 | 576.4787 | 577.1124 |  | 508
[6 empty] | bellow |  |  |  | 0.1035 | 219.0214 | 439.49 | 41.094 | 41.094 | 439.5386 | 219.0037 | 577.1642 | 577.2159 |  | 508.6855
BDS | FFS | BPM | 5772 | BDS_FFS:BPM_D5772 | BPM3 | position monitor | 50 |  | 219.0037 | 0 | 219.0037 | 439.5386 | 41.094 | 41.094 | 439.5386 | 219.0037 | 577.2159 | 577.2159 |  | 508.7373
[6 empty] | diagnostic box |  |  |  | 0.14528 | 218.9788 | 439.6069 | 41.094 | 41.094 | 439.6752 | 218.954 | 577.2886 | 577.3612 |  | 508.8099
BDS | FFS | PM | 5774 | BDS_FFS:PM_D5774 | PM3 | profile monitor | 50 |  |  | 0 | 218.954 | 439.6752 | 41.094 | 41.094 | 439.6752 | 218.954 | 577.3612 | 577.3612 |  | 508.8826
[6 empty] | diagnostic box |  |  |  | 0 | 218.954 | 439.6752 | 41.094 | 41.094 | 439.6752 | 218.954 | 577.3612 | 577.3612 |  | 508.8826
BDS | FFS | IP | 5774 | BDS_FFS:IP_D5774 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 218.954 | 439.6752 | 41.094 | 41.094 | 439.6752 | 218.954 | 577.3612 | 577.3612 |  | 508.8826
[6 empty] | bellow |  | 0 |  | 0.13122 | 218.9315 | 439.7368 | 41.094 | 41.094 | 439.7985 | 218.9091 | 577.4268 | 577.4924 |  | 508.9482
[8 empty] | 0.33668 | 0.33668 | 0.080468 | 218.8953 | 439.8363 | 41.094 | 41.094 | 439.8741 | 218.8816 | 577.5326 | 577.5729 |  | 509.054
BDS | FFS | QH | 5777 | BDS_FFS:QH_D5777 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 218.8388 | 439.9915 | 41.094 | 41.094 | 440.109 | 218.7961 | 577.6979 | 577.8229 |  | 509.2192
[10 empty] | 0.175 | 218.7661 | 440.1912 | 41.094 | 41.094 | 440.2735 | 218.7362 | 577.9104 | 577.9979 |  | 509.4317
BDS | FFS | DCHV | 5781 | BDS_FFS:DCHV_D5781 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 218.702 | 440.3674 | 41.094 | 41.094 | 440.4614 | 218.6678 | 578.0979 | 578.1979 |  | 509.6192
[10 empty] | 0.125 | 218.6464 | 440.5201 | 41.094 | 41.094 | 440.5789 | 218.6251 | 578.2604 | 578.3229 |  | 509.7817
BDS | FFS | QV | 5784 | BDS_FFS:QV_D5784 | Q1 | quadrupole-Y | 50 | 1e-07 |  | 0.25 | 218.5823 | 440.6963 | 41.094 | 41.094 | 440.8138 | 218.5396 | 578.4479 | 578.5729 |  | 509.9692
[8 empty] | 0.31124 | 0.31124 | 0.14874 | 218.5141 | 440.8837 | 41.094 | 41.094 | 440.9535 | 218.4887 | 578.6472 | 578.7216 |  | 510.1686
BDS | FFS | GV | 5788 | BDS_FFS:GV_D5788 | GV_M | gate valve | 63 | 1.249e-16 |  | 0.075 | 218.4759 | 440.9888 | 41.094 | 41.094 | 441.024 | 218.463 | 578.7591 | 578.7966 |  | 510.2805
[8 empty] | 0.12363 | 0.12363 | 0.086132 | 218.4483 | 441.0645 | 41.094 | 41.094 | 441.105 | 218.4336 | 578.8397 | 578.8827 |  | 510.3611
BDS | FFS | BCM | 5789 | BDS_FFS:BCM_D5789 | BCM3 | current monitor |  | 0 |  | 0 | 218.4336 | 441.105 | 41.094 | 41.094 | 441.105 | 218.4336 | 578.8827 | 578.8827 |  | 510.4041
[8 empty] | 0.1235 | 0.1235 | 0.019997 | 218.4302 | 441.1144 | 41.094 | 41.094 | 441.1237 | 218.4267 | 578.8927 | 578.9027 |  | 510.4141
[6 empty] | bellow |  |  |  | 0.1035 | 218.409 | 441.1724 | 41.094 | 41.094 | 441.221 | 218.3913 | 578.9545 | 579.0062 |  | 510.4759
BDS | FFS | BPM | 5790 | BDS_FFS:BPM_D5790 | BPM3 | position monitor | 50 |  |  | 0 | 218.3913 | 441.221 | 41.094 | 41.094 | 441.221 | 218.3913 | 579.0062 | 579.0062 |  | 510.5276
[6 empty] | diagnostic box |  |  |  | 0.14528 | 218.3665 | 441.2893 | 41.094 | 41.094 | 441.3575 | 218.3416 | 579.0789 | 579.1515 |  | 510.6003
BDS | FFS | PM | 5792 | BDS_FFS:PM_D5792 | PM3 | profile monitor | 50 |  |  | 0 | 218.3416 | 441.3575 | 41.094 | 41.094 | 441.3575 | 218.3416 | 579.1515 | 579.1515 |  | 510.6729
[6 empty] | diagnostic box |  |  |  | 0 | 218.3416 | 441.3575 | 41.094 | 41.094 | 441.3575 | 218.3416 | 579.1515 | 579.1515 | 20170217 box (pump,PM), BPM, BCM, GV moved upstream | 510.6729
BDS | FFS | TMP | 5792 | BDS_FFS:TMP_D5792 | TMP_M | turbo pump, 550 l/s |  |  |  | 0 | 218.3416 | 441.3575 | 41.094 | 41.094 | 441.3575 | 218.3416 | 579.1515 | 579.1515 | 20170217 pump 300 -> 550 l/s | 510.6729
[6 empty] | bellow |  |  |  | 0.13122 | 218.3192 | 441.4192 | 41.094 | 41.094 | 441.4808 | 218.2968 | 579.2171 | 579.2827 |  | 510.7385
[10 empty] | 1.0284 | 218.1209 | 441.964 | 41.094 | 41.094 | 442.4472 | 217.945 | 579.797 | 580.3112 |  | 511.3183
BDS | FFS | BPM | 5803 | BDS_FFS:BPM_D5803 | BPM? | position monitor | 75 | 0 |  | 0 | 217.945 | 442.4472 | 41.094 | 41.094 | 442.4472 | 217.945 | 580.3112 | 580.3112 | 20170217 BPM added | 511.8325
[8 empty] | 0.2367 | 0.2367 | 0.1367 | 217.9216 | 442.5115 | 41.094 | 41.094 | 442.5757 | 217.8983 | 580.3795 | 580.4479 |  | 511.9009
BDS | FFS | DCHV | 5805 | BDS_FFS:DCHV_D5805 | C3R | correctors-X,Y | 100 | 0 |  | 0.2 | 217.8641 | 442.6697 | 41.094 | 41.094 | 442.7636 | 217.8299 | 580.5479 | 580.6479 |  | 512.0692
[8 empty] | 0.40506 | 0.40506 | 0.17506 | 217.7999 | 442.8459 | 41.094 | 41.094 | 442.9281 | 217.77 | 580.7354 | 580.8229 |  | 512.2568
BDS | FFS | QH | 5810 | BDS_FFS:QH_D5810 | Q2R | quadrupole-X | 75 |  |  | 0.26 | 217.7255 | 443.0503 | 41.094 | 41.094 | 443.1725 | 217.6811 | 580.9529 | 581.0829 |  | 512.4743
[10 empty] | 0.24 | 217.64 | 443.2852 | 41.094 | 41.094 | 443.398 | 217.599 | 581.2029 | 581.3229 |  | 512.7243
BDS | FFS | QV | 5815 | BDS_FFS:QV_D5815 | Q3R | quadrupole-Y | 75 |  |  | 0.4 | 217.5306 | 443.5859 | 41.094 | 41.094 | 443.7739 | 217.4622 | 581.5229 | 581.7229 |  | 513.0443
[10 empty] | 0.12288 | 217.4412 | 443.8316 | 41.094 | 41.094 | 443.8893 | 217.4201 | 581.7844 | 581.8458 |  | 513.3058
BDS | FFS | BPM | 5818 | BDS_FFS:BPM_D5818 | BPM? | position monitor | 75 | 0 |  | 0 | 217.4201 | 443.8893 | 41.094 | 41.094 | 443.8893 | 217.4201 | 581.8458 | 581.8458 | 20170217 BPM added | 513.3672
[8 empty] | 0.24712 | 0.24712 | 0.11712 | 217.4001 | 443.9444 | 41.094 | 41.094 | 443.9994 | 217.3801 | 581.9044 | 581.9629 |  | 513.4258
BDS | FFS | QH | 5821 | BDS_FFS:QH_D5821 | Q2R | quadrupole-X | 75 |  |  | 0.26 | 217.3356 | 444.1216 | 41.094 | 41.094 | 444.2437 | 217.2912 | 582.0929 | 582.2229 |  | 513.6143
[10 empty] | 0.7264 | 217.1669 | 444.585 | 41.094 | 41.094 | 444.9263 | 217.0427 | 582.5861 | 582.9493 |  | 514.1075
[10 empty] | 1.0236 | 216.8677 | 445.4072 | 41.094 | 41.094 | 445.8882 | 216.6926 | 583.4611 | 583.9729 |  | 514.9825
Target Center | [9 empty] | 0 | 216.6926 | 445.8882 | 41.094 | 41.094 | 445.8882 | 216.6926 | 583.9729 | 583.9729 |  | 515.4943
Note on F34: Author: S0 is an existing solenoid from Artemis B source
Note on C283: Author: to be determined as "QH" or "QV" at a later time
Note on C285: Author: to be determined as "QH" or "QV" at a later time
Note on C287: Author: to be determined as "QH" or "QV" at a later time
Note on C299: Author: to be determined as "QH" or "QV" at a later time
Note on C301: Author: to be determined as "QH" or "QV" at a later time
Note on C303: Author: to be determined as "QH" or "QV" at a later time
Note on C315: Author: to be determined as "QH" or "QV" at a later time
Note on C317: Author: to be determined as "QH" or "QV" at a later time
Note on C319: Author: to be determined as "QH" or "QV" at a later time
Note on C334: Author: to be determined as "QH" or "QV" at a later time
Note on C336: Author: to be determined as "QH" or "QV" at a later time
Note on C338: Author: to be determined as "QH" or "QV" at a later time

## MASTER
Long Run Cable | [6 empty] | Inter-Rack Connections | [9 empty] | Terminating Device
Diagnostics Device | Lattice Beamline Location | Device Type | CableDB Location | Cable Number | Cable Type | Function | Termination Device | Port Number In | Pin Number | CPC pinout | CPC pinout2 | CPC pinout3 | Signal Pin | Shield Pin | Port Number Out | Patch Cable | MTCA Chassis | MTCA Slot | PICO Port | RTM Port | Bias Supply Chassis | Bias Supply Slot | Bias Supply Channel | Notes
PM Flapper | FE_LEBT:PM_D0856 | PM | D0856 | 24F000084 | 4C_22AWG_Multi_006 | X Signal | CPC to LEMO650 Box | CPC-1 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-1A | LEMO650/LEMO650 | 1 | 4 | CH6 - fmc1-3 | 0 | 1 | 1: EBS8005 | Ch0 | 0
PM Flapper | FE_LEBT:PM_D0856 | PM | D0856 | 24F000080 | 4C_22AWG_Multi_006 | Y Signal | CPC to LEMO650 Box | CPC-1 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | LEMO-1B | LEMO650/LEMO650 | 1 | 4 | CH5 - fmc1-2 | 0 | 1 | 1: EBS8005 | Ch0 | 0
PM Flapper | FE_LEBT:PM_D0885 | PM | D0884 | 24F000128 | 4C_22AWG_Multi_006 | X Signal | CPC to LEMO650 Box | CPC-2 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-2A | LEMO650/LEMO650 | 1 | 4 | CH4 - fmc1-1 | 0 | 1 | 1: EBS8005 | Ch0 | 0
PM Flapper | FE_LEBT:PM_D0885 | PM | D0884 | 24F000124 | 4C_22AWG_Multi_006 | Y Signal | CPC to LEMO650 Box | CPC-2 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | LEMO-2B | LEMO650/LEMO650 | 1 | 4 | CH2 - fmc2-3 | 0 | 1 | 1: EBS8005 | Ch0 | 0
PM Flapper | FE_LEBT:PM_D0912 | PM | D0913 | 24F000120 | 4C_22AWG_Multi_006 | X Signal | CPC to LEMO650 Box | CPC-3 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-3A | LEMO650/LEMO650 | 1 | 4 | CH1 - fmc2-2 | 0 | 1 | 1: EBS8005 | Ch0 | 0
PM Flapper | FE_LEBT:PM_D0912 | PM | D0913 | 24F000116 | 4C_22AWG_Multi_006 | Y Signal | CPC to LEMO650 Box | CPC-3 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | LEMO-3B | LEMO650/LEMO650 | 1 | 4 | CH0 - fmc2-1 | 0 | 1 | 1: EBS8005 | Ch0 | 0
PM 6" | FE_SCS2:PM_D0752 | PM | D0752 | 24F000086 | 4C_22AWG_Multi_006 | C Signal | CPC to LEMO650 Box | CPC-4 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-4A | LEMO650/LEMO650 | 1 | 5 | CH6 - fmc1-3 | 0 | 1 | 1: EBS8005 | Ch2 | 0
PM 12" | FE_SCS2:PM_D0752 | PM | D0752 | 24F000140 | 6C_22AWG_Multi_010 | A Signal | CPC to LEMO650 Box | CPC-4 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | LEMO-4B | LEMO650/LEMO650 | 1 | 5 | CH4 - fmc1-1 | 0 | 1 | 1: EBS8005 | Ch2 | 0
PM 12" | FE_SCS2:PM_D0752 | PM | D0752 | 24F000140 | 6C_22AWG_Multi_010 | B Signal | CPC to LEMO650 Box | CPC-4 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-4C | LEMO650/LEMO650 | 1 | 5 | CH5 - fmc1-2 | 0 | 1 | 1: EBS8005 | Ch2 | 0
PM 6" | FE_LEBT:PM_D0783 | PM | D0783 | 24F000060 | 4C_22AWG_Multi_006 | C Signal | CPC to LEMO650 Box | CPC-5 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-5A | LEMO650/LEMO650 | 1 | 5 | CH2 - fmc2-3 | 0 | 1 | 1: EBS8005 | Ch3 | 0
PM 12" | FE_LEBT:PM_D0783 | PM | D0783 | 24F000067 | 6C_22AWG_Multi_010 | A Signal | CPC to LEMO650 Box | CPC-5 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | LEMO-5B | LEMO650/LEMO650 | 1 | 5 | CH0 - fmc2-1 | 0 | 1 | 1: EBS8005 | Ch3 | 0
PM 12" | FE_LEBT:PM_D0783 | PM | D0783 | 24F000067 | 6C_22AWG_Multi_010 | B Signal | CPC to LEMO650 Box | CPC-5 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-5C | LEMO650/LEMO650 | 1 | 5 | CH1 - fmc2-2 | 0 | 1 | 1: EBS8005 | Ch3 | 0
PM 6" | FE_LEBT:PM_D0808 | PM | D0811 | 24F000131 | 4C_22AWG_Multi_006 | C Signal | CPC to LEMO650 Box | CPC-6 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-6A | LEMO650/LEMO650 | 1 | 6 | CH6 - fmc1-3 | 0 | 1 | 1: EBS8005 | Ch4 | 0
PM 12" | FE_LEBT:PM_D0808 | PM | D0811 | 24F000130 | 6C_22AWG_Multi_010 | A Signal | CPC to LEMO650 Box | CPC-6 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | LEMO-6B | LEMO650/LEMO650 | 1 | 6 | CH4 - fmc1-1 | 0 | 1 | 1: EBS8005 | Ch4 | 0
PM 12" | FE_LEBT:PM_D0808 | PM | D0811 | 24F000130 | 6C_22AWG_Multi_010 | B Signal | CPC to LEMO650 Box | CPC-6 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-6C | LEMO650/LEMO650 | 1 | 6 | CH5 - fmc1-2 | 0 | 1 | 1: EBS8005 | Ch4 | 0
PM 6" | FE_LEBT:PM_D0824 | PM | D0824 | 24F000072 | 4C_22AWG_Multi_006 | C Signal | CPC to LEMO650 Box | CPC-7 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-7A | LEMO650/LEMO650 | 1 | 6 | CH2 - fmc2-3 | 0 | 1 | 1: EBS8005 | Ch5 | 0
PM 12" | FE_LEBT:PM_D0824 | PM | D0824 | 24F000066 | 6C_22AWG_Multi_010 | A Signal | CPC to LEMO650 Box | CPC-7 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | LEMO-7B | LEMO650/LEMO650 | 1 | 6 | CH0 - fmc2-1 | 0 | 1 | 1: EBS8005 | Ch5 | 0
PM 12" | FE_LEBT:PM_D0824 | PM | D0824 | 24F000066 | 6C_22AWG_Multi_010 | B Signal | CPC to LEMO650 Box | CPC-7 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-7C | LEMO650/LEMO650 | 1 | 6 | CH1 - fmc2-2 | 0 | 1 | 1: EBS8005 | Ch5 | 0
PM 8" | FE_LEBT:PM_D0998 | PM | D0998 | 24F000142 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-8 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-8A | LEMO650/LEMO650 | 1 | 7 | CH6 - fmc1-3 | 0 | 1 | 1: EBS8005 | Ch6 | 0
PM 8" | FE_LEBT:PM_D0998 | PM | D0998 | 24F000142 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-8 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield-6 | 4 | 6 | LEMO-8B | LEMO650/LEMO650 | 1 | 7 | CH5 - fmc1-2 | 0 | 1 | 1: EBS8005 | Ch6 | 0
PM 8" | FE_LEBT:PM_D0998 | PM | D0998 | 24F000142 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-8 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-8C | LEMO650/LEMO650 | 1 | 7 | CH4 - fmc1-1 | 0 | 1 | 1: EBS8005 | Ch6 | 0
PM 8" | FE_MEBT:PM_D1092 | PM | D1092 | 24F000134 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-9 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-9A | LEMO650/LEMO650 | 1 | 7 | CH2 - fmc2-3 | 0 | 1 | 1: EBS8005 | Ch7 | 0
PM 8" | FE_MEBT:PM_D1092 | PM | D1092 | 24F000134 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-9 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | LEMO-9B | LEMO650/LEMO650 | 1 | 7 | CH1 - fmc2-2 | 0 | 1 | 1: EBS8005 | Ch7 | 0
PM 8" | FE_MEBT:PM_D1092 | PM | D1092 | 24F000134 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-9 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-9C | LEMO650/LEMO650 | 1 | 7 | CH0 - fmc2-1 | 0 | 1 | 1: EBS8005 | Ch7 | 0
Allison Scanner | FE_SCS1:EMS_D0739 | EMS | D0739 | 24E000050 | 2C_75LC_Coax_089 | X Signal | N-Type to SMA Panel | J1 | NA | 0 | 0 | 0 | 0 | 0 | J1 | SMA/LEMO650 | 1 | 3 | CH5 - fmc1-2 | 0 | 0 | 0 | 0 | 0
Allison Scanner | FE_SCS1:EMS_D0739 | EMS | D0739 | 24E000051 | 2C_75LC_Coax_089 | Y Signal | N-Type to SMA Panel | J2 | NA | 0 | 0 | 0 | 0 | 0 | J2 | SMA/LEMO650 | 1 | 3 | CH4 - fmc1-1 | 0 | 0 | 0 | 0 | 0
Allison Scanner | FE_SCS2:EMS_D0718 | EMS | D0718 | 24E000052 | 2C_75LC_Coax_089 | X Signal | N-Type to SMA Panel | J3 | NA | 0 | 0 | 0 | 0 | 0 | J3 | SMA/LEMO650 | 1 | 3 | CH1 - fmc2-2 | 0 | 0 | 0 | 0 | 0
Allison Scanner | FE_SCS2:EMS_D0718 | EMS | D0718 | 24E000053 | 2C_75LC_Coax_089 | Y Signal | N-Type to SMA Panel | J4 | NA | 0 | 0 | 0 | 0 | 0 | J4 | SMA/LEMO650 | 1 | 3 | CH0 - fmc2-1 | 0 | 0 | 0 | 0 | 0
Faraday Cup | FE_SCS2:FC_D0717 | FC | D0718 | 24E000054 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | J5 | NA | 0 | 0 | 0 | 0 | 0 | J5 | SMA/LEMO650 | 1 | 9 | CH0 - fmc2-1 | 0 | 0 | 0 | 0 | 0
Faraday Cup | FE_SCS1:FC_D0738 | FC | D0739 | 24E000049 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | J6 | NA | 0 | 0 | 0 | 0 | 0 | J6 | SMA/LEMO650 | 1 | 9 | CH2 - fmc2-3 | 0 | 0 | 0 | 0 | 0
Faraday Cup | FE_LEBT:FC_D0796 | FC | D0794 | 24E000055 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | J7 | NA | 0 | 0 | 0 | 0 | 0 | J7 | SMA/LEMO650 | 1 | 8 | CH0 - fmc2-1 | 0 | 0 | 0 | 0 | 0
Faraday Cup | FE_LEBT:FC_D0814 | FC | D0811 | 24E000056 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | J8 | NA | 0 | 0 | 0 | 0 | 0 | J8 | SMA/LEMO650 | 1 | 8 | CH1 - fmc2-2 | 0 | 0 | 0 | 0 | 0
Faraday Cup | FE_LEBT:FC_D0977 | FC | D0976 | 24E000057 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | J9 | NA | 0 | 0 | 0 | 0 | 0 | J9 | SMA/LEMO650 | 1 | 8 | CH2 - fmc2-3 | 0 | 0 | 0 | 0 | 0
Faraday Cup | FE_LEBT:FC_D0998 | FC | D0998 | 24E000059 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | J10 | NA | 0 | 0 | 0 | 0 | 0 | J10 | SMA/LEMO650 | 1 | 8 | CH4 - fmc1-1 | 0 | 0 | 0 | 0 | 0
Faraday Cup | FE_BTS:FC_D1104 | FC | D1100 | 24E000063 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | J11 | NA | 0 | 0 | 0 | 0 | 0 | J11 | SMA/LEMO650 | 1 | 8 | CH5 - fmc1-2 | 0 | 0 | 0 | 0 | 0
Faraday Cup | FE_MEBT:FC_D1102 | FC | D1123 | 24E000064 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | J12 | NA | 0 | 0 | 0 | 0 | 0 | J12 | SMA/LEMO650 | 1 | 8 | CH6 - fmc1-3 | 0 | 0 | 0 | 0 | 0
Allison Scanner | FE_SCS1:EMS_D0739 | 0 | D0739 | 24H000011 | 2C_5kVDC_Coax_035 | HV Bias | SHV 6-> 12 Box | 1-A | 0 | 0 | 0 | 0 | 0 | 0 | 1 IN | SHV/SHV | 0 | 0 | 0 | 0 | 1 | 2: EHS8020n | Ch0 | 0
Allison Scanner | FE_SCS1:EMS_D0739 | 0 | D0739 | 24H000012 | 2C_5kVDC_Coax_035 | HV Bias | SHV 6-> 12 Box | 1-B | 0 | 0 | 0 | 0 | 0 | 0 | 1 IN | SHV/SHV | 0 | 0 | 0 | 0 | 1 | 2: EHS8020n | Ch0 | 0
Allison Scanner | FE_SCS2:EMS_D0718 | 0 | D0718 | 24H000002 | 2C_5kVDC_Coax_035 | HV Bias | SHV 6-> 12 Box | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SHV/SHV | 0 | 0 | 0 | 0 | 1 | 2: EHS8020n | Ch1 | 0
Allison Scanner | FE_SCS2:EMS_D0718 | 0 | D0718 | 24H000003 | 2C_5kVDC_Coax_035 | HV Bias | SHV 6-> 12 Box | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SHV/SHV | 0 | 0 | 0 | 0 | 1 | 2: EHS8020n | Ch1 | 0
Faraday Cup | FE_SCS2:FC_D0717 | 0 | D0718 | 24H000007 | 2C_5kVDC_Coax_035 | HV Bias | ISEG - EHS8020n | Ch2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2: EHS8020n | Ch2 | 0
Faraday Cup | FE_SCS2:FC_D0717 | 0 | D0718 | 24H000008 | 2C_5kVDC_Coax_035 | HV Bias | 0 | no connection | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Faraday Cup | FE_SCS1:FC_D0738 | 0 | D0739 | 24H000009 | 2C_5kVDC_Coax_035 | HV Bias | SHV 6-> 12 Box | 4-A | 0 | 0 | 0 | 0 | 0 | 0 | 4 IN | SHV/SHV | 0 | 0 | 0 | 0 | 1 | 2: EHS8020n | Ch3 | 0
Faraday Cup | FE_LEBT:FC_D0814 | 0 | D0811 | 24H000013 | 2C_5kVDC_Coax_035 | HV Bias | SHV 6-> 12 Box | 4-B | 0 | 0 | 0 | 0 | 0 | 0 | 4 IN | SHV/SHV | 0 | 0 | 0 | 0 | 1 | 2: EHS8020n | Ch3 | 0
Faraday Cup | FE_LEBT:FC_D0977 | 0 | D0976 | 24H000015 | 2C_5kVDC_Coax_035 | HV Bias | SHV 6-> 12 Box | 5-A | 0 | 0 | 0 | 0 | 0 | 0 | 5 IN | SHV/SHV | 0 | 0 | 0 | 0 | 1 | 2: EHS8020n | Ch4 | 0
Faraday Cup | FE_LEBT:FC_D0977 | 0 | D0976 | 24H000016 | 2C_5kVDC_Coax_035 | HV Bias | 0 | no connection | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Faraday Cup | FE_LEBT:FC_D0998 | 0 | D0998 | 24H000019 | 2C_5kVDC_Coax_035 | HV Bias | SHV 6-> 12 Box | 3-A | 0 | 0 | 0 | 0 | 0 | 0 | 3 IN | SHV/SHV | 0 | 0 | 0 | 0 | 1 | 2: EHS8020n | Ch5 | 0
Faraday Cup | FE_LEBT:FC_D0998 | 0 | D0998 | 24H000020 | 2C_5kVDC_Coax_035 | HV Bias | 0 | no connection | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Faraday Cup | FE_BTS:FC_D1104 | 0 | D1100 | 24H000021 | 2C_5kVDC_Coax_035 | HV Bias | SHV 6-> 12 Box | 5-B | 0 | 0 | 0 | 0 | 0 | 0 | 5 IN | SHV/SHV | 0 | 0 | 0 | 0 | 1 | 2: EHS8020n | Ch4 | 0
Faraday Cup | FE_BTS:FC_D1104 | 0 | D1100 | 24H000022 | 2C_5kVDC_Coax_035 | HV Bias | 0 | no connection | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Faraday Cup | FE_MEBT:FC_D1102 | 0 | D1123 | 24H000023 | 2C_5kVDC_Coax_035 | HV Bias | SHV 6-> 12 Box | 3-B | 0 | 0 | 0 | 0 | 0 | 0 | 3 IN | SHV/SHV | 0 | 0 | 0 | 0 | 1 | 2: EHS8020n | Ch5 | 0
Faraday Cup | FE_MEBT:FC_D1102 | 0 | D1123 | 24H000026 | 2C_5kVDC_Coax_035 | HV Bias | 0 | no connection | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Fast Faraday Cup | FE_LEBT:FFC_D0998 | 0 | D0998 | 24H000017 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Fast Faraday Cup | FE_LEBT:FFC_D0998 | 0 | D0998 | 24H000018 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Fast Faraday Cup | FE_MEBT:FFC_D1102 | 0 | D1120 | 24H000024 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Fast Faraday Cup | FE_MEBT:FFC_D1102 | 0 | D1120 | 24H000025 | 2C_5kVDC_Coax_035 | HV Bias | SHV 6-> 12 Box | 2-A | 0 | 0 | 0 | 0 | 0 | 0 | 2IN | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 1 | 2: EBS8003 | Ch1 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | SHV 6-> 12 Box (modded) | 6IN | 0 | 0 | 0 | 0 | 0 | 0 | 6-A | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 1 | 2: EBS8003 | Ch0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | SHV 6-> 12 Box (modded) | 6IN | 0 | 0 | 0 | 0 | 0 | 0 | 6-B | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 1 | 2: EBS8003 | Ch0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 6-A | 0 | 0 | 0 | 0 | 0 | 0 | 6-A | SHV/Lemo650 | 1 | 4 | CH7 - fmc1-4 | 0 | 1 | 2: EBS8003 | Ch0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 6-B | 0 | 0 | 0 | 0 | 0 | 0 | 6-B | SHV/Lemo650 | 1 | 4 | CH3 - fmc2-4 | 0 | 1 | 2: EBS8003 | Ch0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 3_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 1 | 2: EBS8003 | Ch2 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 3_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 1 | 5 | CH7 - fmc1-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 4_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 1 | 2: EBS8003 | Ch3 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 4_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 1 | 5 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 5_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 1 | 2: EBS8003 | Ch4 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 5_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 1 | 6 | CH7 - fmc1-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 6_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 1 | 2: EBS8003 | Ch5 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 6_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 1 | 6 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 7_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 1 | 2: EBS8003 | Ch6 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 7_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 1 | 7 | CH7 - fmc1-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 8_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 1 | 2: EBS8003 | Ch7 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 8_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 1 | 7 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Grounding Plug | 0 | ground | 0 | - | Lemo650 Plug | Ground | Pico-8 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Ground Plug | 1 | 3 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Grounding Plug | 0 | ground | 0 | - | Lemo650 Plug | Ground | Pico-8 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Ground Plug | 1 | 3 | CH7 - fmc1-4 | 0 | 0 | 0 | 0 | 0
Grounding Plug | 0 | ground | 0 | - | Lemo650 Plug | Ground | Pico-8 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Ground Plug | 1 | 8 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Grounding Plug | 0 | ground | 0 | - | Lemo650 Plug | Ground | Pico-8 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Ground Plug | 1 | 8 | CH7 - fmc1-4 | 0 | 0 | 0 | 0 | 0
Grounding Plug | 0 | ground | 0 | - | Lemo650 Plug | Ground | Pico-8 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Ground Plug | 1 | 9 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Grounding Plug | 0 | ground | 0 | - | Lemo650 Plug | Ground | Pico-8 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Ground Plug | 1 | 9 | CH7 - fmc1-4 | 0 | 0 | 0 | 0 | 0
BPM | FE_MEBT:BPM_D1056 | BPM | D1056 | 24E000015 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 2 | 3 | 0 | 1 | 0 | 0 | 0 | 0
BPM | FE_MEBT:BPM_D1056 | BPM | D1056 | 24E000016 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 2 | 3 | 0 | 2 | 0 | 0 | 0 | 0
BPM | FE_MEBT:BPM_D1056 | BPM | D1056 | 24E000017 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 2 | 3 | 0 | 3 | 0 | 0 | 0 | 0
BPM | FE_MEBT:BPM_D1056 | BPM | D1056 | 24E000018 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 2 | 3 | 0 | 4 | 0 | 0 | 0 | 0
BPM | FE_MEBT:RFS_D1056 | BPM | D1056 | 24E000019 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 2 | 3 | 0 | Ref | 0 | 0 | 0 | 0
BPM | FE_MEBT:BPM_D1072 | BPM | D1073 | 24E000011 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 2 | 3 | 0 | 5 | 0 | 0 | 0 | 0
BPM | FE_MEBT:BPM_D1072 | BPM | D1073 | 24E000012 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 2 | 3 | 0 | 6 | 0 | 0 | 0 | 0
BPM | FE_MEBT:BPM_D1072 | BPM | D1073 | 24E000013 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 2 | 3 | 0 | 7 | 0 | 0 | 0 | 0
BPM | FE_MEBT:BPM_D1072 | BPM | D1073 | 24E000014 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 2 | 3 | 0 | 8 | 0 | 0 | 0 | 0
BPM | FE_MEBT:RFS_D1072 | BPM | D1073 | 24E000010 | 2C_1/4in_Hardline_036 | Ref | No termination | No termination | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | FE_MEBT:BPM_D1094 | BPM | D1095 | 24E000005 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 2 | 4 | 0 | 1 | 0 | 0 | 0 | 0
BPM | FE_MEBT:BPM_D1094 | BPM | D1095 | 24E000006 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 2 | 4 | 0 | 2 | 0 | 0 | 0 | 0
BPM | FE_MEBT:BPM_D1094 | BPM | D1095 | 24E000007 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 2 | 4 | 0 | 3 | 0 | 0 | 0 | 0
BPM | FE_MEBT:BPM_D1094 | BPM | D1095 | 24E000008 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 2 | 4 | 0 | 4 | 0 | 0 | 0 | 0
BPM | FE_MEBT:RFS_D1094 | BPM | D1095 | 24E000009 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 2 | 4 | 0 | Ref | 0 | 0 | 0 | 0
BPM | FE_MEBT:BPM_D1111 | BPM | D1113 | 24E000001 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 2 | 4 | 0 | 5 | 0 | 0 | 0 | 0
BPM | FE_MEBT:BPM_D1111 | BPM | D1113 | 24E000002 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 2 | 4 | 0 | 6 | 0 | 0 | 0 | 0
BPM | FE_MEBT:BPM_D1111 | BPM | D1113 | 24E000003 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 2 | 4 | 0 | 7 | 0 | 0 | 0 | 0
BPM | FE_MEBT:BPM_D1111 | BPM | D1113 | 24E000004 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 2 | 4 | 0 | 8 | 0 | 0 | 0 | 0
BPM | FE_MEBT:RFS_D1111 | BPM | D1113 | 24E000000 | 2C_1/4in_Hardline_036 | Ref | No termination | No termination | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS1_CA01:BPM_D1129 | BPM | D1129 | 34E000104 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 2 | 5 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS1_CA01:BPM_D1129 | BPM | D1129 | 34E000105 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 2 | 5 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS1_CA01:BPM_D1129 | BPM | D1129 | 34E000106 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 2 | 5 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS1_CA01:BPM_D1129 | BPM | D1129 | 34E000107 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 2 | 5 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS1_CA01:RFS_D1129 | BPM | D1129 | 34E000109 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 2 | 5 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS1_CA01:BPM_D1144 | BPM | D1144 | 34E000100 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 2 | 5 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS1_CA01:BPM_D1144 | BPM | D1144 | 34E000101 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 2 | 5 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS1_CA01:BPM_D1144 | BPM | D1144 | 34E000102 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 2 | 5 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS1_CA01:BPM_D1144 | BPM | D1144 | 34E000103 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 2 | 5 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS1_CA01:RFS_D1144 | BPM | D1144 | 34E000108 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No terminiation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS1_WA01:BPM_D1155 | BPM | D1155 | 34E000391 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 2 | 6 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS1_WA01:BPM_D1155 | BPM | D1155 | 34E000392 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 2 | 6 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS1_WA01:BPM_D1155 | BPM | D1155 | 34E000393 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 2 | 6 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS1_WA01:BPM_D1155 | BPM | D1155 | 34E000394 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 2 | 6 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS1_WA01:RFS_D1155 | BPM | D1155 | 34E000306 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 2 | 6 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS1_CA02:BPM_D1163 | BPM | D1163 | 34E000343 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 2 | 6 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS1_CA02:BPM_D1163 | BPM | D1163 | 34E000344 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 2 | 6 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS1_CA02:BPM_D1163 | BPM | D1163 | 34E000345 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 2 | 6 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS1_CA02:BPM_D1163 | BPM | D1163 | 34E000346 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 2 | 6 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS1_CA02:RFS_D1163 | BPM | D1163 | 34E000305 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No terminiation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS1_CA02:BPM_D1177 | BPM | D1178 | 34E000339 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 2 | 7 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS1_CA02:BPM_D1177 | BPM | D1178 | 34E000340 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 2 | 7 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS1_CA02:BPM_D1177 | BPM | D1178 | 34E000341 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 2 | 7 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS1_CA02:BPM_D1177 | BPM | D1178 | 34E000342 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 2 | 7 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS1_CA02:RFS_D1177 | BPM | D1178 | 34E000304 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No terminiation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS1_WA02:BPM_D1188 | BPM | D1189 | 34E000387 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 2 | 7 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS1_WA02:BPM_D1188 | BPM | D1189 | 34E000388 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 2 | 7 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS1_WA02:BPM_D1188 | BPM | D1189 | 34E000389 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 2 | 7 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS1_WA02:BPM_D1188 | BPM | D1189 | 34E000390 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 2 | 7 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS1_WA02:RFS_D1188 | BPM | D1189 | 34E000303 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 2 | 7 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS1_CA03:BPM_D1196 | BPM | D1196 | 34E000335 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 2 | 8 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS1_CA03:BPM_D1196 | BPM | D1196 | 34E000336 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 2 | 8 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS1_CA03:BPM_D1196 | BPM | D1196 | 34E000337 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 2 | 8 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS1_CA03:BPM_D1196 | BPM | D1196 | 34E000338 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 2 | 8 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS1_CA03:RFS_D1196 | BPM | D1196 | 34E000302 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 2 | 8 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS1_CA03:BPM_D1211 | BPM | D1212 | 34E000331 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 2 | 8 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS1_CA03:BPM_D1211 | BPM | D1212 | 34E000332 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 2 | 8 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS1_CA03:BPM_D1211 | BPM | D1212 | 34E000333 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 2 | 8 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS1_CA03:BPM_D1211 | BPM | D1212 | 34E000334 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 2 | 8 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS1_CA03:RFS_D1211 | BPM | D1212 | 34E000301 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No terminiation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | FS2_BMS:BPM_D4216 | BPM | D4223 | 44E000065 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 2 | 9 | 0 | 1 | 0 | 0 | 0 | 0
BPM | FS2_BMS:BPM_D4216 | BPM | D4223 | 44E000066 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 2 | 9 | 0 | 2 | 0 | 0 | 0 | 0
BPM | FS2_BMS:BPM_D4216 | BPM | D4223 | 44E000067 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 2 | 9 | 0 | 3 | 0 | 0 | 0 | 0
BPM | FS2_BMS:BPM_D4216 | BPM | D4223 | 44E000068 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 2 | 9 | 0 | 4 | 0 | 0 | 0 | 0
BPM | FS2_BMS:RFS_D4216 | BPM | D4223 | 44E000073 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No terminiation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | FS2_BMS:BPM_D4283 | BPM | D4283 | 44E000061 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 2 | 9 | 0 | 5 | 0 | 0 | 0 | 0
BPM | FS2_BMS:BPM_D4283 | BPM | D4283 | 44E000062 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 2 | 9 | 0 | 6 | 0 | 0 | 0 | 0
BPM | FS2_BMS:BPM_D4283 | BPM | D4283 | 44E000063 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 2 | 9 | 0 | 7 | 0 | 0 | 0 | 0
BPM | FS2_BMS:BPM_D4283 | BPM | D4283 | 44E000064 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 2 | 9 | 0 | 8 | 0 | 0 | 0 | 0
BPM | FS2_BMS:RFS_D4283 | BPM | D4283 | 44E000072 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 2 | 9 | 0 | Ref | 0 | 0 | 0 | 0
BPM | FS2_BMS:BPM_D4326 | BPM | D4327 | 44E000059 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 2 | 10 | 0 | 1 | 0 | 0 | 0 | 0
BPM | FS2_BMS:BPM_D4326 | BPM | D4327 | 44E000069 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 2 | 10 | 0 | 2 | 0 | 0 | 0 | 0
BPM | FS2_BMS:BPM_D4326 | BPM | D4327 | 44E000070 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 2 | 10 | 0 | 3 | 0 | 0 | 0 | 0
BPM | FS2_BMS:BPM_D4326 | BPM | D4327 | 44E000071 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 2 | 10 | 0 | 4 | 0 | 0 | 0 | 0
BPM | FS2_BMS:RFS_D4326 | BPM | D4327 | 44E000060 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No terminiation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS3_WD01:BPM_D4389 | BPM | D4389 | 34E000092 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 2 | 10 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS3_WD01:BPM_D4389 | BPM | D4389 | 34E000093 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 2 | 10 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS3_WD01:BPM_D4389 | BPM | D4389 | 34E000094 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 2 | 10 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS3_WD01:BPM_D4389 | BPM | D4389 | 34E000095 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 2 | 10 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS3_WD01:RFS_D4389 | BPM | D4389 | 34E000099 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 2 | 10 | 0 | Ref | 0 | 0 | 0 | 0
Temp Pico | LS1_CA01:HMR_D0001 | HMR | 0 | 0 | 0 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to lemo | 2 | 11 | CH0 - fmc2-1 | 0 | 0 | 0 | 0 | 0
Temp Pico | LS1_CA01:HMR_D0002 | HMR | 0 | 0 | 0 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to lemo | 2 | 11 | CH1 - fmc2-2 | 0 | 0 | 0 | 0 | 0
Temp Pico | LS1_CA01:HMR_D0003 | HMR | 0 | 0 | 0 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to lemo | 2 | 11 | CH2 - fmc2-3 | 0 | 0 | 0 | 0 | 0
Temp Pico | LS1_CA01:HMR_D0004 | HMR | 0 | 0 | 0 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to lemo | 2 | 11 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Temp Pico | LS1_CA01:HMR_D0005 | HMR | 0 | 0 | 0 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to lemo | 2 | 11 | CH4 - fmc1-1 | 0 | 0 | 0 | 0 | 0
Temp Pico | LS1_CA01:HMR_D0006 | HMR | 0 | 0 | 0 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to lemo | 2 | 11 | CH5 - fmc1-2 | 0 | 0 | 0 | 0 | 0
Temp Pico | LS1_CA01:HMR_D0007 | HMR | 0 | 0 | 0 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to lemo | 2 | 11 | CH6 - fmc1-3 | 0 | 0 | 0 | 0 | 0
Temp Pico | LS1_CA01:HMR_D0008 | HMR | 0 | 0 | 0 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to lemo | 2 | 11 | CH7 - fmc1-4 | 0 | 0 | 0 | 0 | 0
Grounding Plug | 0 | ground | 0 | - | Lemo650 Plug | Ground | Pico-8 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Ground Plug | 1 | 11 | CH3 - fmc2-4 | 0
Grounding Plug | 0 | ground | 0 | - | Lemo650 Plug | Ground | Pico-8 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Ground Plug | 1 | 11 | CH7 - fmc1-4 | 0
Ion Chamber | FS2_BMS:IC_D4141 | IC | D4141 | 44E000237 | 2C_1kV75LC_Coax_090 | Signal | N-Type to Lemo Box | J1 | 0 | 0 | 0 | 0 | 0 | 0 | J1 | Lemo to Lemo | 3 | 3 | CH0 - fmc2-1 | 0 | 0 | 0 | 0 | 0
Ion Chamber | FS2_BMS:IC_D4180 | IC | D4180 | 44E000238. | 2C_1kV75LC_Coax_090 | Signal | N-Type to Lemo Box | J2 | 0 | 0 | 0 | 0 | 0 | 0 | J2 | Lemo to Lemo | 3 | 3 | CH1 - fmc2-2 | 0 | 0 | 0 | 0 | 0
Ion Chamber | FS2_BMS:IC_D4221 | IC | D4221 | 44E000239 | 2C_1kV75LC_Coax_090 | Signal | N-Type to Lemo Box | J3 | 0 | 0 | 0 | 0 | 0 | 0 | J3 | Lemo to Lemo | 3 | 3 | CH2 - fmc2-3 | 0 | 0 | 0 | 0 | 0
Ion Chamber | FS2_BMS:IC_D4277 | IC | D4277 | 44E000240 | 2C_1kV75LC_Coax_090 | Signal | N-Type to Lemo Box | J4 | 0 | 0 | 0 | 0 | 0 | 0 | J4 | Lemo to Lemo | 3 | 3 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Ion Chamber | FS2_BMS:IC_D4305 | IC | D4305 | 44E000241 | 2C_1kV75LC_Coax_090 | Signal | N-Type to Lemo Box | J5 | 0 | 0 | 0 | 0 | 0 | 0 | J5 | Lemo to Lemo | 3 | 3 | CH4 - fmc1-1 | 0 | 0 | 0 | 0 | 0
Ion Chamber | LS3_CD01:IC_D4385 | IC | D4385 | 34E000558 | 2C_1kV75LC_Coax_090 | Signal | N-Type to Lemo Box | J6 | 0 | 0 | 0 | 0 | 0 | 0 | J6 | Lemo to Lemo | 3 | 3 | CH5 - fmc1-2 | 0 | 0 | 0 | 0 | 0
Ion Chamber | LS3_CD03:IC_D4503 | IC | D4503 | 34E000559 | 2C_1kV75LC_Coax_090 | Signal | N-Type to Lemo Box | J7 | 0 | 0 | 0 | 0 | 0 | 0 | J7 | Lemo to Lemo | 3 | 3 | CH6 - fmc1-3 | 0 | 0 | 0 | 0 | 0
Ion Chamber | FS2_BMS:IC_D4141 | 0 | D4141 | 44H000021 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | SHV to SHV | 0 | 0 | 0 | 0 | 2 | 1 : EHS8020n | Ch0 | 0
Ion Chamber | FS2_BMS:IC_D4181 | 0 | D4181 | 44H000022 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | SHV to SHV | 0 | 0 | 0 | 0 | 2 | 1 : EHS8020n | Ch1 | 0
Ion Chamber | FS2_BMS:IC_D4221 | 0 | D4221 | 44H000023 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | SHV to SHV | 0 | 0 | 0 | 0 | 2 | 1 : EHS8020n | Ch2 | 0
Ion Chamber | FS2_BMS:HMR_D4277 | 0 | D4277 | 44H000024 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | SHV to SHV | 0 | 0 | 0 | 0 | 2 | 1 : EHS8020n | Ch3 | 0
Ion Chamber | FS2_BMS:HMR_D4305 | 0 | D4305 | 44H000025 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 5 | SHV to SHV | 0 | 0 | 0 | 0 | 2 | 1 : EHS8020n | Ch4 | 0
Ion Chamber | LS3_CD01:HMR_D4385 | 0 | D4385 | 34H000076 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 6 | 0 | 0 | 0 | 0 | 0 | 0 | 6 | SHV to SHV | 0 | 0 | 0 | 0 | 2 | 1 : EHS8020n | Ch5 | 0
Faraday Cup | FS2_BMS:FC_D4253 | 0 | D4253 | 34H000077 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 8 | 0 | 0 | 0 | 0 | 0 | 0 | 8 | SHV to SHV | 0 | 0 | 0 | 0 | 2 | 1 : EHS8020n | Ch7 | 0
Faraday Cup | LS1_WA03:FC_D1224 | FC | D1224 | 34H000007 | 2C_1kV75LC_Coax_090 | Signal | N-Type to SMA Panel | J1 | 0 | 0 | 0 | 0 | 0 | 0 | J1 | Lemo250/Lemo650 | 3 | 4 | CH0 - fmc2-1 | 0 | 2 | 1: EBS8003 | Ch0 | 1:8 current divider installed
Faraday Cup | FS2_BMS:FC_D4253 | FC | D4253 | 44H000001 | 2C_1kV75LC_Coax_090 | Signal | N-Type to SMA Panel | J2 | 0 | 0 | 0 | 0 | 0 | 0 | J2 | Lemo250/Lemo650 | 3 | 5 | CH0 - fmc2-1 | 0 | 2 | 1 : EHS8020n | Ch7 | 0
Halo Monitor Ring,Spare | FS2_BMS:HMR_D4144 | HMR - Unused | D4144 | 44H000000 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | FS2_BMS:HMR_D4144 | 0 | D4144 | 44H000055 | 2C_5kVDC_Coax_035 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring,Spare | FS2_BMS:HMR_D4178 | HMR - Unused | D4179 | 44H000001 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | FS2_BMS:HMR_D4178 | 0 | D4179 | 44H000054 | 2C_5kVDC_Coax_035 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring | FS2_BMS:HMR_D4328 | HMR | D4328 | 44H000002 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J2 | 0 | 0 | 0 | 0 | 0 | 0 | J2 | SHV to Lemo | 3 | 4 | CH4 - fmc1-1 | 0 | 2 | 1: EBS8003 | Ch1 | 0
0 | FS2_BMS:HMR_D4328 | 0 | D4328 | 44H000057 | 2C_5kVDC_Coax_035 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring,Spare | LS3_WD01:HMR_D4390 | HMR - Unused | D4390 | 34H000009 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS3_WD01:HMR_D4390 | 0 | D4390 | 34H000100 | 2C_5kVDC_Coax_035 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring,Spare | LS3_WD02:HMR_D4452 | HMR - Unused | D4452 | 34H000008 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS3_WD02:HMR_D4452 | 0 | D4452 | 34H000101 | 2C_5kVDC_Coax_035 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS3_WD03:HMR_D4514 | 0 | D4514 | 34H000086 | 2C_5kVDC_Coax_035 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring | LS3_WD04:HMR_D4576 | HMR | D4576 | 34H000010 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J3 | 0 | 0 | 0 | 0 | 0 | 0 | J3 | SHV to Lemo | 3 | 4 | CH5 - fmc1-2 | 0 | 2 | 1: EBS8003 | Ch1 | 0
0 | LS3_WD04:HMR_D4576 | 0 | D4576 | 34H000099 | 2C_5kVDC_Coax_035 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring,Spare | LS3_WD05:HMR_D4638 | HMR - Unused | D4638 | 34H000011 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS3_WD05:HMR_D4638 | 0 | D4638 | 34H000095 | 2C_5kVDC_Coax_035 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 1_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 2 | 1: EBS8003 | Ch0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 1_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 3 | 4 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 2_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 2 | 1: EBS8003 | Ch1 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 2_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 3 | 4 | CH7 - fmc1-4 | 0 | 0 | 0 | 0 | 0
Grounding Plug | 0 | ground | 0 | - | Lemo650 Plug | Ground | Pico-8 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Ground Plug | 3 | 5 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Neutron Detector | FE_RFQ:ND_D1025 | ND | D1025 | 24E000066 | 2C_75LC_Coax_089 | Signal | N-type to Lemo Box | J1 | 0 | 0 | 0 | 0 | 0 | 0 | J1 | Lemo to Lemo | 4 | 3 | CH0 - fmc2-1 | 0 | 0 | 0 | 0 | 0
Neutron Detector | FE_RFQ:ND_D1025 | 0 | D1025 | 24E000067 | 2C_75LC_Coax_089 | Test Signal | N-type to SMA Panel | J1 | 0 | 0 | 0 | 0 | 0 | 0 | J1 | SMA to Ferrule | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS1_CA01:ND_D1128 | ND | D1128 | 34E000522 | 2C_75LC_Coax_089 | Signal | N-type to Lemo Box | J2 | 0 | 0 | 0 | 0 | 0 | 0 | J2 | Lemo to Lemo | 4 | 3 | CH1 - fmc2-2 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS1_CA01:ND_D1128 | 0 | D1128 | 34E000521 | 2C_75LC_Coax_089 | Test signal | N-type to SMA Panel | J2 | 0 | 0 | 0 | 0 | 0 | 0 | J2 | SMA to Ferrule | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS1_CA03:ND_D1195 | ND | D1195 | 34E000524 | 2C_75LC_Coax_089 | Signal | N-type to Lemo Box | J3 | 0 | 0 | 0 | 0 | 0 | 0 | J3 | Lemo to Lemo | 4 | 3 | CH2 - fmc2-3 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS1_CA03:ND_D1195 | 0 | D1195 | 34E000523 | 2C_75LC_Coax_089 | Test signal | N-type to SMA Panel | J3 | 0 | 0 | 0 | 0 | 0 | 0 | J3 | SMA to Ferrule | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS1_CB01:ND_D1250 | ND | D1250 | 34E000520 | 2C_75LC_Coax_089 | Signal | N-type to Lemo Box | J4 | 0 | 0 | 0 | 0 | 0 | 0 | J4 | Lemo to Lemo | 4 | 3 | CH4 - fmc1-1 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS1_CB01:ND_D1250 | 0 | D1250 | 34E000519 | 2C_75LC_Coax_089 | Test signal | N-type to SMA Panel | J4 | 0 | 0 | 0 | 0 | 0 | 0 | J4 | SMA to Ferrule | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS1_CB03:ND_D1378 | ND | D1378 | 34E000518 | 2C_75LC_Coax_089 | Signal | N-type to Lemo Box | J5 | 0 | 0 | 0 | 0 | 0 | 0 | J5 | Lemo to Lemo | 4 | 3 | CH5 - fmc1-2 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS1_CB03:ND_D1378 | 0 | D1378 | 34E000517 | 2C_75LC_Coax_089 | Test signal | N-type to SMA Panel | J5 | 0 | 0 | 0 | 0 | 0 | 0 | J5 | SMA to Ferrule | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS3_CD06:ND_D4647 | ND | D4647 | 34E000516 | 2C_75LC_Coax_089 | Signal | N-type to Lemo Box | J6 | 0 | 0 | 0 | 0 | 0 | 0 | J6 | Lemo to Lemo | 4 | 3 | CH6 - fmc1-3 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS3_CD06:ND_D4647 | 0 | D4647 | 34E000515 | 2C_75LC_Coax_089 | Test signal | N-type to SMA Panel | J6 | 0 | 0 | 0 | 0 | 0 | 0 | J6 | SMA to Ferrule | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Neutron Detector | FE_RFQ:ND_D1025 | 0 | D1025 | 24H000006 | 2C_5kVDC_Coax_035 | HV Bias | Direct to ISEG | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 2: EHS8020n | Ch0 | 0
Neutron Detector | LS1_CA01:ND_D1128 | 0 | D1128 | 34H000057 | 2C_5kVDC_Coax_035 | HV Bias | Direct to ISEG | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 2: EHS8020n | Ch1 | 0
Neutron Detector | LS1_CA03:ND_D1195 | 0 | D1195 | 34H000058 | 2C_5kVDC_Coax_035 | HV Bias | Direct to ISEG | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 2: EHS8020n | Ch2 | 0
Neutron Detector | LS1_CB01:ND_D1250 | 0 | D1250 | 34H000056 | 2C_5kVDC_Coax_035 | HV Bias | Direct to ISEG | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 2: EHS8020n | Ch3 | 0
Neutron Detector | LS1_CB03:ND_D1378 | 0 | D1378 | 34H000055 | 2C_5kVDC_Coax_035 | HV Bias | Direct to ISEG | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 2: EHS8020n | Ch4 | 0
Neutron Detector | LS3_CD06:ND_D4647 | 0 | D4647 | 34H000054 | 2C_5kVDC_Coax_035 | HV Bias | Direct to ISEG | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 2: EHS8020n | Ch5 | 0
Halo Monitor Ring | LS1_WA01:HMR_D1156 | HMR | D1156 | 34H000029 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J1 | 0 | 0 | 0 | 0 | 0 | 0 | J1 | SHV to Lemo | 4 | 4 | CH0 - fmc2-1 | 0 | 3 | 1: EBS8003 | Ch0 | 0
0 | LS1_WA01:HMR_D1156 | 0 | D1156 | 34H000110 | 2C_5kVDC_Coax_035 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring | LS1_WA02:HMR_D1190 | HMR | D1190 | 34H000024 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J2 | 0 | 0 | 0 | 0 | 0 | 0 | J2 | SHV to Lemo | 4 | 4 | CH1 - fmc2-2 | 0 | 3 | 1: EBS8003 | Ch0 | 0
0 | LS1_WA02:HMR_D1190 | 0 | D1190 | 34H000114 | 2C_5kVDC_Coax_035 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring | LS1_WA03:HMR_D1224 | HMR | D1223 | 34H000026 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J3 | 0 | 0 | 0 | 0 | 0 | 0 | J3 | SHV to Lemo | 4 | 4 | CH2 - fmc2-3 | 0 | 3 | 1: EBS8003 | Ch0 | Spliter Installed
0 | LS1_WA03:HMR_D1224 | 0 | D1223 | 34H000109 | 2C_5kVDC_Coax_035 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring | LS1_WB01:HMR_D1287 | HMR | D1287 | 34H000027 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J4 | 0 | 0 | 0 | 0 | 0 | 0 | J4 | SHV to Lemo | 4 | 4 | CH4 - fmc1-1 | 0 | 3 | 1: EBS8003 | Ch1 | 0
0 | LS1_WB01:HMR_D1287 | 0 | D1287 | 34H000111 | 2C_5kVDC_Coax_035 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring | LS1_WB02:HMR_D1351 | HMR | D1351 | 34H000028 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J5 | 0 | 0 | 0 | 0 | 0 | 0 | J5 | SHV to Lemo | 4 | 4 | CH5 - fmc1-2 | 0 | 3 | 1: EBS8003 | Ch1 | 0
0 | LS1_WB02:HMR_D1351 | 0 | D1351 | 34H000112 | 2C_5kVDC_Coax_035 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring | LS1_WB03:HMR_D1415 | HMR | D1415 | 34H000025 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J6 | 0 | 0 | 0 | 0 | 0 | 0 | J6 | SHV to Lemo | 4 | 4 | CH6 - fmc1-3 | 0 | 3 | 1: EBS8003 | Ch1 | Spliter Installed
0 | LS1_WB03:HMR_D1415 | 0 | D1415 | 34H000113 | 2C_5kVDC_Coax_035 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS1_WA03:BPM_D1222 | BPM | D1222 | 34E000383 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 4 | 5 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS1_WA03:BPM_D1222 | BPM | D1222 | 34E000384 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 4 | 5 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS1_WA03:BPM_D1222 | BPM | D1222 | 34E000385 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 4 | 5 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS1_WA03:BPM_D1222 | BPM | D1222 | 34E000386 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 4 | 5 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS1_WA03:BPM_D1222 | BPM | D1222 | 34E000300 | 2C_1/4in_Hardline_036 | RF Pick-up | 0 | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 4 | 5 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS1_CB01:BPM_D1231 | BPM | D1231 | 34E000323 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 4 | 5 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS1_CB01:BPM_D1231 | BPM | D1231 | 34E000324 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 4 | 5 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS1_CB01:BPM_D1231 | BPM | D1231 | 34E000325 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 4 | 5 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS1_CB01:BPM_D1231 | BPM | D1231 | 34E000326 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 4 | 5 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS1_CB01:BPM_D1231 | BPM | D1231 | 34E000299 | 2C_1/4in_Hardline_036 | RF Pick-up | 0 | Ref | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS1_CB01:BPM_D1251 | BPM | D1251 | 34E000312 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 4 | 6 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS1_CB01:BPM_D1251 | BPM | D1251 | 34E000313 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 4 | 6 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS1_CB01:BPM_D1251 | BPM | D1251 | 34E000314 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 4 | 6 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS1_CB01:BPM_D1251 | BPM | D1251 | 34E000315 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 4 | 6 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS1_CB01:BPM_D1251 | BPM | D1251 | 34E000298 | 2C_1/4in_Hardline_036 | RF Pick-up | 0 | Ref | 0 | 0 | 0 | 0 | 0 | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0 | 0 | 0 | 0
BPM | LS1_CB01:BPM_D1271 | BPM | D1271 | 34E000308 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 4 | 6 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS1_CB01:BPM_D1271 | BPM | D1271 | 34E000309 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 4 | 6 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS1_CB01:BPM_D1271 | BPM | D1271 | 34E000310 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 4 | 6 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS1_CB01:BPM_D1271 | BPM | D1271 | 34E000311 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 4 | 6 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS1_CB01:BPM_D1271 | BPM | D1271 | 34E000297 | 2C_1/4in_Hardline_036 | RF Pick-up | 0 | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 4 | 6 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS1_WB01:BPM_D1286 | BPM | D1286 | 34E000379 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 4 | 7 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS1_WB01:BPM_D1286 | BPM | D1286 | 34E000380 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 4 | 7 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS1_WB01:BPM_D1286 | BPM | D1286 | 34E000381 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 4 | 7 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS1_WB01:BPM_D1286 | BPM | D1286 | 34E000382 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 4 | 7 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS1_WB01:BPM_D1286 | BPM | D1286 | 34E000296 | 2C_1/4in_Hardline_036 | RF Pick-up | 0 | Ref | 0 | 0 | 0 | 0 | 0 | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0 | 0 | 0 | 0
BPM | LS1_CB02:BPM_D1295 | BPM | D1295 | 34E000316 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 4 | 7 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS1_CB02:BPM_D1295 | BPM | D1295 | 34E000317 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 4 | 7 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS1_CB02:BPM_D1295 | BPM | D1295 | 34E000318 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 4 | 7 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS1_CB02:BPM_D1295 | BPM | D1295 | 34E000319 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 4 | 7 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS1_CB02:BPM_D1295 | BPM | D1295 | 34E000295 | 2C_1/4in_Hardline_036 | RF Pick-up | 0 | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 4 | 7 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS1_CB02:BPM_D1315 | BPM | D1315 | 34E000327 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 4 | 8 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS1_CB02:BPM_D1315 | BPM | D1315 | 34E000328 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 4 | 8 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS1_CB02:BPM_D1315 | BPM | D1315 | 34E000329 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 4 | 8 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS1_CB02:BPM_D1315 | BPM | D1315 | 34E000330 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 4 | 8 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS1_CB02:BPM_D1315 | BPM | D1315 | 34E000294 | 2C_1/4in_Hardline_036 | RF Pick-up | 0 | Ref | 0 | 0 | 0 | 0 | 0 | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0 | 0 | 0 | 0
BPM | LS1_CB02:BPM_D1335 | BPM | D1335 | 34E000307 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 4 | 8 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS1_CB02:BPM_D1335 | BPM | D1335 | 34E000320 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 4 | 8 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS1_CB02:BPM_D1335 | BPM | D1335 | 34E000321 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 4 | 8 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS1_CB02:BPM_D1335 | BPM | D1335 | 34E000322 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 4 | 8 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS1_CB02:BPM_D1335 | BPM | D1335 | 34E000293 | 2C_1/4in_Hardline_036 | RF Pick-up | 0 | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 4 | 8 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS1_WB02:BPM_D1350 | BPM | D1350 | 34E000375 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 4 | 9 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS1_WB02:BPM_D1350 | BPM | D1350 | 34E000376 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 4 | 9 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS1_WB02:BPM_D1350 | BPM | D1350 | 34E000377 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 4 | 9 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS1_WB02:BPM_D1350 | BPM | D1350 | 34E000378 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 4 | 9 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS1_WB02:BPM_D1350 | BPM | D1350 | 34E000292 | 2C_1/4in_Hardline_036 | RF Pick-up | 0 | Ref | 0 | 0 | 0 | 0 | 0 | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0 | 0 | 0 | 0
BPM | LS1_CB03:BPM_D1359 | BPM | D1359 | 34E000371 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 4 | 9 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS1_CB03:BPM_D1359 | BPM | D1359 | 34E000372 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 4 | 9 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS1_CB03:BPM_D1359 | BPM | D1359 | 34E000373 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 4 | 9 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS1_CB03:BPM_D1359 | BPM | D1359 | 34E000374 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 4 | 9 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS1_CB03:BPM_D1359 | BPM | D1359 | 34E000291 | 2C_1/4in_Hardline_036 | RF Pick-up | 0 | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 4 | 9 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS1_CB03:BPM_D1379 | BPM | D1379 | 34E000352 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 4 | 10 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS1_CB03:BPM_D1379 | BPM | D1379 | 34E000353 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 4 | 10 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS1_CB03:BPM_D1379 | BPM | D1379 | 34E000354 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 4 | 10 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS1_CB03:BPM_D1379 | BPM | D1379 | 34E000355 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 4 | 10 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS1_CB03:BPM_D1379 | BPM | D1379 | 34E000287 | 2C_1/4in_Hardline_036 | RF Pick-up | 0 | Ref | 0 | 0 | 0 | 0 | 0 | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0 | 0 | 0 | 0
BPM | LS1_CB03:BPM_D1399 | BPM | D1399 | 34E000348 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 4 | 10 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS1_CB03:BPM_D1399 | BPM | D1399 | 34E000349 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 4 | 10 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS1_CB03:BPM_D1399 | BPM | D1399 | 34E000350 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 4 | 10 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS1_CB03:BPM_D1399 | BPM | D1399 | 34E000351 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 4 | 10 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS1_CB03:BPM_D1399 | BPM | D1399 | 34E000286 | 2C_1/4in_Hardline_036 | RF Pick-up | 0 | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 4 | 10 | 0 | Ref | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 1_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 3 | 1: EBS8003 | Ch0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 1_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 4 | 4 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 2_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 3 | 1: EBS8003 | Ch1 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 2_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 4 | 4 | CH7 - fmc1-4 | 0 | 0 | 0 | 0 | 0
BPM | LS1_WB03:BPM_D1413 | BPM | D1413 | 34E000367 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 5 | 3 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS1_WB03:BPM_D1413 | BPM | D1413 | 34E000368 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 5 | 3 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS1_WB03:BPM_D1413 | BPM | D1413 | 34E000369 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 5 | 3 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS1_WB03:BPM_D1413 | BPM | D1413 | 34E000370 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 5 | 3 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS1_WB03:BPM_D1413 | BPM | D1413 | 34E000289 | 2C_1/4in_Hardline_036 | Ref | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 5 | 3 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS1_CB04:BPM_D1423 | BPM | D1423 | 34E000356 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 5 | 3 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS1_CB04:BPM_D1423 | BPM | D1423 | 34E000357 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 5 | 3 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS1_CB04:BPM_D1423 | BPM | D1423 | 34E000358 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 5 | 3 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS1_CB04:BPM_D1423 | BPM | D1423 | 34E000359 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 5 | 3 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS1_CB04:BPM_D1423 | BPM | D1423 | 34E000288 | 2C_1/4in_Hardline_036 | Ref | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS1_CB04:BPM_D1442 | BPM | D1442 | 34E000360 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 5 | 4 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS1_CB04:BPM_D1442 | BPM | D1442 | 34E000361 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 5 | 4 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS1_CB04:BPM_D1442 | BPM | D1442 | 34E000362 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 5 | 4 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS1_CB04:BPM_D1442 | BPM | D1442 | 34E000363 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 5 | 4 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS1_CB04:BPM_D1442 | BPM | D1442 | 34E000285 | 2C_1/4in_Hardline_036 | Ref | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 5 | 4 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS1_CB04:BPM_D1462 | BPM | D1462 | 34E000347 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 5 | 4 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS1_CB04:BPM_D1462 | BPM | D1462 | 34E000364 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 5 | 4 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS1_CB04:BPM_D1462 | BPM | D1462 | 34E000365 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 5 | 4 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS1_CB04:BPM_D1462 | BPM | D1462 | 34E000366 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 5 | 4 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS1_CB04:BPM_D1462 | BPM | D1462 | 34E000290 | 2C_1/4in_Hardline_036 | Ref | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS1_WB04:BPM_D1477 | BPM | D1477 | 34E000216 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 5 | 5 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS1_WB04:BPM_D1477 | BPM | D1477 | 34E000217 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 5 | 5 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS1_WB04:BPM_D1477 | BPM | D1477 | 34E000218 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 5 | 5 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS1_WB04:BPM_D1477 | BPM | D1477 | 34E000219 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 5 | 5 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS1_WB04:BPM_D1477 | BPM | D1477 | 34E000119 | 2C_1/4in_Hardline_036 | Ref | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 5 | 5 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS1_CB05:BPM_D1486 | BPM | D1486 | 34E000212 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 5 | 5 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS1_CB05:BPM_D1486 | BPM | D1486 | 34E000213 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 5 | 5 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS1_CB05:BPM_D1486 | BPM | D1486 | 34E000214 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 5 | 5 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS1_CB05:BPM_D1486 | BPM | D1486 | 34E000215 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 5 | 5 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS1_CB05:BPM_D1486 | BPM | D1486 | 34E000118 | 2C_1/4in_Hardline_036 | Ref | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS1_CB05:BPM_D1506 | BPM | D1506 | 34E000208 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 5 | 6 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS1_CB05:BPM_D1506 | BPM | D1506 | 34E000209 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 5 | 6 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS1_CB05:BPM_D1506 | BPM | D1506 | 34E000210 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 5 | 6 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS1_CB05:BPM_D1506 | BPM | D1506 | 34E000211 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 5 | 6 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS1_CB05:BPM_D1506 | BPM | D1506 | 34E000117 | 2C_1/4in_Hardline_036 | Ref | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 5 | 6 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS1_CB05:BPM_D1526 | BPM | D1526 | 34E000204 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 5 | 6 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS1_CB05:BPM_D1526 | BPM | D1526 | 34E000205 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 5 | 6 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS1_CB05:BPM_D1526 | BPM | D1526 | 34E000206 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 5 | 6 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS1_CB05:BPM_D1526 | BPM | D1526 | 34E000207 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 5 | 6 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS1_CB05:BPM_D1526 | BPM | D1526 | 34E000116 | 2C_1/4in_Hardline_036 | Ref | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS1_WB05:BPM_D1541 | BPM | D1541 | 34E000200 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 5 | 7 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS1_WB05:BPM_D1541 | BPM | D1541 | 34E000201 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 5 | 7 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS1_WB05:BPM_D1541 | BPM | D1541 | 34E000202 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 5 | 7 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS1_WB05:BPM_D1541 | BPM | D1541 | 34E000203 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 5 | 7 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS1_WB05:BPM_D1541 | BPM | D1541 | 34E000112 | 2C_1/4in_Hardline_036 | Ref | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 5 | 7 | 0 | Ref | 0 | 0 | 0 | 0
Halo Monitor Ring | LS1_WB04:HMR_D1478 | HMR | D1478 | 34H000018 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J1 | 0 | 0 | 0 | 0 | 0 | 0 | J1 | SHV to Lemo | 6 | 3 | CH0 - fmc2-1 | 0 | 4 | 1: EBS8003 | Ch0 | 0
Halo Monitor Ring | LS1_WB05:HMR_D1542 | HMR | D1542 | 34H000020 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J2 | 0 | 0 | 0 | 0 | 0 | 0 | J2 | SHV to Lemo | 6 | 3 | CH1 - fmc2-2 | 0 | 4 | 1: EBS8003 | Ch0 | 0
Halo Monitor Ring | LS1_WB06:HMR_D1606 | HMR | D1606 | 34H000021 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J3 | 0 | 0 | 0 | 0 | 0 | 0 | J3 | SHV to Lemo | 6 | 3 | CH2 - fmc2-3 | 0 | 4 | 1: EBS8003 | Ch0 | 0
Halo Monitor Ring | LS1_WB07:HMR_D1670 | HMR | D1670 | 34H000023 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J4 | 0 | 0 | 0 | 0 | 0 | 0 | J4 | SHV to Lemo | 6 | 3 | CH4 - fmc1-1 | 0 | 4 | 1: EBS8003 | Ch1 | 0
Halo Monitor Ring | LS1_WB08:HMR_D1733 | HMR | D1733 | 34H000022 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J5 | 0 | 0 | 0 | 0 | 0 | 0 | J5 | SHV to Lemo | 6 | 3 | CH5 - fmc1-2 | 0 | 4 | 1: EBS8003 | Ch1 | 0
Halo Monitor Ring | LS1_WB09:HMR_D1797 | HMR | D1797 | 34H000019 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J6 | 0 | 0 | 0 | 0 | 0 | 0 | J6 | SHV to Lemo | 6 | 3 | CH6 - fmc1-3 | 0 | 4 | 1: EBS8003 | Ch1 | 0
0 | LS1_WB04:HMR_D1478 | 0 | D1478 | 34H000105 | 2C_5kVDC_Coax_035 | 0 | SHV to SHV Panel | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS1_WB05:HMR_D1542 | 0 | D1542 | 34H000103 | 2C_5kVDC_Coax_035 | 0 | SHV to SHV Panel | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS1_WB06:HMR_D1606 | 0 | D1606 | 34H000102 | 2C_5kVDC_Coax_035 | 0 | SHV to SHV Panel | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS1_WB07:HMR_D1670 | 0 | D1670 | 34H000115 | 2C_5kVDC_Coax_035 | 0 | SHV to SHV Panel | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS1_WB08:HMR_D1733 | 0 | D1733 | 34H000116 | 2C_5kVDC_Coax_035 | 0 | SHV to SHV Panel | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS1_WB09:HMR_D1797 | 0 | D1797 | 34H000104 | 2C_5kVDC_Coax_035 | 0 | SHV to SHV Panel | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS1_CB06:BPM_D1550 | BPM | D1550 | 34E000196 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 6 | 7 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS1_CB06:BPM_D1550 | BPM | D1550 | 34E000197 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 6 | 7 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS1_CB06:BPM_D1550 | BPM | D1550 | 34E000198 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 6 | 7 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS1_CB06:BPM_D1550 | BPM | D1550 | 34E000199 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 6 | 7 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS1_CB06:BPM_D1550 | BPM | D1550 | 34E000111 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No terminiation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS1_CB06:BPM_D1570 | BPM | D1570 | 34E000192 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 6 | 7 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS1_CB06:BPM_D1570 | BPM | D1570 | 34E000193 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 6 | 7 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS1_CB06:BPM_D1570 | BPM | D1570 | 34E000194 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 6 | 7 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS1_CB06:BPM_D1570 | BPM | D1570 | 34E000195 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 6 | 7 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS1_CB06:BPM_D1570 | BPM | D1570 | 34E000113 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 6 | 7 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS1_CB06:BPM_D1590 | BPM | D1590 | 34E000188 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 6 | 8 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS1_CB06:BPM_D1590 | BPM | D1590 | 34E000189 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 6 | 8 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS1_CB06:BPM_D1590 | BPM | D1590 | 34E000190 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 6 | 8 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS1_CB06:BPM_D1590 | BPM | D1590 | 34E000191 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 6 | 8 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS1_CB06:BPM_D1590 | BPM | D1590 | 34E000115 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No terminiation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS1_WB06:BPM_D1604 | BPM | D1605 | 34E000184 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 6 | 8 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS1_WB06:BPM_D1604 | BPM | D1605 | 34E000185 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 6 | 8 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS1_WB06:BPM_D1604 | BPM | D1605 | 34E000186 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 6 | 8 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS1_WB06:BPM_D1604 | BPM | D1605 | 34E000187 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 6 | 8 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS1_WB06:BPM_D1604 | BPM | D1605 | 34E000114 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 6 | 8 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS1_CB07:BPM_D1614 | BPM | D1614 | 34E000180 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 6 | 9 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS1_CB07:BPM_D1614 | BPM | D1614 | 34E000181 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 6 | 9 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS1_CB07:BPM_D1614 | BPM | D1614 | 34E000182 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 6 | 9 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS1_CB07:BPM_D1614 | BPM | D1614 | 34E000183 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 6 | 9 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS1_CB07:BPM_D1614 | BPM | D1614 | 34E000110 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No terminiation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS1_CB07:BPM_D1634 | BPM | D1634 | 34E000176 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 6 | 9 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS1_CB07:BPM_D1634 | BPM | D1634 | 34E000177 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 6 | 9 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS1_CB07:BPM_D1634 | BPM | D1634 | 34E000178 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 6 | 9 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS1_CB07:BPM_D1634 | BPM | D1634 | 34E000179 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 6 | 9 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS1_CB07:BPM_D1634 | BPM | D1634 | 34E000129 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 6 | 9 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS1_CB07:BPM_D1654 | BPM | D1654 | 34E000172 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 2 | 10 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS1_CB07:BPM_D1654 | BPM | D1654 | 34E000173 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 6 | 10 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS1_CB07:BPM_D1654 | BPM | D1654 | 34E000174 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 6 | 10 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS1_CB07:BPM_D1654 | BPM | D1654 | 34E000175 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 6 | 10 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS1_CB07:BPM_D1654 | BPM | D1654 | 34E000128 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 6 | 10 | 0 | Ref | 0 | 0 | 0 | 0
BCM | FE_LEBT:BCM_D0989 | BCM | D0998 | 24E000044 | 2C_22AWG_Multi_063 | Signal | Amplifier | FE_LEBT:AIM_D0998 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Amplifier in FE_N0404
BCM | FE_MEBT:BCM_D1055 | BCM | D1055 | 24E000042 | 2C_22AWG_Multi_063 | Signal | Amplifier | FE_MEBT:AIM_D1055 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Amplifier in FE_N0404
BCM | FE_MEBT:BCM_D1120 | BCM | D1121 | 24E000043 | 2C_22AWG_Multi_063 | Signal | Amplifier | FE_MEBT:AIM_D1121 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Amplifier in FE_N0404
BCM | FS1_CSS:BCM_D2183 | BCM | D2183 | 44E000222 | 2C_22AWG_Multi_063 | Signal | Amplifier | FS1_CSS:AIM_D2183 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Amplifier in BDS_N0305
BCM | FS1_CSS:BCM_D2264 | BCM | D2264 | 44E000221 | 2C_22AWG_Multi_063 | Signal | Amplifier | FS1_CSS:AIM_D2264 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Amplifier in BDS_N0305
BCM | FS1_CSS:BCM_D2353 | BCM | D2353 | 44E000223 | 2C_22AWG_Multi_063 | Signal | Amplifier | FS1_CSS:AIM_D2359 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Amplifier in BDS_N0305
BCM | FS1_BTS:BCM_D2449 | BCM | D2464 | 44E000219 | 2C_22AWG_Multi_063 | Signal | Amplifier | BDS_FFS:AIM_D5798 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Amplifier in BDS_N0305
BCM | FS1_BMS:BCM_D2519 | BCM | D2519 | 44E000224 | 2C_22AWG_Multi_063 | Signal | Amplifier | FS1_BMS:AIM_D2519 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Amplifier in FS1_N0506
BCM | FS1_BMS:BCM_D2675 | BCM | D2675 | 44E000225 | 2C_22AWG_Multi_063 | Signal | Amplifier | FS1_BMS:AIM_D2675 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Amplifier in FS1_N0506
BCM | FS2_BTS:BCM_D3936 | BCM | D3952 | 44E000227 | 2C_22AWG_Multi_063 | Signal | Amplifier | FS2_BTS:AIM_D3952 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Amplifier in LS2_N4201
BCM | FS2_BBS:BCM_D4169 | BCM | D4175 | 44E000226 | 2C_22AWG_Multi_063 | Signal | Amplifier | FS2_BMS:AIM_D4175 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Amplifier in FE_N0404
BCM | BDS_BTS:BCM_D5521 | BCM | D5550 | 44E000220 | 2C_22AWG_Multi_063 | Signal | Amplifier | BDS_BTS:AIM_D5550 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Amplifier in BDS_N0305
BCM | BDS_FFS:BCM_D5789 | BCM | D5798 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | CABLE MOVED TO: FS1_BTS:BCM_D2449
BCM | FE_LEBT:BCM_D0989 | BCM | D0998 | 24E000041 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | J1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to SMB | 6 | 4 | 0 | 1 | 0 | 0 | 0 | FE_LEBT:AIM_D0998 in FE_N0404
BCM | FE_MEBT:BCM_D1055 | BCM | D1055 | 24E000039 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | J2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to SMB | 6 | 4 | 0 | 2 | 0 | 0 | 0 | FE_MEBT:AIM_D1055 in FE_N0404
BCM | FE_MEBT:BCM_D1120 | BCM | D1121 | 24E000040 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | J3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to SMB | 6 | 4 | 0 | 3 | 0 | 0 | 0 | FE_MEBT:AIM_D1121 in FE_N0404
BCM | FS1_CSS:BCM_D2183 | BCM | D2228 | 44E000214 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | J5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to SMB | 6 | 4 | 0 | 4 | 0 | 0 | 0 | FS1_CSS:AIM_D2228 in BDS_N0305
BCM | FS1_CSS:BCM_D2264 | BCM | D2265 | 44E000213 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | J4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to SMB | 6 | 5 | 0 | 1 | 0 | 0 | 0 | FS1_CSS:AIM_D2265 in BDS_N0305
BCM | FS1_BTS:BCM_D2449 | BCM | D2464 | 44E000210 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | J12 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to SMB | 6 | 5 | 0 | 2 | 0 | 0 | 0 | BDS_FFS:AIM_D5798 in BDS_N0305
BCM | FS1_CSS:BCM_D2353 | BCM | D2359 | 44E000212 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | J6 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to SMB | 6 | 5 | 0 | 3 | 0 | 0 | 0 | FS1_CSS:AIM_D2359 in BDS_N0305
BCM | FS1_BMS:BCM_D2519 | BCM | D2525 | 44E000215 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | J7 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to SMB | 6 | 5 | 0 | 4 | 0 | 0 | 0 | FS1_BMS:AIM_D2525 in FS1_N0506
BCM | FS1_BMS:BCM_D2675 | BCM | D2686 | 44E000216 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | J8 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to SMB | 6 | 6 | 0 | 1 | 0 | 0 | 0 | FS1_BMS:AIM_D2686 in FS1_N0506
BCM | FS2_BTS:BCM_D3936 | BCM | D3952 | 44E000209 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | J9 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to SMB | 6 | 6 | 0 | 2 | 0 | 0 | 0 | FS2_BTS:AIM_D3952 in LS2_N4201
BCM | FS2_BBS:BCM_D4169 | BCM | D4175 | 44E000217 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | J10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to SMB | 6 | 6 | 0 | 3 | 0 | 0 | 0 | FS2_BMS:AIM_D4175 in FE_N0404
BCM | BDS_BTS:BCM_D5521 | BCM | D5550 | 44E000211 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | J11 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to SMB | 6 | 6 | 0 | 4 | 0 | 0 | 0 | BDS_BTS:AIM_D5550 in BDS_N0305
BCM | BDS_FFS:BCM_D5789 | BCM | D5798 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | CABLE MOVED TO: FS1_BTS:BCM_D2449
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 1_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 4 | 1: EBS8003 | Ch0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 1_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 6 | 3 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 2_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 4 | 1: EBS8003 | Ch1 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 2_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 6 | 3 | CH7 - fmc1-4 | 0 | 0 | 0 | 0 | 0
BPM | LS1_WB07:BPM_D1668 | BPM | D1668 | 34E000168 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 7 | 3 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS1_WB07:BPM_D1668 | BPM | D1668 | 34E000169 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 7 | 3 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS1_WB07:BPM_D1668 | BPM | D1668 | 34E000170 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 7 | 3 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS1_WB07:BPM_D1668 | BPM | D1668 | 34E000171 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 7 | 3 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS1_WB07:RFS_D1668 | BPM | D1668 | 34E000127 | 2C_1/4in_Hardline_036 | Ref | No terminiation | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0 | 0 | 0 | 0
BPM | LS1_CB08:BPM_D1677 | BPM | D1677 | 34E000164 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 7 | 3 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS1_CB08:BPM_D1677 | BPM | D1677 | 34E000165 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 7 | 3 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS1_CB08:BPM_D1677 | BPM | D1677 | 34E000166 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 7 | 3 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS1_CB08:BPM_D1677 | BPM | D1677 | 34E000167 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 7 | 3 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS1_CB08:RFS_D1677 | BPM | D1677 | 34E000126 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 7 | 3 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS1_CB08:BPM_D1697 | BPM | D1697 | 34E000160 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 7 | 4 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS1_CB08:BPM_D1697 | BPM | D1697 | 34E000161 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 7 | 4 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS1_CB08:BPM_D1697 | BPM | D1697 | 34E000162 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 7 | 4 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS1_CB08:BPM_D1697 | BPM | D1697 | 34E000163 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 7 | 4 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS1_CB08:RFS_D1697 | BPM | D1697 | 34E000124 | 2C_1/4in_Hardline_036 | Ref | No terminiation | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0 | 0 | 0 | 0
BPM | LS1_CB08:BPM_D1717 | BPM | D1717 | 34E000156 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 7 | 4 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS1_CB08:BPM_D1717 | BPM | D1717 | 34E000157 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 7 | 4 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS1_CB08:BPM_D1717 | BPM | D1717 | 34E000158 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 7 | 4 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS1_CB08:BPM_D1717 | BPM | D1717 | 34E000159 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 7 | 4 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS1_CB08:RFS_D1717 | BPM | D1717 | 34E000121 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 7 | 4 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS1_WB08:BPM_D1732 | BPM | D1732 | 34E000140 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 7 | 5 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS1_WB08:BPM_D1732 | BPM | D1732 | 34E000141 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 7 | 5 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS1_WB08:BPM_D1732 | BPM | D1732 | 34E000142 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 7 | 5 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS1_WB08:BPM_D1732 | BPM | D1732 | 34E000143 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 7 | 5 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS1_WB08:RFS_D1732 | BPM | D1732 | 34E000122 | 2C_1/4in_Hardline_036 | Ref | No terminiation | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0 | 0 | 0 | 0
BPM | LS1_CB09:BPM_D1741 | BPM | D1741 | 34E000136 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 7 | 5 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS1_CB09:BPM_D1741 | BPM | D1741 | 34E000137 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 7 | 5 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS1_CB09:BPM_D1741 | BPM | D1741 | 34E000138 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 7 | 5 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS1_CB09:BPM_D1741 | BPM | D1741 | 34E000139 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 7 | 5 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS1_CB09:RFS_D1741 | BPM | D1741 | 34E000123 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 7 | 5 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS1_CB09:BPM_D1761 | BPM | D1761 | 34E000152 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 7 | 6 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS1_CB09:BPM_D1761 | BPM | D1761 | 34E000153 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 7 | 6 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS1_CB09:BPM_D1761 | BPM | D1761 | 34E000154 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 7 | 6 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS1_CB09:BPM_D1761 | BPM | D1761 | 34E000155 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 7 | 6 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS1_CB09:RFS_D1761 | BPM | D1761 | 34E000125 | 2C_1/4in_Hardline_036 | Ref | No terminiation | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0 | 0 | 0 | 0
BPM | LS1_CB09:BPM_D1781 | BPM | D1781 | 34E000148 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 7 | 6 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS1_CB09:BPM_D1781 | BPM | D1781 | 34E000149 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 7 | 6 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS1_CB09:BPM_D1781 | BPM | D1781 | 34E000150 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 7 | 6 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS1_CB09:BPM_D1781 | BPM | D1781 | 34E000151 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 7 | 6 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS1_CB09:RFS_D1781 | BPM | D1781 | 34E000120 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 7 | 6 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS1_WB09:BPM_D1796 | BPM | D1796 | 34E000132 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 7 | 7 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS1_WB09:BPM_D1796 | BPM | D1796 | 34E000133 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 7 | 7 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS1_WB09:BPM_D1796 | BPM | D1796 | 34E000134 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 7 | 7 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS1_WB09:BPM_D1796 | BPM | D1796 | 34E000135 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 7 | 7 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS1_WB09:RFS_D1796 | BPM | D1796 | 34E000131 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 7 | 7 | 0 | Ref | 0 | 0 | 0 | 0
Neutron Detector | LS1_CB05:ND_D1505 | ND | D1505 | 34E000526 | 2C_75LC_Coax_089 | Signal | N-Type to Lemo Box | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | SMA to Lemo650 | 7 | 8 | CH0 - fmc2-1 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS1_CB07:ND_D1633 | ND | D1633 | 34E000528 | 2C_75LC_Coax_089 | Signal | N-Type to Lemo Box | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | SMA to Lemo650 | 7 | 8 | CH1 - fmc2-2 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS1_CB09:ND_D1760 | ND | D1760 | 34E000530 | 2C_75LC_Coax_089 | Signal | N-Type to Lemo Box | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | SMA to Lemo650 | 7 | 8 | CH2 - fmc2-3 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS1_CB11:ND_D1908 | ND | D1908 | 34E000532 | 2C_75LC_Coax_089 | Signal | N-Type to Lemo Box | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | SMA to Lemo650 | 7 | 8 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS1_BTS:ND_D2076 | ND | D2076 | 34E000534 | 2C_75LC_Coax_089 | Signal | N-Type to Lemo Box | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 5 | SMA to Lemo650 | 7 | 8 | CH4 - fmc1-1 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS1_CB05:ND_D1505 | 0 | D1505 | 34E000525 | 2C_75LC_Coax_089 | Test Signal | N-Type to SMA Panel | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to Ferrule | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS1_CB07:ND_D1633 | 0 | D1633 | 34E000527 | 2C_75LC_Coax_089 | Test Signal | N-Type to SMA Panel | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to Ferrule | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS1_CB09:ND_D1760 | 0 | D1760 | 34E000529 | 2C_75LC_Coax_089 | Test Signal | N-Type to SMA Panel | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to Ferrule | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS1_CB11:ND_D1908 | 0 | D1908 | 34E000531 | 2C_75LC_Coax_089 | Test Signal | N-Type to SMA Panel | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to Ferrule | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS1_BTS:ND_D2076 | 0 | D2076 | 34E000533 | 2C_75LC_Coax_089 | Test Signal | N-Type to SMA Panel | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to Ferrule | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS1_CB05:ND_D1505 | 0 | D1505 | 34H000059 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | SHV to SHV | 0 | 0 | 0 | 0 | 5 | 1:EHS8020n | Ch0 | 0
Neutron Detector | LS1_CB07:ND_D1633 | 0 | D1633 | 34H000060 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | SHV to SHV | 0 | 0 | 0 | 0 | 5 | 1:EHS8020n | Ch1 | 0
Neutron Detector | LS1_CB09:ND_D1760 | 0 | D1760 | 34H000061 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | SHV to SHV | 0 | 0 | 0 | 0 | 5 | 1:EHS8020n | Ch2 | 0
Neutron Detector | LS1_CB11:ND_D1908 | 0 | D1908 | 34H000062 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | SHV to SHV | 0 | 0 | 0 | 0 | 5 | 1:EHS8020n | Ch3 | 0
Neutron Detector | LS1_BTS:ND_D2076 | 0 | D2076 | 34H000063 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 5 | SHV to SHV | 0 | 0 | 0 | 0 | 5 | 1:EHS8020n | Ch4 | 0
BPM | LS1_CB10:BPM_D1805 | BPM | D1805 | 34E000144 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 8 | 3 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS1_CB10:BPM_D1805 | BPM | D1805 | 34E000145 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 8 | 3 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS1_CB10:BPM_D1805 | BPM | D1805 | 34E000146 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 8 | 3 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS1_CB10:BPM_D1805 | BPM | D1805 | 34E000147 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 8 | 3 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS1_CB10:RFS_D1805 | BPM | D1805 | 34E000130 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS1_CB10:BPM_D1825 | BPM | D1825 | 34E000233 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 8 | 3 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS1_CB10:BPM_D1825 | BPM | D1825 | 34E000250 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 8 | 3 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS1_CB10:BPM_D1825 | BPM | D1825 | 34E000251 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 8 | 3 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS1_CB10:BPM_D1825 | BPM | D1825 | 34E000252 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 8 | 3 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS1_CB10:RFS_D1825 | BPM | D1825 | 34E000224 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 8 | 3 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS1_CB10:BPM_D1845 | BPM | D1845 | 34E000246 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 8 | 4 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS1_CB10:BPM_D1845 | BPM | D1845 | 34E000247 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 8 | 4 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS1_CB10:BPM_D1845 | BPM | D1845 | 34E000248 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 8 | 4 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS1_CB10:BPM_D1845 | BPM | D1845 | 34E000249 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 8 | 4 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS1_CB10:RFS_D1845 | BPM | D1845 | 34E000225 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS1_WB10:BPM_D1859 | BPM | D1860 | 34E000253 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 8 | 4 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS1_WB10:BPM_D1859 | BPM | D1860 | 34E000254 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 8 | 4 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS1_WB10:BPM_D1859 | BPM | D1860 | 34E000255 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 8 | 4 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS1_WB10:BPM_D1859 | BPM | D1860 | 34E000256 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 8 | 4 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS1_WB10:RFS_D1859 | BPM | D1860 | 34E000220 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 8 | 4 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS1_CB11:BPM_D1869 | BPM | D1869 | 34E000234 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 8 | 5 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS1_CB11:BPM_D1869 | BPM | D1869 | 34E000235 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 8 | 5 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS1_CB11:BPM_D1869 | BPM | D1869 | 34E000236 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 8 | 5 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS1_CB11:BPM_D1869 | BPM | D1869 | 34E000237 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 8 | 5 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS1_CB11:RFS_D1869 | BPM | D1869 | 34E000222 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS1_CB11:BPM_D1889 | BPM | D1889 | 34E000242 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 8 | 5 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS1_CB11:BPM_D1889 | BPM | D1889 | 34E000243 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 8 | 5 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS1_CB11:BPM_D1889 | BPM | D1889 | 34E000244 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 8 | 5 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS1_CB11:BPM_D1889 | BPM | D1889 | 34E000245 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 8 | 5 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS1_CB11:RFS_D1889 | BPM | D1889 | 34E000223 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 8 | 5 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS1_CB11:BPM_D1909 | BPM | D1909 | 34E000238 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 8 | 6 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS1_CB11:BPM_D1909 | BPM | D1909 | 34E000239 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 8 | 6 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS1_CB11:BPM_D1909 | BPM | D1909 | 34E000240 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 8 | 6 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS1_CB11:BPM_D1909 | BPM | D1909 | 34E000241 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 8 | 6 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS1_CB11:RFS_D1909 | BPM | D1909 | 34E000221 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS1_WB11:BPM_D1923 | BPM | D1923 | 34E000257 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 8 | 6 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS1_WB11:BPM_D1923 | BPM | D1923 | 34E000258 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 8 | 6 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS1_WB11:BPM_D1923 | BPM | D1923 | 34E000259 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 8 | 6 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS1_WB11:BPM_D1923 | BPM | D1923 | 34E000260 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 8 | 6 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS1_WB11:RFS_D1923 | BPM | D1923 | 34E000226 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 8 | 6 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS1_BTS:BPM_D1967 | BPM | D1966 | 34E000261 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 8 | 7 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS1_BTS:BPM_D1967 | BPM | D1966 | 34E000262 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 8 | 7 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS1_BTS:BPM_D1967 | BPM | D1966 | 34E000263 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 8 | 7 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS1_BTS:BPM_D1967 | BPM | D1966 | 34E000264 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 8 | 7 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS1_BTS:RFS_D1967 | BPM | D1966 | 34E000227 | 2C_1/4in_Hardline_036 | Ref | No terminiation | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0 | 0 | 0 | 0
BPM | LS1_BTS:BPM_D1980 | BPM | D1980 | 34E000265 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 8 | 7 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS1_BTS:BPM_D1980 | BPM | D1980 | 34E000266 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 8 | 7 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS1_BTS:BPM_D1980 | BPM | D1980 | 34E000267 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 8 | 7 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS1_BTS:BPM_D1980 | BPM | D1980 | 34E000268 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 8 | 7 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS1_BTS:RFS_D1980 | BPM | D1980 | 34E000228 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 8 | 7 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS1_BTS:BPM_D2027 | BPM | D2033 | 34E000269 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 8 | 8 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS1_BTS:BPM_D2027 | BPM | D2033 | 34E000270 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 8 | 8 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS1_BTS:BPM_D2027 | BPM | D2033 | 34E000271 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 8 | 8 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS1_BTS:BPM_D2027 | BPM | D2033 | 34E000272 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 8 | 8 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS1_BTS:RFS_D2027 | BPM | D2033 | 34E000229 | 2C_1/4in_Hardline_036 | Ref | No terminiation | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0 | 0 | 0 | 0
BPM | LS1_BTS:BPM_D2054 | BPM | D2056 | 34E000273 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 8 | 8 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS1_BTS:BPM_D2054 | BPM | D2056 | 34E000282 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 8 | 8 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS1_BTS:BPM_D2054 | BPM | D2056 | 34E000283 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 8 | 8 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS1_BTS:BPM_D2054 | BPM | D2056 | 34E000284 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 8 | 8 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS1_BTS:RFS_D2054 | BPM | D2056 | 34E000230 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 8 | 8 | 0 | Ref | 0 | 0 | 0 | 0
PM 8" | LS1_BTS:PM_D2056 | PM | D2056 | 34F000019 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-1 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-1A | LEMO650/LEMO650 | 9 | 3 | CH0 - fmc2-1 | 0 | 6 | 1: EBS8005 | Ch0 | 0
PM 8" | LS1_BTS:PM_D2056 | PM | D2056 | 34F000019 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-1 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield-6 | 4 | 6 | LEMO-1B | LEMO650/LEMO650 | 9 | 3 | CH1 - fmc2-2 | 0 | 6 | 1: EBS8005 | Ch0 | 0
PM 8" | LS1_BTS:PM_D2056 | PM | D2056 | 34F000019 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-1 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-1C | LEMO650/LEMO650 | 9 | 3 | CH2 - fmc2-3 | 0 | 6 | 1: EBS8005 | Ch0 | 0
PM 8" | LS1_BTS:PM_D2131 | PM | D2131 | 34F000012 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-2 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-2A | LEMO650/LEMO650 | 9 | 3 | CH4 - fmc1-1 | 0 | 6 | 1: EBS8005 | Ch1 | 0
PM 8" | LS1_BTS:PM_D2131 | PM | D2131 | 34F000012 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-2 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | LEMO-2B | LEMO650/LEMO650 | 9 | 3 | CH5 - fmc1-2 | 0 | 6 | 1: EBS8005 | Ch1 | 0
PM 8" | LS1_BTS:PM_D2131 | PM | D2131 | 34F000012 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-2 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-2C | LEMO650/LEMO650 | 9 | 3 | CH6 - fmc1-3 | 0 | 6 | 1: EBS8005 | Ch1 | 0
PM 8" | FS1_CSS:PM_D2198 | PM | D2198 | 44F000056 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-3 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-3A | LEMO650/LEMO650 | 9 | 4 | CH0 - fmc2-1 | 0 | 6 | 1: EBS8005 | Ch2 | 0
PM 8" | FS1_CSS:PM_D2198 | PM | D2198 | 44F000056 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-3 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield-6 | 4 | 6 | LEMO-3B | LEMO650/LEMO650 | 9 | 4 | CH1 - fmc2-2 | 0 | 6 | 1: EBS8005 | Ch2 | 0
PM 8" | FS1_CSS:PM_D2198 | PM | D2198 | 44F000056 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-3 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-3C | LEMO650/LEMO650 | 9 | 4 | CH2 - fmc2-3 | 0 | 6 | 1: EBS8005 | Ch2 | 0
PM 8" | FS1_CSS:PM_D2225 | PM | D2225 | 44F000052 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-4 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-4A | LEMO650/LEMO650 | 9 | 4 | CH4 - fmc1-1 | 0 | 6 | 1: EBS8005 | Ch3 | 0
PM 8" | FS1_CSS:PM_D2225 | PM | D2225 | 44F000052 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-4 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | LEMO-4B | LEMO650/LEMO650 | 9 | 4 | CH5 - fmc1-2 | 0 | 6 | 1: EBS8005 | Ch3 | 0
PM 8" | FS1_CSS:PM_D2225 | PM | D2225 | 44F000052 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-4 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-4C | LEMO650/LEMO650 | 9 | 4 | CH6 - fmc1-3 | 0 | 6 | 1: EBS8005 | Ch3 | 0
PM 8" | FS1_CSS:PM_D2249 | PM | D2249 | 44F000048 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-5 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-5A | LEMO650/LEMO650 | 9 | 5 | CH0 - fmc2-1 | 0 | 6 | 1: EBS8005 | Ch4 | 0
PM 8" | FS1_CSS:PM_D2249 | PM | D2249 | 44F000048 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-5 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield-6 | 4 | 6 | LEMO-5B | LEMO650/LEMO650 | 9 | 5 | CH1 - fmc2-2 | 0 | 6 | 1: EBS8005 | Ch4 | 0
PM 8" | FS1_CSS:PM_D2249 | PM | D2249 | 44F000048 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-5 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-5C | LEMO650/LEMO650 | 9 | 5 | CH2 - fmc2-3 | 0 | 6 | 1: EBS8005 | Ch4 | 0
PM 8" | FS1_CSS:PM_D2385 | PM | D2385 | 44F000044 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-6 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-6A | LEMO650/LEMO650 | 9 | 5 | CH4 - fmc1-1 | 0 | 6 | 1: EBS8005 | Ch5 | 0
PM 8" | FS1_CSS:PM_D2385 | PM | D2385 | 44F000044 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-6 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | LEMO-6B | LEMO650/LEMO650 | 9 | 5 | CH5 - fmc1-2 | 0 | 6 | 1: EBS8005 | Ch5 | 0
PM 8" | FS1_CSS:PM_D2385 | PM | D2385 | 44F000044 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-6 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-6C | LEMO650/LEMO650 | 9 | 5 | CH6 - fmc1-3 | 0 | 6 | 1: EBS8005 | Ch5 | 0
PM 12" | FS1_BTS:PM_D2439 | PM | D2439 | 44F000036 | 6C_22AWG_Multi_010 | A Signal | CPC to LEMO650 Box | CPC-7 |  | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-7A | LEMO650/LEMO650 | 9 | 6 | CH0 - fmc2-1 | 0 | 6 | 1: EBS8005 | Ch6 | 0
PM 12" | FS1_BTS:PM_D2439 | PM | D2439 | 44F000036 | 6C_22AWG_Multi_010 | B Signal | CPC to LEMO650 Box | CPC-7 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | LEMO-7B | LEMO650/LEMO650 | 9 | 6 | CH1 - fmc2-2 | 0 | 6 | 1: EBS8005 | Ch6 | 0
PM 6" | FS1_BTS:PM_D2439 | PM | D2439 | 44F000040 | 4C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-7 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-7C | LEMO650/LEMO650 | 9 | 6 | CH2 - fmc2-3 | 0 | 6 | 1: EBS8005 | Ch6 | 0
Ion Chamber | FS1_BTS:IC_D2448 | IC | D2448 | 44E000256 | 2C_1kV75LC_Coax_090 | Signal | N-type to Lemo Box | J1 | 0 | 0 | 0 | 0 | 0 | 0 | J1 | Lemo to Lemo | 9 | 7 | CH0 - fmc2-1 | 0 | 0 | 0 | 0 | 0
Ion Chamber | FS1_BTS:IC_D2449 | IC | D2449 | 44E000257 | 2C_1kV75LC_Coax_090 | Signal | N-type to Lemo Box | J2 | 0 | 0 | 0 | 0 | 0 | 0 | J2 | Lemo to Lemo | 9 | 7 | CH1 - fmc2-2 | 0 | 0 | 0 | 0 | 0
Ion Chamber | FS1_BTS:IC_D2478 | IC | D2478 | 44E000254 | 2C_1kV75LC_Coax_090 | Signal | N-type to Lemo Box | J3 | 0 | 0 | 0 | 0 | 0 | 0 | J3 | Lemo to Lemo | 9 | 7 | CH2 - fmc2-3 | 0 | 0 | 0 | 0 | 0
Ion Chamber | FS1_BTS:IC_D2479 | IC | D2479 | 44E000255 | 2C_1kV75LC_Coax_090 | Signal | N-type to Lemo Box | J4 | 0 | 0 | 0 | 0 | 0 | 0 | J4 | Lemo to Lemo | 9 | 7 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Ion Chamber | FS1_STRL:IC_D2233 | IC | D2233 | 44E000258 | 2C_1kV75LC_Coax_090 | Signal | N-type to Lemo Box | J5 | 0 | 0 | 0 | 0 | 0 | 0 | J5 | Lemo to Lemo | 9 | 7 | CH4 - fmc1-1 | 0 | 0 | 0 | 0 | 0
Ion Chamber | FS1_STRL:IC_D2234 | IC | D2234 | 44E000259 | 2C_1kV75LC_Coax_090 | Signal | N-type to Lemo Box | J6 | 0 | 0 | 0 | 0 | 0 | 0 | J6 | Lemo to Lemo | 9 | 7 | CH5 - fmc1-2 | 0 | 0 | 0 | 0 | 0
Ion Chamber | FS1_STRL:IC_D2253 | IC | D2253 | 44E000261 | 2C_1kV75LC_Coax_090 | Signal | N-type to Lemo Box | J7 | 0 | 0 | 0 | 0 | 0 | 0 | J7 | Lemo to Lemo | 9 | 7 | CH6 - fmc1-3 | 0 | 0 | 0 | 0 | 0
Ion Chamber | FS1_STRL:IC_D2254 | IC | D2254 | 44E000260 | 2C_1kV75LC_Coax_090 | Signal | N-type to Lemo Box | J8 | 0 | 0 | 0 | 0 | 0 | 0 | J8 | Lemo to Lemo | 9 | 7 | CH7 - fmc1-4 | 0 | 0 | 0 | 0 | 0
Ion Chamber | FS1_BTS:IC_D2448 | 0 | D2448 | 44H000041 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | SHV to SHV | 0 | 0 | 0 | 0 | 6 | 2: EHS8020n | Ch0 | 0
Ion Chamber | FS1_BTS:IC_D2449 | 0 | D2449 | 44H000040 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | SHV to SHV | 0 | 0 | 0 | 0 | 6 | 2: EHS8020n | Ch1 | 0
Ion Chamber | FS1_BTS:IC_D2478 | 0 | D2478 | 44H000039 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | SHV to SHV | 0 | 0 | 0 | 0 | 6 | 2: EHS8020n | Ch2 | 0
Ion Chamber | FS1_BTS:IC_D2479 | 0 | D2479 | 44H000038 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | SHV to SHV | 0 | 0 | 0 | 0 | 6 | 2: EHS8020n | Ch3 | 0
Ion Chamber | FS1_STRL:IC_D2233 | 0 | D2233 | 44H000044 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 5 | SHV to SHV | 0 | 0 | 0 | 0 | 6 | 2: EHS8020n | Ch4 | 0
Ion Chamber | FS1_STRL:IC_D2234 | 0 | D2234 | 44H000045 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 6 | 0 | 0 | 0 | 0 | 0 | 0 | 6 | SHV to SHV | 0 | 0 | 0 | 0 | 6 | 2: EHS8020n | Ch5 | 0
Ion Chamber | FS1_STRL:IC_D2253 | 0 | D2253 | 44H000042 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 7 | 0 | 0 | 0 | 0 | 0 | 0 | 7 | SHV to SHV | 0 | 0 | 0 | 0 | 6 | 2: EHS8020n | Ch6 | 0
Ion Chamber | FS1_STRL:IC_D2254 | 0 | D2254 | 44H000043 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 8 | 0 | 0 | 0 | 0 | 0 | 0 | 8 | SHV to SHV | 0 | 0 | 0 | 0 | 6 | 2: EHS8020n | Ch7 | 0
Halo Monitor Ring | LS1_WB10:HMR_D1861 | HMR | D1861 | 34H000013 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J1 | 0 | 0 | 0 | 0 | 0 | 0 | J1 | SHV to Lemo | 9 | 9 | CH0 - fmc2-1 | 0 | 6 | 1: EBS8005 | Ch7 | 0
Halo Monitor Ring | LS1_WB11:HMR_D1925 | HMR | D1924 | 34H000017 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J2 | 0 | 0 | 0 | 0 | 0 | 0 | J2 | SHV to Lemo | 9 | 9 | CH1 - fmc2-2 | 0 | 6 | 1: EBS8005 | Ch7 | 0
0 | LS1_BTS:HMR_D1981 | 0 | D1981 | 34H000014 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS1_BTS:HMR_D2057 | 0 | D2057 | 34H000016 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS1_BTS:HMR_D2150 | 0 | D2150 | 34H000015 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | FS1_CSS:HMR_D2204 | 0 | D2204 | 44H000003 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | FS1_CSS:HMR_D2232 | 0 | D2232 | 44H000004 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | FS1_CSS:HMR_D2255 | 0 | D2255 | 44H000005 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | FS1_CSS:HMR_D2318 | 0 | D2318 | 44H000006 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | FS1_CSS:HMR_D2385 | 0 | D2385 | 44H000007 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | FS1_BMS:HMR_D2504 | 0 | D2504 | 44H000008 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS1_WB10:HMR_D1861 | 0 | D1861 | 34H000098 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS1_WB11:HMR_D1924 | 0 | D1924 | 34H000106 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS1_BTS:HMR_D1981 | 0 | D1981 | 34H000096 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS1_BTS:HMR_D2057 | 0 | D2057 | 34H000107 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS1_BTS:HMR_D2150 | 0 | D2150 | 34H000094 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | FS1_CSS:HMR_D2204 | 0 | D2204 | 44H000056 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | FS1_CSS:HMR_D2232 | 0 | D2232 | 44H000061 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | FS1_CSS:HMR_D2255 | 0 | D2255 | 44H000060 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | FS1_CSS:HMR_D2318 | 0 | D2318 | 44H000058 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | FS1_CSS:HMR_D2385 | 0 | D2385 | 44H000059 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | FS1_BMS:HMR_D2504 | 0 | D2504 | 44H000062 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS1_BTS:BPM_D2116 | BPM | D2116 | 34E000274 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 9 | 10 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS1_BTS:BPM_D2116 | BPM | D2116 | 34E000275 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 9 | 10 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS1_BTS:BPM_D2116 | BPM | D2116 | 34E000276 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 9 | 10 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS1_BTS:BPM_D2116 | BPM | D2116 | 34E000277 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 9 | 10 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS1_BTS:BPM_D2133 | BPM | D2133 | 34E000232 | 2C_1/4in_Hardline_036 | Ref | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS1_BTS:BPM_D2130 | BPM | D2130 | 34E000278 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 9 | 10 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS1_BTS:BPM_D2130 | BPM | D2130 | 34E000279 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 9 | 10 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS1_BTS:BPM_D2130 | BPM | D2130 | 34E000280 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 9 | 10 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS1_BTS:BPM_D2130 | BPM | D2130 | 34E000281 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 9 | 10 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS1_BTS:RFS_D2149 | BPM | D2149 | 34E000275 | 2C_1/4in_Hardline_036 | No Term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 8 | 0 | 0 | 0 | 0
BSM | FS1_CSS:LPM_D2263 | BSM | D2263 | 44F000181 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 9 | 9 | CH4 - fmc1-1 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 1_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 6 | 2: EBS8003 | Ch0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 1_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 9 | 3 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 2_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 6 | 2: EBS8003 | Ch1 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 2_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 9 | 3 | CH7 - fmc1-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 3_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 6 | 2: EBS8003 | Ch2 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 3_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 9 | 4 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 4_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 6 | 2: EBS8003 | Ch3 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 4_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 9 | 4 | CH7 - fmc1-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 5_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 6 | 2: EBS8003 | Ch4 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 5_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 9 | 5 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 6_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 6 | 2: EBS8003 | Ch5 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 6_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 9 | 5 | CH7 - fmc1-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 7_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 6 | 2: EBS8003 | Ch6 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 7_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 9 | 6 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 8_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 6 | 2: EBS8003 | Ch7 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 8_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 9 | 9 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Grounding Plug | 0 | ground | 0 | - | Lemo650 Plug | Ground | Pico-8 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Ground Plug | 9 | 9 | CH7 - fmc1-4 | 0 | 0 | 0 | 0 | 0
BPM | FS1_BBS:BPM_D2466 | BPM | D2466 | 44E000134 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 10 | 3 | 0 | 1 | 0 | 0 | 0 | 0
BPM | FS1_BBS:BPM_D2466 | BPM | D2466 | 44E000140 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 10 | 3 | 0 | 2 | 0 | 0 | 0 | 0
BPM | FS1_BBS:BPM_D2466 | BPM | D2466 | 44E000141 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 10 | 3 | 0 | 3 | 0 | 0 | 0 | 0
BPM | FS1_BBS:BPM_D2466 | BPM | D2466 | 44E000142 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 10 | 3 | 0 | 4 | 0 | 0 | 0 | 0
BPM | FS1_BBS:BPM_D2466 | BPM | D2466 | 44E000139 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | FS1_BTS:BPM_D2467 | BPM | D2467 | 44E000174 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 10 | 3 | 0 | 5 | 0 | 0 | 0 | 0
BPM | FS1_BTS:BPM_D2467 | BPM | D2467 | 44E000175 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 10 | 3 | 0 | 6 | 0 | 0 | 0 | 0
BPM | FS1_BTS:BPM_D2467 | BPM | D2467 | 44E000176 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 10 | 3 | 0 | 7 | 0 | 0 | 0 | 0
BPM | FS1_BTS:BPM_D2467 | BPM | D2467 | 44E000177 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 10 | 3 | 0 | 8 | 0 | 0 | 0 | 0
BPM | FS1_BTS:BPM_D2467 | BPM | D2467 | 44E000178 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 10 | 3 | 0 | Ref | 0 | 0 | 0 | 0
BPM | FS1_BMS:BPM_D2502 | BPM | D2502 | 44E000147 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 10 | 4 | 0 | 1 | 0 | 0 | 0 | 0
BPM | FS1_BMS:BPM_D2502 | BPM | D2502 | 44E000148 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 10 | 4 | 0 | 2 | 0 | 0 | 0 | 0
BPM | FS1_BMS:BPM_D2502 | BPM | D2502 | 44E000149 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 10 | 4 | 0 | 3 | 0 | 0 | 0 | 0
BPM | FS1_BMS:BPM_D2502 | BPM | D2502 | 44E000150 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 10 | 4 | 0 | 4 | 0 | 0 | 0 | 0
BPM | FS1_BMS:BPM_D2502 | BPM | D2502 | 44E000155 | 2C_1/4in_Hardline_036 | Ref | No terminiation | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 10 | 4 | 0 | Ref | 0 | 0 | 0 | 0
BPM | FS1_BMS:BPM_D2537 | BPM | D2537 | 44E000151 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 10 | 4 | 0 | 5 | 0 | 0 | 0 | 0
BPM | FS1_BMS:BPM_D2537 | BPM | D2537 | 44E000152 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 10 | 4 | 0 | 6 | 0 | 0 | 0 | 0
BPM | FS1_BMS:BPM_D2537 | BPM | D2537 | 44E000153 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 10 | 4 | 0 | 7 | 0 | 0 | 0 | 0
BPM | FS1_BMS:BPM_D2537 | BPM | D2537 | 44E000154 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 10 | 4 | 0 | 8 | 0 | 0 | 0 | 0
BPM | FS1_BMS:BPM_D2537 | BPM | D2537 | 44E000156 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | FS1_BMS:BPM_D2587 | BPM | D2587 | 44E000143 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 10 | 5 | 0 | 1 | 0 | 0 | 0 | 0
BPM | FS1_BMS:BPM_D2587 | BPM | D2587 | 44E000144 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 10 | 5 | 0 | 2 | 0 | 0 | 0 | 0
BPM | FS1_BMS:BPM_D2587 | BPM | D2587 | 44E000145 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 10 | 5 | 0 | 3 | 0 | 0 | 0 | 0
BPM | FS1_BMS:BPM_D2587 | BPM | D2587 | 44E000146 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 10 | 5 | 0 | 4 | 0 | 0 | 0 | 0
BPM | FS1_BMS:BPM_D2587 | BPM | D2587 | 44E000157 | 2C_1/4in_Hardline_036 | Ref | No terminiation | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 10 | 5 | 0 | Ref | 0 | 0 | 0 | 0
BPM | FS1_BMS:BPM_D2600 | BPM | D2600 | 44E000135 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 10 | 5 | 0 | 5 | 0 | 0 | 0 | 0
BPM | FS1_BMS:BPM_D2600 | BPM | D2600 | 44E000136 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 10 | 5 | 0 | 6 | 0 | 0 | 0 | 0
BPM | FS1_BMS:BPM_D2600 | BPM | D2600 | 44E000137 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 10 | 5 | 0 | 7 | 0 | 0 | 0 | 0
BPM | FS1_BMS:BPM_D2600 | BPM | D2600 | 44E000138 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 10 | 5 | 0 | 8 | 0 | 0 | 0 | 0
BPM | FS1_BMS:BPM_D2600 | BPM | D2600 | 44E000158 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | FS1_BMS:BPM_D2665 | BPM | D2665 | 44E000159 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 10 | 6 | 0 | 1 | 0 | 0 | 0 | 0
BPM | FS1_BMS:BPM_D2665 | BPM | D2665 | 44E000170 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 10 | 6 | 0 | 2 | 0 | 0 | 0 | 0
BPM | FS1_BMS:BPM_D2665 | BPM | D2665 | 44E000171 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 10 | 6 | 0 | 3 | 0 | 0 | 0 | 0
BPM | FS1_BMS:BPM_D2665 | BPM | D2665 | 44E000172 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 10 | 6 | 0 | 4 | 0 | 0 | 0 | 0
BPM | FS1_BMS:BPM_D2665 | BPM | D2665 | 44E000173 | 2C_1/4in_Hardline_036 | Ref | No terminiation | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 10 | 6 | 0 | Ref | 0 | 0 | 0 | 0
BPM | FS1_BMS:BPM_D2690 | BPM | D2690 | 44E000166 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 10 | 6 | 0 | 5 | 0 | 0 | 0 | 0
BPM | FS1_BMS:BPM_D2690 | BPM | D2690 | 44E000167 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 10 | 6 | 0 | 6 | 0 | 0 | 0 | 0
BPM | FS1_BMS:BPM_D2690 | BPM | D2690 | 44E000168 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 10 | 6 | 0 | 7 | 0 | 0 | 0 | 0
BPM | FS1_BMS:BPM_D2690 | BPM | D2690 | 44E000169 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 10 | 6 | 0 | 8 | 0 | 0 | 0 | 0
BPM | FS1_BMS:BPM_D2690 | BPM | D2690 | 44E000160 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | FS1_BMS:BPM_D2702 | BPM | D2702 | 44E000162 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 10 | 7 | 0 | 1 | 0 | 0 | 0 | 0
BPM | FS1_BMS:BPM_D2702 | BPM | D2702 | 44E000163 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 10 | 7 | 0 | 2 | 0 | 0 | 0 | 0
BPM | FS1_BMS:BPM_D2702 | BPM | D2702 | 44E000164 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 10 | 7 | 0 | 3 | 0 | 0 | 0 | 0
BPM | FS1_BMS:BPM_D2702 | BPM | D2702 | 44E000165 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 10 | 7 | 0 | 4 | 0 | 0 | 0 | 0
BPM | FS1_BMS:BPM_D2702 | BPM | D2702 | 44E000161 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS2_WC01:BPM_D2742 | BPM | D2742 | 34E000395 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 10 | 7 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS2_WC01:BPM_D2742 | BPM | D2742 | 34E000396 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 10 | 7 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS2_WC01:BPM_D2742 | BPM | D2742 | 34E000397 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 10 | 7 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS2_WC01:BPM_D2742 | BPM | D2742 | 34E000398 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 10 | 7 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS2_WC01:BPM_D2742 | BPM | D2742 | 34E000411 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 10 | 7 | 0 | Ref | 0 | 0 | 0 | 0
PM 12" | FS1_BBS:PM_D2444 | PM | D2444 | 44F000084 | 6C_22AWG_Multi_010 | A Signal | CPC to LEMO650 Box | CPC-1 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-1A | LEMO650/LEMO650 | 10 | 8 | CH0 - fmc2-1 | 0 | 7 | 1: EBS8005 | Ch0 | 0
PM 12" | FS1_BBS:PM_D2444 | PM | D2444 | 44F000084 | 6C_22AWG_Multi_010 | B Signal | CPC to LEMO650 Box | CPC-1 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | LEMO-1B | LEMO650/LEMO650 | 10 | 8 | CH1 - fmc2-2 | 0 | 7 | 1: EBS8005 | Ch0 | 0
PM 6" | FS1_BBS:PM_D2444 | PM | D2444 | 44F000080 | 4C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-1 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-1C | LEMO650/LEMO650 | 10 | 8 | CH2 - fmc2-3 | 0 | 7 | 1: EBS8005 | Ch0 | 0
PM 12" | FS1_BBS:PM_D2482 | PM | D2482 | 44F000088 | 6C_22AWG_Multi_010 | A Signal | CPC to LEMO650 Box | CPC-2 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-2A | LEMO650/LEMO650 | 10 | 8 | CH4 - fmc1-1 | 0 | 7 | 1: EBS8005 | Ch1 | 0
PM 12" | FS1_BBS:PM_D2482 | PM | D2482 | 44F000088 | 6C_22AWG_Multi_010 | B Signal | CPC to LEMO650 Box | CPC-2 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | LEMO-2B | LEMO650/LEMO650 | 10 | 8 | CH5 - fmc1-2 | 0 | 7 | 1: EBS8005 | Ch1 | 0
PM 6" | FS1_BBS:PM_D2482 | PM | D2482 | 44F000091 | 4C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-2 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-2C | LEMO650/LEMO650 | 10 | 8 | CH6 - fmc1-3 | 0 | 7 | 1: EBS8005 | Ch1 | 0
PM 12" | FS1_BTS:PM_D2470 | PM | D2470 | 44F000072 | 6C_22AWG_Multi_010 | A Signal | CPC to LEMO650 Box | CPC-3 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-3A | LEMO650/LEMO650 | 10 | 9 | CH0 - fmc2-1 | 0 | 7 | 1: EBS8005 | Ch2 | 0
PM 12" | FS1_BTS:PM_D2470 | PM | D2470 | 44F000072 | 6C_22AWG_Multi_010 | B Signal | CPC to LEMO650 Box | CPC-3 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | LEMO-3B | LEMO650/LEMO650 | 10 | 9 | CH1 - fmc2-2 | 0 | 7 | 1: EBS8005 | Ch2 | 0
PM 6" | FS1_BTS:PM_D2470 | PM | D2470 | 44F000076 | 4C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-3 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-3C | LEMO650/LEMO650 | 10 | 9 | CH2 - fmc2-3 | 0 | 7 | 1: EBS8005 | Ch2 | 0
PM 8" | FS1_BMS:PM_D2552 | PM | D2504 | 44F000063 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-4 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-4A | LEMO650/LEMO650 | 10 | 9 | CH4 - fmc1-1 | 0 | 7 | 1: EBS8005 | Ch3 | 0
PM 8" | FS1_BMS:PM_D2552 | PM | D2504 | 44F000063 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-4 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | LEMO-4B | LEMO650/LEMO650 | 10 | 9 | CH5 - fmc1-2 | 0 | 7 | 1: EBS8005 | Ch3 | 0
PM 8" | FS1_BMS:PM_D2552 | PM | D2504 | 44F000063 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-4 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-4C | LEMO650/LEMO650 | 10 | 9 | CH6 - fmc1-3 | 0 | 7 | 1: EBS8005 | Ch3 | 0
PM 8" | FS1_BMS:PM_D2602 | PM | D2602 | 44F000064 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-5 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-5A | LEMO650/LEMO650 | 10 | 10 | CH0 - fmc2-1 | 0 | 7 | 1: EBS8005 | Ch4 | 0
PM 8" | FS1_BMS:PM_D2602 | PM | D2602 | 44F000064 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-5 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield-6 | 4 | 6 | LEMO-5B | LEMO650/LEMO650 | 10 | 10 | CH1 - fmc2-2 | 0 | 7 | 1: EBS8005 | Ch4 | 0
PM 8" | FS1_BMS:PM_D2602 | PM | D2602 | 44F000064 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-5 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-6C | LEMO650/LEMO650 | 10 | 10 | CH2 - fmc2-3 | 0 | 7 | 1: EBS8005 | Ch4 | 0
PM 8" | FS1_BMS:PM_D2703 | PM | D2703 | 44F000068 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-6 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-7A | LEMO650/LEMO650 | 10 | 10 | CH4 - fmc1-1 | 0 | 7 | 1: EBS8005 | Ch5 | 0
PM 8" | FS1_BMS:PM_D2703 | PM | D2703 | 44F000068 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-6 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | LEMO-7B | LEMO650/LEMO650 | 10 | 10 | CH5 - fmc1-2 | 0 | 7 | 1: EBS8005 | Ch5 | 0
PM 8" | FS1_BMS:PM_D2703 | PM | D2703 | 44F000068 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-6 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-7C | LEMO650/LEMO650 | 10 | 10 | CH6 - fmc1-3 | 0 | 7 | 1: EBS8005 | Ch5 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 1_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 7 | 2: EBS8003 | Ch0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 1_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 10 | 8 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 2_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 7 | 2: EBS8003 | Ch1 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 2_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 10 | 8 | CH7 - fmc1-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 3_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 7 | 2: EBS8003 | Ch2 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 3_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 10 | 9 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 4_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 7 | 2: EBS8003 | Ch3 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 4_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 10 | 9 | CH7 - fmc1-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 5_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 7 | 2: EBS8003 | Ch4 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 5_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 10 | 10 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 6_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 1 | 2: EBS8003 | Ch5 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 6_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 10 | 10 | CH7 - fmc1-4 | 0 | 0 | 0 | 0 | 0
BPM | LS2_WC02:BPM_D2782 | BPM | D2782 | 34E000399 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 11 | 3 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS2_WC02:BPM_D2782 | BPM | D2782 | 34E000400 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 11 | 3 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS2_WC02:BPM_D2782 | BPM | D2782 | 34E000401 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 11 | 3 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS2_WC02:BPM_D2782 | BPM | D2782 | 34E000402 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 11 | 3 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS2_WC02:BPM_D2782 | BPM | D2782 | 34E000412 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS2_WC03:BPM_D2821 | BPM | D2821 | 34E000407 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 11 | 3 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS2_WC03:BPM_D2821 | BPM | D2821 | 34E000408 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 11 | 3 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS2_WC03:BPM_D2821 | BPM | D2821 | 34E000409 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 11 | 3 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS2_WC03:BPM_D2821 | BPM | D2821 | 34E000410 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 11 | 3 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS2_WC03:BPM_D2821 | BPM | D2821 | 34E000413 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 11 | 3 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS2_WC04:BPM_D2861 | BPM | D2861 | 34E000403 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 11 | 4 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS2_WC04:BPM_D2861 | BPM | D2861 | 34E000404 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 11 | 4 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS2_WC04:BPM_D2861 | BPM | D2861 | 34E000405 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 11 | 4 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS2_WC04:BPM_D2861 | BPM | D2861 | 34E000406 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 11 | 4 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS2_WC04:BPM_D2861 | BPM | D2861 | 34E000414 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS2_WC05:BPM_D2901 | BPM | D2901 | 34E000415 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 11 | 4 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS2_WC05:BPM_D2901 | BPM | D2901 | 34E000416 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 11 | 4 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS2_WC05:BPM_D2901 | BPM | D2901 | 34E000417 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 11 | 4 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS2_WC05:BPM_D2901 | BPM | D2901 | 34E000418 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 11 | 4 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS2_WC05:BPM_D2901 | BPM | D2901 | 34E000420 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 11 | 4 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS2_WC06:BPM_D2941 | BPM | D2941 | 34E000421 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 11 | 5 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS2_WC06:BPM_D2941 | BPM | D2941 | 34E000422 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 11 | 5 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS2_WC06:BPM_D2941 | BPM | D2941 | 34E000423 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 11 | 5 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS2_WC06:BPM_D2941 | BPM | D2941 | 34E000424 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 11 | 5 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS2_WC06:BPM_D2941 | BPM | D2941 | 34E000419 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS2_WC07:BPM_D2981 | BPM | D2981 | 34E000425 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 11 | 5 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS2_WC07:BPM_D2981 | BPM | D2981 | 34E000426 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 11 | 5 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS2_WC07:BPM_D2981 | BPM | D2981 | 34E000427 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 11 | 5 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS2_WC07:BPM_D2981 | BPM | D2981 | 34E000428 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 11 | 5 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS2_WC07:BPM_D2981 | BPM | D2981 | 34E000469 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 11 | 5 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS2_WC08:BPM_D3020 | BPM | D3020 | 34E000433 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 11 | 6 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS2_WC08:BPM_D3020 | BPM | D3020 | 34E000434 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 11 | 6 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS2_WC08:BPM_D3020 | BPM | D3020 | 34E000435 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 11 | 6 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS2_WC08:BPM_D3020 | BPM | D3020 | 34E000436 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 11 | 6 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS2_WC08:BPM_D3020 | BPM | D3020 | 34E000470 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS2_WC09:BPM_D3060 | BPM | D3060 | 34E000437 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 11 | 6 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS2_WC09:BPM_D3060 | BPM | D3060 | 34E000438 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 11 | 6 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS2_WC09:BPM_D3060 | BPM | D3060 | 34E000439 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 11 | 6 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS2_WC09:BPM_D3060 | BPM | D3060 | 34E000440 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 11 | 6 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS2_WC09:BPM_D3060 | BPM | D3060 | 34E000471 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 11 | 6 | 0 | Ref | 0 | 0 | 0 | 0
Neutron Detector | FS1_BMS:ND_D2588 | ND | D2588 | 44E000229 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | SMA to Lemo650 | 11 | 9 | CH0 - fmc2-1 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS2_CC02:ND_D2748 | ND | D2748 | 34E000536 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | SMA to Lemo650 | 11 | 9 | CH1 - fmc2-2 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS2_CC04:ND_D2828 | ND | D2828 | 34E000538 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | SMA to Lemo650 | 11 | 9 | CH2 - fmc2-3 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS2_CC06:ND_D2916 | ND | D2916 | 34E000540 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | SMA to Lemo650 | 11 | 9 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Neutron Detector | FS1_BMS:ND_D2588 | 0 | D2588 | 44E000228 | 2C_75LC_Coax_089 | Test Signal | N-Type to SMA Panel | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to Ferrule | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS2_CC02:ND_D2748 | 0 | D2748 | 34E000535 | 2C_75LC_Coax_089 | Test Signal | N-Type to SMA Panel |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to Ferrule | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS2_CC04:ND_D2828 | 0 | D2828 | 34E000537 | 2C_75LC_Coax_089 | Test Signal | N-Type to SMA Panel | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to Ferrule | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS2_CC06:ND_D2916 | 0 | D2916 | 34E000539 | 2C_75LC_Coax_089 | Test Signal | N-Type to SMA Panel | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to Ferrule | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Neutron Detector | FS1_BMS:ND_D2588 | 0 | D2588 | 44H000013 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | SHV to SHV | 0 | 0 | 0 | 0 | 8 | 2 : EHS8020n | Ch0 | 0
Neutron Detector | LS2_CC02:ND_D2748 | 0 | D2748 | 34H000064 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | SHV to SHV | 0 | 0 | 0 | 0 | 8 | 2 : EHS8020n | Ch1 | 0
Neutron Detector | LS2_CC04:ND_D2828 | 0 | D2828 | 34H000065 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | SHV to SHV | 0 | 0 | 0 | 0 | 8 | 2 : EHS8020n | Ch2 | 0
Neutron Detector | LS2_CC06:ND_D2916 | 0 | D2916 | 34H000066 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | SHV to SHV | 0 | 0 | 0 | 0 | 8 | 2 : EHS8020n | Ch3 | 0
Faraday Cup | FS1_BMS:FC_D2634 | FC | D2605 | 44H000011 | 2C_1kV75LC_Coax_090 | Signal | N-Type to SMA Panel | J1 | 0 | 0 | 0 | 0 | 0 | 0 | J1 | SMA/LEMO650 | 11 | 7 | CH4 - fmc1-1 | 0 | 0 | 0 | 0 | 0
Faraday Cup | FS1_BMS:FC_D2634 | 0 | D2605 | 44H000063 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 5 | SHV to SHV | 0 | 0 | 0 | 0 | 8 | 2 : EHS8020n | Ch4 | 0
Halo Monitor Ring | FS1_BMS:HMR_D2703 | HMR | D2703 | 44H000009 | 2C_1kV75LC_Coax_090 | Signal | SHV to Lemo | J1 | 0 | 0 | 0 | 0 | 0 | 0 | J1 | Lemo to Lemo | 11 | 7 | CH0 - fmc2-1 | 0 | 8 | 1: EBS8003 | Ch0 | 0
0 | FS1_BMS:HMR_D2703 | 0 | D2703 | 44H000065 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring,Spare | LS2_WC01:HMR_D2743 | HMR - Unused | D2743 | 34H000030 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS2_WC01:HMR_D2743 | 0 | D2743 | 34H000108 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring,Spare | LS2_WC02:HMR_D2783 | HMR - Unused | D2783 | 34H000031 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS2_WC02:HMR_D2783 | 0 | D2783 | 34H000123 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring | LS2_WC03:HMR_D2823 | HMR | D2823 | 34H000032 | 2C_1kV75LC_Coax_090 | Signal | SHV to Lemo | J2 | 0 | 0 | 0 | 0 | 0 | 0 | J2 | Lemo to Lemo | 11 | 7 | CH1 - fmc2-2 | 0 | 8 | 1: EBS8003 | Ch0 | 0
0 | LS2_WC03:HMR_D2823 | 0 | D2823 | 34H000118 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring,Spare | LS2_WC04:HMR_D2862 | HMR - Unused | D2862 | 34H000033 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS2_WC04:HMR_D2862 | 0 | D2862 | 34H000121 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring,Spare | LS2_WC05:HMR_D2902 | HMR - Unused | D2902 | 34H000034 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS2_WC05:HMR_D2902 | 0 | D2902 | 34H000122 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring | LS2_WC06:HMR_D2942 | HMR | D2942 | 34H000035 | 2C_1kV75LC_Coax_090 | Signal | SHV to Lemo | J3 | 0 | 0 | 0 | 0 | 0 | 0 | J3 | Lemo to Lemo | 11 | 7 | CH2 - fmc2-3 | 0 | 8 | 1: EBS8003 | Ch0 | 0
0 | LS2_WC06:HMR_D2942 | 0 | D2942 | 34H000120 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring,Spare | LS2_WC07:HMR_D2982 | HMR - Unused | D2982 | 34H000036 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS2_WC07:HMR_D2982 | 0 | D2982 | 34H000119 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Grounding Plug | 0 | ground | 0 | - | Lemo650 Plug | Ground | Pico-8 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Ground Plug | 11 | 7 | CH7 - fmc1-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 1_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 8 | 2: EBS8003 | Ch0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 1_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 11 | 7 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
BPM | LS2_WC10:BPM_D3100 | BPM | D3100 | 34E000441 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 12 | 6 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS2_WC10:BPM_D3100 | BPM | D3100 | 34E000442 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 12 | 6 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS2_WC10:BPM_D3100 | BPM | D3100 | 34E000443 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 12 | 6 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS2_WC10:BPM_D3100 | BPM | D3100 | 34E000444 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 12 | 6 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS2_WC10:BPM_D3100 | BPM | D3100 | 34E000472 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS2_WC11:BPM_D3140 | BPM | D3140 | 34E000445 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 12 | 6 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS2_WC11:BPM_D3140 | BPM | D3140 | 34E000446 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 12 | 6 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS2_WC11:BPM_D3140 | BPM | D3140 | 34E000447 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 12 | 6 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS2_WC11:BPM_D3140 | BPM | D3140 | 34E000448 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 12 | 6 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS2_WC11:BPM_D3140 | BPM | D3140 | 34E000473 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 12 | 6 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS2_WC12:BPM_D3180 | BPM | D3180 | 34E000429 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 12 | 7 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS2_WC12:BPM_D3180 | BPM | D3180 | 34E000430 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 12 | 7 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS2_WC12:BPM_D3180 | BPM | D3180 | 34E000431 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 12 | 7 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS2_WC12:BPM_D3180 | BPM | D3180 | 34E000432 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 12 | 7 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS2_WC12:BPM_D3180 | BPM | D3180 | 34E000474 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS2_WD01:BPM_D3242 | BPM | D3242 | 34E000449 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 12 | 7 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS2_WD01:BPM_D3242 | BPM | D3242 | 34E000450 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 12 | 7 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS2_WD01:BPM_D3242 | BPM | D3242 | 34E000451 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 12 | 7 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS2_WD01:BPM_D3242 | BPM | D3242 | 34E000452 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 12 | 7 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS2_WD01:BPM_D3242 | BPM | D3242 | 34E000475 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 12 | 7 | 0 | Ref | 0 | 0 | 0 | 0
Neutron Detector | LS2_CC08:ND_D2987 | ND | D2987 | 34E000542 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | SMA to Lemo650 | 12 | 5 | CH0 - fmc2-1 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS2_CC10:ND_D3075 | ND | D3075 | 34E000543 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | SMA to Lemo650 | 12 | 5 | CH1 - fmc2-2 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS2_CC12:ND_D3155 | ND | D3155 | 34E000619 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | SMA to Lemo650 | 12 | 5 | CH2 - fmc2-3 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS2_CD02:ND_D3252 | ND | D3252 | 34E000547 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | SMA to Lemo650 | 12 | 5 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS2_CD04:ND_D3376 | ND | D3376 | 34E000549 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 5 | SMA to Lemo650 | 12 | 5 | CH4 - fmc1-1 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS2_CD06:ND_D3500 | ND | D3500 | 34E000551 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | 6 | 0 | 0 | 0 | 0 | 0 | 0 | 6 | SMA to Lemo650 | 12 | 5 | CH5 - fmc1-2 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS2_CD08:ND_D3624 | ND | D3624 | 34E000553 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | 7 | 0 | 0 | 0 | 0 | 0 | 0 | 7 | SMA to Lemo650 | 12 | 5 | CH6 - fmc1-3 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS2_CC08:ND_D2987 | 0 | D2987 | 34E000541 | 2C_75LC_Coax_089 | Test Signal | N-Type to SMA Panel | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to Ferrule | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS2_CC10:ND_D3075 | 0 | D3075 | 34E000544 | 2C_75LC_Coax_089 | Test Signal | N-Type to SMA Panel | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to Ferrule | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS2_CC12:ND_D3155 | 0 | D3155 | 34E000545 | 2C_75LC_Coax_089 | Test Signal | N-Type to SMA Panel | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to Ferrule | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS2_CD02:ND_D3252 | 0 | D3252 | 34E000546 | 2C_75LC_Coax_089 | Test Signal | N-Type to SMA Panel | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to Ferrule | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS2_CD04:ND_D3376 | 0 | D3376 | 34E000548 | 2C_75LC_Coax_089 | Test Signal | N-Type to SMA Panel | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to Ferrule | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS2_CD06:ND_D3500 | 0 | D3500 | 34E000550 | 2C_75LC_Coax_089 | Test Signal | N-Type to SMA Panel | 6 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to Ferrule | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS2_CD08:ND_D3624 | 0 | D3624 | 34E000552 | 2C_75LC_Coax_089 | Test Signal | N-Type to SMA Panel | 7 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to Ferrule | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS2_CC08:ND_D2987 | 0 | D2987 | 34H000067 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | SHV to SHV | 0 | 0 | 0 | 0 | 9 | 2 : EHS8020n | Ch0 | 0
Neutron Detector | LS2_CC10:ND_D3075 | 0 | D3075 | 34H000068 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | SHV to SHV | 0 | 0 | 0 | 0 | 9 | 2 : EHS8020n | Ch1 | 0
Neutron Detector | LS2_CC12:ND_D3155 | 0 | D3155 | 34H000069 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | SHV to SHV | 0 | 0 | 0 | 0 | 9 | 2 : EHS8020n | Ch2 | 0
Neutron Detector | LS2_CD02:ND_D3252 | 0 | D3252 | 34H000070 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | SHV to SHV | 0 | 0 | 0 | 0 | 9 | 2 : EHS8020n | Ch3 | 0
Neutron Detector | LS2_CD04:ND_D3376 | 0 | D3376 | 34H000071 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 5 | SHV to SHV | 0 | 0 | 0 | 0 | 9 | 2 : EHS8020n | Ch4 | 0
Neutron Detector | LS2_CD06:ND_D3500 | 0 | D3500 | 34H000072 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 6 | 0 | 0 | 0 | 0 | 0 | 0 | 6 | SHV to SHV | 0 | 0 | 0 | 0 | 9 | 2 : EHS8020n | Ch5 | 0
Neutron Detector | LS2_CD08:ND_D3624 | 0 | D3624 | 34H000073 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 7 | 0 | 0 | 0 | 0 | 0 | 0 | 7 | SHV to SHV | 0 | 0 | 0 | 0 | 9 | 2 : EHS8020n | Ch6 | 0
Halo Monitor Ring,Spare | LS2_WC08:HMR_D3022 | HMR - Unused | D3022 | 34H000037 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS2_WC08:HMR_D3022 | 0 | D3022 | 34H000117 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring | LS2_WC09:HMR_D3062 | HMR | D3061 | 34H000039 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J1 | 0 | 0 | 0 | 0 | 0 | 0 | J1 | SHV to Lemo | 12 | 3 | CH0 - fmc2-1 | 0 | 9 | 1:EHS8003 | Ch0 | 0
0 | LS2_WC09:HMR_D3062 | 0 | D3061 | 34H000132 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring,Spare | LS2_WC10:HMR_D3101 | HMR - Unused | D3101 | 34H000040 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS2_WC10:HMR_D3101 | 0 | D3101 | 34H000131 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring,Spare | LS2_WC11:HMR_D3141 | HMR - Unused | D3141 | 34H000041 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS2_WC11:HMR_D3141 | 0 | D3141 | 34H000130 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring | LS2_WC12:HMR_D3181 | HMR | D3181 | 34H000042 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J2 | 0 | 0 | 0 | 0 | 0 | 0 | J2 | SHV to Lemo | 12 | 3 | CH1 - fmc2-2 | 0 | 9 | 1:EHS8003 | Ch0 | 0
0 | LS2_WC12:HMR_D3181 | 0 | D3181 | 34H000126 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring,Spare | LS2_WD01:HMR_D3243 | HMR - Unused | D3243 | 34H000038 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS2_WD01:HMR_D3243 | 0 | D3243 | 34H000133 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring,Spare | LS2_WD02:HMR_D3305 | HMR - Unused | D3305 | 34H000043 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS2_WD02:HMR_D3305 | 0 | D3305 | 34H000128 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring | LS2_WD03:HMR_D3367 | HMR | D3367 | 34H000044 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J3 | 0 | 0 | 0 | 0 | 0 | 0 | J3 | SHV to Lemo | 12 | 3 | CH4 - fmc1-1 | 0 | 9 | 1:EHS8003 | Ch1 | 0
0 | LS2_WD03:HMR_D3367 | 0 | D3367 | 34H000127 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring,Spare | LS2_WD04:HMR_D3429 | HMR - Unused | D3429 | 34H000045 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS2_WD04:HMR_D3429 | 0 | D3429 | 34H000125 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring,Spare | LS2_WD05:HMR_D3491 | HMR - Unused | D3491 | 34H000046 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS2_WD05:HMR_D3491 | 0 | D3491 | 34H000129 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring | LS2_WD06:HMR_D3553 | HMR | D3553 | 34H000047 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J4 | 0 | 0 | 0 | 0 | 0 | 0 | J4 | SHV to Lemo | 12 | 3 | CH5 - fmc1-2 | 0 | 9 | 1:EHS8003 | Ch1 | 0
0 | LS2_WD06:HMR_D3553 | 0 | D3553 | 34H000124 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring,Spare | LS2_WD07:HMR_D3615 | HMR - Unused | D3615 | 34H000049 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS2_WD07:HMR_D3615 | 0 | D3615 | 34H000134 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 1_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 9 | 2: EBS8003 | Ch0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 1_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 12 | 3 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 2_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 9 | 2: EBS8003 | Ch1 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 2_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 12 | 3 | CH7 - fmc1-4 | 0 | 0 | 0 | 0 | 0
BPM | LS2_WD02:BPM_D3304 | BPM | D3304 | 34E000453 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 13 | 3 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS2_WD02:BPM_D3304 | BPM | D3304 | 34E000454 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 13 | 3 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS2_WD02:BPM_D3304 | BPM | D3304 | 34E000455 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 13 | 3 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS2_WD02:BPM_D3304 | BPM | D3304 | 34E000456 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 13 | 3 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS2_WD02:BPM_D3304 | BPM | D3304 | 34E000476 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS2_WD03:BPM_D3366 | BPM | D3366 | 34E000457 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 13 | 3 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS2_WD03:BPM_D3366 | BPM | D3366 | 34E000458 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 13 | 3 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS2_WD03:BPM_D3366 | BPM | D3366 | 34E000459 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 13 | 3 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS2_WD03:BPM_D3366 | BPM | D3366 | 34E000460 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 13 | 3 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS2_WD03:BPM_D3366 | BPM | D3366 | 34E000477 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 13 | 3 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS2_WD04:BPM_D3428 | BPM | D3428 | 34E000461 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 13 | 4 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS2_WD04:BPM_D3428 | BPM | D3428 | 34E000462 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 13 | 4 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS2_WD04:BPM_D3428 | BPM | D3428 | 34E000463 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 13 | 4 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS2_WD04:BPM_D3428 | BPM | D3428 | 34E000464 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 13 | 4 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS2_WD04:BPM_D3428 | BPM | D3428 | 34E000478 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS2_WD05:BPM_D3490 | BPM | D3490 | 34E000465 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 13 | 4 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS2_WD05:BPM_D3490 | BPM | D3490 | 34E000466 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 13 | 4 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS2_WD05:BPM_D3490 | BPM | D3490 | 34E000467 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 13 | 4 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS2_WD05:BPM_D3490 | BPM | D3490 | 34E000468 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 13 | 4 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS2_WD05:BPM_D3490 | BPM | D3490 | 34E000479 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 13 | 4 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS2_WD06:BPM_D3552 | BPM | D3552 | 34E000485 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 13 | 5 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS2_WD06:BPM_D3552 | BPM | D3552 | 34E000486 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 13 | 5 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS2_WD06:BPM_D3552 | BPM | D3552 | 34E000487 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 13 | 5 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS2_WD06:BPM_D3552 | BPM | D3552 | 34E000488 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 13 | 5 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS2_WD06:BPM_D3552 | BPM | D3552 | 34E000484 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS2_WD07:BPM_D3614 | BPM | D3614 | 34E000480 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 13 | 5 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS2_WD07:BPM_D3614 | BPM | D3614 | 34E000481 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 13 | 5 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS2_WD07:BPM_D3614 | BPM | D3614 | 34E000482 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 13 | 5 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS2_WD07:BPM_D3614 | BPM | D3614 | 34E000483 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 13 | 5 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS2_WD07:BPM_D3614 | BPM | D3614 | 34E000497 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 13 | 5 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS2_WD08:BPM_D3676 | BPM | D3676 | 34E000489 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 13 | 6 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS2_WD08:BPM_D3676 | BPM | D3676 | 34E000490 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 13 | 6 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS2_WD08:BPM_D3676 | BPM | D3676 | 34E000491 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 13 | 6 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS2_WD08:BPM_D3676 | BPM | D3676 | 34E000492 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 13 | 6 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS2_WD08:BPM_D3676 | BPM | D3676 | 34E000498 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS2_WD09:BPM_D3738 | BPM | D3738 | 34E000493 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 13 | 6 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS2_WD09:BPM_D3738 | BPM | D3738 | 34E000494 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 13 | 6 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS2_WD09:BPM_D3738 | BPM | D3738 | 34E000495 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 13 | 6 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS2_WD09:BPM_D3738 | BPM | D3738 | 34E000496 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 13 | 6 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS2_WD09:BPM_D3738 | BPM | D3738 | 34E000499 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 13 | 6 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS2_WD10:BPM_D3800 | BPM | D3800 | 34E000505 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 14 | 3 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS2_WD10:BPM_D3800 | BPM | D3800 | 34E000506 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 14 | 3 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS2_WD10:BPM_D3800 | BPM | D3800 | 34E000507 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 14 | 3 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS2_WD10:BPM_D3800 | BPM | D3800 | 34E000508 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 14 | 3 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS2_WD10:BPM_D3800 | BPM | D3800 | 34E000514 | 2C_1/4in_Hardline_036 | Ref | No terminiation | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 14 | 3 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS2_WD11:BPM_D3862 | BPM | D3862 | 34E000501 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 14 | 3 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS2_WD11:BPM_D3862 | BPM | D3862 | 34E000502 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 14 | 3 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS2_WD11:BPM_D3862 | BPM | D3862 | 34E000503 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 14 | 3 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS2_WD11:BPM_D3862 | BPM | D3862 | 34E000504 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 14 | 3 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS2_WD11:BPM_D3862 | BPM | D3862 | 34E000500 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS2_WD12:BPM_D3924 | BPM | D3924 | 34E000509 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 14 | 4 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS2_WD12:BPM_D3924 | BPM | D3924 | 34E000510 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 14 | 4 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS2_WD12:BPM_D3924 | BPM | D3924 | 34E000511 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 14 | 4 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS2_WD12:BPM_D3924 | BPM | D3924 | 34E000512 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 14 | 4 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS2_WD12:BPM_D3924 | BPM | D3924 | 34E000513 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | FS2_BTS:BPM_D3943 | BPM | D3947 | 44E000192 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 14 | 4 | 0 | 5 | 0 | 0 | 0 | 0
BPM | FS2_BTS:BPM_D3943 | BPM | D3947 | 44E000193 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 14 | 4 | 0 | 6 | 0 | 0 | 0 | 0
BPM | FS2_BTS:BPM_D3943 | BPM | D3947 | 44E000194 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 14 | 4 | 0 | 7 | 0 | 0 | 0 | 0
BPM | FS2_BTS:BPM_D3943 | BPM | D3947 | 44E000195 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 14 | 4 | 0 | 8 | 0 | 0 | 0 | 0
BPM | FS2_BTS:BPM_D3943 | BPM | D3947 | 44E000184 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 14 | 4 | 0 | Ref | 0 | 0 | 0 | 0
BPM | FS2_BTS:BPM_D3958 | BPM | D3958 | 44E000179 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 14 | 5 | 0 | 1 | 0 | 0 | 0 | 0
BPM | FS2_BTS:BPM_D3958 | BPM | D3958 | 44E000185 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 14 | 5 | 0 | 2 | 0 | 0 | 0 | 0
BPM | FS2_BTS:BPM_D3958 | BPM | D3958 | 44E000186 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 14 | 5 | 0 | 3 | 0 | 0 | 0 | 0
BPM | FS2_BTS:BPM_D3958 | BPM | D3958 | 44E000187 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 14 | 5 | 0 | 4 | 0 | 0 | 0 | 0
BPM | FS2_BTS:BPM_D3958 | BPM | D3958 | 44E000200 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | FS2_BTS:BPM_D4006 | BPM | D4012 | 44E000204 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 14 | 5 | 0 | 5 | 0 | 0 | 0 | 0
BPM | FS2_BTS:BPM_D4006 | BPM | D4012 | 44E000205 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 14 | 5 | 0 | 6 | 0 | 0 | 0 | 0
BPM | FS2_BTS:BPM_D4006 | BPM | D4012 | 44E000206 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 14 | 5 | 0 | 7 | 0 | 0 | 0 | 0
BPM | FS2_BTS:BPM_D4006 | BPM | D4012 | 44E000207 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 14 | 5 | 0 | 8 | 0 | 0 | 0 | 0
BPM | FS2_BTS:BPM_D4006 | BPM | D4012 | 44E000208 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 14 | 5 | 0 | Ref | 0 | 0 | 0 | 0
BPM | FS2_BBS:BPM_D4019 | BPM | D4017 | 44E000196 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 14 | 6 | 0 | 1 | 0 | 0 | 0 | 0
BPM | FS2_BBS:BPM_D4019 | BPM | D4017 | 44E000197 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 14 | 6 | 0 | 2 | 0 | 0 | 0 | 0
BPM | FS2_BBS:BPM_D4019 | BPM | D4017 | 44E000198 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 14 | 6 | 0 | 3 | 0 | 0 | 0 | 0
BPM | FS2_BBS:BPM_D4019 | BPM | D4017 | 44E000199 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 14 | 6 | 0 | 4 | 0 | 0 | 0 | 0
BPM | FS2_BBS:BPM_D4019 | BPM | D4017 | 44E000201 | 2C_1/4in_Hardline_036 | Ref | No terminiation | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 14 | 6 | 0 | Ref | 0 | 0 | 0 | 0
BPM | FS2_BBS:BPM_D4054 | BPM | D4054 | 44E000188 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 14 | 6 | 0 | 5 | 0 | 0 | 0 | 0
BPM | FS2_BBS:BPM_D4054 | BPM | D4054 | 44E000189 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 14 | 6 | 0 | 6 | 0 | 0 | 0 | 0
BPM | FS2_BBS:BPM_D4054 | BPM | D4054 | 44E000190 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 14 | 6 | 0 | 7 | 0 | 0 | 0 | 0
BPM | FS2_BBS:BPM_D4054 | BPM | D4054 | 44E000191 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 14 | 6 | 0 | 8 | 0 | 0 | 0 | 0
BPM | FS2_BBS:BPM_D4054 | BPM | D4054 | 44E000202 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS2_CD10:ND_D3748 | ND | D3748 | 34E000555 | 2C_75LC_Coax_089 | Signal | N-Type to Lemo Box | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | SMA to Lemo650 | 14 | 8 | CH0 - fmc2-1 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS2_CD12:ND_D3872 | ND | D3872 | 34E000557 | 2C_75LC_Coax_089 | Signal | N-Type to Lemo Box | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | SMA to Lemo650 | 14 | 8 | CH1 - fmc2-2 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS2_CD10:BLM_D3748 | 0 | D3748 | 34E000554 | 2C_75LC_Coax_089 | Test Signal | N-Type to SMA Panel | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to Ferrule | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS2_CD12:BLM_D3872 | 0 | D3872 | 34E000556 | 2C_75LC_Coax_089 | Test Signal | N-Type to SMA Panel | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | SMA to Ferrule | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Neutron Detector | LS2_CD10:ND_D3748 | 0 | D3748 | 34H000074 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | SHV to SHV | 0 | 0 | 0 | 0 | 10 | 2 : EHS8020n | Ch0 | 0
Neutron Detector | LS2_CD12:ND_D3872 | 0 | D3872 | 34H000075 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | SHV to SHV | 0 | 0 | 0 | 0 | 10 | 2 : EHS8020n | Ch1 | 0
Halo Monitor Ring,Spare | FS2_BTS:HMR_D3959 | HMR - Unused | D3959 | 44H000012 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | FS2_BTS:HMR_D3959 | 0 | D3959 | 44H000066 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring,Spare | LS2_WD08:HMR_D3677 | HMR - Unused | D3677 | 34H000050 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS2_WD08:HMR_D3677 | 0 | D3677 | 34H000137 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring | LS2_WD09:HMR_D3739 | HMR | D3739 | 34H000052 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J1 | 0 | 0 | 0 | 0 | 0 | 0 | J1 | SHV to Lemo | 14 | 7 | CH0 - fmc2-1 | 0 | 10 | 1:EHS8003 | Ch0 | 0
0 | LS2_WD09:HMR_D3739 | 0 | D3739 | 34H000136 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring,Spare | LS2_WD10:HMR_D3801 | HMR - Unused | D3801 | 34H000051 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS2_WD10:HMR_D3801 | 0 | D3801 | 34H000138 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring,Spare | LS2_WD11:HMR_D3863 | HMR - Unused | D3863 | 34H000048 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS2_WD11:HMR_D3863 | 0 | D3863 | 34H000139 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring | LS2_WD12:HMR_D3925 | HMR | D3925 | 34H000053 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J2 | 0 | 0 | 0 | 0 | 0 | 0 | J2 | SHV to Lemo | 14 | 7 | CH1 - fmc2-2 | 0 | 10 | 1:EHS8003 | Ch0 | 0
0 | LS2_WD12:HMR_D3925 | 0 | D3925 | 34H000135 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 1_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 10 | 2: EBS8003 | Ch0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 1_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 14 | 7 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
BPM | FS2_BBS:BPM_D4087 | BPM | D4088 | 44E000180 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 15 | 9 | 0 | 1 | 0 | 0 | 0 | 0
BPM | FS2_BBS:BPM_D4087 | BPM | D4088 | 44E000181 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 15 | 9 | 0 | 2 | 0 | 0 | 0 | 0
BPM | FS2_BBS:BPM_D4087 | BPM | D4088 | 44E000182 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 15 | 9 | 0 | 3 | 0 | 0 | 0 | 0
BPM | FS2_BBS:BPM_D4087 | BPM | D4088 | 44E000183 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 15 | 9 | 0 | 4 | 0 | 0 | 0 | 0
BPM | FS2_BBS:RFS_D4087 | BPM | D4088 | 44E000203 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | FS2_BMS:BPM_D4142 | BPM | D4142 | 44E000083 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 15 | 9 | 0 | 5 | 0 | 0 | 0 | 0
BPM | FS2_BMS:BPM_D4142 | BPM | D4142 | 44E000084 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 15 | 9 | 0 | 6 | 0 | 0 | 0 | 0
BPM | FS2_BMS:BPM_D4142 | BPM | D4142 | 44E000085 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 15 | 9 | 0 | 7 | 0 | 0 | 0 | 0
BPM | FS2_BMS:BPM_D4142 | BPM | D4142 | 44E000086 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 15 | 9 | 0 | 8 | 0 | 0 | 0 | 0
BPM | FS2_BMS:RFS_D4142 | BPM | D4142 | 44E000078 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 15 | 9 | 0 | Ref | 0 | 0 | 0 | 0
BPM | FS2_BMS:BPM_D4164 | BPM | D4164 | 44E000079 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 15 | 10 | 0 | 1 | 0 | 0 | 0 | 0
BPM | FS2_BMS:BPM_D4164 | BPM | D4164 | 44E000080 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 15 | 10 | 0 | 2 | 0 | 0 | 0 | 0
BPM | FS2_BMS:BPM_D4164 | BPM | D4164 | 44E000081 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 15 | 10 | 0 | 3 | 0 | 0 | 0 | 0
BPM | FS2_BMS:BPM_D4164 | BPM | D4164 | 44E000082 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 15 | 10 | 0 | 4 | 0 | 0 | 0 | 0
BPM | FS2_BMS:RFS_D4164 | BPM | D4164 | 44E000087 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | FS2_BMS:BPM_D4177 | BPM | D4178 | 44E000074 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 15 | 10 | 0 | 5 | 0 | 0 | 0 | 0
BPM | FS2_BMS:BPM_D4177 | BPM | D4178 | 44E000075 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 15 | 10 | 0 | 6 | 0 | 0 | 0 | 0
BPM | FS2_BMS:BPM_D4177 | BPM | D4178 | 44E000076 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 15 | 10 | 0 | 7 | 0 | 0 | 0 | 0
BPM | FS2_BMS:BPM_D4177 | BPM | D4178 | 44E000077 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 15 | 10 | 0 | 8 | 0 | 0 | 0 | 0
BPM | FS2_BMS:RFS_D4177 | BPM | D4178 | 44E000088 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 15 | 10 | 0 | Ref | 0 | 0 | 0 | 0
Ion Chamber | FS2_BBS:IC_D3979 | IC | D3979 | 44E000231 | 2C_1kV75LC_Coax_090 | Signal | 0 | J1 | 0 | 0 | 0 | 0 | 0 | 0 | J1 | Lemo to Lemo | 15 | 7 | CH0 - fmc2-1 | 0 | 0 | 0 | 0 | 0
Ion Chamber | FS2_BBS:IC_D3979 | 0 | D3979 | 44E000015 | 2C_5kVDC_Coax_035 | HV Bias | SHV | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | SHV to SHV | 0 | 0 | 0 | 0 | 11 | 2: EHS8020n | Ch0 | 0
Ion Chamber | FS2_BBS:IC_D4018 | IC | D4018 | 44E000232 | 2C_1kV75LC_Coax_090 | Signal | 0 | J2 | 0 | 0 | 0 | 0 | 0 | 0 | J2 | Lemo to Lemo | 15 | 7 | CH1 - fmc2-2 | 0 | 0 | 0 | 0 | 0
Ion Chamber | FS2_BBS:IC_D4018 | 0 | D4018 | 44E000017 | 2C_5kVDC_Coax_035 | HV Bias | SHV | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | SHV to SHV | 0 | 0 | 0 | 0 | 11 | 2: EHS8020n | Ch1 | 0
Ion Chamber | FS2_BBS:IC_D4055 | IC | D4055 | 44E000235 | 2C_1kV75LC_Coax_090 | Signal | 0 | J3 | 0 | 0 | 0 | 0 | 0 | 0 | J3 | Lemo to Lemo | 15 | 7 | CH2 - fmc2-3 | 0 | 0 | 0 | 0 | 0
Ion Chamber | FS2_BBS:IC_D4055 | 0 | D4055 | 44E000019 | 2C_5kVDC_Coax_035 | HV Bias | SHV | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | SHV to SHV | 0 | 0 | 0 | 0 | 11 | 2: EHS8020n | Ch2 | 0
Ion Chamber | FS2_BBS:IC_D4102 | IC | D4102 | 44E000236 | 2C_1kV75LC_Coax_090 | Signal | 0 | J4 | 0 | 0 | 0 | 0 | 0 | 0 | J4 | Lemo to Lemo | 15 | 7 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Ion Chamber | FS2_BBS:IC_D4102 | 0 | D4102 | 44E000020 | 2C_5kVDC_Coax_035 | HV Bias | SHV | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | SHV to SHV | 0 | 0 | 0 | 0 | 11 | 2: EHS8020n | Ch3 | 0
Ion Chamber | FS2_BTS:IC_D3940 | IC | D3940 | 44E000230 | 2C_1kV75LC_Coax_090 | Signal | 0 | J5 | 0 | 0 | 0 | 0 | 0 | 0 | J5 | Lemo to Lemo | 15 | 7 | CH4 - fmc1-1 | 0 | 0 | 0 | 0 | 0
Ion Chamber | FS2_BTS:IC_D3940 | 0 | D3940 | 44E000014 | 2C_5kVDC_Coax_035 | HV Bias | SHV | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 5 | SHV to SHV | 0 | 0 | 0 | 0 | 11 | 2: EHS8020n | Ch4 | 0
Ion Chamber | FS2_BTS:IC_D4018 | IC | D4018 | 44E000233 | 2C_1kV75LC_Coax_090 | Signal | 0 | J6 | 0 | 0 | 0 | 0 | 0 | 0 | J6 | Lemo to Lemo | 15 | 7 | CH5 - fmc1-2 | 0 | 0 | 0 | 0 | 0
Ion Chamber | FS2_BTS:IC_D4018 | 0 | D4018 | 44E000016 | 2C_5kVDC_Coax_035 | HV Bias | SHV | 6 | 0 | 0 | 0 | 0 | 0 | 0 | 6 | SHV to SHV | 0 | 0 | 0 | 0 | 11 | 2: EHS8020n | Ch5 | 0
Ion Chamber | FS2_BTS:IC_D4019 | IC | D4019 | 44E000234 | 2C_1kV75LC_Coax_090 | Signal | 0 | J7 | 0 | 0 | 0 | 0 | 0 | 0 | J7 | Lemo to Lemo | 15 | 7 | CH6 - fmc1-3 | 0 | 0 | 0 | 0 | 0
Ion Chamber | FS2_BTS:IC_D4019 | 0 | D4019 | 44E000018 | 2C_5kVDC_Coax_035 | HV Bias | SHV | 7 | 0 | 0 | 0 | 0 | 0 | 0 | 7 | SHV to SHV | 0 | 0 | 0 | 0 | 11 | 2: EHS8020n | Ch6 | 0
PM 6" | FE_LEBT:PM_D0945 | PM | D0945 | 24F000107 | 4C_22AWG_Multi_006 | C Signal | CPC to LEMO650 Box | CPC-1 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | Lemo-1A | LEMO650/LEMO650 | 15 | 3 | CH2 - fmc2-3 | 0 | 11 | 1: EBS8005 | Ch0 | 0
PM 12" | FE_LEBT:PM_D0945 | PM | D0945 | 24F000110 | 6C_22AWG_Multi_010 | A Signal | CPC to LEMO650 Box | CPC-1 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | Lemo-1B | LEMO650/LEMO650 | 15 | 3 | CH0 - fmc2-1 | 0 | 11 | 1: EBS8005 | Ch0 | 0
PM 12" | FE_LEBT:PM_D0945 | PM | D0945 | 24F000110 | 6C_22AWG_Multi_010 | B Signal | CPC to LEMO650 Box | CPC-1 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | Lemo-1C | LEMO650/LEMO650 | 15 | 3 | CH1 - fmc2-2 | 0 | 11 | 1: EBS8005 | Ch0 | 0
PM 6" | FE_LEBT:PM_D0961 | PM | D0962 | 24F000102 | 4C_22AWG_Multi_006 | C Signal | CPC to LEMO650 Box | CPC-2 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | Lemo-2A | LEMO650/LEMO650 | 15 | 3 | CH6 - fmc1-3 | 0 | 11 | 1: EBS8005 | Ch1 | 0
PM 12" | FE_LEBT:PM_D0961 | PM | D0962 | 24F000053 | 6C_22AWG_Multi_010 | A Signal | CPC to LEMO650 Box | CPC-2 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | Lemo-2B | LEMO650/LEMO650 | 15 | 3 | CH4 - fmc1-1 | 0 | 11 | 1: EBS8005 | Ch1 | 0
PM 12" | FE_LEBT:PM_D0961 | PM | D0962 | 24F000053 | 6C_22AWG_Multi_010 | B Signal | CPC to LEMO650 Box | CPC-2 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | Lemo-2C | LEMO650/LEMO650 | 15 | 3 | CH5 - fmc1-2 | 0 | 11 | 1: EBS8005 | Ch1 | 0
PM 6" | FE_LEBT:PM_D0972 | PM | D0976 | 24F000094 | 4C_22AWG_Multi_006 | C Signal | CPC to LEMO650 Box | CPC-3 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | Lemo-3A | LEMO650/LEMO650 | 15 | 4 | CH2 - fmc2-3 | 0 | 11 | 1: EBS8005 | Ch2 | 0
PM 12" | FE_LEBT:PM_D0972 | PM | D0976 | 24F000098 | 6C_22AWG_Multi_010 | A Signal | CPC to LEMO650 Box | CPC-3 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | Lemo-3B | LEMO650/LEMO650 | 15 | 4 | CH0 - fmc2-1 | 0 | 11 | 1: EBS8005 | Ch2 | 0
PM 12" | FE_LEBT:PM_D0972 | PM | D0976 | 24F000098 | 6C_22AWG_Multi_010 | B Signal | CPC to LEMO650 Box | CPC-3 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | Lemo-3C | LEMO650/LEMO650 | 15 | 4 | CH1 - fmc2-2 | 0 | 11 | 1: EBS8005 | Ch2 | 0
PM 8" | FE_LEBT:PM_D0986 | PM | D0986 | 24F000091 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-4 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | Lemo-4A | LEMO650/LEMO650 | 15 | 4 | CH6 - fmc1-3 | 0 | 11 | 1: EBS8005 | Ch3 | 0
PM 8" | FE_LEBT:PM_D0986 | PM | D0986 | 24F000091 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-4 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | Lemo-4B | LEMO650/LEMO650 | 15 | 4 | CH5 - fmc1-2 | 0 | 11 | 1: EBS8005 | Ch3 | 0
PM 8" | FE_LEBT:PM_D0986 | PM | D0986 | 24F000091 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-4 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | Lemo-4C | LEMO650/LEMO650 | 15 | 4 | CH4 - fmc1-1 | 0 | 11 | 1: EBS8005 | Ch3 | 0
PM 6" | FS2_BBS:PM_D4051 | PM | D4051 | 44F000013 | 4C_22AWG_Multi_006 | C Signal | CPC to LEMO650 Box | CPC-5 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | Lemo-5A | LEMO650/LEMO650 | 15 | 5 | CH2 - fmc2-3 | 0 | 11 | 1: EBS8005 | Ch4 | 0
PM 12" | FS2_BBS:PM_D4051 | PM | D4051 | 44F000012 | 6C_22AWG_Multi_010 | A Signal | CPC to LEMO650 Box | CPC-5 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | Lemo-5B | LEMO650/LEMO650 | 15 | 5 | CH0 - fmc2-1 | 0 | 11 | 1: EBS8005 | Ch4 | 0
PM 12" | FS2_BBS:PM_D4051 | PM | D4051 | 44F000012 | 6C_22AWG_Multi_010 | B Signal | CPC to LEMO650 Box | CPC-5 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | Lemo-5C | LEMO650/LEMO650 | 15 | 5 | CH1 - fmc2-2 | 0 | 11 | 1: EBS8005 | Ch4 | 0
PM 8" | FS2_BMS:PM_D4144 | PM | D4143 | 44F000096 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-6 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | Lemo-6A | LEMO650/LEMO650 | 15 | 5 | CH6 - fmc1-3 | 0 | 11 | 1: EBS8005 | Ch5 | 0
PM 8" | FS2_BMS:PM_D4144 | PM | D4143 | 44F000096 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-6 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | Lemo-6B | LEMO650/LEMO650 | 15 | 5 | CH5 - fmc1-2 | 0 | 11 | 1: EBS8005 | Ch5 | 0
PM 8" | FS2_BMS:PM_D4144 | PM | D4143 | 44F000096 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-6 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | Lemo-6C | LEMO650/LEMO650 | 15 | 5 | CH4 - fmc1-1 | 0 | 11 | 1: EBS8005 | Ch5 | 0
PM 8" | FS2_BTS:PM_D3959 | PM | D3958 | 44F000003 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-7 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | Lemo-7A | LEMO650/LEMO650 | 15 | 6 | CH2 - fmc2-3 | 0 | 11 | 1: EBS8005 | Ch6 | 0
PM 8" | FS2_BTS:PM_D3959 | PM | D3958 | 44F000003 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-7 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | Lemo-7B | LEMO650/LEMO650 | 15 | 6 | CH1 - fmc2-2 | 0 | 11 | 1: EBS8005 | Ch6 | 0
PM 8" | FS2_BTS:PM_D3959 | PM | D3958 | 44F000003 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-7 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | Lemo-7C | LEMO650/LEMO650 | 15 | 6 | CH0 - fmc2-1 | 0 | 11 | 1: EBS8005 | Ch6 | 0
PM 6" | FS2_BTS:PM_D4009 | PM | D4013 | 44F000006 | 4C_22AWG_Multi_006 | C Signal | CPC to LEMO650 Box | CPC-8 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | Lemo-8A | LEMO650/LEMO650 | 15 | 6 | CH6 - fmc1-3 | 0 | 11 | 1: EBS8005 | Ch7 | 0
PM 12" | FS2_BTS:PM_D4009 | PM | D4013 | 44F000004 | 6C_22AWG_Multi_010 | A Signal | CPC to LEMO650 Box | CPC-8 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | Lemo-8B | LEMO650/LEMO650 | 15 | 6 | CH4 - fmc1-1 | 0 | 11 | 1: EBS8005 | Ch7 | 0
PM 12" | FS2_BTS:PM_D4009 | PM | D4013 | 44F000004 | 6C_22AWG_Multi_010 | B Signal | CPC to LEMO650 Box | CPC-8 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | Lemo-8C | LEMO650/LEMO650 | 15 | 6 | CH5 - fmc1-2 | 0 | 11 | 1: EBS8005 | Ch7 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 1_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 11 | 2: EBS8003 | Ch0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 1_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 15 | 3 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 2_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 11 | 2: EBS8003 | Ch1 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 2_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 15 | 3 | CH7 - fmc1-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 3_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 11 | 2: EBS8003 | Ch2 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 3_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 15 | 4 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 4_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 11 | 2: EBS8003 | Ch3 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 4_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 15 | 4 | CH7 - fmc1-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 5_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 11 | 2: EBS8003 | Ch4 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 5_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 15 | 5 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 6_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 11 | 2: EBS8003 | Ch5 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 6_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 15 | 5 | CH7 - fmc1-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 7_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 11 | 2: EBS8003 | Ch6 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 7_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 15 | 6 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 8_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 11 | 2: EBS8003 | Ch7 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 8_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 15 | 6 | CH7 - fmc1-4 | 0 | 0 | 0 | 0 | 0
BPM | LS3_WD02:BPM_D4451 | BPM | D4451 | 34E000088 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 16 | 3 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS3_WD02:BPM_D4451 | BPM | D4451 | 34E000089 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 16 | 3 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS3_WD02:BPM_D4451 | BPM | D4451 | 34E000090 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 16 | 3 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS3_WD02:BPM_D4451 | BPM | D4451 | 34E000091 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 16 | 3 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS3_WD02:BPM_D4451 | BPM | D4451 | 34E000098 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS3_WD03:BPM_D4513 | BPM | D4513 | 34E000084 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 16 | 3 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS3_WD03:BPM_D4513 | BPM | D4513 | 34E000085 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 16 | 3 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS3_WD03:BPM_D4513 | BPM | D4513 | 34E000086 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 16 | 3 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS3_WD03:BPM_D4513 | BPM | D4513 | 34E000087 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 16 | 3 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS3_WD03:BPM_D4513 | BPM | D4513 | 34E000097 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 16 | 3 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS3_WD04:BPM_D4575 | BPM | D4575 | 34E000080 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 16 | 4 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS3_WD04:BPM_D4575 | BPM | D4575 | 34E000081 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 16 | 4 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS3_WD04:BPM_D4575 | BPM | D4575 | 34E000082 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 16 | 4 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS3_WD04:BPM_D4575 | BPM | D4575 | 34E000083 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 16 | 4 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS3_WD04:BPM_D4575 | BPM | D4575 | 34E000096 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS3_WD05:BPM_D4637 | BPM | D4637 | 34E000075 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 16 | 4 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS3_WD05:BPM_D4637 | BPM | D4637 | 34E000076 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 16 | 4 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS3_WD05:BPM_D4637 | BPM | D4637 | 34E000077 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 16 | 4 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS3_WD05:BPM_D4637 | BPM | D4637 | 34E000078 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 16 | 4 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS3_WD05:BPM_D4637 | BPM | D4637 | 34E000079 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 16 | 4 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS3_WD06:BPM_D4699 | BPM | D4699 | 34E000025 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 16 | 5 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS3_WD06:BPM_D4699 | BPM | D4699 | 34E000026 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 16 | 5 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS3_WD06:BPM_D4699 | BPM | D4699 | 34E000027 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 16 | 5 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS3_WD06:BPM_D4699 | BPM | D4699 | 34E000028 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 16 | 5 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS3_WD06:BPM_D4699 | BPM | D4699 | 34E000029 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D4753 | BPM | D4753 | 34E000069 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 16 | 5 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D4753 | BPM | D4753 | 34E000070 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 16 | 5 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D4753 | BPM | D4753 | 34E000071 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 16 | 5 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D4753 | BPM | D4753 | 34E000072 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 16 | 5 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D4753 | BPM | D4753 | 34E000074 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 16 | 5 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D4769 | BPM | D4769 | 34E000064 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 16 | 6 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D4769 | BPM | D4769 | 34E000065 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 16 | 6 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D4769 | BPM | D4769 | 34E000066 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 16 | 6 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D4769 | BPM | D4769 | 34E000067 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 16 | 6 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D4769 | BPM | D4769 | 34E000073 | 2C_1/4in_Hardline_036 | Ref | No terminiation | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 16 | 6 | 0 | Ref | 0 | 0 | 0 | 0
Ion Chamber | LS3_CD05:IC_D4602 | IC | D4602 | 34E000560 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J1 | 0 | 0 | 0 | 0 | 0 | 0 | J1 | Lemo to Lemo | 16 | 7 | CH0 - fmc2-1 | 0 | 0 | 0 | 0 | 0
Ion Chamber | LS3_BTS:IC_D4701 | IC | D4701 | 34E000563 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J2 | 0 | 0 | 0 | 0 | 0 | 0 | J2 | Lemo to Lemo | 16 | 7 | CH1 - fmc2-2 | 0 | 0 | 0 | 0 | 0
Ion Chamber | LS3_CD05:IC_D4602 | 0 | D4602 | 34H000078 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | SHV to SHV | 0 | 0 | 0 | 0 | 12 | 1 : EHS8020n | Ch0 | 0
Ion Chamber | LS3_BTS:IC_D4701 | 0 | D4701 | 34H000081 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | SHV to SHV | 0 | 0 | 0 | 0 | 12 | 1 : EHS8020n | Ch1 | 0
BPM | LS3_BTS:BPM_D4843 | BPM | D4843 | 34E000060 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 17 | 3 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D4843 | BPM | D4843 | 34E000061 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 17 | 3 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D4843 | BPM | D4843 | 34E000062 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 17 | 3 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D4843 | BPM | D4843 | 34E000063 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 17 | 3 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D4843 | BPM | D4843 | 34E000059 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D4886 | BPM | D4886 | 34E000055 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 17 | 3 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D4886 | BPM | D4886 | 34E000056 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 17 | 3 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D4886 | BPM | D4886 | 34E000057 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 17 | 3 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D4886 | BPM | D4886 | 34E000058 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 17 | 3 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D4886 | BPM | D4886 | 34E000068 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 17 | 3 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D4968 | BPM | D4967 | 34E000032 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 17 | 4 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D4968 | BPM | D4967 | 34E000033 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 17 | 4 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D4968 | BPM | D4967 | 34E000034 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 17 | 4 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D4968 | BPM | D4967 | 34E000035 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 17 | 4 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D4968 | BPM | D4967 | 34E000054 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5010 | BPM | D5010 | 34E000047 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 17 | 4 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5010 | BPM | D5010 | 34E000048 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 17 | 4 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5010 | BPM | D5010 | 34E000049 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 17 | 4 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5010 | BPM | D5010 | 34E000050 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 17 | 4 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5010 | BPM | D5010 | 34E000053 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 17 | 4 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5092 | BPM | D5091 | 34E000043 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 17 | 5 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5092 | BPM | D5091 | 34E000044 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 17 | 5 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5092 | BPM | D5091 | 34E000045 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 17 | 5 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5092 | BPM | D5091 | 34E000046 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 17 | 5 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5092 | BPM | D5091 | 34E000052 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5134 | BPM | D5134 | 34E000039 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 17 | 5 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5134 | BPM | D5134 | 34E000040 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 17 | 5 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5134 | BPM | D5134 | 34E000041 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 17 | 5 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5134 | BPM | D5134 | 34E000042 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 17 | 5 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5134 | BPM | D5134 | 34E000051 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 17 | 5 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5216 | BPM | D5215 | 34E000030 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 17 | 6 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5216 | BPM | D5215 | 34E000031 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 17 | 6 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5216 | BPM | D5215 | 34E000037 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 17 | 6 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5216 | BPM | D5215 | 34E000038 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 17 | 6 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5216 | BPM | D5215 | 34E000036 | 2C_1/4in_Hardline_036 | Ref | No terminiation | Ref | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5259 | BPM | D5258 | 34E000009 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 17 | 6 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5259 | BPM | D5258 | 34E000010 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 17 | 6 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5259 | BPM | D5258 | 34E000011 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 17 | 6 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5259 | BPM | D5258 | 34E000012 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 17 | 6 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5259 | BPM | D5258 | 34E000024 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 17 | 6 | 0 | Ref | 0 | 0 | 0 | 0
Ion Chamber | LS3_BTS:IC_D4799 | IC | D4799 | 34E000562 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J1 | 0 | 0 | 0 | 0 | 0 | 0 | J1 | Lemo to Lemo | 17 | 9 | CH0 - fmc2-1 | 0 | 0 | 0 | 0 | 0
Ion Chamber | LS3_BTS:IC_D4895 | IC | D4895 | 34E000561 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J2 | 0 | 0 | 0 | 0 | 0 | 0 | J2 | Lemo to Lemo | 17 | 9 | CH1 - fmc2-2 | 0 | 0 | 0 | 0 | 0
Ion Chamber | LS3_BTS:IC_D5001 | IC | D5001 | 34E000567 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J3 | 0 | 0 | 0 | 0 | 0 | 0 | J3 | Lemo to Lemo | 17 | 9 | CH2 - fmc2-3 | 0 | 0 | 0 | 0 | 0
Ion Chamber | LS3_BTS:IC_D5105 | IC | D5105 | 34E000566 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J4 | 0 | 0 | 0 | 0 | 0 | 0 | J4 | Lemo to Lemo | 17 | 9 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Ion Chamber | LS3_BTS:IC_D4799 | 0 | D4799 | 34H000080 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | SHV to SHV | 0 | 0 | 0 | 0 | 12 | 2 : EHS8020n | Ch0 | 0
Ion Chamber | LS3_BTS:IC_D4895 | 0 | D4895 | 34H000079 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | SHV to SHV | 0 | 0 | 0 | 0 | 12 | 2 : EHS8020n | Ch1 | 0
Ion Chamber | LS3_BTS:IC_D5001 | 0 | D5001 | 34H000085 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | SHV to SHV | 0 | 0 | 0 | 0 | 12 | 2 : EHS8020n | Ch2 | 0
Ion Chamber | LS3_BTS:IC_D5105 | 0 | D5105 | 34H000084 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | SHV to SHV | 0 | 0 | 0 | 0 | 12 | 2 : EHS8020n | Ch3 | 0
PM 8" | LS3_BTS:PM_D4771 | PM | D4769 | 34F000023 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-1 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | Lemo-1A | LEMO650/LEMO650 | 17 | 7 | CH0 - fmc2-1 | 0 | 12 | 1: EBS8005 | Ch0 | 0
PM 8" | LS3_BTS:PM_D4771 | PM | D4769 | 34F000023 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-1 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | Lemo-1B | LEMO650/LEMO650 | 17 | 7 | CH1 - fmc2-2 | 0 | 12 | 1: EBS8005 | Ch0 | 0
PM 8" | LS3_BTS:PM_D4771 | PM | D4769 | 34F000023 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-1 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | Lemo-1C | LEMO650/LEMO650 | 17 | 7 | CH2 - fmc2-3 | 0 | 12 | 1: EBS8005 | Ch0 | 0
PM 8" | LS3_BTS:PM_D4797 | PM | D4794 | 34F000001 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-2 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | Lemo-2A | LEMO650/LEMO650 | 17 | 7 | CH4 - fmc1-1 | 0 | 12 | 1: EBS8005 | Ch1 | 0
PM 8" | LS3_BTS:PM_D4797 | PM | D4794 | 34F000001 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-2 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield-6 | 4 | 6 | Lemo-2B | LEMO650/LEMO650 | 17 | 7 | CH5 - fmc1-2 | 0 | 12 | 1: EBS8005 | Ch1 | 0
PM 8" | LS3_BTS:PM_D4797 | PM | D4794 | 34F000001 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-2 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | Lemo-2C | LEMO650/LEMO650 | 17 | 7 | CH6 - fmc1-3 | 0 | 12 | 1: EBS8005 | Ch1 | 0
PM 8" | LS3_BTS:PM_D4827 | PM | D4824 | 34F000004 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-3 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | Lemo-3A | LEMO650/LEMO650 | 17 | 8 | CH0 - fmc2-1 | 0 | 12 | 1: EBS8005 | Ch2 | 0
PM 8" | LS3_BTS:PM_D4827 | PM | D4824 | 34F000004 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-3 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | Lemo-3B | LEMO650/LEMO650 | 17 | 8 | CH1 - fmc2-2 | 0 | 12 | 1: EBS8005 | Ch2 | 0
PM 8" | LS3_BTS:PM_D4827 | PM | D4824 | 34F000004 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-3 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | Lemo-3C | LEMO650/LEMO650 | 17 | 8 | CH2 - fmc2-3 | 0 | 12 | 1: EBS8005 | Ch2 | 0
PM 8" | LS3_BTS:PM_D4862 | PM | D4862 | 34F000008 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-4 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | Lemo-4A | LEMO650/LEMO650 | 17 | 8 | CH4 - fmc1-1 | 0 | 12 | 1: EBS8005 | Ch3 | 0
PM 8" | LS3_BTS:PM_D4862 | PM | D4862 | 34F000008 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-4 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | Lemo-4B | LEMO650/LEMO650 | 17 | 8 | CH5 - fmc1-2 | 0 | 12 | 1: EBS8005 | Ch3 | 0
PM 8" | LS3_BTS:PM_D4862 | PM | D4862 | 34F000008 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-4 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | Lemo-4C | LEMO650/LEMO650 | 17 | 8 | CH6 - fmc1-3 | 0 | 12 | 1: EBS8005 | Ch3 | 0
Halo Monitor Ring | LS3_WD06:HMR_D4700 | HMR | D4700 | 34H000012 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV | J1 | 0 | 0 | 0 | 0 | 0 | 0 | J1 | SHV to Lemo | 17 | 10 | CH0 - fmc2-1 | 0 | 12 | 1: EBS8003 | Ch4 | 0
0 | LS3_WD06:HMR_D4700 | 0 | D4700 | 34H000097 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring,Spare | LS3_BTS:HMR_D4770 | HMR - Unused | D4770 | 34H000000 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS3_BTS:HMR_D4770 | 0 | D4770 | 34H000093 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring,Spare | LS3_BTS:HMR_D4887 | HMR - Unused | D4887 | 34H000002 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS3_BTS:HMR_D4887 | 0 | D4887 | 34H000091 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring,Spare | LS3_BTS:HMR_D5011 | HMR - Unused | D5011 | 34H000003 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS3_BTS:HMR_D5011 | 0 | D5011 | 34H000090 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring,Spare | LS3_BTS:HMR_D5135 | HMR - Unused | D5135 | 34H000005 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS3_BTS:HMR_D5135 | 0 | D5135 | 34H000087 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring,Spare | LS3_BTS:HMR_D5259 | HMR - Unused | D5259 | 34H000004 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS3_BTS:HMR_D5259 | 0 | D5259 | 34H000088 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 1_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 12 | 2: EBS8003 | Ch0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 1_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 17 | 7 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 2_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 12 | 2: EBS8003 | Ch1 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 2_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 17 | 7 | CH7 - fmc1-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 3_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 12 | 2: EBS8003 | Ch2 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 3_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 17 | 8 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 4_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 12 | 2: EBS8003 | Ch3 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 4_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 17 | 8 | CH7 - fmc1-4 | 0 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5340 | BPM | D5339 | 34E000013 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 18 | 3 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5340 | BPM | D5339 | 34E000014 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 18 | 3 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5340 | BPM | D5339 | 34E000015 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 18 | 3 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5340 | BPM | D5339 | 34E000016 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 18 | 3 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5340 | BPM | D5339 | 34E000023 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5381 | BPM | D5381 | 34E000002 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 18 | 3 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5381 | BPM | D5381 | 34E000003 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 18 | 3 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5381 | BPM | D5381 | 34E000004 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 18 | 3 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5381 | BPM | D5381 | 34E000005 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 18 | 3 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5381 | BPM | D5381 | 34E000022 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 18 | 3 | 0 | Ref | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5430 | BPM | D5430 | 34E000017 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 18 | 4 | 0 | 1 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5430 | BPM | D5430 | 34E000018 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 18 | 4 | 0 | 2 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5430 | BPM | D5430 | 34E000019 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 18 | 4 | 0 | 3 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5430 | BPM | D5430 | 34E000020 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 18 | 4 | 0 | 4 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5430 | BPM | D5430 | 34E000021 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5445 | BPM | D5444 | 34E000000 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 18 | 4 | 0 | 5 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5445 | BPM | D5444 | 34E000006 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 18 | 4 | 0 | 6 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5445 | BPM | D5444 | 34E000007 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 18 | 4 | 0 | 7 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5445 | BPM | D5444 | 34E000008 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 18 | 4 | 0 | 8 | 0 | 0 | 0 | 0
BPM | LS3_BTS:BPM_D5445 | BPM | D5444 | 34E000001 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 18 | 4 | 0 | Ref | 0 | 0 | 0 | 0
BPM | BDS_BTS:BPM_D5499 | BPM | D5499 | 44E000025 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 18 | 5 | 0 | 1 | 0 | 0 | 0 | 0
BPM | BDS_BTS:BPM_D5499 | BPM | D5499 | 44E000026 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 18 | 5 | 0 | 2 | 0 | 0 | 0 | 0
BPM | BDS_BTS:BPM_D5499 | BPM | D5499 | 44E000027 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 18 | 5 | 0 | 3 | 0 | 0 | 0 | 0
BPM | BDS_BTS:BPM_D5499 | BPM | D5499 | 44E000028 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 18 | 5 | 0 | 4 | 0 | 0 | 0 | 0
BPM | BDS_BTS:BPM_D5499 | BPM | D5499 | 44E000048 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | BDS_BTS:BPM_D5513 | BPM | D5512 | 44E000033 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 18 | 5 | 0 | 5 | 0 | 0 | 0 | 0
BPM | BDS_BTS:BPM_D5513 | BPM | D5512 | 44E000034 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 18 | 5 | 0 | 6 | 0 | 0 | 0 | 0
BPM | BDS_BTS:BPM_D5513 | BPM | D5512 | 44E000035 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 18 | 5 | 0 | 7 | 0 | 0 | 0 | 0
BPM | BDS_BTS:BPM_D5513 | BPM | D5512 | 44E000036 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 18 | 5 | 0 | 8 | 0 | 0 | 0 | 0
BPM | BDS_BTS:BPM_D5513 | BPM | D5512 | 44E000047 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 18 | 5 | 0 | Ref | 0 | 0 | 0 | 0
BPM | BDS_BTS:BPM_D5565 | BPM | D5566 | 44E000041 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 18 | 6 | 0 | 1 | 0 | 0 | 0 | 0
BPM | BDS_BTS:BPM_D5565 | BPM | D5566 | 44E000042 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 18 | 6 | 0 | 2 | 0 | 0 | 0 | 0
BPM | BDS_BTS:BPM_D5565 | BPM | D5566 | 44E000043 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 18 | 6 | 0 | 3 | 0 | 0 | 0 | 0
BPM | BDS_BTS:BPM_D5565 | BPM | D5566 | 44E000044 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 18 | 6 | 0 | 4 | 0 | 0 | 0 | 0
BPM | BDS_BTS:BPM_D5565 | BPM | D5566 | 44E000046 | 2C_1/4in_Hardline_036 | Ref | No terminiation | Ref | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | BDS_BBS:BPM_D5625 | BPM | D5626 | 44E000053 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 18 | 6 | 0 | 5 | 0 | 0 | 0 | 0
BPM | BDS_BBS:BPM_D5625 | BPM | D5626 | 44E000054 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 18 | 6 | 0 | 6 | 0 | 0 | 0 | 0
BPM | BDS_BBS:BPM_D5625 | BPM | D5626 | 44E000055 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 18 | 6 | 0 | 7 | 0 | 0 | 0 | 0
BPM | BDS_BBS:BPM_D5625 | BPM | D5626 | 44E000056 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 18 | 6 | 0 | 8 | 0 | 0 | 0 | 0
BPM | BDS_BBS:BPM_D5625 | BPM | D5626 | 44E000057 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 18 | 6 | 0 | Ref | 0 | 0 | 0 | 0
BPM | BDS_BBS:BPM_D5653 | BPM | D5653 | 44E000049 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 18 | 7 | 0 | 1 | 0 | 0 | 0 | 0
BPM | BDS_BBS:BPM_D5653 | BPM | D5653 | 44E000050 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 18 | 7 | 0 | 2 | 0 | 0 | 0 | 0
BPM | BDS_BBS:BPM_D5653 | BPM | D5653 | 44E000051 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 18 | 7 | 0 | 3 | 0 | 0 | 0 | 0
BPM | BDS_BBS:BPM_D5653 | BPM | D5653 | 44E000052 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 18 | 7 | 0 | 4 | 0 | 0 | 0 | 0
BPM | BDS_BBS:BPM_D5653 | BPM | D5653 | 44E000058 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | BDS_BBS:BPM_D5653 | BPM | D5655 | 44E000037 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 18 | 7 | 0 | 5 | 0 | 0 | 0 | 0
BPM | BDS_BBS:BPM_D5653 | BPM | D5655 | 44E000038 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 18 | 7 | 0 | 6 | 0 | 0 | 0 | 0
BPM | BDS_BBS:BPM_D5653 | BPM | D5655 | 44E000039 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 18 | 7 | 0 | 7 | 0 | 0 | 0 | 0
BPM | BDS_BBS:BPM_D5653 | BPM | D5655 | 44E000040 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 18 | 7 | 0 | 8 | 0 | 0 | 0 | 0
BPM | BDS_BBS:BPM_D5653 | BPM | D5655 | 44E000045 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 18 | 7 | 0 | Ref | 0 | 0 | 0 | 0
BPM | BDS_BBS:BPM_D5680 | BPM | D5681 | 44E000024 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 18 | 8 | 0 | 1 | 0 | 0 | 0 | 0
BPM | BDS_BBS:BPM_D5680 | BPM | D5681 | 44E000030 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 18 | 8 | 0 | 2 | 0 | 0 | 0 | 0
BPM | BDS_BBS:BPM_D5680 | BPM | D5681 | 44E000031 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 18 | 8 | 0 | 3 | 0 | 0 | 0 | 0
BPM | BDS_BBS:BPM_D5680 | BPM | D5681 | 44E000032 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 18 | 8 | 0 | 4 | 0 | 0 | 0 | 0
BPM | BDS_BBS:BPM_D5680 | BPM | D5681 | 44E000029 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | BDS_FFS:BPM_D5742 | BPM | D5742 | 44E000009 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 18 | 8 | 0 | 5 | 0 | 0 | 0 | 0
BPM | BDS_FFS:BPM_D5742 | BPM | D5742 | 44E000010 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 18 | 8 | 0 | 6 | 0 | 0 | 0 | 0
BPM | BDS_FFS:BPM_D5742 | BPM | D5742 | 44E000011 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 18 | 8 | 0 | 7 | 0 | 0 | 0 | 0
BPM | BDS_FFS:BPM_D5742 | BPM | D5742 | 44E000012 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 18 | 8 | 0 | 8 | 0 | 0 | 0 | 0
BPM | BDS_FFS:BPM_D5742 | BPM | D5742 | 44E000013 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 18 | 8 | 0 | Ref | 0 | 0 | 0 | 0
BPM | BDS_FFS:BPM_D5772 | BPM | D5772 | 44E000020 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 18 | 9 | 0 | 1 | 0 | 0 | 0 | 0
BPM | BDS_FFS:BPM_D5772 | BPM | D5772 | 44E000021 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 18 | 9 | 0 | 2 | 0 | 0 | 0 | 0
BPM | BDS_FFS:BPM_D5772 | BPM | D5772 | 44E000022 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 18 | 9 | 0 | 3 | 0 | 0 | 0 | 0
BPM | BDS_FFS:BPM_D5772 | BPM | D5772 | 44E000023 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 18 | 9 | 0 | 4 | 0 | 0 | 0 | 0
BPM | BDS_FFS:BPM_D5772 | BPM | D5772 | 44E000015 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | BDS_FFS:BPM_D5790 | BPM | D5799 | 44E000016 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 18 | 9 | 0 | 5 | 0 | 0 | 0 | 0
BPM | BDS_FFS:BPM_D5790 | BPM | D5799 | 44E000017 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 18 | 9 | 0 | 6 | 0 | 0 | 0 | 0
BPM | BDS_FFS:BPM_D5790 | BPM | D5799 | 44E000018 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 18 | 9 | 0 | 7 | 0 | 0 | 0 | 0
BPM | BDS_FFS:BPM_D5790 | BPM | D5799 | 44E000019 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 18 | 9 | 0 | 8 | 0 | 0 | 0 | 0
BPM | BDS_FFS:BPM_D5790 | BPM | D5799 | 44E000014 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 18 | 9 | 0 | Ref | 0 | 0 | 0 | 0
BPM | BDS_FFS:BPM_D5803 | BPM | D5803 | 44E000271 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 18 | 10 | 0 | 1 | 0 | 0 | 0 | 0
BPM | BDS_FFS:BPM_D5803 | BPM | D5803 | 44E000270 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 18 | 10 | 0 | 2 | 0 | 0 | 0 | 0
BPM | BDS_FFS:BPM_D5803 | BPM | D5803 | 44E000269 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 18 | 10 | 0 | 3 | 0 | 0 | 0 | 0
BPM | BDS_FFS:BPM_D5803 | BPM | D5803 | 44E000268 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 18 | 10 | 0 | 4 | 0 | 0 | 0 | 0
BPM | BDS_FFS:BPM_D5803 | BPM | D5803 | 44E000266 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
BPM | BDS_FFS:BPM_D5818 | BPM | D5818 | 44E000265 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 18 | 10 | 0 | 5 | 0 | 0 | 0 | 0
BPM | BDS_FFS:BPM_D5818 | BPM | D5818 | 44E000264 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 18 | 10 | 0 | 6 | 0 | 0 | 0 | 0
BPM | BDS_FFS:BPM_D5818 | BPM | D5818 | 44E000263 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 18 | 10 | 0 | 7 | 0 | 0 | 0 | 0
BPM | BDS_FFS:BPM_D5818 | BPM | D5818 | 44E000262 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 18 | 10 | 0 | 8 | 0 | 0 | 0 | 0
BPM | BDS_FFS:BPM_D5818 | BPM | D5818 | 44E000267 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 18 | 10 | 0 | Ref | 0 | 0 | 0 | 0
PM 8" | BDS_BTS:PM_D5514 | PM | D5513 | 44F000020 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-1 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | Lemo-1A | LEMO650/LEMO650 | 19 | 3 | CH0 - fmc2-1 | 0 | 14 | 1: EBS8005 | Ch2 | 0
PM 8" | BDS_BTS:PM_D5514 | PM | D5513 | 44F000020 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-1 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | Lemo-1B | LEMO650/LEMO650 | 19 | 3 | CH1 - fmc2-2 | 0 | 14 | 1: EBS8005 | Ch2 | 0
PM 8" | BDS_BTS:PM_D5514 | PM | D5513 | 44F000020 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-1 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | Lemo-1C | LEMO650/LEMO650 | 19 | 3 | CH2 - fmc2-3 | 0 | 14 | 1: EBS8005 | Ch2 | 0
PM 8" | BDS_BTS:PM_D5567 | PM | D5567 | 44F000024 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-2 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | Lemo-2A | LEMO650/LEMO650 | 19 | 3 | CH4 - fmc1-1 | 0 | 14 | 1: EBS8005 | Ch3 | 0
PM 8" | BDS_BTS:PM_D5567 | PM | D5567 | 44F000024 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-2 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield-6 | 4 | 6 | Lemo-2B | LEMO650/LEMO650 | 19 | 3 | CH5 - fmc1-2 | 0 | 14 | 1: EBS8005 | Ch3 | 0
PM 8" | BDS_BTS:PM_D5567 | PM | D5567 | 44F000024 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-2 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | Lemo-2C | LEMO650/LEMO650 | 19 | 3 | CH6 - fmc1-3 | 0 | 14 | 1: EBS8005 | Ch3 | 0
PM 12" | BDS_BBS:PM_D5628 | PM | D5628 | 44F000108 | 6C_22AWG_Multi_010 | A Signal | CPC to LEMO650 Box | CPC-3 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo-3A | LEMO650/LEMO650 | 19 | 4 | CH0 - fmc2-1 | 0 | 14 | 1: EBS8005 | Ch0 | 0
PM 12" | BDS_BBS:PM_D5628 | PM | D5628 | 44F000108 | 6C_22AWG_Multi_010 | B Signal | CPC to LEMO650 Box | CPC-3 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo-3B | LEMO650/LEMO650 | 19 | 4 | CH1 - fmc2-2 | 0 | 14 | 1: EBS8005 | Ch0 | 0
PM 6" | BDS_BBS:PM_D5628 | PM | D5628 | 44F000112 | 4C_22AWG_Multi_006 | C Signal | CPC to LEMO650 Box | CPC-3 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo-3C | LEMO650/LEMO650 | 19 | 4 | CH2 - fmc2-3 | 0 | 14 | 1: EBS8005 | Ch0 | 0
PM 12" | BDS_BBS:PM_D5653 | PM | D5655 | 44F000101 | 6C_22AWG_Multi_010 | A Signal | CPC to LEMO650 Box | CPC-4 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo-4A | LEMO650/LEMO650 | 19 | 4 | CH4 - fmc1-1 | 0 | 14 | 1: EBS8005 | Ch1 | 0
PM 12" | BDS_BBS:PM_D5653 | PM | D5655 | 44F000101 | 6C_22AWG_Multi_010 | B Signal | CPC to LEMO650 Box | CPC-4 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo-4B | LEMO650/LEMO650 | 19 | 4 | CH5 - fmc1-2 | 0 | 14 | 1: EBS8005 | Ch1 | 0
PM 6" | BDS_BBS:PM_D5653 | PM | D5655 | 44F000100 | 4C_22AWG_Multi_006 | C Signal | CPC to LEMO650 Box | CPC-4 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo-4C | LEMO650/LEMO650 | 19 | 4 | CH6 - fmc1-3 | 0 | 14 | 1: EBS8005 | Ch1 | 0
PM 8" | BDS_FFS:PM_D5743 | PM | D5742 | 44F000028 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-5 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | Lemo-5A | LEMO650/LEMO650 | 19 | 5 | CH0 - fmc2-1 | 0 | 14 | 1: EBS8005 | Ch4 | 0
PM 8" | BDS_FFS:PM_D5743 | PM | D5742 | 44F000028 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-5 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | Lemo-5B | LEMO650/LEMO650 | 19 | 5 | CH1 - fmc2-2 | 0 | 14 | 1: EBS8005 | Ch4 | 0
PM 8" | BDS_FFS:PM_D5743 | PM | D5742 | 44F000028 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-5 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | Lemo-5C | LEMO650/LEMO650 | 19 | 5 | CH2 - fmc2-3 | 0 | 14 | 1: EBS8005 | Ch4 | 0
PM 8" | BDS_FFS:PM_D5774 | PM | D5773 | 44F000034 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-6 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | Lemo-6A | LEMO650/LEMO650 | 19 | 5 | CH4 - fmc1-1 | 0 | 14 | 1: EBS8005 | Ch5 | 0
PM 8" | BDS_FFS:PM_D5774 | PM | D5773 | 44F000034 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-6 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | Lemo-6B | LEMO650/LEMO650 | 19 | 5 | CH5 - fmc1-2 | 0 | 14 | 1: EBS8005 | Ch5 | 0
PM 8" | BDS_FFS:PM_D5774 | PM | D5773 | 44F000034 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-6 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | Lemo-6C | LEMO650/LEMO650 | 19 | 5 | CH6 - fmc1-3 | 0 | 14 | 1: EBS8005 | Ch5 | 0
PM 8" | BDS_FFS:PM_D5792 | PM | D5800 | 44F000116 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-7 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | Lemo-7A | LEMO650/LEMO650 | 19 | 6 | CH0 - fmc2-1 | 0 | 14 | 1: EBS8005 | Ch6 | 0
PM 8" | BDS_FFS:PM_D5792 | PM | D5800 | 44F000116 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-7 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | Lemo-7B | LEMO650/LEMO650 | 19 | 6 | CH1 - fmc2-2 | 0 | 14 | 1: EBS8005 | Ch6 | 0
PM 8" | BDS_FFS:PM_D5792 | PM | D5800 | 44F000116 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-7 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | Lemo-7C | LEMO650/LEMO650 | 19 | 6 | CH2 - fmc2-3 | 0 | 14 | 1: EBS8005 | Ch6 | 0
Ion Chamber | LS3_BTS:IC_D5205 | IC | D5205 | 34E000565 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J1 | 0 | 0 | 0 | 0 | 0 | 0 | J1 | Lemo to Lemo | 19 | 7 | CH0 - fmc2-1 | 0 | 0 | 0 | 0 | 0
Ion Chamber | LS3_BTS:IC_D5291 | IC | D5291 | 34E000564 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J2 | 0 | 0 | 0 | 0 | 0 | 0 | J2 | Lemo to Lemo | 19 | 7 | CH1 - fmc2-2 | 0 | 0 | 0 | 0 | 0
Ion Chamber | BDS_BTS:IC_D5496 | IC | D5496 | 44E000253 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J3 | 0 | 0 | 0 | 0 | 0 | 0 | J3 | Lemo to Lemo | 19 | 7 | CH2 - fmc2-3 | 0 | 0 | 0 | 0 | 0
Ion Chamber | BDS_BTS:IC_D5515 | IC | D5515 | 44E000250 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J4 | 0 | 0 | 0 | 0 | 0 | 0 | J4 | Lemo to Lemo | 19 | 7 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Ion Chamber | BDS_BTS:IC_D5550 | IC | D5550 | 44E000249 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J5 | 0 | 0 | 0 | 0 | 0 | 0 | J5 | Lemo to Lemo | 19 | 7 | CH4 - fmc1-1 | 0 | 0 | 0 | 0 | 0
Ion Chamber | BDS_BBS:IC_D5578 | IC | D5578 | 44E000248 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J6 | 0 | 0 | 0 | 0 | 0 | 0 | J6 | Lemo to Lemo | 19 | 7 | CH5 - fmc1-2 | 0 | 0 | 0 | 0 | 0
Ion Chamber | BDS_BBS:IC_D5611 | IC | D5611 | 44E000252 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J7 | 0 | 0 | 0 | 0 | 0 | 0 | J7 | Lemo to Lemo | 19 | 7 | CH6 - fmc1-3 | 0 | 0 | 0 | 0 | 0
Ion Chamber | BDS_BBS:IC_D5641 | IC | D5641 | 44E000251 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J8 | 0 | 0 | 0 | 0 | 0 | 0 | J8 | Lemo to Lemo | 19 | 7 | CH7 - fmc1-4 | 0 | 0 | 0 | 0 | 0
Ion Chamber | BDS_BBS:IC_D5679 | IC | D5679 | 44E000247 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J1 | 0 | 0 | 0 | 0 | 0 | 0 | J1 | Lemo to Lemo | 19 | 8 | CH0 - fmc2-1 | 0 | 0 | 0 | 0 | 0
Ion Chamber | BDS_BBS:IC_D5711 | IC | D5711 | 44E000246 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J2 | 0 | 0 | 0 | 0 | 0 | 0 | J2 | Lemo to Lemo | 19 | 8 | CH1 - fmc2-2 | 0 | 0 | 0 | 0 | 0
Ion Chamber | BDS_FFS:IC_D5740 | IC | D5740 | 44E000245 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J3 | 0 | 0 | 0 | 0 | 0 | 0 | J3 | Lemo to Lemo | 19 | 8 | CH2 - fmc2-3 | 0 | 0 | 0 | 0 | 0
Ion Chamber | BDS_FFS:IC_D5765 | IC | D5765 | 44E000244 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J4 | 0 | 0 | 0 | 0 | 0 | 0 | J4 | Lemo to Lemo | 19 | 8 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Ion Chamber | BDS_FFS:IC_D5800 | IC | D5800 | 44E000243 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J5 | 0 | 0 | 0 | 0 | 0 | 0 | J5 | Lemo to Lemo | 19 | 8 | CH4 - fmc1-1 | 0 | 0 | 0 | 0 | 0
Ion Chamber | BDS_FFS:IC_D5822 | IC | D5822 | 44E000242 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J6 | 0 | 0 | 0 | 0 | 0 | 0 | J6 | Lemo to Lemo | 19 | 8 | CH5 - fmc1-2 | 0 | 0 | 0 | 0 | 0
Ion Chamber | BDS_BTS:IC_D5661 | IC | D5661 | 44E000277 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J7 | 0 | 0 | 0 | 0 | 0 | 0 | J7 | Lemo to Lemo | 19 | 8 | CH6 - fmc1-3 | 0 | 0 | 0 | 0 | 0
Ion Chamber | BDS_BTS:IC_D5662 | IC | D5662 | 44E000278 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J8 | 0 | 0 | 0 | 0 | 0 | 0 | J8 | Lemo to Lemo | 19 | 8 | CH7 - fmc1-4 | 0 | 0 | 0 | 0 | 0
Ion Chamber | LS3_BTS:IC_D5205 | 0 | D5205 | 34H000083 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | SHV to SHV | 0 | 0 | 0 | 0 | 14 | 2 : EHS8020n | Ch0 | 0
Ion Chamber | LS3_BTS:IC_D5291 | 0 | D5291 | 34H000082 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | SHV to SHV | 0 | 0 | 0 | 0 | 14 | 2 : EHS8020n | Ch0 | 0
Ion Chamber | BDS_BTS:IC_D5496 | 0 | D5496 | 44H000037 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | SHV to SHV | 0 | 0 | 0 | 0 | 14 | 2 : EHS8020n | Ch1 | 0
Ion Chamber | BDS_BTS:IC_D5515 | 0 | D5515 | 44H000034 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | SHV to SHV | 0 | 0 | 0 | 0 | 14 | 2 : EHS8020n | Ch1 | 0
Ion Chamber | BDS_BTS:IC_D5550 | 0 | D5550 | 44H000033 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 5 | SHV to SHV | 0 | 0 | 0 | 0 | 14 | 2 : EHS8020n | Ch2 | 0
Ion Chamber | BDS_BBS:IC_D5578 | 0 | D5578 | 44H000032 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 6 | 0 | 0 | 0 | 0 | 0 | 0 | 6 | SHV to SHV | 0 | 0 | 0 | 0 | 14 | 2 : EHS8020n | Ch2 | 0
Ion Chamber | BDS_BBS:IC_D5611 | 0 | D5611 | 44H000035 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 7 | 0 | 0 | 0 | 0 | 0 | 0 | 7 | SHV to SHV | 0 | 0 | 0 | 0 | 14 | 2 : EHS8020n | Ch3 | 0
Ion Chamber | BDS_BBS:IC_D5641 | 0 | D5641 | 44H000036 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 8 | 0 | 0 | 0 | 0 | 0 | 0 | 8 | SHV to SHV | 0 | 0 | 0 | 0 | 14 | 2 : EHS8020n | Ch3 | 0
Ion Chamber | BDS_BBS:IC_D5679 | 0 | D5679 | 44H000031 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | SHV to SHV | 0 | 0 | 0 | 0 | 14 | 2 : EHS8020n | Ch4 | 0
Ion Chamber | BDS_BBS:IC_D5711 | 0 | D5711 | 44H000030 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | SHV to SHV | 0 | 0 | 0 | 0 | 14 | 2 : EHS8020n | Ch4 | 0
Ion Chamber | BDS_FFS:IC_D5740 | 0 | D5740 | 44H000029 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | SHV to SHV | 0 | 0 | 0 | 0 | 14 | 2 : EHS8020n | Ch5 | 0
Ion Chamber | BDS_FFS:IC_D5765 | 0 | D5765 | 44H000028 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | SHV to SHV | 0 | 0 | 0 | 0 | 14 | 2 : EHS8020n | Ch5 | 0
Ion Chamber | BDS_FFS:IC_D5800 | 0 | D5800 | 44H000027 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 5 | SHV to SHV | 0 | 0 | 0 | 0 | 14 | 2 : EHS8020n | Ch6 | 0
Ion Chamber | BDS_FFS:IC_D5822 | 0 | D5822 | 44H000026 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 6 | 0 | 0 | 0 | 0 | 0 | 0 | 6 | SHV to SHV | 0 | 0 | 0 | 0 | 14 | 2 : EHS8020n | Ch6 | 0
Ion Chamber | BDS_BTS:IC_D5661 | 0 | D5661 | 44H000068 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 7 | 0 | 0 | 0 | 0 | 0 | 0 | 7 | SHV to SHV | 0 | 0 | 0 | 0 | 14 | 2 : EHS8020n | Ch7 | 0
Ion Chamber | BDS_BTS:IC_D5662 | 0 | D5662 | 44H000069 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 8 | 0 | 0 | 0 | 0 | 0 | 0 | 8 | SHV to SHV | 0 | 0 | 0 | 0 | 14 | 2 : EHS8020n | Ch7 | 0
Halo Monitor Ring,Spare | LS3_BTS:HMR_D5382 | HMR - Unused | D5382 | 34H000001 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS3_BTS:HMR_D5382 | 0 | D5382 | 34H000092 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Halo Monitor Ring,Spare | LS3_BTS:HMR_D5445 | HMR - Unused | D5445 | 34H000006 | 2C_1kV75LC_Coax_090 | Signal | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | LS3_BTS:HMR_D5445 | 0 | D5445 | 34H000089 | 2C_5kVDC_Coax_035 | HV Bias | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 1_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 14 | 2: EBS8003 | Ch0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 1_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 19 | 4 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 2_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 14 | 2: EBS8003 | Ch1 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 2_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 19 | 4 | CH7 - fmc1-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 3_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 14 | 2: EBS8003 | Ch2 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 3_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 19 | 3 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 4_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 14 | 2: EBS8003 | Ch3 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 4_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 19 | 3 | CH7 - fmc1-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 5_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 14 | 2: EBS8003 | Ch4 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 5_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 19 | 5 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 6_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 14 | 2: EBS8003 | Ch5 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 6_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 19 | 5 | CH7 - fmc1-4 | 0 | 0 | 0 | 0 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 7_IN Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo250/Lemo250 | 0 | 0 | 0 | 0 | 14 | 2: EBS8003 | Ch6 | 0
Interconnect | PM bias for PICO | bias | 0 | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 7_OUT Port | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Lemo650/Lemo650 | 19 | 6 | CH3 - fmc2-4 | 0 | 0 | 0 | 0 | 0
BPM | FS1_CSS:BPM_D2212 | BPM | D2212 | 44E000102 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 20 | 3 | 0 | 1 | 0 | 0 | 0 | 0
BPM | FS1_CSS:BPM_D2212 | BPM | D2212 | 44E000103 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 20 | 3 | 0 | 2 | 0 | 0 | 0 | 0
BPM | FS1_CSS:BPM_D2212 | BPM | D2212 | 44E000104 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 20 | 3 | 0 | 3 | 0 | 0 | 0 | 0
BPM | FS1_CSS:BPM_D2212 | BPM | D2212 | 44E000105 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 20 | 3 | 0 | 4 | 0 | 0 | 0 | 0
BPM | FS1_CSS:BPM_D2212 | BPM | D2212 | 44E000090 | 2C_1/4in_Hardline_036 | Ref | No terminiation | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0 | 0 | 0 | 0
BPM | FS1_CSS:BPM_D2223 | BPM | D2223 | 44E000106 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 20 | 3 | 0 | 5 | 0 | 0 | 0 | 0
BPM | FS1_CSS:BPM_D2223 | BPM | D2223 | 44E000107 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 20 | 3 | 0 | 6 | 0 | 0 | 0 | 0
BPM | FS1_CSS:BPM_D2223 | BPM | D2223 | 44E000108 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 20 | 3 | 0 | 7 | 0 | 0 | 0 | 0
BPM | FS1_CSS:BPM_D2223 | BPM | D2223 | 44E000109 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 20 | 3 | 0 | 8 | 0 | 0 | 0 | 0
BPM | FS1_CSS:BPM_D2223 | BPM | D2223 | 44E000091 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 20 | 3 | 0 | Ref | 0 | 0 | 0 | 0
BPM | FS1_CSS:BPM_D2248 | BPM | D2248 | 44E000098 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 20 | 4 | 0 | 1 | 0 | 0 | 0 | 0
BPM | FS1_CSS:BPM_D2248 | BPM | D2248 | 44E000099 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 20 | 4 | 0 | 2 | 0 | 0 | 0 | 0
BPM | FS1_CSS:BPM_D2248 | BPM | D2248 | 44E000100 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 20 | 4 | 0 | 3 | 0 | 0 | 0 | 0
BPM | FS1_CSS:BPM_D2248 | BPM | D2248 | 44E000101 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 20 | 4 | 0 | 4 | 0 | 0 | 0 | 0
BPM | FS1_CSS:BPM_D2248 | BPM | D2248 | 44E000089 | 2C_1/4in_Hardline_036 | Ref | No terminiation | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0 | 0 | 0 | 0
BPM | FS1_CSS:BPM_D2278 | BPM | D2278 | 44E000110 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 20 | 4 | 0 | 5 | 0 | 0 | 0 | 0
BPM | FS1_CSS:BPM_D2278 | BPM | D2278 | 44E000111 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 20 | 4 | 0 | 6 | 0 | 0 | 0 | 0
BPM | FS1_CSS:BPM_D2278 | BPM | D2278 | 44E000112 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 20 | 4 | 0 | 7 | 0 | 0 | 0 | 0
BPM | FS1_CSS:BPM_D2278 | BPM | D2278 | 44E000117 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 20 | 4 | 0 | 8 | 0 | 0 | 0 | 0
BPM | FS1_CSS:BPM_D2278 | BPM | D2278 | 44E000092 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 20 | 4 | 0 | Ref | 0 | 0 | 0 | 0
BPM | FS1_CSS:BPM_D2313 | BPM | D2313 | 44E000113 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 20 | 5 | 0 | 1 | 0 | 0 | 0 | 0
BPM | FS1_CSS:BPM_D2313 | BPM | D2313 | 44E000114 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 20 | 5 | 0 | 2 | 0 | 0 | 0 | 0
BPM | FS1_CSS:BPM_D2313 | BPM | D2313 | 44E000115 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 20 | 5 | 0 | 3 | 0 | 0 | 0 | 0
BPM | FS1_CSS:BPM_D2313 | BPM | D2313 | 44E000116 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 20 | 5 | 0 | 4 | 0 | 0 | 0 | 0
BPM | FS1_CSS:BPM_D2313 | BPM | D2313 | 44E000093 | 2C_1/4in_Hardline_036 | Ref | No terminiation | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0 | 0 | 0 | 0
BPM | FS1_CSS:BPM_D2369 | BPM | D2369 | 44E000118 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 20 | 5 | 0 | 5 | 0 | 0 | 0 | 0
BPM | FS1_CSS:BPM_D2369 | BPM | D2369 | 44E000119 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 20 | 5 | 0 | 6 | 0 | 0 | 0 | 0
BPM | FS1_CSS:BPM_D2369 | BPM | D2369 | 44E000120 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 20 | 5 | 0 | 7 | 0 | 0 | 0 | 0
BPM | FS1_CSS:BPM_D2369 | BPM | D2369 | 44E000121 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 20 | 5 | 0 | 8 | 0 | 0 | 0 | 0
BPM | FS1_CSS:BPM_D2369 | BPM | D2369 | 44E000095 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 20 | 5 | 0 | Ref | 0 | 0 | 0 | 0
BPM | FS1_CSS:BPM_D2383 | BPM | D2383 | 44E000122 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 20 | 6 | 0 | 1 | 0 | 0 | 0 | 0
BPM | FS1_CSS:BPM_D2383 | BPM | D2383 | 44E000123 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 20 | 6 | 0 | 2 | 0 | 0 | 0 | 0
BPM | FS1_CSS:BPM_D2383 | BPM | D2383 | 44E000124 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 20 | 6 | 0 | 3 | 0 | 0 | 0 | 0
BPM | FS1_CSS:BPM_D2383 | BPM | D2383 | 44E000125 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 20 | 6 | 0 | 4 | 0 | 0 | 0 | 0
BPM | FS1_CSS:BPM_D2383 | BPM | D2383 | 44E000094 | 2C_1/4in_Hardline_036 | Ref | No terminiation | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0 | 0 | 0 | 0
BPM | FS1_BBS:BPM_D2421 | BPM | D2421 | 44E000130 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 20 | 6 | 0 | 5 | 0 | 0 | 0 | 0
BPM | FS1_BBS:BPM_D2421 | BPM | D2421 | 44E000131 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 20 | 6 | 0 | 6 | 0 | 0 | 0 | 0
BPM | FS1_BBS:BPM_D2421 | BPM | D2421 | 44E000132 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 20 | 6 | 0 | 7 | 0 | 0 | 0 | 0
BPM | FS1_BBS:BPM_D2421 | BPM | D2421 | 44E000133 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 20 | 6 | 0 | 8 | 0 | 0 | 0 | 0
BPM | FS1_BBS:BPM_D2421 | BPM | D2421 | 44E000097 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 20 | 6 | 0 | Ref | 0 | 0 | 0 | 0
BPM | FS1_BTS:BPM_D2436 | BPM | D2436 | 44E000126 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | 0 | 0 | 0 | 0 | 0 | 0 | 1T | SMA to SMA | 20 | 7 | 0 | 1 | 0 | 0 | 0 | 0
BPM | FS1_BTS:BPM_D2436 | BPM | D2436 | 44E000127 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | 0 | 0 | 0 | 0 | 0 | 0 | 2L | SMA to SMA | 20 | 7 | 0 | 2 | 0 | 0 | 0 | 0
BPM | FS1_BTS:BPM_D2436 | BPM | D2436 | 44E000128 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | 0 | 0 | 0 | 0 | 0 | 0 | 3R | SMA to SMA | 20 | 7 | 0 | 3 | 0 | 0 | 0 | 0
BPM | FS1_BTS:BPM_D2436 | BPM | D2436 | 44E000129 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | 0 | 0 | 0 | 0 | 0 | 0 | 4B | SMA to SMA | 20 | 7 | 0 | 4 | 0 | 0 | 0 | 0
BPM | FS1_BTS:BPM_D2436 | BPM | D2436 | 44E000096 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | 0 | 0 | 0 | 0 | 0 | 0 | Ref | SMA to SMA | 20 | 7 | 0 | Ref | 0 | 0 | 0 | 0

## Diag Lattice
BRANCH 1 Artemis_B Source Line
BEAM Branch | System | Sub- system | Device | Beam line position (dm) | Name | Device Type | Element Name | Minimum Beam Pipe Inner Diameter (mm) | Element Flange to Flange Length (m) | PLACE HOLDER | Element Effective Length (m) | Global Coordinate element center Xc (m) | Global Coordinate element center Yc (m) | Global Coordinate element center Zc (m) | Global Coordinate Z (m) | Global Coordinate Y (m) | Global Coordinate X (m) | Accumulated center-to-center Length (m) | Accumulated end-to-end Length (m) | Comment | Distance from Artemis source (m)
68.4786
0 | FE | ISRC1 | ELC1 | 685 | FE_ISRC1:ELC1_D0685 |  | extraction electrode inner wall |  |  |  | 0 | 370.7066 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 370.7066 | 68.4786 | 68.4786 |  | 0
0 | [6 empty] | 8 mm extraction hole |  |  |  | 0.0023657 | 370.7054 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 370.7042 | 68.4798 | 68.481 |  | 0.0011829
0 | [6 empty] | artemis_b extraction conical wall |  |  |  | 0.011632 | 370.6984 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 370.6926 | 68.4868 | 68.4926 |  | 0.0081818
0 | [6 empty] | extraction mounting plate |  |  |  | 0 | 370.6926 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 370.6926 | 68.4926 | 68.4926 |  | 0.013998
0 | [6 empty] | gap (puller & extraction hole) |  |  |  | 0.038348 | 370.6734 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 370.6543 | 68.5118 | 68.531 |  | 0.033172
0 | FE | ISRC1 | ELC2 | 685 | FE_ISRC1:ELC2_D0685 | PULL_B | puller electrode (bias) |  |  |  | 0.0044341 | 370.652 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 370.6498 | 68.5332 | 68.5354 |  | 0.054563
0 | [6 empty] | gap (puller main & bias) |  |  |  | 0.0047569 | 370.6475 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 370.6451 | 68.5378 | 68.5402 |  | 0.059159
0 | FE | ISRC1 | ELC3 | 685 | FE_ISRC1:ELC3_D0685 | PULL_M | puller electrode (main) |  |  |  | 0.0044341 | 370.6429 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 370.6406 | 68.5424 | 68.5446 |  | 0.063754
0 | [6 empty] | puller tube |  |  |  | 0.13195 | 370.5747 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 370.5087 | 68.6106 | 68.6766 |  | 0.13195
0 | [6 empty] | extraction box |  |  |  | 0.036269 | 370.4905 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 370.4724 | 68.6947 | 68.7128 |  | 0.21606
0 | FE | ISRC1 | TMP | 687 | FE_ISRC1:TMP_D0687 | TMP_M | turbo pump, 700 l/s |  |  |  | 0 | 370.4724 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 370.4724 | 68.7128 | 68.7128 |  | 0.23419
0 | [6 empty] | extraction box |  |  |  | 0.1768 | 370.384 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 370.2956 | 68.8012 | 68.8896 |  | 0.32259
0 | FE | ISRC1 | SOLR | 690 | FE_ISRC1:SOLR_D0690 | S0 | solenoid | 150 |  |  | 0.26032 | 370.1655 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 370.0353 | 69.0198 | 69.1499 | 40cm effective length | 0.54116
0 | [6 empty] | bellow | 150 |  |  | 0.23421 | 369.9182 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 369.8011 | 69.2671 | 69.3842 |  | 0.78843
0 | FE | ISRC1 | DCHV | 695 | FE_ISRC1:DCHV_D0695 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 369.7058 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 369.6106 | 69.4794 | 69.5747 |  | 1.0008
0 | [6 empty] | tube | 150 |  |  | 0.02686 | 369.5971 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 369.5837 | 69.5881 | 69.6015 |  | 1.1095
0 | FE | ISRC1 | GV | 696 | FE_ISRC1:GV_D0696 | GV_L | gate valve | 150 |  |  | 0.080264 | 369.5436 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 369.5035 | 69.6417 | 69.6818 |  | 1.163
0 | [7 empty] | 150 |  |  | 0 | 369.5035 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 369.5035 | 69.6818 | 69.6818 | extra length due to a 10" CFF on GV_D0696 | 1.2032
0 | FE | ISRC1 | ACC | 698 | FE_ISRC1:ACC_D0698 |  | acceleration column | 150 |  |  | 0.26772 | 369.3696 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 369.2357 | 69.8156 | 69.9495 |  | 1.337
0 | FE | ISRC1 | ACC | 700 | FE_ISRC1:ACC_D0700 |  | suppression electrode | 150 |  |  | 0.024638 | 369.2234 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 369.2111 | 69.9618 | 69.9741 |  | 1.4832
0 | [10 empty] | 0.016854 | 369.2027 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 369.1942 | 69.9826 | 69.991 |  | 1.5039
0 | [10 empty] | 0.1905 | 369.099 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 369.0037 | 70.0862 | 70.1815 | Moved to D0695 location | 1.6076
0 | [7 empty] | 150 |  |  | 0.0254 | 368.991 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 368.9783 | 70.1942 | 70.2069 |  | 1.7156
0 | FE | SCS1 | SOLR | 704 | FE_SCS1:SOLR_D0704 | S4b | solenoid | 150 |  |  | 0.4 | 368.7783 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 368.5783 | 70.4069 | 70.6069 | 40cm effective length | 1.9283
0 | [7 empty] | 150 |  |  | 0.125 | 368.5158 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 368.4533 | 70.6694 | 70.7319 |  | 2.1908
0 | FE | SCS1 | TMP | 707 | FE_SCS1:TMP_D0707 | TMP_M | turbo pump, 700 l/s |  |  |  | 0 | 368.4533 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 368.4533 | 70.7319 | 70.7319 |  | 2.2533
0 | [7 empty] | 150 |  |  | 0.11975 | 368.3935 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 368.3336 | 70.7918 | 70.8516 |  | 2.3131
0 | [10 empty] | 0 | 368.3336 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 368.3336 | 70.8516 | 70.8516 |  | 2.373
0 | FE | SCS1 | DCHV | 709 | FE_SCS1:DCHV_D0709 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 368.2383 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 368.1431 | 70.9469 | 71.0421 | dimension to yoke not flange | 2.4683
0 | [7 empty] | 150 |  |  | 0.13975 | 368.0732 | 427.5728 | 51.6647 | 51.6647 | 427.5728 | 368.0033 | 71.112 | 71.1819 |  | 2.6334
0 | FE | SCS1 | DH | 717 | FE_SCS1:DH_D0717 | D5 | dipole-90deg, D5-1 | 120 x 150 |  |  | 1 | 367.685 | 427.2545 | 51.6647 | 51.6647 | 426.9362 | 367.3667 | 71.6819 | 72.1819 |  | 3.2033
0 | FE | SCS1 | TMP | 717 | FE_SCS1:TMP_D0717 | TMP_M | turbo pump, 700 l/s |  |  |  | 0 | 367.3667 | 426.9362 | 51.6647 | 51.6647 | 426.9362 | 367.3667 | 72.1819 | 72.1819 | special: pump port on the side of magnet | 3.7033
0 | [7 empty] | 150 |  |  | 0 | 367.3667 | 426.9362 | 51.6647 | 51.6647 | 426.9362 | 367.3667 | 72.1819 | 72.1819 | straight port | 3.7033
0 | [7 empty] | 150 |  |  | 0 | 367.3667 | 426.9362 | 51.6647 | 51.6647 | 426.9362 | 367.3667 | 72.1819 | 72.1819 |  | 3.7033
0 | [7 empty] | 150 |  |  | 0.06705 | 367.3667 | 426.9027 | 51.6647 | 51.6647 | 426.8692 | 367.3667 | 72.2154 | 72.2489 | dimension to yoke not flange | 3.7368
0 | FE | SCS1 | DCHV | 723 | FE_SCS1:DCHV_D0723 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 367.3667 | 426.7739 | 51.6647 | 51.6647 | 426.6787 | 367.3667 | 72.3442 | 72.4394 |  | 3.8656
0 | [7 empty] | 150 |  |  | 0.09245 | 367.3667 | 426.6324 | 51.6647 | 51.6647 | 426.5862 | 367.3667 | 72.4857 | 72.5319 |  | 4.007
0 | FE | SCS1 | QHE | 726 | FE_SCS1:QHE_D0726 | Q7 | ES_quad_X | 150 |  |  | 0.2 | 367.3667 | 426.4862 | 51.6647 | 51.6647 | 426.3862 | 367.3667 | 72.6319 | 72.7319 | TRIPLET HOUSING | 4.1533
0 | [7 empty] | 150 |  |  | 0.135 | 367.3667 | 426.3187 | 51.6647 | 51.6647 | 426.2512 | 367.3667 | 72.7994 | 72.8669 | TRIPLET HOUSING | 4.3208
0 | FE | SCS1 | QVE | 730 | FE_SCS1:QVE_D0730 | Q7 | ES_quad_Y | 150 |  |  | 0.2 | 367.3667 | 426.1512 | 51.6647 | 51.6647 | 426.0512 | 367.3667 | 72.9669 | 73.0669 | TRIPLET HOUSING | 4.4883
0 | [7 empty] | 150 |  |  | 0.135 | 367.3667 | 425.9837 | 51.6647 | 51.6647 | 425.9162 | 367.3667 | 73.1344 | 73.2019 | TRIPLET HOUSING | 4.6558
0 | FE | SCS1 | QHE | 733 | FE_SCS1:QHE_D0733 | Q7 | ES_quad_X | 150 |  |  | 0.2 | 367.3667 | 425.8162 | 51.6647 | 51.6647 | 425.7162 | 367.3667 | 73.3019 | 73.4019 | TRIPLET HOUSING | 4.8233
0 | [6 empty] | bellow+box | 150 |  |  | 0.22821 | 367.3667 | 425.6021 | 51.6647 | 51.6647 | 425.488 | 367.3667 | 73.516 | 73.6301 |  | 5.0374
0 | FE | SCS1 | SLT | 736 | FE_SCS1:SLT_D0736 |  | four jaw collimator | 150 |  |  | 0 | 367.3667 | 425.488 | 51.6647 | 51.6647 | 425.488 | 367.3667 | 73.6301 | 73.6301 |  | 5.1515
0 | [7 empty] | 150 |  |  | 0.021 | 367.3667 | 425.4775 | 51.6647 | 51.6647 | 425.467 | 367.3667 | 73.6406 | 73.6511 |  | 5.162
0 | FE | SCS1 | TMP | 737 | FE_SCS1:TMP_D0737 | TMP_M | turbo pump, 700 l/s |  |  |  | 0 | 367.3667 | 425.467 | 51.6647 | 51.6647 | 425.467 | 367.3667 | 73.6511 | 73.6511 |  | 5.1725
0 | [6 empty] | box | 150 |  |  | 0.11811 | 367.3667 | 425.408 | 51.6647 | 51.6647 | 425.3489 | 367.3667 | 73.7102 | 73.7692 |  | 5.2315
0 | FE | SCS1 | FC | 738 | FE_SCS1:FC_D0738 | FC2 | Faraday cup | 150 |  |  | 0 | 367.3667 | 425.3489 | 51.6647 | 51.6647 | 425.3489 | 367.3667 | 73.7692 | 73.7692 |  | 5.2906
0 | [7 empty] | 150 |  |  | 0.18034 | 367.3667 | 425.2587 | 51.6647 | 51.6647 | 425.1686 | 367.3667 | 73.8594 | 73.9496 |  | 5.3808
0 | FE | SCS1 | VD | 739 | FE_SCS1:VD_D0739 | V1 | viewer | 150 |  |  | 0 | 367.3667 | 425.1686 | 51.6647 | 51.6647 | 425.1686 | 367.3667 | 73.9496 | 73.9496 |  | 5.4709
0 | FE | SCS1 | EMS | 739 | FE_SCS1:EMS_D0739 | EM1 | emittance scanner (x & y) | 150 |  |  | 0 | 367.3667 | 425.1686 | 51.6647 | 51.6647 | 425.1686 | 367.3667 | 73.9496 | 73.9496 |  | 5.4709
0 | [6 empty] | bellow+box | 150 |  |  | 0.21234 | 367.3667 | 425.0624 | 51.6647 | 51.6647 | 424.9562 | 367.3667 | 74.0557 | 74.1619 |  | 5.5771
0 | FE | SCS1 | QHE | 743 | FE_SCS1:QHE_D0743 | Q7 | ES_quad_X | 150 |  |  | 0.2 | 367.3667 | 424.8562 | 51.6647 | 51.6647 | 424.7562 | 367.3667 | 74.2619 | 74.3619 | TRIPLET HOUSING | 5.7833
0 | [7 empty] | 150 |  |  | 0.135 | 367.3667 | 424.6887 | 51.6647 | 51.6647 | 424.6212 | 367.3667 | 74.4294 | 74.4969 | TRIPLET HOUSING | 5.9508
0 | FE | SCS1 | QVE | 746 | FE_SCS1:QVE_D0746 | Q7 | ES_quad_Y | 150 |  |  | 0.2 | 367.3667 | 424.5212 | 51.6647 | 51.6647 | 424.4212 | 367.3667 | 74.5969 | 74.6969 | TRIPLET HOUSING | 6.1183
0 | [7 empty] | 150 |  |  | 0.135 | 367.3667 | 424.3537 | 51.6647 | 51.6647 | 424.2862 | 367.3667 | 74.7644 | 74.8319 | TRIPLET HOUSING | 6.2858
0 | FE | SCS1 | QHE | 749 | FE_SCS1:QHE_D0749 | Q7 | ES_quad_X | 150 |  |  | 0.2 | 367.3667 | 424.1862 | 51.6647 | 51.6647 | 424.0862 | 367.3667 | 74.9319 | 75.0319 | TRIPLET HOUSING | 6.4533
0 | [6 empty] | bellow | 150 |  |  | 0.14959 | 367.3667 | 424.0114 | 51.6647 | 51.6647 | 423.9366 | 367.3667 | 75.1067 | 75.1815 |  | 6.6281
0 | FE | SCS1 | GV | 752 | FE_SCS1:GV_D0752 | GV_L | gate valve | 150 |  |  | 0.080264 | 367.3667 | 423.8965 | 51.6647 | 51.6647 | 423.8564 | 367.3667 | 75.2216 | 75.2618 |  | 6.743
0 | [7 empty] | 150 |  |  | 0.12014 | 367.3667 | 423.7963 | 51.6647 | 51.6647 | 423.7362 | 367.3667 | 75.3218 | 75.3819 | dimension to yoke not flange | 6.8432
0 | FE | LEBT | DH | 759 | FE_LEBT:DH_D0759 | D5 | dipole-90deg, D5-2 | 120 x 150 |  |  | 1 | 367.685 | 423.4179 | 51.6647 | 51.6647 | 423.0996 | 368.0033 | 75.8819 | 76.3819 |  | 7.4033
0 | [7 empty] | 150 |  |  | 0.225 | 368.1158 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 368.2283 | 76.4944 | 76.6069 | dimension to yoke not flange | 8.0158
0 | FE | LEBT | QVE | 767 | FE_LEBT:QVE_D0767 | Q7 | ES_quad_Y | 150 |  |  | 0.2 | 368.3283 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 368.4283 | 76.7069 | 76.8069 | DOUBLET HOUSING | 8.2283
0 | [7 empty] | 150 |  |  | 0.135 | 368.4958 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 368.5633 | 76.8744 | 76.9419 | DOUBLET HOUSING | 8.3958
0 | FE | LEBT | QHE | 770 | FE_LEBT:QHE_D0770 | Q7 | ES_quad_X | 150 |  |  | 0.2 | 368.6633 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 368.7633 | 77.0419 | 77.1419 | DOUBLET HOUSING | 8.5633
0 | [7 empty] | 150 |  |  | 0.10225 | 368.8145 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 368.8656 | 77.193 | 77.2441 |  | 8.7144
0 | [7 empty] | 150 |  |  | 0 | 368.8656 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 368.8656 | 77.2441 | 77.2441 |  | 8.7655
0 | FE | LEBT | DCHV | 773 | FE_LEBT:DCHV_D0773 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 368.9608 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 369.0561 | 77.3394 | 77.4346 |  | 8.8608
0 | [7 empty] | 150 |  |  | 0.10225 | 369.1072 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 369.1583 | 77.4858 | 77.5369 |  | 9.0071
0 | FE | LEBT | QHE | 776 | FE_LEBT:QHE_D0776 | Q7 | ES_quad_X | 150 |  |  | 0.2 | 369.2583 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 369.3583 | 77.6369 | 77.7369 | DOUBLET HOUSING | 9.1583
0 | [7 empty] | 150 |  |  | 0.135 | 369.4258 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 369.4933 | 77.8044 | 77.8719 | DOUBLET HOUSING | 9.3258
0 | FE | LEBT | QVE | 780 | FE_LEBT:QVE_D0780 | Q7 | ES_quad_Y | 150 |  |  | 0.2 | 369.5933 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 369.6933 | 77.9719 | 78.0719 | DOUBLET HOUSING | 9.4933
0 | [7 empty] | 150 |  |  | 0.19159 | 369.7891 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 369.8849 | 78.1677 | 78.2635 |  | 9.6891
0 | FE | LEBT | PM | 783 | FE_LEBT:PM_D0783 | PM1 | profile monitor (large fork) | 150 |  |  | 0 | 369.8849 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 369.8849 | 78.2635 | 78.2635 |  | 9.7849
0 | FE | LEBT | TMP | 783 | FE_LEBT:TMP_D0783 | TMP_M | pump, 700 l/s |  |  |  | 0 | 369.8849 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 369.8849 | 78.2635 | 78.2635 |  | 9.7849
0 | [7 empty] | 150 |  |  | 0.21841 | 369.9941 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 370.1033 | 78.3727 | 78.4819 |  | 9.8941
0 | FE | LEBT | SOLR | 787 | FE_LEBT:SOLR_D0787 | S4 | solenoid | 150 |  |  | 0.4 | 370.3033 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 370.5033 | 78.6819 | 78.8819 | 40cm effective length | 10.2033
0 | [7 empty] | 150 |  |  | 0.0254 | 370.516 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 370.5287 | 78.8946 | 78.9073 |  | 10.416
0 | FE | LEBT | DCHV | 790 | FE_LEBT:DCHV_D0790 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 370.624 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 370.7192 | 79.0025 | 79.0978 |  | 10.5239
0 | [7 empty] | 150 |  |  | 0.060413 | 370.7495 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 370.7797 | 79.128 | 79.1582 |  | 10.6494
0 | FE | LEBT | CHP | 792 | FE_LEBT:CHP_D0792 |  | chopper | 70 x 150 |  |  | 0.1 | 370.8297 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 370.8797 | 79.2082 | 79.2582 | 10cm deflector plates | 10.7296
0 | [6 empty] | box+tube | 150 |  |  | 0.20349 | 370.9814 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 371.0831 | 79.36 | 79.4617 |  | 10.8813
0 | [7 empty] | 150 |  |  | 0 | 371.0831 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 371.0831 | 79.4617 | 79.4617 |  | 10.9831
0 | FE | LEBT | NEGP | 795 | FE_LEBT:NEGP_D0795 | NEG-0 | getter pump, 3500 l/s |  |  |  | 0 | 371.0831 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 371.0831 | 79.4617 | 79.4617 |  | 10.9831
0 | [7 empty] | 150 |  |  | 0.11717 | 371.1417 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 371.2003 | 79.5203 | 79.5789 |  | 11.0417
0 | FE | LEBT | AP | 796 | FE_LEBT:AP_D0796 |  | apertures | 150 |  |  | 0 | 371.2003 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 371.2003 | 79.5789 | 79.5789 |  | 11.1002
0 | [6 empty] | bellow+tube/box | 150 |  |  | 0.025375 | 371.213 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 371.2257 | 79.5916 | 79.6042 |  | 11.1129
0 | FE | LEBT | FC | 796 | FE_LEBT:FC_D0796 | FC1 | Faraday cup | 150 |  |  | 0 | 371.2257 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 371.2257 | 79.6042 | 79.6042 |  | 11.1256
0 | [7 empty] | 150 |  |  | 0.3528 | 371.4021 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 371.5785 | 79.7806 | 79.957 |  | 11.302
0 | FE | LEBT | SOLR | 802 | FE_LEBT:SOLR_D0802 | S4 | solenoid | 150 |  |  | 0.4 | 371.7785 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 371.9785 | 80.157 | 80.357 | 40cm effective length | 11.6784
0 | [7 empty] | 150 |  |  | 0.0254 | 371.9912 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 372.0039 | 80.3697 | 80.3824 |  | 11.8911
0 | FE | LEBT | DCHV | 805 | FE_LEBT:DCHV_D0805 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 372.0991 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 372.1944 | 80.4777 | 80.5729 |  | 11.9991
0 | [7 empty] | 150 |  |  | 0.16071 | 372.2748 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 372.3551 | 80.6533 | 80.7337 |  | 12.1747
0 | FE | LEBT | AP | 807 | FE_LEBT:AP_D0807 |  | apertures | 150 |  |  | 0 | 372.3551 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 372.3551 | 80.7337 | 80.7337 |  | 12.255
0 | [7 empty] | 150 |  |  | 0.042926 | 372.3766 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 372.398 | 80.7551 | 80.7766 |  | 12.2765
0 | FE | LEBT | PM | 808 | FE_LEBT:PM_D0808 | PM1 | profile monitor (large fork) | 150 |  |  | 0 | 372.398 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 372.398 | 80.7766 | 80.7766 |  | 12.298
0 | [7 empty] | 150 |  |  | 0.3954 | 372.5957 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 372.7934 | 80.9743 | 81.172 |  | 12.4957
0 | FE | LEBT | EMS | 812 | FE_LEBT:EMS_D0812 | EM2 | Pepper Pot | 150 |  |  | 0 | 372.7934 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 372.7934 | 81.172 | 81.172 |  | 12.6934
0 | FE | LEBT | VD | 812 | FE_LEBT:VD_D0812 | V1 | viewer | 150 |  |  | 0 | 372.7934 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 372.7934 | 81.172 | 81.172 |  | 12.6934
0 | [7 empty] | 150 |  |  | 0.20828 | 372.8976 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 373.0017 | 81.2761 | 81.3803 |  | 12.7975
0 | FE | LEBT | NEGP | 814 | FE_LEBT:NEGP_D0814 | NEG-0 | getter pump, 3500 l/s |  |  |  | 0 | 373.0017 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 373.0017 | 81.3803 | 81.3803 |  | 12.9016
0 | FE | LEBT | FC | 814 | FE_LEBT:FC_D0814 | FC1 | Faraday cup | 150 |  |  | 0 | 373.0017 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 373.0017 | 81.3803 | 81.3803 |  | 12.9016
0 | [7 empty] | 150 |  |  | 0.21693 | 373.1102 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 373.2186 | 81.4887 | 81.5972 | 40cm effective length | 13.0101
0 | FE | LEBT | SOLR | 818 | FE_LEBT:SOLR_D0818 | S4 | solenoid | 150 |  |  | 0.4 | 373.4186 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 373.6186 | 81.7972 | 81.9972 |  | 13.3186
0 | [7 empty] | 150 |  |  | 0.0254 | 373.6313 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 373.644 | 82.0099 | 82.0226 |  | 13.5313
0 | FE | LEBT | DCHV | 821 | FE_LEBT:DCHV_D0821 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 373.7393 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 373.8345 | 82.1178 | 82.2131 |  | 13.6392
0 | [7 empty] | 150 |  |  | 0.13973 | 373.9044 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 373.9743 | 82.283 | 82.3528 |  | 13.8043
0 | FE | LEBT | BD | 824 | FE_LEBT:BD_D0824 | dump | beam plug (low power) |  |  |  | 0 | 373.9743 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 373.9743 | 82.3528 | 82.3528 |  | 13.8742
0 | [7 empty] | 150 |  |  | 0.038798 | 373.9937 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 374.0131 | 82.3722 | 82.3916 |  | 13.8936
0 | FE | LEBT | PM | 824 | FE_LEBT:PM_D0824 | PM1 | profile monitor (large fork) | 150 |  |  | 0 | 374.0131 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 374.0131 | 82.3916 | 82.3916 |  | 13.913
0 | [7 empty] | 150 |  |  | 0.2018 | 374.114 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 374.2149 | 82.4925 | 82.5934 |  | 14.0139
0 | FE | LEBT | GV | 826 | FE_LEBT:GV_D0826 | GV_L | gate valve | 150 |  |  | 0.080264 | 374.255 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 374.2951 | 82.6336 | 82.6737 |  | 14.1549
0 | [7 empty] | 150 |  |  | 0.12351 | 374.3569 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 374.4186 | 82.7354 | 82.7972 | dimension to yoke not flange | 14.2568
0 | FE | LEBT | DVE | 833 | FE_LEBT:DVE_D0833 | D7 | ES_bender 1 | 100 x 150 |  |  | 1 | 374.737 | 423.0996 | 51.3464 | 51.0281 | 423.0996 | 375.0553 | 83.2972 | 83.7972 |  | 14.8186
0 | FE | LEBT | TMP | 833 | FE_LEBT:TMP_D0833 | TMP_M | pump, 700 l/s |  |  |  | 0 | 375.0553 | 423.0996 | 51.0281 | 51.0281 | 423.0996 | 375.0553 | 83.7972 | 83.7972 | special: pump port on the side of e-bender | 15.3186
0 | [7 empty] | 150 |  |  | 0 | 375.0553 | 423.0996 | 50.9682 | 50.9084 | 423.0996 | 375.0553 | 83.7972 | 83.7972 |  | 15.3186
0 | [7 empty] | 150 |  |  | 0.11975 | 375.0553 | 423.0996 | 50.9084 | 50.9084 | 423.0996 | 375.0553 | 83.8571 | 83.9169 | dimension to yoke not flange | 15.3784
0 | FE | LEBT | DCHV | 840 | FE_LEBT:DCHV_D0840 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 375.0553 | 423.0996 | 50.8131 | 50.7179 | 423.0996 | 375.0553 | 84.0122 | 84.1074 |  | 15.5336
0 | [7 empty] | 150 |  |  | 0.23975 | 375.0553 | 423.0996 | 50.598 | 50.4781 | 423.0996 | 375.0553 | 84.2273 | 84.3472 |  | 15.7487
0 | FE | LEBT | QHE | 844 | FE_LEBT:QHE_D0844 | Q7 | ES_quad_X | 150 |  |  | 0.2 | 375.0553 | 423.0996 | 50.3781 | 50.2781 | 423.0996 | 375.0553 | 84.4472 | 84.5472 | TRIPLET HOUSING | 15.9686
0 | [7 empty] | 150 |  |  | 0.135 | 375.0553 | 423.0996 | 50.2106 | 50.1431 | 423.0996 | 375.0553 | 84.6147 | 84.6822 | TRIPLET HOUSING | 16.1361
0 | FE | LEBT | QVE | 848 | FE_LEBT:QVE_D0848 | Q7 | ES_quad_Y | 150 |  |  | 0.2 | 375.0553 | 423.0996 | 50.0431 | 49.9431 | 423.0996 | 375.0553 | 84.7822 | 84.8822 | TRIPLET HOUSING | 16.3036
0 | [7 empty] | 150 |  |  | 0.135 | 375.0553 | 423.0996 | 49.8756 | 49.8081 | 423.0996 | 375.0553 | 84.9497 | 85.0172 | TRIPLET HOUSING | 16.4711
0 | FE | LEBT | QHE | 851 | FE_LEBT:QHE_D0851 | Q7 | ES_quad_X | 150 |  |  | 0.2 | 375.0553 | 423.0996 | 49.7081 | 49.6081 | 423.0996 | 375.0553 | 85.1172 | 85.2172 | TRIPLET HOUSING | 16.6386
0 | [6 empty] | bellow ? | 150 |  |  | 0.34302 | 375.0553 | 423.0996 | 49.4366 | 49.2651 | 423.0996 | 375.0553 | 85.3887 | 85.5602 |  | 16.9101
0 | FE | LEBT | PM | 856 | FE_LEBT:PM_D0856 | PM1 | profile monitor (large fork) | 150 |  |  | 0 | 375.0553 | 423.0996 | 49.2651 | 49.2651 | 423.0996 | 375.0553 | 85.5602 | 85.5602 |  | 17.0816
0 | [6 empty] | bellow+tube | 150 |  |  | 0.73973 | 375.0553 | 423.0996 | 48.8952 | 48.5254 | 423.0996 | 375.0553 | 85.9301 | 86.2999 |  | 17.4514
0 | [6 empty] | 4 way cross | 150 |  |  | 0 | 375.0553 | 423.0996 | 48.5254 | 48.5254 | 423.0996 | 375.0553 | 86.2999 | 86.2999 |  | 17.8213
0 | FE | LEBT | NEGP | 863 | FE_LEBT:NEGP_D0863 | NEG-0 | getter pump, 3500 l/s |  |  |  | 0 | 375.0553 | 423.0996 | 48.5254 | 48.5254 | 423.0996 | 375.0553 | 86.2999 | 86.2999 |  | 17.8213
0 | [7 empty] | 150 |  |  | 0.42075 | 375.0553 | 423.0996 | 48.315 | 48.1046 | 423.0996 | 375.0553 | 86.5103 | 86.7207 |  | 18.0317
0 | FE | LEBT | DCHV | 868 | FE_LEBT:DCHV_D0868 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 375.0553 | 423.0996 | 48.0094 | 47.9141 | 423.0996 | 375.0553 | 86.8159 | 86.9112 |  | 18.3373
0 | [7 empty] | 150 |  |  | 0.08475 | 375.0553 | 423.0996 | 47.8717 | 47.8294 | 423.0996 | 375.0553 | 86.9536 | 86.9959 |  | 18.4749
0 | FE | LEBT | QVE | 871 | FE_LEBT:QVE_D0871 | Q7 | ES_quad_Y | 150 |  |  | 0.2 | 375.0553 | 423.0996 | 47.7294 | 47.6294 | 423.0996 | 375.0553 | 87.0959 | 87.1959 | TRIPLET HOUSING | 18.6173
0 | [7 empty] | 150 |  |  | 0.135 | 375.0553 | 423.0996 | 47.5619 | 47.4944 | 423.0996 | 375.0553 | 87.2634 | 87.3309 | TRIPLET HOUSING | 18.7848
0 | FE | LEBT | QHE | 874 | FE_LEBT:QHE_D0874 | Q7 | ES_quad_X | 150 |  |  | 0.2 | 375.0553 | 423.0996 | 47.3944 | 47.2944 | 423.0996 | 375.0553 | 87.4309 | 87.5309 | TRIPLET HOUSING | 18.9523
0 | [7 empty] | 150 |  |  | 0.135 | 375.0553 | 423.0996 | 47.2269 | 47.1594 | 423.0996 | 375.0553 | 87.5984 | 87.6659 | TRIPLET HOUSING | 19.1198
0 | FE | LEBT | QVE | 878 | FE_LEBT:QVE_D0878 | Q7 | ES_quad_Y | 150 |  |  | 0.2 | 375.0553 | 423.0996 | 47.0594 | 46.9594 | 423.0996 | 375.0553 | 87.7659 | 87.8659 | TRIPLET HOUSING | 19.2873
0 | [7 empty] | 150 |  |  | 0.08475 | 375.0553 | 423.0996 | 46.917 | 46.8746 | 423.0996 | 375.0553 | 87.9083 | 87.9507 |  | 19.4297
0 | FE | LEBT | DCHV | 880 | FE_LEBT:DCHV_D0880 | C4 | corrector+bellow+tube | 150 |  |  | 0.1905 | 375.0553 | 423.0996 | 46.7794 | 46.6841 | 423.0996 | 375.0553 | 88.0459 | 88.1412 |  | 19.5673
0 | [7 empty] | 150 |  |  | 0.37948 | 375.0553 | 423.0996 | 46.4944 | 46.3046 | 423.0996 | 375.0553 | 88.3309 | 88.5207 |  | 19.8523
0 | FE | LEBT | PM | 885 | FE_LEBT:PM_D0885 | PM1 | profile monitor (large fork) | 150 |  |  | 0 | 375.0553 | 423.0996 | 46.3046 | 46.3046 | 423.0996 | 375.0553 | 88.5207 | 88.5207 |  | 20.042
0 | [7 empty] | 150 |  |  | 0.058903 | 375.0553 | 423.0996 | 46.2752 | 46.2457 | 423.0996 | 375.0553 | 88.5501 | 88.5796 |  | 20.0715
0 | [6 empty] | 4 way cross ?? | 150 |  |  | 0 | 375.0553 | 423.0996 | 46.2457 | 46.2457 | 423.0996 | 375.0553 | 88.5796 | 88.5796 |  | 20.1009
0 | FE | LEBT | NEGP | 886 | FE_LEBT:NEGP_D0886 | NEG-0 | getter pump, 3500 l/s |  |  |  | 0 | 375.0553 | 423.0996 | 46.2457 | 46.2457 | 423.0996 | 375.0553 | 88.5796 | 88.5796 |  | 20.1009
0 | [6 empty] | tube | 150 |  |  | 0.44637 | 375.0553 | 423.0996 | 46.0226 | 45.7994 | 423.0996 | 375.0553 | 88.8028 | 89.0259 |  | 20.3241
0 | FE | LEBT | QVE | 891 | FE_LEBT:QVE_D0891 | Q7 | ES_quad_Y | 150 |  |  | 0.2 | 375.0553 | 423.0996 | 45.6994 | 45.5994 | 423.0996 | 375.0553 | 89.1259 | 89.2259 | TRIPLET HOUSING | 20.6473
0 | [7 empty] | 150 |  |  | 0.135 | 375.0553 | 423.0996 | 45.5319 | 45.4644 | 423.0996 | 375.0553 | 89.2934 | 89.3609 | TRIPLET HOUSING | 20.8148
0 | FE | LEBT | QHE | 895 | FE_LEBT:QHE_D0895 | Q7 | ES_quad_X | 150 |  |  | 0.2 | 375.0553 | 423.0996 | 45.3644 | 45.2644 | 423.0996 | 375.0553 | 89.4609 | 89.5609 | TRIPLET HOUSING | 20.9823
0 | [7 empty] | 150 |  |  | 0.135 | 375.0553 | 423.0996 | 45.1969 | 45.1294 | 423.0996 | 375.0553 | 89.6284 | 89.6959 | TRIPLET HOUSING | 21.1498
0 | FE | LEBT | QVE | 898 | FE_LEBT:QVE_D0898 | Q7 | ES_quad_Y | 150 |  |  | 0.2 | 375.0553 | 423.0996 | 45.0294 | 44.9294 | 423.0996 | 375.0553 | 89.7959 | 89.8959 | TRIPLET HOUSING | 21.3173
0 | [7 empty] | 150 |  |  | 0.08475 | 375.0553 | 423.0996 | 44.887 | 44.8446 | 423.0996 | 375.0553 | 89.9383 | 89.9807 |  | 21.4597
0 | FE | LEBT | DCHV | 901 | FE_LEBT:DCHV_D0901 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 375.0553 | 423.0996 | 44.7494 | 44.6541 | 423.0996 | 375.0553 | 90.0759 | 90.1712 |  | 21.5973
0 | [7 empty] | 150 |  |  | 0.42077 | 375.0553 | 423.0996 | 44.4437 | 44.2334 | 423.0996 | 375.0553 | 90.3816 | 90.592 |  | 21.903
0 | [6 empty] | 4 way cross ?? | 150 |  |  | 0 | 375.0553 | 423.0996 | 44.2334 | 44.2334 | 423.0996 | 375.0553 | 90.592 | 90.592 |  | 22.1133
0 | FE | LEBT | NEGP | 906 | FE_LEBT:NEGP_D0906 | NEG-0 | getter pump, 3500 l/s |  |  |  | 0 | 375.0553 | 423.0996 | 44.2334 | 44.2334 | 423.0996 | 375.0553 | 90.592 | 90.592 |  | 22.1133
0 | [6 empty] | tube | 150 |  |  | 0.57746 | 375.0553 | 423.0996 | 43.9446 | 43.6559 | 423.0996 | 375.0553 | 90.8807 | 91.1694 |  | 22.4021
0 | [6 empty] | 4 way cross | 150 |  |  | 0 | 375.0553 | 423.0996 | 43.6559 | 43.6559 | 423.0996 | 375.0553 | 91.1694 | 91.1694 |  | 22.6908
0 | [7 empty] | 150 |  |  | 0 | 375.0553 | 423.0996 | 43.6559 | 43.6559 | 423.0996 | 375.0553 | 91.1694 | 91.1694 |  | 22.6908
0 | FE | LEBT | PM | 912 | FE_LEBT:PM_D0912 | PM1 | profile monitor (large fork) | 150 |  |  | 0 | 375.0553 | 423.0996 | 43.6559 | 43.6559 | 423.0996 | 375.0553 | 91.1694 | 91.1694 |  | 22.6908
0 | [6 empty] | bellow | 150 |  |  | 0.50527 | 375.0553 | 423.0996 | 43.4033 | 43.1506 | 423.0996 | 375.0553 | 91.4221 | 91.6747 |  | 22.9434
0 | FE | LEBT | QHE | 918 | FE_LEBT:QHE_D0918 | Q7 | ES_quad_X | 150 |  |  | 0.2 | 375.0553 | 423.0996 | 43.0506 | 42.9506 | 423.0996 | 375.0553 | 91.7747 | 91.8747 | TRIPLET HOUSING | 23.2961
0 | [7 empty] | 150 |  |  | 0.135 | 375.0553 | 423.0996 | 42.8831 | 42.8156 | 423.0996 | 375.0553 | 91.9422 | 92.0097 | TRIPLET HOUSING | 23.4636
0 | FE | LEBT | QVE | 921 | FE_LEBT:QVE_D0921 | Q7 | ES_quad_Y | 150 |  |  | 0.2 | 375.0553 | 423.0996 | 42.7156 | 42.6156 | 423.0996 | 375.0553 | 92.1097 | 92.2097 | TRIPLET HOUSING | 23.6311
0 | [7 empty] | 150 |  |  | 0.135 | 375.0553 | 423.0996 | 42.5481 | 42.4806 | 423.0996 | 375.0553 | 92.2772 | 92.3447 | TRIPLET HOUSING | 23.7986
0 | FE | LEBT | QHE | 924 | FE_LEBT:QHE_D0924 | Q7 | ES_quad_X | 150 |  |  | 0.2 | 375.0553 | 423.0996 | 42.3806 | 42.2806 | 423.0996 | 375.0553 | 92.4447 | 92.5447 | TRIPLET HOUSING | 23.9661
0 | [7 empty] | 150 |  |  | 0.23975 | 375.0553 | 423.0996 | 42.1607 | 42.0409 | 423.0996 | 375.0553 | 92.6646 | 92.7844 |  | 24.1859
0 | FE | LEBT | DCHV | 929 | FE_LEBT:DCHV_D0929 | C4 | corrector+bellow+tube | 150 |  |  | 0.1905 | 375.0553 | 423.0996 | 41.9456 | 41.8504 | 423.0996 | 375.0553 | 92.8797 | 92.9749 |  | 24.4011
0 | [7 empty] | 150 |  |  | 0.11975 | 375.0553 | 423.0996 | 41.7905 | 41.7306 | 423.0996 | 375.0553 | 93.0348 | 93.0947 | dimension to yoke not flange | 24.5562
0 | FE | LEBT | DVE | 936 | FE_LEBT:DVE_D0936 | D7 | ES_bender 2 | 100 x 150 |  |  | 1 | 374.737 | 423.0996 | 41.4123 | 41.094 | 423.0996 | 374.4186 | 93.5947 | 94.0947 |  | 25.1161
0 | FE | LEBT | NEGP | 936 | FE_LEBT:NEGP_D0936 | NEG-0 | getter pump, 3500 l/s |  |  |  | 0 | 374.737 | 423.0996 | 41.094 | 41.094 | 423.0996 | 375.0553 | 94.0947 | 94.0947 | special: pump port on the side of e-bender | 25.6161
0 | [7 empty] | 150 |  |  | 3e-06 | 374.737 | 423.0996 | 41.094 | 41.094 | 423.0996 | 374.4186 | 94.0947 | 94.0947 | adjustment for residuals | 25.6161
0 | [7 empty] | 150 |  |  | 0.12351 | 374.3569 | 423.0996 | 41.094 | 41.094 | 423.0996 | 374.2951 | 94.1565 | 94.2182 | dimension to yoke not flange | 25.6778
0 | FE | LEBT | GV | 943 | FE_LEBT:GV_D0943 | GV_L | gate valve | 150 |  |  | 0.080264 | 374.255 | 423.0996 | 41.094 | 41.094 | 423.0996 | 374.2149 | 94.2583 | 94.2985 |  | 25.7797
0 | [7 empty] | 150 |  |  | 0.2018 | 374.114 | 423.0996 | 41.094 | 41.094 | 423.0996 | 374.0131 | 94.3994 | 94.5003 |  | 25.9207
0 | FE | LEBT | PM | 945 | FE_LEBT:PM_D0945 | PM1 | profile monitor (large fork) | 150 |  |  | 0 | 374.0131 | 423.0996 | 41.094 | 41.094 | 423.0996 | 374.0131 | 94.5003 | 94.5003 |  | 26.0216
0 | [7 empty] | 150 |  |  | 0.17853 | 373.9238 | 423.0996 | 41.094 | 41.094 | 423.0996 | 373.8345 | 94.5895 | 94.6788 |  | 26.1109
0 | FE | LEBT | DCHV | 948 | FE_LEBT:DCHV_D0948 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 373.7393 | 423.0996 | 41.094 | 41.094 | 423.0996 | 373.644 | 94.774 | 94.8693 |  | 26.2954
0 | [7 empty] | 150 |  |  | 0.0254 | 373.6313 | 423.0996 | 41.094 | 41.094 | 423.0996 | 373.6186 | 94.882 | 94.8947 |  | 26.4034
0 | FE | LEBT | SOLR | 951 | FE_LEBT:SOLR_D0951 | S4 | solenoid | 150 |  |  | 0.4 | 373.4186 | 423.0996 | 41.094 | 41.094 | 423.0996 | 373.2186 | 95.0947 | 95.2947 | 40cm effective length | 26.6161
0 | [6 empty] | bellow+tube | 150 |  |  | 0.21693 | 373.1102 | 423.0996 | 41.094 | 41.094 | 423.0996 | 373.0017 | 95.4032 | 95.5116 |  | 26.9245
0 | FE | LEBT | NEGP | 955 | FE_LEBT:NEGP_D0955 | NEG-0 | getter pump, 3500 l/s |  |  |  | 0 | 373.0017 | 423.0996 | 41.094 | 41.094 | 423.0996 | 373.0017 | 95.5116 | 95.5116 |  | 27.033
0 | [7 empty] | 150 |  |  | 0.18546 | 372.909 | 423.0996 | 41.094 | 41.094 | 423.0996 | 372.8163 | 95.6044 | 95.6971 |  | 27.1257
0 | FE | LEBT | ATT | 957 | FE_LEBT:ATT_D0957 | ATT1 | beam attenuator | 150 |  |  | 0 | 372.8163 | 423.0996 | 41.094 | 41.094 | 423.0996 | 372.8163 | 95.6971 | 95.6971 |  | 27.2185
0 | [7 empty] | 150 |  |  | 0.073025 | 372.7797 | 423.0996 | 41.094 | 41.094 | 423.0996 | 372.7432 | 95.7336 | 95.7701 |  | 27.255
0 | FE | LEBT | ATT | 958 | FE_LEBT:ATT_D0958 | ATT1 | beam attenuator | 150 |  |  | 0 | 372.7432 | 423.0996 | 41.094 | 41.094 | 423.0996 | 372.7432 | 95.7701 | 95.7701 |  | 27.2915
0 | [7 empty] | 150 |  |  | 0.34527 | 372.5706 | 423.0996 | 41.094 | 41.094 | 423.0996 | 372.398 | 95.9427 | 96.1154 |  | 27.4641
0 | FE | LEBT | PM | 961 | FE_LEBT:PM_D0961 | PM1 | profile monitor (large fork) | 150 |  |  | 0 | 372.398 | 423.0996 | 41.094 | 41.094 | 423.0996 | 372.398 | 96.1154 | 96.1154 |  | 27.6368
0 | [7 empty] | 150 |  |  | 0.20357 | 372.2962 | 423.0996 | 41.094 | 41.094 | 423.0996 | 372.1944 | 96.2172 | 96.3189 |  | 27.7385
0 | FE | LEBT | DCHV | 964 | FE_LEBT:DCHV_D0964 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 372.0991 | 423.0996 | 41.094 | 41.094 | 423.0996 | 372.0039 | 96.4142 | 96.5094 |  | 27.9356
0 | [7 empty] | 150 |  |  | 0.0254 | 371.9912 | 423.0996 | 41.094 | 41.094 | 423.0996 | 371.9785 | 96.5221 | 96.5348 |  | 28.0435
0 | [7 empty] | 150 |  |  | 0 | 371.9785 | 423.0996 | 41.094 | 41.094 | 423.0996 | 371.9785 | 96.5348 | 96.5348 |  | 28.0562
0 | FE | LEBT | SOLR | 967 | FE_LEBT:SOLR_D0967 | S4 | solenoid | 150 |  |  | 0.4 | 371.7785 | 423.0996 | 41.094 | 41.094 | 423.0996 | 371.5785 | 96.7348 | 96.9348 | 40cm effective length | 28.2562
0 | [6 empty] | bellow | 150 |  |  | 0.21855 | 371.4692 | 423.0996 | 41.094 | 41.094 | 423.0996 | 371.3599 | 97.0441 | 97.1534 |  | 28.5655
0 | FE | LEBT | PM | 972 | FE_LEBT:PM_D0972 | PM1 | profile monitor (large fork) | 150 |  |  | 0 | 371.3599 | 423.0996 | 41.094 | 41.094 | 423.0996 | 371.3599 | 97.1534 | 97.1534 | no longer has aperture | 28.6748
0 | [7 empty] | 150 |  |  | 0.28661 | 371.2166 | 423.0996 | 41.094 | 41.094 | 423.0996 | 371.0733 | 97.2967 | 97.44 |  | 28.8181
0 | FE | LEBT | ATT | 974 | FE_LEBT:ATT_D0974 | ATT1 | beam attenuator | 150 |  |  | 0 | 371.0733 | 423.0996 | 41.094 | 41.094 | 423.0996 | 371.0733 | 97.44 | 97.44 |  | 28.9614
0 | [7 empty] | 150 |  |  | 0.073025 | 371.0368 | 423.0996 | 41.094 | 41.094 | 423.0996 | 371.0003 | 97.4765 | 97.513 |  | 28.9979
0 | FE | LEBT | ATT | 975 | FE_LEBT:ATT_D0975 | ATT1 | beam attenuator | 150 |  |  | 0 | 371.0003 | 423.0996 | 41.094 | 41.094 | 423.0996 | 371.0003 | 97.513 | 97.513 |  | 29.0344
0 | [7 empty] | 150 |  |  | 0.1563 | 370.9222 | 423.0996 | 41.094 | 41.094 | 423.0996 | 370.844 | 97.5912 | 97.6693 |  | 29.1126
0 | FE | LEBT | VD | 977 | FE_LEBT:VD_D0977 | V2 | viewer | 150 |  |  | 0 | 370.844 | 423.0996 | 41.094 | 41.094 | 423.0996 | 370.844 | 97.6693 | 97.6693 |  | 29.1907
0 | [7 empty] | 150 |  |  | 0 | 370.844 | 423.0996 | 41.094 | 41.094 | 423.0996 | 370.844 | 97.6693 | 97.6693 |  | 29.1907
0 | FE | LEBT | IP | 977 | FE_LEBT:IP_D0977 | IP_L | ion pump, 300 L/s | 150 |  |  | 0 | 370.844 | 423.0996 | 41.094 | 41.094 | 423.0996 | 370.844 | 97.6693 | 97.6693 |  | 29.1907
0 | [7 empty] | 150 |  |  | 0 | 370.844 | 423.0996 | 41.094 | 41.094 | 423.0996 | 370.844 | 97.6693 | 97.6693 |  | 29.1907
0 | FE | LEBT | FC | 977 | FE_LEBT:FC_D0977 | FC1 | faraday cup | 150 |  |  | 0 | 370.844 | 423.0996 | 41.094 | 41.094 | 423.0996 | 370.844 | 97.6693 | 97.6693 |  | 29.1907
0 | [7 empty] | 150 |  |  | 0.12477 | 370.7816 | 423.0996 | 41.094 | 41.094 | 423.0996 | 370.7192 | 97.7317 | 97.7941 |  | 29.2531
0 | FE | LEBT | DCHV | 979 | FE_LEBT:DCHV_D0979 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 370.624 | 423.0996 | 41.094 | 41.094 | 423.0996 | 370.5287 | 97.8893 | 97.9846 |  | 29.4107
0 | [7 empty] | 150 |  |  | 0.0254 | 370.516 | 423.0996 | 41.094 | 41.094 | 423.0996 | 370.5033 | 97.9973 | 98.01 |  | 29.5187
0 | FE | LEBT | SOLR | 982 | FE_LEBT:SOLR_D0982 | S4 | solenoid | 150 |  |  | 0.4 | 370.3033 | 423.0996 | 41.094 | 41.094 | 423.0996 | 370.1033 | 98.21 | 98.41 | 40cm effective length | 29.7314
0 | [6 empty] | mhb box & bellows | 150 |  |  | 0.22 | 369.9933 | 423.0996 | 41.094 | 41.094 | 423.0996 | 369.8833 | 98.52 | 98.63 |  | 30.0414
0 | [7 empty] | 150 |  |  | 0 | 369.8833 | 423.0996 | 41.094 | 41.094 | 423.0996 | 369.8833 | 98.63 | 98.63 |  | 30.1514
0 | FE | LEBT | PM | 986 | FE_LEBT:PM_D0986 | PM1 | profile monitor (small fork) | 150 |  |  | 0 | 369.8833 | 423.0996 | 41.094 | 41.094 | 423.0996 | 369.8833 | 98.63 | 98.63 |  | 30.1514
0 | [6 empty] | mhb box | 150 |  |  | 0.105 | 369.8308 | 423.0996 | 41.094 | 41.094 | 423.0996 | 369.7783 | 98.6825 | 98.735 |  | 30.2039
0 | FE | LEBT | CAV | 987 | FE_LEBT:CAV_D0987 |  | multi harmonic buncher | 150 |  |  | 0 | 369.7783 | 423.0996 | 41.094 | 41.094 | 423.0996 | 369.7783 | 98.735 | 98.735 |  | 30.2564
0 | [7 empty] | 150 |  |  | 0 | 369.7783 | 423.0996 | 41.094 | 41.094 | 423.0996 | 369.7783 | 98.735 | 98.735 |  | 30.2564
0 | [6 empty] | mhb box | 150 |  |  | 0 | 369.7783 | 423.0996 | 41.094 | 41.094 | 423.0996 | 369.7783 | 98.735 | 98.735 |  | 30.2564
0 | FE | LEBT | NEGP | 987 | FE_LEBT:NEGP_D0987 | NEG-0 | getter pump, 3500 l/s |  |  |  | 0 | 369.7783 | 423.0996 | 41.094 | 41.094 | 423.0996 | 369.7783 | 98.735 | 98.735 |  | 30.2564
0 | [7 empty] | 150 |  |  | 0 | 369.7783 | 423.0996 | 41.094 | 41.094 | 423.0996 | 369.7783 | 98.735 | 98.735 |  | 30.2564
0 | [7 empty] | 150 |  |  | 0 | 369.7783 | 423.0996 | 41.094 | 41.094 | 423.0996 | 369.7783 | 98.735 | 98.735 | Possible HV tube starts here | 30.2564
0 | [7 empty] | 150 |  |  | 0.14649 | 369.7051 | 423.0996 | 41.094 | 41.094 | 423.0996 | 369.6318 | 98.8082 | 98.8815 |  | 30.3296
0 | FE | LEBT | BCM | 989 | FE_LEBT:BCM_D0989 | BCM1 | AC current transformer | 150 | 50mm |  | 0 | 369.6318 | 423.0996 | 41.094 | 41.094 | 423.0996 | 369.6318 | 98.8815 | 98.8815 |  | 30.4029
0 | [7 empty] | 150 |  |  | 0.22261 | 369.5205 | 423.0996 | 41.094 | 41.094 | 423.0996 | 369.4092 | 98.9928 | 99.1041 |  | 30.5142
0 | FE | LEBT | DCHV | 992 | FE_LEBT:DCHV_D0992 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 369.314 | 423.0996 | 41.094 | 41.094 | 423.0996 | 369.2187 | 99.1993 | 99.2946 |  | 30.7207
0 | [7 empty] | 150 |  |  | 0.0254 | 369.206 | 423.0996 | 41.094 | 41.094 | 423.0996 | 369.1933 | 99.3073 | 99.32 | Possible HV tube ends in solenoid | 30.8287
0 | FE | LEBT | SOLR | 995 | FE_LEBT:SOLR_D0995 | S4b | solenoid | 150 |  |  | 0.4 | 368.9933 | 423.0996 | 41.094 | 41.094 | 423.0996 | 368.7933 | 99.52 | 99.72 | one side ISO flange connection | 31.0414
0 | [7 empty] | 40 |  |  | 0.075524 | 368.7556 | 423.0996 | 41.094 | 41.094 | 423.0996 | 368.7178 | 99.7578 | 99.7955 |  | 31.2791
0 | FE | LEBT | FC | 998 | FE_LEBT:FC_D0998 | FC1 | faraday cup | 40 |  |  | 0 | 368.7178 | 423.0996 | 41.094 | 41.094 | 423.0996 | 368.7178 | 99.7955 | 99.7955 | Not used currently | 31.3169
0 | FE | LEBT | FFC | 998 | FE_LEBT:FFC_D0998 | LPM1 | fast faraday cup | 40 |  |  | 0 | 368.7178 | 423.0996 | 41.094 | 41.094 | 423.0996 | 368.7178 | 99.7955 | 99.7955 |  | 31.3169
0 | [7 empty] | 40 |  |  | 0.023317 | 368.7062 | 423.0996 | 41.094 | 41.094 | 423.0996 | 368.6945 | 99.8072 | 99.8188 |  | 31.3286
0 | FE | LEBT | PM | 998 | FE_LEBT:PM_D0998 | PM1 | profile monitor (small fork) | 40 |  |  | 0 | 368.6945 | 423.0996 | 41.094 | 41.094 | 423.0996 | 368.6945 | 99.8188 | 99.8188 |  | 31.3402
0 | [7 empty] | 40 |  |  | 0.076159 | 368.6564 | 423.0996 | 41.094 | 41.094 | 423.0996 | 368.6183 | 99.8569 | 99.895 |  | 31.3783
0 | FE | LEBT | CAV | 999 | FE_LEBT:CAV_D0999 |  | velocity equalizer resonator | 40 |  |  | 0 | 368.6183 | 423.0996 | 41.094 | 41.094 | 423.0996 | 368.6183 | 99.895 | 99.895 | port available | 31.4164
0 | [7 empty] | 40 |  |  | 0.025336 | 368.6057 | 423.0996 | 41.094 | 41.094 | 423.0996 | 368.593 | 99.9077 | 99.9203 |  | 31.429
0 | FE | LEBT | GV | 999 | FE_LEBT:GV_D0999 | GV_C | gate valve w/ special clamp | 40 |  |  | 0.0347 | 368.5757 | 423.0996 | 41.094 | 41.094 | 423.0996 | 368.5583 | 99.9377 | 99.955 | insertable gate valve | 31.4591
0 | [10 empty] | 0.013215 | 368.5517 | 423.0996 | 41.094 | 41.094 | 423.0996 | 368.5451 | 99.9616 | 99.9682 |  | 31.483
0 | [10 empty] | 0 | 368.5451 | 423.0996 | 41.094 | 41.094 | 423.0996 | 368.5451 | 99.9682 | 99.9682 | moved | 31.4896
0 | FE | LEBT | end | [7 empty] | 0 | 368.5451 | 423.0996 | 41.094 | 41.094 | 423.0996 | 368.5451 | 99.9682 | 99.9682 |  | 31.4896
0 | [6 empty] | RFQ end wall |  |  |  | 0.03175 | 368.5292 | 423.0996 | 41.094 | 41.094 | 423.0996 | 368.5133 | 99.9841 | 100 |  | 31.5055
0 | [6 empty] | RFQ inn-wall (match point) |  |  |  | 0 | 368.5133 | 423.0996 | 41.094 | 41.094 | 423.0996 | 368.5133 | 100 | 100 |  | 31.5214
0 | FE | RFQ | CAV | 1005 | FE_RFQ:CAV_D1005 |  | RFQ tank -1 |  |  |  | 1.0082 | 368.0093 | 423.0996 | 41.094 | 41.094 | 423.0996 | 367.5052 | 100.5041 | 101.0082 |  | 32.0254
0 | FE | RFQ | IP | 1005 | FE_RFQ:IP_D1005 | NEG-I | getter-ion pump 2000-10 l/s |  |  |  | 0 | 367.5052 | 423.0996 | 41.094 | 41.094 | 423.0996 | 367.5052 | 101.0082 | 101.0082 |  | 32.5295
0 | FE | RFQ | CAV | 1015 | FE_RFQ:CAV_D1015 |  | RFQ tank -2 |  |  |  | 1.0082 | 367.0011 | 423.0996 | 41.094 | 41.094 | 423.0996 | 366.497 | 101.5122 | 102.0163 |  | 33.0336
0 | FE | RFQ | IP | 1015 | FE_RFQ:IP_D1015 | NEG-I | getter-ion pump 2000-10 l/s |  |  |  | 0 | 366.497 | 423.0996 | 41.094 | 41.094 | 423.0996 | 366.497 | 102.0163 | 102.0163 |  | 33.5377
0 | FE | RFQ | CAV | 1025 | FE_RFQ:CAV_D1025 |  | RFQ tank -3 |  |  |  | 1.0082 | 365.993 | 423.0996 | 41.094 | 41.094 | 423.0996 | 365.4889 | 102.5204 | 103.0245 |  | 34.0418
0 | FE | RFQ | TMP | 1025 | FE_RFQ:TMP_D1025 | TMP_M | pump, 700 l/s |  |  |  | 0 | 365.4889 | 423.0996 | 41.094 | 41.094 | 423.0996 | 365.4889 | 103.0245 | 103.0245 |  | 34.5458
0 | FE | RFQ | CAV | 1035 | FE_RFQ:CAV_D1035 |  | RFQ tank -4 |  |  |  | 1.0082 | 364.9848 | 423.0996 | 41.094 | 41.094 | 423.0996 | 364.4807 | 103.5285 | 104.0326 |  | 35.0499
0 | FE | RFQ | IP | 1035 | FE_RFQ:IP_D1035 | NEG-I | getter-ion pump 2000-10 l/s |  |  |  | 0 | 364.4807 | 423.0996 | 41.094 | 41.094 | 423.0996 | 364.4807 | 104.0326 | 104.0326 |  | 35.554
0 | FE | RFQ | CAV | 1045 | FE_RFQ:CAV_D1045 |  | RFQ tank -5 |  |  |  | 1.0082 | 363.9766 | 423.0996 | 41.094 | 41.094 | 423.0996 | 363.4726 | 104.5367 | 105.0408 |  | 36.0581
0 | FE | RFQ | IP | 1045 | FE_RFQ:IP_D1045 | NEG-I | getter-ion pump 2000-10 l/s |  |  |  | 0 | 363.4726 | 423.0996 | 41.094 | 41.094 | 423.0996 | 363.4726 | 105.0408 | 105.0408 |  | 36.5621
0 | [6 empty] | RFQ inn-wall |  |  |  | 0 | 363.4726 | 423.0996 | 41.094 | 41.094 | 423.0996 | 363.4726 | 105.0408 | 105.0408 |  | 36.5621
0 | [6 empty] | RFQ end wall |  |  |  | 0.03175 | 363.4567 | 423.0996 | 41.094 | 41.094 | 423.0996 | 363.4408 | 105.0567 | 105.0725 |  | 36.578
0 | FE | MEBT | start | [7 empty] | 0 | 363.4408 | 423.0996 | 41.094 | 41.094 | 423.0996 | 363.4408 | 105.0725 | 105.0725 |  | 36.5939
0 | [6 empty] | bellow | 30 |  |  | 0.057907 | 363.4119 | 423.0996 | 41.094 | 41.094 | 423.0996 | 363.3829 | 105.1015 | 105.1304 |  | 36.6228
0 | FE | MEBT | GV | 1052 | FE_MEBT:GV_D1052 | GV_S | gate valve | 30 |  |  | 0.060706 | 363.3526 | 423.0996 | 41.094 | 41.094 | 423.0996 | 363.3222 | 105.1608 | 105.1911 |  | 36.6822
0 | [10 empty] | 0.13118 | 363.2566 | 423.0996 | 41.094 | 41.094 | 423.0996 | 363.191 | 105.2567 | 105.3223 |  | 36.7781
0 | FE | MEBT | PM | 1053 | FE_MEBT:PM_D1053 | PM1 | profile monitor (small fork) | 30 |  |  | 0 | 363.191 | 423.0996 | 41.094 | 41.094 | 423.0996 | 363.191 | 105.3223 | 105.3223 | this PM was removed per Gerlind K. 1/10/2017 | 36.8437
0 | FE | MEBT | IP | 1053 | FE_MEBT:IP_D1053 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 363.191 | 423.0996 | 41.094 | 41.094 | 423.0996 | 363.191 | 105.3223 | 105.3223 |  | 36.8437
0 | [7 empty] | 30 |  |  | 0.12291 | 363.1296 | 423.0996 | 41.094 | 41.094 | 423.0996 | 363.0681 | 105.3838 | 105.4452 |  | 36.9051
0 | FE | MEBT | BCM | 1055 | FE_MEBT:BCM_D1055 | BCM1 | AC current transformer | 30 |  |  | 0.04 | 363.0481 | 423.0996 | 41.094 | 41.094 | 423.0996 | 363.0281 | 105.4652 | 105.4852 |  | 36.9866
0 | [6 empty] | BPM bellow | 30 |  |  | 0.1035 | 362.9764 | 423.0996 | 41.094 | 41.094 | 423.0996 | 362.9246 | 105.537 | 105.5887 |  | 37.0584
0 | FE | MEBT | BPM | 1056 | FE_MEBT:BPM_D1056 | BPM1 | position monitor | 40 |  |  | 0 | 362.9246 | 423.0996 | 41.094 | 41.094 | 423.0996 | 362.9246 | 105.5887 | 105.5887 |  | 37.1101
0 | [7 empty] | 30 |  |  | 0.10663 | 362.8713 | 423.0996 | 41.094 | 41.094 | 423.0996 | 362.818 | 105.642 | 105.6954 |  | 37.1634
0 | FE | MEBT | Q | 1057 | FE_MEBT:Q_D1057 | Q8 | quad, triplet 1-1 | 47.5 |  |  | 0.1 | 362.768 | 423.0996 | 41.094 | 41.094 | 423.0996 | 362.718 | 105.7454 | 105.7954 |  | 37.2667
0 | [7 empty] | 47.5 |  |  | 0.1 | 362.668 | 423.0996 | 41.094 | 41.094 | 423.0996 | 362.618 | 105.8454 | 105.8954 |  | 37.3667
0 | FE | MEBT | Q | 1060 | FE_MEBT:Q_D1060 | Q9 | quad, triplet 1-2 | 47.5 |  |  | 0.15 | 362.543 | 423.0996 | 41.094 | 41.094 | 423.0996 | 362.468 | 105.9704 | 106.0454 |  | 37.4917
0 | [7 empty] | 47.5 |  |  | 0.1 | 362.418 | 423.0996 | 41.094 | 41.094 | 423.0996 | 362.368 | 106.0954 | 106.1454 |  | 37.6167
0 | FE | MEBT | Q | 1062 | FE_MEBT:Q_D1062 | Q10 | quad+steerers, triplet 1-3 | 47.5 |  |  | 0.1 | 362.318 | 423.0996 | 41.094 | 41.094 | 423.0996 | 362.268 | 106.1954 | 106.2454 |  | 37.7167
0 | FE | MEBT | DCHV | 1062 | FE_MEBT:DCHV_D1062 | Q10 | quad+steerers, triplet 1-3 | 47.5 |  |  | 0 | 362.268 | 423.0996 | 41.094 | 41.094 | 423.0996 | 362.268 | 106.2454 | 106.2454 |  | 37.7667
0 | [6 empty] | bellow? | 47.5/30 |  |  | 0.24092 | 362.1475 | 423.0996 | 41.094 | 41.094 | 423.0996 | 362.0271 | 106.3658 | 106.4863 |  | 37.8872
0 | FE | MEBT | CAV | 1066 | FE_MEBT:CAV_D1066 |  | rebuncher | 30 |  |  | 0.32 | 361.8671 | 423.0996 | 41.094 | 41.094 | 423.0996 | 361.7071 | 106.6463 | 106.8063 |  | 38.1676
0 | FE | MEBT | IP | 1068 | FE_MEBT:IP_D1066 | IP_L | ion pump, 300 l/s |  |  |  | 0 | 361.7071 | 423.0996 | 41.094 | 41.094 | 423.0996 | 361.7071 | 106.8063 | 106.8063 |  | 38.3276
0 | [6 empty] | bellow+tube ?? | 30 |  |  | 0.22222 | 361.596 | 423.0996 | 41.094 | 41.094 | 423.0996 | 361.4848 | 106.9174 | 107.0285 |  | 38.4388
0 | [6 empty] | 6 way cross ?? | 30 |  |  | 0 | 361.4848 | 423.0996 | 41.094 | 41.094 | 423.0996 | 361.4848 | 107.0285 | 107.0285 |  | 38.5499
0 | FE | MEBT | IP | 1070 | FE_MEBT:IP_D1070 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 361.4848 | 423.0996 | 41.094 | 41.094 | 423.0996 | 361.4848 | 107.0285 | 107.0285 |  | 38.5499
0 | FE | MEBT | FAVS | 1070 | FE_MEBT:FAVS_D1070 | FVS | fast valve sensor |  |  |  | 0 | 361.4848 | 423.0996 | 41.094 | 41.094 | 423.0996 | 361.4848 | 107.0285 | 107.0285 |  | 38.5499
0 | [10 empty] | 0.21444 | 361.3776 | 423.0996 | 41.094 | 41.094 | 423.0996 | 361.2704 | 107.1357 | 107.2429 |  | 38.6571
0 | FE | MEBT | BPM | 1072 | FE_MEBT:BPM_D1072 | BPM1 | position monitor | 40 |  |  | 0 | 361.2704 | 423.0996 | 41.094 | 41.094 | 423.0996 | 361.2704 | 107.2429 | 107.2429 |  | 38.7643
0 | [7 empty] | 30 |  |  | 0.10663 | 361.2171 | 423.0996 | 41.094 | 41.094 | 423.0996 | 361.1638 | 107.2963 | 107.3496 |  | 38.8176
0 | FE | MEBT | Q | 1074 | FE_MEBT:Q_D1074 | Q8 | quad, triplet 2-1 | 47.5 |  |  | 0.1 | 361.1138 | 423.0996 | 41.094 | 41.094 | 423.0996 | 361.0638 | 107.3996 | 107.4496 |  | 38.9209
0 | [7 empty] | 47.5 |  |  | 0.1 | 361.0138 | 423.0996 | 41.094 | 41.094 | 423.0996 | 360.9638 | 107.4996 | 107.5496 |  | 39.0209
0 | FE | MEBT | Q | 1076 | FE_MEBT:Q_D1076 | Q9 | quad, triplet 2-2 | 47.5 |  |  | 0.15 | 360.8888 | 423.0996 | 41.094 | 41.094 | 423.0996 | 360.8138 | 107.6246 | 107.6996 |  | 39.1459
0 | [7 empty] | 47.5 |  |  | 0.1 | 360.7638 | 423.0996 | 41.094 | 41.094 | 423.0996 | 360.7138 | 107.7496 | 107.7996 |  | 39.2709
0 | FE | MEBT | Q | 1078 | FE_MEBT:Q_D1078 | Q10 | quad+steerers, triplet 2-3 | 47.5 |  |  | 0.1 | 360.6638 | 423.0996 | 41.094 | 41.094 | 423.0996 | 360.6138 | 107.8496 | 107.8996 |  | 39.3709
0 | FE | MEBT | DCHV | 1078 | FE_MEBT:DCHV_D1078 | Q10 | quad+steerers, triplet 2-3 | 47.5 |  |  | 0 | 360.6138 | 423.0996 | 41.094 | 41.094 | 423.0996 | 360.6138 | 107.8996 | 107.8996 |  | 39.4209
0 | [10 empty] | 0 | 360.6138 | 423.0996 | 41.094 | 41.094 | 423.0996 | 360.6138 | 107.8996 | 107.8996 |  | 39.4209
0 | [10 empty] | 0.69723 | 360.2652 | 423.0996 | 41.094 | 41.094 | 423.0996 | 359.9165 | 108.2482 | 108.5968 |  | 39.7696
0 | FE | MEBT | DH | 1086 | FE_MEBT:DH_D1086 | D6 | dipole-45deg, D6 (straight leg) | 40 |  |  | 0 | 359.9165 | 423.0996 | 41.094 | 41.094 | 423.0996 | 359.9165 | 108.5968 | 108.5968 |  | 40.1182
0 | [10 empty] | 0.562 | 359.6355 | 423.0996 | 41.094 | 41.094 | 423.0996 | 359.3545 | 108.8778 | 109.1588 |  | 40.3992
0 | FE | MEBT | PM | 1092 | FE_MEBT:PM_D1092 | PM1 | profile monitor (small fork) | 30 ? |  |  | 0 | 359.3545 | 423.0996 | 41.094 | 41.094 | 423.0996 | 359.3545 | 109.1588 | 109.1588 |  | 40.6802
0 | FE | MEBT | IP | 1092 | FE_MEBT:IP_D1092 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 359.3545 | 423.0996 | 41.094 | 41.094 | 423.0996 | 359.3545 | 109.1588 | 109.1588 |  | 40.6802
0 | [6 empty] | bellow ? |  |  |  | 0.22641 | 359.2413 | 423.0996 | 41.094 | 41.094 | 423.0996 | 359.1281 | 109.272 | 109.3852 |  | 40.7934
0 | FE | MEBT | BPM | 1094 | FE_MEBT:BPM_D1094 | BPM1 | position monitor | 40 |  |  | 0 | 359.1281 | 423.0996 | 41.094 | 41.094 | 423.0996 | 359.1281 | 109.3852 | 109.3852 |  | 40.9066
0 | [7 empty] | 40 |  |  | 0.10663 | 359.0748 | 423.0996 | 41.094 | 41.094 | 423.0996 | 359.0215 | 109.4385 | 109.4918 |  | 40.9599
0 | [10 empty] | 0 | 359.0215 | 423.0996 | 41.094 | 41.094 | 423.0996 | 359.0215 | 109.4918 | 109.4918 |  | 41.0132
0 | FE | MEBT | Q | 1095 | FE_MEBT:Q_D1095 | Q8 | quad, triplet 3-1 | 47.5 |  |  | 0.1 | 358.9715 | 423.0996 | 41.094 | 41.094 | 423.0996 | 358.9215 | 109.5418 | 109.5918 |  | 41.0632
0 | [7 empty] | 47.5 |  |  | 0.1 | 358.8715 | 423.0996 | 41.094 | 41.094 | 423.0996 | 358.8215 | 109.6418 | 109.6918 |  | 41.1632
0 | FE | MEBT | Q | 1098 | FE_MEBT:Q_D1098 | Q9 | quad, triplet 3-2 | 47.5 |  |  | 0.15 | 358.7465 | 423.0996 | 41.094 | 41.094 | 423.0996 | 358.6715 | 109.7668 | 109.8418 |  | 41.2882
0 | [7 empty] | 47.5 |  |  | 0.1 | 358.6215 | 423.0996 | 41.094 | 41.094 | 423.0996 | 358.5715 | 109.8918 | 109.9418 |  | 41.4132
0 | FE | MEBT | Q | 1100 | FE_MEBT:Q_D1100 | Q10 | quad+steerers, triplet 3-3 | 47.5 |  |  | 0.1 | 358.5215 | 423.0996 | 41.094 | 41.094 | 423.0996 | 358.4715 | 109.9918 | 110.0418 |  | 41.5132
0 | FE | MEBT | DCHV | 1100 | FE_MEBT:DCHV_D1100 | Q10 | quad+steerers, triplet 3-3 | 47.5 |  |  | 0 | 358.4715 | 423.0996 | 41.094 | 41.094 | 423.0996 | 358.4715 | 110.0418 | 110.0418 |  | 41.5632
0 | [10 empty] | 0.17064 | 358.3862 | 423.0996 | 41.094 | 41.094 | 423.0996 | 358.3009 | 110.1272 | 110.2125 |  | 41.6485
0 | FE | MEBT | FC | 1102 | FE_MEBT:FC_D1102 | FC3 | beam stopper / faraday cup | 40 |  |  | 0 | 358.3009 | 423.0996 | 41.094 | 41.094 | 423.0996 | 358.3009 | 110.2125 | 110.2125 |  | 41.7339
0 | FE | MEBT | FFC | 1102 | FE_MEBT:FFC_D1102 | LPM1 | fast faraday cup | 40 |  |  | 0 | 358.3009 | 423.0996 | 41.094 | 41.094 | 423.0996 | 358.3009 | 110.2125 | 110.2125 |  | 41.7339
0 | FE | MEBT | IP | 1102 | FE_MEBT:IP_D1102 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 358.3009 | 423.0996 | 41.094 | 41.094 | 423.0996 | 358.3009 | 110.2125 | 110.2125 |  | 41.7339
0 | [10 empty] | 0.11303 | 358.2443 | 423.0996 | 41.094 | 41.094 | 423.0996 | 358.1878 | 110.269 | 110.3255 |  | 41.7904
0 | FE | MEBT | GV | 1104 | FE_MEBT:GV_D1104 | GV_S | gate valve | 30 ? |  |  | 0.060711 | 358.1575 | 423.0996 | 41.094 | 41.094 | 423.0996 | 358.1271 | 110.3559 | 110.3862 |  | 41.8772
0 | [10 empty] | 0.18331 | 358.0355 | 423.0996 | 41.094 | 41.094 | 423.0996 | 357.9438 | 110.4779 | 110.5695 |  | 41.9992
0 | FE | MEBT | CAV | 1107 | FE_MEBT:CAV_D1107 |  | rebuncher | 30 |  |  | 0.32 | 357.7838 | 423.0996 | 41.094 | 41.094 | 423.0996 | 357.6238 | 110.7295 | 110.8895 |  | 42.2509
0 | FE | MEBT | IP | 1109 | FE_MEBT:IP_D1107 | IP_L | ion pump, 300 l/s |  |  |  | 0 | 357.6238 | 423.0996 | 41.094 | 41.094 | 423.0996 | 357.6238 | 110.8895 | 110.8895 |  | 42.4109
0 | [6 empty] | bellow+tube | 40 |  |  | 0.21062 | 357.5185 | 423.0996 | 41.094 | 41.094 | 423.0996 | 357.4132 | 110.9948 | 111.1001 |  | 42.5162
0 | FE | MEBT | BPM | 1111 | FE_MEBT:BPM_D1111 | BPM1 | position monitor | 40 |  |  | 0 | 357.4132 | 423.0996 | 41.094 | 41.094 | 423.0996 | 357.4132 | 111.1001 | 111.1001 |  | 42.6215
0 | [7 empty] | 40 |  |  | 0.10663 | 357.3599 | 423.0996 | 41.094 | 41.094 | 423.0996 | 357.3066 | 111.1535 | 111.2068 |  | 42.6748
0 | [10 empty] | 0 | 357.3066 | 423.0996 | 41.094 | 41.094 | 423.0996 | 357.3066 | 111.2068 | 111.2068 |  | 42.7281
0 | FE | MEBT | Q | 1113 | FE_MEBT:Q_D1113 | Q8 | quad, triplet 4-1 | 47.5 |  |  | 0.1 | 357.2566 | 423.0996 | 41.094 | 41.094 | 423.0996 | 357.2066 | 111.2568 | 111.3068 |  | 42.7781
0 | [7 empty] | 47.5 |  |  | 0.1 | 357.1566 | 423.0996 | 41.094 | 41.094 | 423.0996 | 357.1066 | 111.3568 | 111.4068 |  | 42.8781
0 | FE | MEBT | Q | 1115 | FE_MEBT:Q_D1115 | Q9 | quad, triplet 4-2 | 47.5 |  |  | 0.15 | 357.0316 | 423.0996 | 41.094 | 41.094 | 423.0996 | 356.9566 | 111.4818 | 111.5568 |  | 43.0031
0 | [7 empty] | 47.5 |  |  | 0.1 | 356.9066 | 423.0996 | 41.094 | 41.094 | 423.0996 | 356.8566 | 111.6068 | 111.6568 |  | 43.1281
0 | FE | MEBT | Q | 1117 | FE_MEBT:Q_D1117 | Q10 | quad+steerers, triplet 4-3 | 47.5 |  |  | 0.1 | 356.8066 | 423.0996 | 41.094 | 41.094 | 423.0996 | 356.7566 | 111.7068 | 111.7568 |  | 43.2281
0 | FE | MEBT | DCHV | 1117 | FE_MEBT:DCHV_D1117 | Q10 | quad+steerers, triplet 4-3 | 47.5 |  |  | 0 | 356.7566 | 423.0996 | 41.094 | 41.094 | 423.0996 | 356.7566 | 111.7568 | 111.7568 |  | 43.2781
0 | [6 empty] | bellow | 40 |  |  | 0.05761 | 356.7278 | 423.0996 | 41.094 | 41.094 | 423.0996 | 356.699 | 111.7856 | 111.8144 |  | 43.307
0 | FE | MEBT | FAV | 1118 | FE_MEBT:FAV_D1118 | FV | fast valve | 40 |  |  | 0.056602 | 356.6707 | 423.0996 | 41.094 | 41.094 | 423.0996 | 356.6424 | 111.8427 | 111.871 |  | 43.3641
0 | [10 empty] | 0 | 356.6424 | 423.0996 | 41.094 | 41.094 | 423.0996 | 356.6424 | 111.871 | 111.871 |  | 43.3924
0 | [10 empty] | 0.11748 | 356.5836 | 423.0996 | 41.094 | 41.094 | 423.0996 | 356.5249 | 111.9297 | 111.9885 |  | 43.4511
0 | [10 empty] | 0 | 356.5249 | 423.0996 | 41.094 | 41.094 | 423.0996 | 356.5249 | 111.9885 | 111.9885 |  | 43.5098
0 | FE | MEBT | BCM | 1120 | FE_MEBT:BCM_D1120 | BCM1 | AC current transformer | 40 |  |  | 0.04 | 356.5049 | 423.0996 | 41.094 | 41.094 | 423.0996 | 356.4849 | 112.0085 | 112.0285 |  | 43.5298
0 | [10 empty] | 0 | 356.4849 | 423.0996 | 41.094 | 41.094 | 423.0996 | 356.4849 | 112.0285 | 112.0285 |  | 43.5498
0 | [10 empty] | 0.11303 | 356.4284 | 423.0996 | 41.094 | 41.094 | 423.0996 | 356.3718 | 112.085 | 112.1415 |  | 43.6064
0 | FE | MEBT | IP | 1121 | FE_MEBT:IP_D1121 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 356.3718 | 423.0996 | 41.094 | 41.094 | 423.0996 | 356.3718 | 112.1415 | 112.1415 |  | 43.6629
0 | [7 empty] | 40 |  |  | 0.2305 | 356.2566 | 423.0996 | 41.094 | 41.094 | 423.0996 | 356.1413 | 112.2567 | 112.372 |  | 43.7781
0 | [10 empty] | -2.63e-06 | 356.1413 | 423.0996 | 41.094 | 41.094 | 423.0996 | 356.1413 | 112.372 | 112.372 | adjustment for residuals | 43.8934
0 | FE | MEBT | END | [7 empty] | 0 | 356.1413 | 423.0996 | 41.094 | 41.094 | 423.0996 | 356.1413 | 112.372 | 112.372 |  | 43.8934
0 | SEGMENT 1 START (begin with b=0.041 QWR cryomodules) | [9 empty] | 0 | 356.1413 | 423.0996 | 41.094 | 41.094 | 423.0996 | 356.1413 | 112.372 | 112.372 |  | 43.8934
0 | LS1 | CA01 | GV | 1124 | LS1_CA01:GV_D1124 |  | gate valve | 40 |  |  | 0.057404 | 356.1126 | 423.0996 | 41.094 | 41.094 | 423.0996 | 356.0839 | 112.4007 | 112.4294 |  | 43.9221
0 | [6 empty] | bellows | 40 |  |  | 0.189 | 355.9894 | 423.0996 | 41.094 | 41.094 | 423.0996 | 355.8949 | 112.5239 | 112.6184 |  | 44.0453
0 | LS1 | CA01 | CAV1 | 1127 | LS1_CA01:CAV1_D1127 |  | b04 resonator | 36 |  |  | 0.18336 | 355.8033 | 423.0996 | 41.094 | 41.094 | 423.0996 | 355.7116 | 112.7101 | 112.8018 |  | 44.2315
0 | [10 empty] | 0 | 355.7116 | 423.0996 | 41.094 | 41.094 | 423.0996 | 355.7116 | 112.8018 | 112.8018 |  | 44.3231
0 | LS1 | CA01 | BLM | 1128 | LS1_CA01:BLM_D1128 |  | outside loss monitor |  |  |  | 0 | 355.7116 | 423.0996 | 41.094 | 41.094 | 423.0996 | 355.7116 | 112.8018 | 112.8018 |  | 44.3231
0 | [6 empty] | bellow | 40 |  |  | 0.09651 | 355.6633 | 423.0996 | 41.094 | 41.094 | 423.0996 | 355.6151 | 112.85 | 112.8983 |  | 44.3714
0 | LS1 | CA01 | BPM | 1129 | LS1_CA01:BPM_D1129 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 355.6151 | 423.0996 | 41.094 | 41.094 | 423.0996 | 355.6151 | 112.8983 | 112.8983 |  | 44.4196
0 | [6 empty] | BPM-box |  |  |  | 0.05588 | 355.5871 | 423.0996 | 41.094 | 41.094 | 423.0996 | 355.5592 | 112.9262 | 112.9542 |  | 44.4476
0 | [6 empty] | solenoid-entry | 40 |  |  | 0.078206 | 355.5201 | 423.0996 | 41.094 | 41.094 | 423.0996 | 355.481 | 112.9933 | 113.0324 | 20170302 space adjusted to match Jlab design | 44.5146
0 | LS1 | CA01 | SOL1 | 1132 | LS1_CA01:SOL1_D1132 | S1 | solenoid+correctors | 40 |  |  | 0.25 | 355.356 | 423.0996 | 41.094 | 41.094 | 423.0996 | 355.231 | 113.1574 | 113.2824 |  | 44.6787
0 | [6 empty] | solenoid-exit | 40 |  |  | 0.078206 | 355.1919 | 423.0996 | 41.094 | 41.094 | 423.0996 | 355.1528 | 113.3215 | 113.3606 | 20170302 space adjusted to match Jlab design | 44.8428
0 | [6 empty] | bellow | 40 |  |  | 0.099517 | 355.103 | 423.0996 | 41.094 | 41.094 | 423.0996 | 355.0533 | 113.4103 | 113.4601 |  | 44.9317
0 | LS1 | CA01 | CAV2 | 1136 | LS1_CA01:CAV2_D1136 |  | b04 resonator | 36 |  |  | 0.18336 | 354.9616 | 423.0996 | 41.094 | 41.094 | 423.0996 | 354.8699 | 113.5518 | 113.6434 |  | 45.0731
0 | [10 empty] | 0 | 354.8699 | 423.0996 | 41.094 | 41.094 | 423.0996 | 354.8699 | 113.6434 | 113.6434 |  | 45.1648
0 | [10 empty] | 0 | 354.8699 | 423.0996 | 41.094 | 41.094 | 423.0996 | 354.8699 | 113.6434 | 113.6434 |  | 45.1648
0 | [6 empty] | 2 bellows + tube | 36 |  |  | 0.45019 | 354.6448 | 423.0996 | 41.094 | 41.094 | 423.0996 | 354.4197 | 113.8685 | 114.0936 | 20170302 space adjusted to match Jlab design | 45.3899
0 | [10 empty] | 0 | 354.4197 | 423.0996 | 41.094 | 41.094 | 423.0996 | 354.4197 | 114.0936 | 114.0936 |  | 45.615
0 | LS1 | CA01 | CAV3 | 1142 | LS1_CA01:CAV3_D1142 |  | b04 resonator | 36 |  |  | 0.18336 | 354.328 | 423.0996 | 41.094 | 41.094 | 423.0996 | 354.2363 | 114.1853 | 114.277 |  | 45.7067
0 | [10 empty] | 0 | 354.2363 | 423.0996 | 41.094 | 41.094 | 423.0996 | 354.2363 | 114.277 | 114.277 |  | 45.7984
0 | LS1 | CA01 | BLM | 1143 | LS1_CA01:BLM_D1143 |  | outside loss monitor |  |  |  | 0 | 354.2363 | 423.0996 | 41.094 | 41.094 | 423.0996 | 354.2363 | 114.277 | 114.277 |  | 45.7984
0 | [6 empty] | bellow |  |  |  | 0.09651 | 354.1881 | 423.0996 | 41.094 | 41.094 | 423.0996 | 354.1398 | 114.3253 | 114.3735 |  | 45.8466
0 | LS1 | CA01 | BPM | 1144 | LS1_CA01:BPM_D1144 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 354.1398 | 423.0996 | 41.094 | 41.094 | 423.0996 | 354.1398 | 114.3735 | 114.3735 |  | 45.8949
0 | [6 empty] | BPM-box |  |  |  | 0.05588 | 354.1119 | 423.0996 | 41.094 | 41.094 | 423.0996 | 354.0839 | 114.4015 | 114.4294 |  | 45.9228
0 | [6 empty] | solenoid-entry |  |  |  | 0.078206 | 354.0448 | 423.0996 | 41.094 | 41.094 | 423.0996 | 354.0057 | 114.4685 | 114.5076 | 20170302 space adjusted to match Jlab design | 45.9899
0 | LS1 | CA01 | SOL2 | 1146 | LS1_CA01:SOL2_D1146 | S1 | solenoid+correctors | 40 |  |  | 0.25 | 353.8807 | 423.0996 | 41.094 | 41.094 | 423.0996 | 353.7557 | 114.6326 | 114.7576 |  | 46.154
0 | [6 empty] | solenoid-exit |  |  |  | 0.078206 | 353.7166 | 423.0996 | 41.094 | 41.094 | 423.0996 | 353.6775 | 114.7967 | 114.8358 | 20170302 space adjusted to match Jlab design | 46.3181
0 | [6 empty] | bellow |  |  |  | 0.099517 | 353.6278 | 423.0996 | 41.094 | 41.094 | 423.0996 | 353.578 | 114.8856 | 114.9353 |  | 46.4069
0 | LS1 | CA01 | CAV4 | 1150 | LS1_CA01:CAV4_D1150 |  | b04 resonator | 36 |  |  | 0.18336 | 353.4863 | 423.0996 | 41.094 | 41.094 | 423.0996 | 353.3947 | 115.027 | 115.1187 |  | 46.5484
0 | [10 empty] | 0 | 353.3947 | 423.0996 | 41.094 | 41.094 | 423.0996 | 353.3947 | 115.1187 | 115.1187 |  | 46.6401
0 | [10 empty] | 0 | 353.3947 | 423.0996 | 41.094 | 41.094 | 423.0996 | 353.3947 | 115.1187 | 115.1187 |  | 46.6401
0 | [6 empty] | bellows |  |  |  | 0.189 | 353.3002 | 423.0996 | 41.094 | 41.094 | 423.0996 | 353.2057 | 115.2132 | 115.3077 |  | 46.7346
0 | LS1 | CA01 | GV | 1153 | LS1_CA01:GV_D1153 |  | gate valve | 40 |  |  | 0.057404 | 353.177 | 423.0996 | 41.094 | 41.094 | 423.0996 | 353.1483 | 115.3364 | 115.3651 |  | 46.8578
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 353.0965 | 423.0996 | 41.094 | 41.094 | 423.0996 | 353.0447 | 115.4168 | 115.4686 |  | 46.9382
0 | LS1 | WA01 | BPM | 1155 | LS1_WA01:BPM_D1155 | BPM2 | position monitor |  |  |  | 0 | 353.0447 | 423.0996 | 41.094 | 41.094 | 423.0996 | 353.0447 | 115.4686 | 115.4686 |  | 46.99
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 352.9721 | 423.0996 | 41.094 | 41.094 | 423.0996 | 352.8995 | 115.5412 | 115.6139 |  | 47.0626
0 | LS1 | WA01 | HMR | 1156 | LS1_WA01:HMR_D1156 | HR_L | Halo ring, large aperture | 33 |  |  | 0 | 352.8995 | 423.0996 | 41.094 | 41.094 | 423.0996 | 352.8995 | 115.6139 | 115.6139 |  | 47.1352
0 | [6 empty] | diagnostic box |  |  |  | 0 | 352.8995 | 423.0996 | 41.094 | 41.094 | 423.0996 | 352.8995 | 115.6139 | 115.6139 |  | 47.1352
0 | LS1 | WA01 | IP | 1156 | LS1_WA01:IP_D1156 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 352.8995 | 423.0996 | 41.094 | 41.094 | 423.0996 | 352.8995 | 115.6139 | 115.6139 |  | 47.1352
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 352.8339 | 423.0996 | 41.094 | 41.094 | 423.0996 | 352.7683 | 115.6795 | 115.7451 |  | 47.2009
0 | LS1 | CA02 | GV | 1158 | LS1_CA02:GV_D1158 |  | gate valve | 40 |  |  | 0.057404 | 352.7395 | 423.0996 | 41.094 | 41.094 | 423.0996 | 352.7108 | 115.7738 | 115.8025 |  | 47.2952
0 | [6 empty] | bellows | 40 |  |  | 0.189 | 352.6163 | 423.0996 | 41.094 | 41.094 | 423.0996 | 352.5218 | 115.897 | 115.9915 |  | 47.4184
0 | LS1 | CA02 | CAV1 | 1161 | LS1_CA02:CAV1_D1161 |  | b04 resonator | 36 |  |  | 0.18336 | 352.4302 | 423.0996 | 41.094 | 41.094 | 423.0996 | 352.3385 | 116.0832 | 116.1749 |  | 47.6045
0 | [10 empty] | 0 | 352.3385 | 423.0996 | 41.094 | 41.094 | 423.0996 | 352.3385 | 116.1749 | 116.1749 |  | 47.6962
0 | LS1 | CA02 | BLM | 1162 | LS1_CA02:BLM_D1162 |  | outside loss monitor |  |  |  | 0 | 352.3385 | 423.0996 | 41.094 | 41.094 | 423.0996 | 352.3385 | 116.1749 | 116.1749 |  | 47.6962
0 | [6 empty] | bellow | 40 |  |  | 0.09651 | 352.2902 | 423.0996 | 41.094 | 41.094 | 423.0996 | 352.242 | 116.2231 | 116.2714 |  | 47.7445
0 | LS1 | CA02 | BPM | 1163 | LS1_CA02:BPM_D1163 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 352.242 | 423.0996 | 41.094 | 41.094 | 423.0996 | 352.242 | 116.2714 | 116.2714 |  | 47.7927
0 | [6 empty] | BPM-box |  |  |  | 0.05588 | 352.214 | 423.0996 | 41.094 | 41.094 | 423.0996 | 352.1861 | 116.2993 | 116.3272 |  | 47.8207
0 | [6 empty] | solenoid-entry | 40 |  |  | 0.078206 | 352.147 | 423.0996 | 41.094 | 41.094 | 423.0996 | 352.1079 | 116.3664 | 116.4055 | 20170302 space adjusted to match Jlab design | 47.8877
0 | LS1 | CA02 | SOL1 | 1165 | LS1_CA02:SOL1_D1165 | S1 | solenoid+correctors | 40 |  |  | 0.25 | 351.9829 | 423.0996 | 41.094 | 41.094 | 423.0996 | 351.8579 | 116.5305 | 116.6555 |  | 48.0518
0 | [6 empty] | solenoid-exit | 40 |  |  | 0.078206 | 351.8188 | 423.0996 | 41.094 | 41.094 | 423.0996 | 351.7797 | 116.6946 | 116.7337 | 20170302 space adjusted to match Jlab design | 48.2159
0 | [6 empty] | bellow | 40 |  |  | 0.099517 | 351.7299 | 423.0996 | 41.094 | 41.094 | 423.0996 | 351.6802 | 116.7834 | 116.8332 |  | 48.3048
0 | LS1 | CA02 | CAV2 | 1169 | LS1_CA02:CAV2_D1169 |  | b04 resonator | 36 |  |  | 0.18336 | 351.5885 | 423.0996 | 41.094 | 41.094 | 423.0996 | 351.4968 | 116.9249 | 117.0165 |  | 48.4462
0 | [10 empty] | 0 | 351.4968 | 423.0996 | 41.094 | 41.094 | 423.0996 | 351.4968 | 117.0165 | 117.0165 |  | 48.5379
0 | [10 empty] | 0 | 351.4968 | 423.0996 | 41.094 | 41.094 | 423.0996 | 351.4968 | 117.0165 | 117.0165 |  | 48.5379
0 | [6 empty] | 2 bellows + tube | 36 |  |  | 0.45019 | 351.2717 | 423.0996 | 41.094 | 41.094 | 423.0996 | 351.0466 | 117.2416 | 117.4667 | 20170302 space adjusted to match Jlab design | 48.763
0 | [10 empty] | 0 | 351.0466 | 423.0996 | 41.094 | 41.094 | 423.0996 | 351.0466 | 117.4667 | 117.4667 |  | 48.9881
0 | LS1 | CA02 | CAV3 | 1176 | LS1_CA02:CAV3_D1176 |  | b04 resonator | 36 |  |  | 0.18336 | 350.9549 | 423.0996 | 41.094 | 41.094 | 423.0996 | 350.8632 | 117.5584 | 117.6501 |  | 49.0798
0 | [10 empty] | 0 | 350.8632 | 423.0996 | 41.094 | 41.094 | 423.0996 | 350.8632 | 117.6501 | 117.6501 |  | 49.1715
0 | LS1 | CA02 | BLM | 1177 | LS1_CA02:BLM_D1177 |  | outside loss monitor |  |  |  | 0 | 350.8632 | 423.0996 | 41.094 | 41.094 | 423.0996 | 350.8632 | 117.6501 | 117.6501 |  | 49.1715
0 | [6 empty] | bellow |  |  |  | 0.09651 | 350.815 | 423.0996 | 41.094 | 41.094 | 423.0996 | 350.7667 | 117.6984 | 117.7466 |  | 49.2197
0 | LS1 | CA02 | BPM | 1177 | LS1_CA02:BPM_D1177 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 350.7667 | 423.0996 | 41.094 | 41.094 | 423.0996 | 350.7667 | 117.7466 | 117.7466 |  | 49.268
0 | [6 empty] | BPM-box |  |  |  | 0.05588 | 350.7388 | 423.0996 | 41.094 | 41.094 | 423.0996 | 350.7109 | 117.7745 | 117.8025 |  | 49.2959
0 | [6 empty] | solenoid-entry |  |  |  | 0.078206 | 350.6718 | 423.0996 | 41.094 | 41.094 | 423.0996 | 350.6326 | 117.8416 | 117.8807 | 20170302 space adjusted to match Jlab design | 49.363
0 | LS1 | CA02 | SOL2 | 1180 | LS1_CA02:SOL2_D1180 | S1 | solenoid+correctors | 40 |  |  | 0.25 | 350.5076 | 423.0996 | 41.094 | 41.094 | 423.0996 | 350.3826 | 118.0057 | 118.1307 |  | 49.5271
0 | [6 empty] | solenoid-exit |  |  |  | 0.078206 | 350.3435 | 423.0996 | 41.094 | 41.094 | 423.0996 | 350.3044 | 118.1698 | 118.2089 | 20170302 space adjusted to match Jlab design | 49.6912
0 | [6 empty] | bellow |  |  |  | 0.099517 | 350.2547 | 423.0996 | 41.094 | 41.094 | 423.0996 | 350.2049 | 118.2587 | 118.3084 |  | 49.78
0 | LS1 | CA02 | CAV4 | 1184 | LS1_CA02:CAV4_D1184 |  | b04 resonator | 36 |  |  | 0.18336 | 350.1132 | 423.0996 | 41.094 | 41.094 | 423.0996 | 350.0216 | 118.4001 | 118.4918 |  | 49.9215
0 | [10 empty] | 0 | 350.0216 | 423.0996 | 41.094 | 41.094 | 423.0996 | 350.0216 | 118.4918 | 118.4918 |  | 50.0132
0 | [10 empty] | 0 | 350.0216 | 423.0996 | 41.094 | 41.094 | 423.0996 | 350.0216 | 118.4918 | 118.4918 |  | 50.0132
0 | [6 empty] | bellows |  |  |  | 0.189 | 349.9271 | 423.0996 | 41.094 | 41.094 | 423.0996 | 349.8326 | 118.5863 | 118.6808 |  | 50.1077
0 | LS1 | CA02 | GV | 1187 | LS1_CA02:GV_D1187 |  | gate valve | 40 |  |  | 0.057404 | 349.8039 | 423.0996 | 41.094 | 41.094 | 423.0996 | 349.7752 | 118.7095 | 118.7382 |  | 50.2309
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 349.7234 | 423.0996 | 41.094 | 41.094 | 423.0996 | 349.6717 | 118.7899 | 118.8417 |  | 50.3113
0 | LS1 | WA02 | BPM | 1188 | LS1_WA02:BPM_D1188 | BPM2 | position monitor |  |  |  | 0 | 349.6717 | 423.0996 | 41.094 | 41.094 | 423.0996 | 349.6717 | 118.8417 | 118.8417 |  | 50.3631
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 349.599 | 423.0996 | 41.094 | 41.094 | 423.0996 | 349.5264 | 118.9143 | 118.987 |  | 50.4357
0 | LS1 | WA02 | HMR | 1190 | LS1_WA02:HMR_D1190 | HR_L | Halo ring, large aperture | 33 |  |  | 0 | 349.5264 | 423.0996 | 41.094 | 41.094 | 423.0996 | 349.5264 | 118.987 | 118.987 |  | 50.5083
0 | [6 empty] | diagnostic box |  |  |  | 0 | 349.5264 | 423.0996 | 41.094 | 41.094 | 423.0996 | 349.5264 | 118.987 | 118.987 |  | 50.5083
0 | LS1 | WA02 | IP | 1190 | LS1_WA02:IP_D1190 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 349.5264 | 423.0996 | 41.094 | 41.094 | 423.0996 | 349.5264 | 118.987 | 118.987 |  | 50.5083
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 349.4608 | 423.0996 | 41.094 | 41.094 | 423.0996 | 349.3952 | 119.0526 | 119.1182 |  | 50.5739
0 | LS1 | CA03 | GV | 1191 | LS1_CA03:GV_D1191 |  | gate valve | 40 |  |  | 0.057404 | 349.3665 | 423.0996 | 41.094 | 41.094 | 423.0996 | 349.3378 | 119.1469 | 119.1756 |  | 50.6683
0 | [6 empty] | bellows | 40 |  |  | 0.189 | 349.2433 | 423.0996 | 41.094 | 41.094 | 423.0996 | 349.1488 | 119.2701 | 119.3646 |  | 50.7915
0 | LS1 | CA03 | CAV1 | 1195 | LS1_CA03:CAV1_D1195 |  | b04 resonator | 36 |  |  | 0.18336 | 349.0571 | 423.0996 | 41.094 | 41.094 | 423.0996 | 348.9654 | 119.4563 | 119.548 |  | 50.9776
0 | [10 empty] | 0 | 348.9654 | 423.0996 | 41.094 | 41.094 | 423.0996 | 348.9654 | 119.548 | 119.548 |  | 51.0693
0 | LS1 | CA03 | BLM | 1195 | LS1_CA03:BLM_D1195 |  | outside loss monitor |  |  |  | 0 | 348.9654 | 423.0996 | 41.094 | 41.094 | 423.0996 | 348.9654 | 119.548 | 119.548 |  | 51.0693
0 | [6 empty] | bellow | 40 |  |  | 0.09651 | 348.9171 | 423.0996 | 41.094 | 41.094 | 423.0996 | 348.8689 | 119.5962 | 119.6445 |  | 51.1176
0 | LS1 | CA03 | BPM | 1196 | LS1_CA03:BPM_D1196 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 348.8689 | 423.0996 | 41.094 | 41.094 | 423.0996 | 348.8689 | 119.6445 | 119.6445 |  | 51.1658
0 | [6 empty] | BPM-box |  |  |  | 0.05588 | 348.8409 | 423.0996 | 41.094 | 41.094 | 423.0996 | 348.813 | 119.6724 | 119.7003 |  | 51.1938
0 | [6 empty] | solenoid-entry | 40 |  |  | 0.078206 | 348.7739 | 423.0996 | 41.094 | 41.094 | 423.0996 | 348.7348 | 119.7394 | 119.7785 | 20170302 space adjusted to match Jlab design | 51.2608
0 | LS1 | CA03 | SOL1 | 1199 | LS1_CA03:SOL1_D1199 | S1 | solenoid+correctors | 40 |  |  | 0.25 | 348.6098 | 423.0996 | 41.094 | 41.094 | 423.0996 | 348.4848 | 119.9035 | 120.0285 |  | 51.4249
0 | [6 empty] | solenoid-exit | 40 |  |  | 0.078206 | 348.4457 | 423.0996 | 41.094 | 41.094 | 423.0996 | 348.4066 | 120.0677 | 120.1068 | 20170302 space adjusted to match Jlab design | 51.589
0 | [6 empty] | bellow | 40 |  |  | 0.099517 | 348.3568 | 423.0996 | 41.094 | 41.094 | 423.0996 | 348.3071 | 120.1565 | 120.2063 |  | 51.6779
0 | LS1 | CA03 | CAV2 | 1203 | LS1_CA03:CAV2_D1203 |  | b04 resonator | 36 |  |  | 0.18336 | 348.2154 | 423.0996 | 41.094 | 41.094 | 423.0996 | 348.1237 | 120.298 | 120.3896 |  | 51.8193
0 | [10 empty] | 0 | 348.1237 | 423.0996 | 41.094 | 41.094 | 423.0996 | 348.1237 | 120.3896 | 120.3896 |  | 51.911
0 | [10 empty] | 0 | 348.1237 | 423.0996 | 41.094 | 41.094 | 423.0996 | 348.1237 | 120.3896 | 120.3896 |  | 51.911
0 | [6 empty] | 2 bellows + tube | 36 |  |  | 0.45019 | 347.8986 | 423.0996 | 41.094 | 41.094 | 423.0996 | 347.6735 | 120.6147 | 120.8398 | 20170302 space adjusted to match Jlab design | 52.1361
0 | [10 empty] | 0 | 347.6735 | 423.0996 | 41.094 | 41.094 | 423.0996 | 347.6735 | 120.8398 | 120.8398 |  | 52.3612
0 | LS1 | CA03 | CAV3 | 1209 | LS1_CA03:CAV3_D1209 |  | b04 resonator | 36 |  |  | 0.18336 | 347.5818 | 423.0996 | 41.094 | 41.094 | 423.0996 | 347.4901 | 120.9315 | 121.0232 |  | 52.4529
0 | [10 empty] | 0 | 347.4901 | 423.0996 | 41.094 | 41.094 | 423.0996 | 347.4901 | 121.0232 | 121.0232 |  | 52.5446
0 | LS1 | CA03 | BLM | 1210 | LS1_CA03:BLM_D1210 |  | outside loss monitor |  |  |  | 0 | 347.4901 | 423.0996 | 41.094 | 41.094 | 423.0996 | 347.4901 | 121.0232 | 121.0232 |  | 52.5446
0 | [6 empty] | bellow |  |  |  | 0.09651 | 347.4419 | 423.0996 | 41.094 | 41.094 | 423.0996 | 347.3936 | 121.0714 | 121.1197 |  | 52.5928
0 | LS1 | CA03 | BPM | 1211 | LS1_CA03:BPM_D1211 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 347.3936 | 423.0996 | 41.094 | 41.094 | 423.0996 | 347.3936 | 121.1197 | 121.1197 |  | 52.6411
0 | [6 empty] | BPM-box |  |  |  | 0.05588 | 347.3657 | 423.0996 | 41.094 | 41.094 | 423.0996 | 347.3378 | 121.1476 | 121.1756 |  | 52.669
0 | [6 empty] | solenoid-entry |  |  |  | 0.078206 | 347.2987 | 423.0996 | 41.094 | 41.094 | 423.0996 | 347.2596 | 121.2147 | 121.2538 | 20170302 space adjusted to match Jlab design | 52.7361
0 | LS1 | CA03 | SOL2 | 1214 | LS1_CA03:SOL2_D1214 | S1 | solenoid+correctors | 40 |  |  | 0.25 | 347.1346 | 423.0996 | 41.094 | 41.094 | 423.0996 | 347.0096 | 121.3788 | 121.5038 |  | 52.9002
0 | [6 empty] | solenoid-exit |  |  |  | 0.078206 | 346.9704 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.9313 | 121.5429 | 121.582 | 20170302 space adjusted to match Jlab design | 53.0643
0 | [6 empty] | bellow |  |  |  | 0.099517 | 346.8816 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.8318 | 121.6318 | 121.6815 |  | 53.1531
0 | LS1 | CA03 | CAV4 | 1218 | LS1_CA03:CAV4_D1218 |  | b04 resonator | 36 |  |  | 0.18336 | 346.7401 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.6485 | 121.7732 | 121.8649 |  | 53.2946
0 | [10 empty] | 0 | 346.6485 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.6485 | 121.8649 | 121.8649 |  | 53.3862
0 | [10 empty] | 0 | 346.6485 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.6485 | 121.8649 | 121.8649 |  | 53.3862
0 | [6 empty] | bellows |  |  |  | 0.189 | 346.554 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.4595 | 121.9594 | 122.0539 |  | 53.4807
0 | LS1 | CA03 | GV | 1221 | LS1_CA03:GV_D1221 |  | gate valve | 40 |  |  | 0.057404 | 346.4308 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.4021 | 122.0826 | 122.1113 |  | 53.6039
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 346.3503 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.2986 | 122.163 | 122.2148 |  | 53.6844
0 | LS1 | WA03 | BPM | 1222 | LS1_WA03:BPM_D1222 | BPM2 | position monitor |  |  |  | 0 | 346.2986 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.2986 | 122.2148 | 122.2148 |  | 53.7362
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 346.2259 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.1533 | 122.2874 | 122.3601 |  | 53.8088
0 | LS1 | WA03 | HMR | 1224 | LS1_WA03:HMR_D1224 | HR_L | Halo ring, large aperture | 33 |  |  | 0 | 346.1533 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.1533 | 122.3601 | 122.3601 | 346.1533 | 53.8814
0 | LS1 | WA03 | BD | 1224 | LS1_WA03:BD_D1224 | dump | Beam dump with nibium plate |  |  |  | 0 | 346.1533 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.1533 | 122.3601 | 122.3601 |  | 53.8814
0 | [6 empty] | diagnostic box |  |  |  | 0 | 346.1533 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.1533 | 122.3601 | 122.3601 |  | 53.8814
0 | LS1 | WA03 | IP | 1224 | LS1_WA03:IP_D1224 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 346.1533 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.1533 | 122.3601 | 122.3601 |  | 53.8814
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 346.0877 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.0221 | 122.4257 | 122.4913 |  | 53.947
0 | D-station starts in "V" column
0 | beta=0.085 QWR cryomodules START | [9 empty] | 0 | 346.0221 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.0221 | 122.4913 | 122.4913 |  | 54.0127
0 | LS1 | CB01 | GV | 1225 | LS1_CB01:GV_D1225 |  | gate valve | 40 |  |  | 0.057404 | 345.9934 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.9647 | 122.52 | 122.5487 |  | 54.0414
0 | [6 empty] | bellows |  |  |  | 0.19082 | 345.8692 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.7738 | 122.6441 | 122.7395 |  | 54.1655
0 | LS1 | CB01 | CAV1 | 1229 | LS1_CB01:CAV1_D1229 |  | b08 resonator | 34 |  |  | 0.29817 | 345.6247 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.4757 | 122.8886 | 123.0377 |  | 54.41
0 | [10 empty] | 0 | 345.4757 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.4757 | 123.0377 | 123.0377 |  | 54.559
0 | LS1 | CB01 | BLM | 1230 | LS1_CB01:BLM_D1230 |  | outside loss monitor |  |  |  | 0 | 345.4757 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.4757 | 123.0377 | 123.0377 |  | 54.559
0 | [6 empty] | bellow | 40 |  |  | 0.09651 | 345.4274 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.3792 | 123.0859 | 123.1342 |  | 54.6073
0 | LS1 | CB01 | BPM | 1231 | LS1_CB01:BPM_D1231 | BPM6 | position monitor, BPMc |  |  |  | 0 | 345.3792 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.3792 | 123.1342 | 123.1342 |  | 54.6556
0 | [6 empty] | BPM-box | 40 |  |  | 0.05588 | 345.3512 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.3233 | 123.1621 | 123.1901 |  | 54.6835
0 | [6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 345.2852 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.2471 | 123.2281 | 123.2662 |  | 54.7495
0 | LS1 | CB01 | SOL1 | 1235 | LS1_CB01:SOL1_D1235 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 344.9971 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.7471 | 123.5162 | 123.7662 |  | 55.0376
0 | [6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 344.7091 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.671 | 123.8043 | 123.8423 |  | 55.3256
0 | [6 empty] | bellow |  |  |  | 0.099517 | 344.6212 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.5715 | 123.8921 | 123.9419 |  | 55.4135
0 | LS1 | CB01 | CAV2 | 1241 | LS1_CB01:CAV2_D1241 |  | b08 resonator | 34 |  |  | 0.29817 | 344.4224 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.2733 | 124.0909 | 124.24 |  | 55.6123
0 | [10 empty] | 0 | 344.2733 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.2733 | 124.24 | 124.24 |  | 55.7614
0 | [10 empty] | 0 | 344.2733 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.2733 | 124.24 | 124.24 |  | 55.7614
0 | [6 empty] | bellow |  |  |  | 0.099517 | 344.2236 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.1738 | 124.2898 | 124.3395 |  | 55.8112
0 | LS1 | CB01 | CAV3 | 1245 | LS1_CB01:CAV3_D1245 |  | b08 resonator | 34 |  |  | 0.29817 | 344.0247 | 423.0996 | 41.094 | 41.094 | 423.0996 | 343.8756 | 124.4886 | 124.6377 |  | 56.01
0 | [10 empty] | 0 | 343.8756 | 423.0996 | 41.094 | 41.094 | 423.0996 | 343.8756 | 124.6377 | 124.6377 |  | 56.1591
0 | [10 empty] | 0 | 343.8756 | 423.0996 | 41.094 | 41.094 | 423.0996 | 343.8756 | 124.6377 | 124.6377 |  | 56.1591
0 | [6 empty] | bellow |  |  |  | 0.099517 | 343.8259 | 423.0996 | 41.094 | 41.094 | 423.0996 | 343.7761 | 124.6875 | 124.7372 |  | 56.2088
0 | LS1 | CB01 | CAV4 | 1249 | LS1_CB01:CAV4_D1249 |  | b08 resonator | 34 |  |  | 0.29817 | 343.627 | 423.0996 | 41.094 | 41.094 | 423.0996 | 343.4779 | 124.8863 | 125.0354 |  | 56.4077
0 | [10 empty] | 0 | 343.4779 | 423.0996 | 41.094 | 41.094 | 423.0996 | 343.4779 | 125.0354 | 125.0354 |  | 56.5568
0 | LS1 | CB01 | BLM | 1250 | LS1_CB01:BLM_D1250 |  | outside loss monitor |  |  |  | 0 | 343.4779 | 423.0996 | 41.094 | 41.094 | 423.0996 | 343.4779 | 125.0354 | 125.0354 |  | 56.5568
0 | [6 empty] | bellow |  |  |  | 0.09651 | 343.4297 | 423.0996 | 41.094 | 41.094 | 423.0996 | 343.3814 | 125.0837 | 125.1319 |  | 56.605
0 | LS1 | CB01 | BPM | 1251 | LS1_CB01:BPM_D1251 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 343.3814 | 423.0996 | 41.094 | 41.094 | 423.0996 | 343.3814 | 125.1319 | 125.1319 |  | 56.6533
0 | [6 empty] | BPM-box |  |  |  | 0.05588 | 343.3535 | 423.0996 | 41.094 | 41.094 | 423.0996 | 343.3255 | 125.1599 | 125.1878 |  | 56.6812
0 | [6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 343.2875 | 423.0996 | 41.094 | 41.094 | 423.0996 | 343.2494 | 125.2259 | 125.2639 |  | 56.7472
0 | LS1 | CB01 | SOL2 | 1255 | LS1_CB01:SOL2_D1255 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 342.9994 | 423.0996 | 41.094 | 41.094 | 423.0996 | 342.7494 | 125.5139 | 125.7639 |  | 57.0353
0 | [6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 342.7113 | 423.0996 | 41.094 | 41.094 | 423.0996 | 342.6733 | 125.802 | 125.8401 |  | 57.3234
0 | [6 empty] | bellow |  |  |  | 0.099517 | 342.6235 | 423.0996 | 41.094 | 41.094 | 423.0996 | 342.5738 | 125.8898 | 125.9396 |  | 57.4112
0 | LS1 | CB01 | CAV5 | 1261 | LS1_CB01:CAV5_D1261 |  | b08 resonator | 34 |  |  | 0.29817 | 342.4247 | 423.0996 | 41.094 | 41.094 | 423.0996 | 342.2756 | 126.0887 | 126.2378 |  | 57.61
0 | [10 empty] | 0 | 342.2756 | 423.0996 | 41.094 | 41.094 | 423.0996 | 342.2756 | 126.2378 | 126.2378 |  | 57.7591
0 | [10 empty] | 0 | 342.2756 | 423.0996 | 41.094 | 41.094 | 423.0996 | 342.2756 | 126.2378 | 126.2378 |  | 57.7591
0 | [6 empty] | bellow |  |  |  | 0.099517 | 342.2258 | 423.0996 | 41.094 | 41.094 | 423.0996 | 342.1761 | 126.2875 | 126.3373 |  | 57.8089
0 | LS1 | CB01 | CAV6 | 1265 | LS1_CB01:CAV6_D1265 |  | b08 resonator | 34 |  |  | 0.29817 | 342.027 | 423.0996 | 41.094 | 41.094 | 423.0996 | 341.8779 | 126.4864 | 126.6354 |  | 58.0077
0 | [10 empty] | 0 | 341.8779 | 423.0996 | 41.094 | 41.094 | 423.0996 | 341.8779 | 126.6354 | 126.6354 |  | 58.1568
0 | [10 empty] | 0 | 341.8779 | 423.0996 | 41.094 | 41.094 | 423.0996 | 341.8779 | 126.6354 | 126.6354 |  | 58.1568
0 | [6 empty] | bellow |  |  |  | 0.099517 | 341.8281 | 423.0996 | 41.094 | 41.094 | 423.0996 | 341.7784 | 126.6852 | 126.735 |  | 58.2066
0 | LS1 | CB01 | CAV7 | 1269 | LS1_CB01:CAV7_D1269 |  | b08 resonator | 34 |  |  | 0.29817 | 341.6293 | 423.0996 | 41.094 | 41.094 | 423.0996 | 341.4802 | 126.884 | 127.0331 |  | 58.4054
0 | [10 empty] | 0 | 341.4802 | 423.0996 | 41.094 | 41.094 | 423.0996 | 341.4802 | 127.0331 | 127.0331 |  | 58.5545
0 | LS1 | CB01 | BLM | 1270 | LS1_CB01:BLM_D1270 |  | outside loss monitor |  |  |  | 0 | 341.4802 | 423.0996 | 41.094 | 41.094 | 423.0996 | 341.4802 | 127.0331 | 127.0331 |  | 58.5545
0 | [6 empty] | bellow |  |  |  | 0.09651 | 341.432 | 423.0996 | 41.094 | 41.094 | 423.0996 | 341.3837 | 127.0814 | 127.1296 |  | 58.6028
0 | LS1 | CB01 | BPM | 1271 | LS1_CB01:BPM_D1271 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 341.3837 | 423.0996 | 41.094 | 41.094 | 423.0996 | 341.3837 | 127.1296 | 127.1296 |  | 58.651
0 | [6 empty] | BPM-box |  |  |  | 0.05588 | 341.3558 | 423.0996 | 41.094 | 41.094 | 423.0996 | 341.3278 | 127.1576 | 127.1855 |  | 58.6789
0 | [6 empty] | solenoid-entry |  |  |  | 0.076136 | 341.2898 | 423.0996 | 41.094 | 41.094 | 423.0996 | 341.2517 | 127.2236 | 127.2617 |  | 58.745
0 | LS1 | CB01 | SOL3 | 1275 | LS1_CB01:SOL3_D1275 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 341.0017 | 423.0996 | 41.094 | 41.094 | 423.0996 | 340.7517 | 127.5117 | 127.7617 |  | 59.033
0 | [6 empty] | solenoid-exit |  |  |  | 0.076136 | 340.7136 | 423.0996 | 41.094 | 41.094 | 423.0996 | 340.6756 | 127.7997 | 127.8378 |  | 59.3211
0 | [6 empty] | bellow |  |  |  | 0.099517 | 340.6258 | 423.0996 | 41.094 | 41.094 | 423.0996 | 340.576 | 127.8875 | 127.9373 |  | 59.4089
0 | LS1 | CB01 | CAV8 | 1281 | LS1_CB01:CAV8_D1281 |  | b08 resonator | 34 |  |  | 0.29817 | 340.4269 | 423.0996 | 41.094 | 41.094 | 423.0996 | 340.2779 | 128.0864 | 128.2355 |  | 59.6078
0 | [10 empty] | 0 | 340.2779 | 423.0996 | 41.094 | 41.094 | 423.0996 | 340.2779 | 128.2355 | 128.2355 |  | 59.7568
0 | [10 empty] | 0 | 340.2779 | 423.0996 | 41.094 | 41.094 | 423.0996 | 340.2779 | 128.2355 | 128.2355 |  | 59.7568
0 | [6 empty] | bellows |  |  |  | 0.19082 | 340.1825 | 423.0996 | 41.094 | 41.094 | 423.0996 | 340.087 | 128.3309 | 128.4263 |  | 59.8523
0 | LS1 | CB01 | GV | 1285 | LS1_CB01:GV_D1285 |  | gate valve | 40 |  |  | 0.057404 | 340.0583 | 423.0996 | 41.094 | 41.094 | 423.0996 | 340.0296 | 128.455 | 128.4837 |  | 59.9764
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 339.9779 | 423.0996 | 41.094 | 41.094 | 423.0996 | 339.9261 | 128.5355 | 128.5872 |  | 60.0568
0 | LS1 | WB01 | BPM | 1286 | LS1_WB01:BPM_D1286 | BPM2 | position monitor |  |  |  | 0 | 339.9261 | 423.0996 | 41.094 | 41.094 | 423.0996 | 339.9261 | 128.5872 | 128.5872 |  | 60.1086
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 339.8535 | 423.0996 | 41.094 | 41.094 | 423.0996 | 339.7809 | 128.6598 | 128.7325 |  | 60.1812
0 | LS1 | WB01 | HMR | 1287 | LS1_WB01:HMR_D1287 | HR_L | Halo ring, large aperture | 33 |  |  | 0 | 339.7809 | 423.0996 | 41.094 | 41.094 | 423.0996 | 339.7809 | 128.7325 | 128.7325 |  | 60.2539
0 | [6 empty] | diagnostic box |  |  |  | 0 | 339.7809 | 423.0996 | 41.094 | 41.094 | 423.0996 | 339.7809 | 128.7325 | 128.7325 |  | 60.2539
0 | LS1 | WB01 | IP | 1287 | LS1_WB01:IP_D1287 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 339.7809 | 423.0996 | 41.094 | 41.094 | 423.0996 | 339.7809 | 128.7325 | 128.7325 |  | 60.2539
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 339.7152 | 423.0996 | 41.094 | 41.094 | 423.0996 | 339.6496 | 128.7981 | 128.8637 |  | 60.3195
0 | LS1 | CB02 | GV | 1289 | LS1_CB02:GV_D1289 |  | gate valve | 40 |  |  | 0.057404 | 339.6209 | 423.0996 | 41.094 | 41.094 | 423.0996 | 339.5922 | 128.8924 | 128.9211 |  | 60.4138
0 | [6 empty] | bellows |  |  |  | 0.19082 | 339.4968 | 423.0996 | 41.094 | 41.094 | 423.0996 | 339.4014 | 129.0165 | 129.1119 |  | 60.5379
0 | LS1 | CB02 | CAV1 | 1293 | LS1_CB02:CAV1_D1293 |  | b08 resonator | 34 |  |  | 0.29817 | 339.2523 | 423.0996 | 41.094 | 41.094 | 423.0996 | 339.1032 | 129.261 | 129.4101 |  | 60.7824
0 | [10 empty] | 0 | 339.1032 | 423.0996 | 41.094 | 41.094 | 423.0996 | 339.1032 | 129.4101 | 129.4101 |  | 60.9315
0 | LS1 | CB02 | BLM | 1294 | LS1_CB02:BLM_D1294 |  | outside loss monitor |  |  |  | 0 | 339.1032 | 423.0996 | 41.094 | 41.094 | 423.0996 | 339.1032 | 129.4101 | 129.4101 |  | 60.9315
0 | [6 empty] | bellow | 40 |  |  | 0.09651 | 339.055 | 423.0996 | 41.094 | 41.094 | 423.0996 | 339.0067 | 129.4584 | 129.5066 |  | 60.9797
0 | LS1 | CB02 | BPM | 1295 | LS1_CB02:BPM_D1295 | BPM6 | position monitor, BPMc |  |  |  | 0 | 339.0067 | 423.0996 | 41.094 | 41.094 | 423.0996 | 339.0067 | 129.5066 | 129.5066 |  | 61.028
0 | [6 empty] | BPM-box | 40 |  |  | 0.05588 | 338.9788 | 423.0996 | 41.094 | 41.094 | 423.0996 | 338.9508 | 129.5346 | 129.5625 |  | 61.0559
0 | [6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 338.9128 | 423.0996 | 41.094 | 41.094 | 423.0996 | 338.8747 | 129.6006 | 129.6386 |  | 61.1219
0 | LS1 | CB02 | SOL1 | 1299 | LS1_CB02:SOL1_D1299 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 338.6247 | 423.0996 | 41.094 | 41.094 | 423.0996 | 338.3747 | 129.8886 | 130.1386 |  | 61.41
0 | [6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 338.3366 | 423.0996 | 41.094 | 41.094 | 423.0996 | 338.2986 | 130.1767 | 130.2148 |  | 61.6981
0 | [6 empty] | bellow |  |  |  | 0.099517 | 338.2488 | 423.0996 | 41.094 | 41.094 | 423.0996 | 338.1991 | 130.2645 | 130.3143 |  | 61.7859
0 | LS1 | CB02 | CAV2 | 1305 | LS1_CB02:CAV2_D1305 |  | b08 resonator | 34 |  |  | 0.29817 | 338.05 | 423.0996 | 41.094 | 41.094 | 423.0996 | 337.9009 | 130.4634 | 130.6125 |  | 61.9847
0 | [10 empty] | 0 | 337.9009 | 423.0996 | 41.094 | 41.094 | 423.0996 | 337.9009 | 130.6125 | 130.6125 |  | 62.1338
0 | [10 empty] | 0 | 337.9009 | 423.0996 | 41.094 | 41.094 | 423.0996 | 337.9009 | 130.6125 | 130.6125 |  | 62.1338
0 | [6 empty] | bellow |  |  |  | 0.099517 | 337.8511 | 423.0996 | 41.094 | 41.094 | 423.0996 | 337.8014 | 130.6622 | 130.712 |  | 62.1836
0 | LS1 | CB02 | CAV3 | 1309 | LS1_CB02:CAV3_D1309 |  | b08 resonator | 34 |  |  | 0.29817 | 337.6523 | 423.0996 | 41.094 | 41.094 | 423.0996 | 337.5032 | 130.8611 | 131.0101 |  | 62.3824
0 | [10 empty] | 0 | 337.5032 | 423.0996 | 41.094 | 41.094 | 423.0996 | 337.5032 | 131.0101 | 131.0101 |  | 62.5315
0 | [10 empty] | 0 | 337.5032 | 423.0996 | 41.094 | 41.094 | 423.0996 | 337.5032 | 131.0101 | 131.0101 |  | 62.5315
0 | [6 empty] | bellow |  |  |  | 0.099517 | 337.4534 | 423.0996 | 41.094 | 41.094 | 423.0996 | 337.4037 | 131.0599 | 131.1097 |  | 62.5813
0 | LS1 | CB02 | CAV4 | 1313 | LS1_CB02:CAV4_D1313 |  | b08 resonator | 34 |  |  | 0.29817 | 337.2546 | 423.0996 | 41.094 | 41.094 | 423.0996 | 337.1055 | 131.2587 | 131.4078 |  | 62.7801
0 | [10 empty] | 0 | 337.1055 | 423.0996 | 41.094 | 41.094 | 423.0996 | 337.1055 | 131.4078 | 131.4078 |  | 62.9292
0 | LS1 | CB02 | BLM | 1314 | LS1_CB02:BLM_D1314 |  | outside loss monitor |  |  |  | 0 | 337.1055 | 423.0996 | 41.094 | 41.094 | 423.0996 | 337.1055 | 131.4078 | 131.4078 |  | 62.9292
0 | [6 empty] | bellow |  |  |  | 0.09651 | 337.0573 | 423.0996 | 41.094 | 41.094 | 423.0996 | 337.009 | 131.4561 | 131.5043 |  | 62.9775
0 | LS1 | CB02 | BPM | 1315 | LS1_CB02:BPM_D1315 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 337.009 | 423.0996 | 41.094 | 41.094 | 423.0996 | 337.009 | 131.5043 | 131.5043 |  | 63.0257
0 | [6 empty] | BPM-box |  |  |  | 0.05588 | 336.9811 | 423.0996 | 41.094 | 41.094 | 423.0996 | 336.9531 | 131.5323 | 131.5602 |  | 63.0536
0 | [6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 336.9151 | 423.0996 | 41.094 | 41.094 | 423.0996 | 336.877 | 131.5983 | 131.6364 |  | 63.1197
0 | LS1 | CB02 | SOL2 | 1319 | LS1_CB02:SOL2_D1319 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 336.627 | 423.0996 | 41.094 | 41.094 | 423.0996 | 336.377 | 131.8864 | 132.1364 |  | 63.4077
0 | [6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 336.3389 | 423.0996 | 41.094 | 41.094 | 423.0996 | 336.3009 | 132.1744 | 132.2125 |  | 63.6958
0 | [6 empty] | bellow |  |  |  | 0.099517 | 336.2511 | 423.0996 | 41.094 | 41.094 | 423.0996 | 336.2013 | 132.2622 | 132.312 |  | 63.7836
0 | LS1 | CB02 | CAV5 | 1325 | LS1_CB02:CAV5_D1325 |  | b08 resonator | 34 |  |  | 0.29817 | 336.0522 | 423.0996 | 41.094 | 41.094 | 423.0996 | 335.9032 | 132.4611 | 132.6102 |  | 63.9825
0 | [10 empty] | 0 | 335.9032 | 423.0996 | 41.094 | 41.094 | 423.0996 | 335.9032 | 132.6102 | 132.6102 |  | 64.1315
0 | [10 empty] | 0 | 335.9032 | 423.0996 | 41.094 | 41.094 | 423.0996 | 335.9032 | 132.6102 | 132.6102 |  | 64.1315
0 | [6 empty] | bellow |  |  |  | 0.099517 | 335.8534 | 423.0996 | 41.094 | 41.094 | 423.0996 | 335.8036 | 132.6599 | 132.7097 |  | 64.1813
0 | LS1 | CB02 | CAV6 | 1329 | LS1_CB02:CAV6_D1329 |  | b08 resonator | 34 |  |  | 0.29817 | 335.6546 | 423.0996 | 41.094 | 41.094 | 423.0996 | 335.5055 | 132.8588 | 133.0079 |  | 64.3802
0 | [10 empty] | 0 | 335.5055 | 423.0996 | 41.094 | 41.094 | 423.0996 | 335.5055 | 133.0079 | 133.0079 |  | 64.5292
0 | [10 empty] | 0 | 335.5055 | 423.0996 | 41.094 | 41.094 | 423.0996 | 335.5055 | 133.0079 | 133.0079 |  | 64.5292
0 | [6 empty] | bellow |  |  |  | 0.099517 | 335.4557 | 423.0996 | 41.094 | 41.094 | 423.0996 | 335.406 | 133.0576 | 133.1074 |  | 64.579
0 | LS1 | CB02 | CAV7 | 1333 | LS1_CB02:CAV7_D1333 |  | b08 resonator | 34 |  |  | 0.29817 | 335.2569 | 423.0996 | 41.094 | 41.094 | 423.0996 | 335.1078 | 133.2565 | 133.4056 |  | 64.7778
0 | [10 empty] | 0 | 335.1078 | 423.0996 | 41.094 | 41.094 | 423.0996 | 335.1078 | 133.4056 | 133.4056 |  | 64.9269
0 | LS1 | CB02 | BLM | 1334 | LS1_CB02:BLM_D1334 |  | outside loss monitor |  |  |  | 0 | 335.1078 | 423.0996 | 41.094 | 41.094 | 423.0996 | 335.1078 | 133.4056 | 133.4056 |  | 64.9269
0 | [6 empty] | bellow |  |  |  | 0.09651 | 335.0595 | 423.0996 | 41.094 | 41.094 | 423.0996 | 335.0113 | 133.4538 | 133.5021 |  | 64.9752
0 | LS1 | CB02 | BPM | 1335 | LS1_CB02:BPM_D1335 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 335.0113 | 423.0996 | 41.094 | 41.094 | 423.0996 | 335.0113 | 133.5021 | 133.5021 |  | 65.0234
0 | [6 empty] | BPM-box |  |  |  | 0.05588 | 334.9833 | 423.0996 | 41.094 | 41.094 | 423.0996 | 334.9554 | 133.53 | 133.5579 |  | 65.0514
0 | [6 empty] | solenoid-entry |  |  |  | 0.076136 | 334.9173 | 423.0996 | 41.094 | 41.094 | 423.0996 | 334.8793 | 133.596 | 133.6341 |  | 65.1174
0 | LS1 | CB02 | SOL3 | 1339 | LS1_CB02:SOL3_D1339 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 334.6293 | 423.0996 | 41.094 | 41.094 | 423.0996 | 334.3793 | 133.8841 | 134.1341 |  | 65.4054
0 | [6 empty] | solenoid-exit |  |  |  | 0.076136 | 334.3412 | 423.0996 | 41.094 | 41.094 | 423.0996 | 334.3031 | 134.1721 | 134.2102 |  | 65.6935
0 | [6 empty] | bellow |  |  |  | 0.099517 | 334.2534 | 423.0996 | 41.094 | 41.094 | 423.0996 | 334.2036 | 134.26 | 134.3097 |  | 65.7813
0 | LS1 | CB02 | CAV8 | 1345 | LS1_CB02:CAV8_D1345 |  | b08 resonator | 34 |  |  | 0.29817 | 334.0545 | 423.0996 | 41.094 | 41.094 | 423.0996 | 333.9054 | 134.4588 | 134.6079 |  | 65.9802
0 | [10 empty] | 0 | 333.9054 | 423.0996 | 41.094 | 41.094 | 423.0996 | 333.9054 | 134.6079 | 134.6079 |  | 66.1293
0 | [10 empty] | 0 | 333.9054 | 423.0996 | 41.094 | 41.094 | 423.0996 | 333.9054 | 134.6079 | 134.6079 |  | 66.1293
0 | [6 empty] | bellows |  |  |  | 0.19082 | 333.81 | 423.0996 | 41.094 | 41.094 | 423.0996 | 333.7146 | 134.7033 | 134.7987 |  | 66.2247
0 | LS1 | CB02 | GV | 1348 | LS1_CB02:GV_D1348 |  | gate valve | 40 |  |  | 0.057404 | 333.6859 | 423.0996 | 41.094 | 41.094 | 423.0996 | 333.6572 | 134.8274 | 134.8561 |  | 66.3488
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 333.6055 | 423.0996 | 41.094 | 41.094 | 423.0996 | 333.5537 | 134.9079 | 134.9596 |  | 66.4293
0 | LS1 | WB02 | BPM | 1350 | LS1_WB02:BPM_D1350 | BPM2 | position monitor |  |  |  | 0 | 333.5537 | 423.0996 | 41.094 | 41.094 | 423.0996 | 333.5537 | 134.9596 | 134.9596 |  | 66.481
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 333.4811 | 423.0996 | 41.094 | 41.094 | 423.0996 | 333.4084 | 135.0323 | 135.1049 |  | 66.5536
0 | LS1 | WB02 | HMR | 1351 | LS1_WB02:HMR_D1351 | HR_L | Halo ring, large aperture | 33 |  |  | 0 | 333.4084 | 423.0996 | 41.094 | 41.094 | 423.0996 | 333.4084 | 135.1049 | 135.1049 |  | 66.6263
0 | [6 empty] | diagnostic box |  |  |  | 0 | 333.4084 | 423.0996 | 41.094 | 41.094 | 423.0996 | 333.4084 | 135.1049 | 135.1049 |  | 66.6263
0 | LS1 | WB02 | IP | 1351 | LS1_WB02:IP_D1351 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 333.4084 | 423.0996 | 41.094 | 41.094 | 423.0996 | 333.4084 | 135.1049 | 135.1049 |  | 66.6263
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 333.3428 | 423.0996 | 41.094 | 41.094 | 423.0996 | 333.2772 | 135.1705 | 135.2361 |  | 66.6919
0 | LS1 | CB03 | GV | 1353 | LS1_CB03:GV_D1353 |  | gate valve | 40 |  |  | 0.057404 | 333.2485 | 423.0996 | 41.094 | 41.094 | 423.0996 | 333.2198 | 135.2648 | 135.2935 |  | 66.7862
0 | [6 empty] | bellows |  |  |  | 0.19082 | 333.1244 | 423.0996 | 41.094 | 41.094 | 423.0996 | 333.029 | 135.3889 | 135.4844 |  | 66.9103
0 | LS1 | CB03 | CAV1 | 1356 | LS1_CB03:CAV1_D1356 |  | b08 resonator | 34 |  |  | 0.29817 | 332.8799 | 423.0996 | 41.094 | 41.094 | 423.0996 | 332.7308 | 135.6334 | 135.7825 |  | 67.1548
0 | [10 empty] | 0 | 332.7308 | 423.0996 | 41.094 | 41.094 | 423.0996 | 332.7308 | 135.7825 | 135.7825 |  | 67.3039
0 | LS1 | CB03 | BLM | 1358 | LS1_CB03:BLM_D1358 |  | outside loss monitor |  |  |  | 0 | 332.7308 | 423.0996 | 41.094 | 41.094 | 423.0996 | 332.7308 | 135.7825 | 135.7825 |  | 67.3039
0 | [6 empty] | bellow | 40 |  |  | 0.09651 | 332.6826 | 423.0996 | 41.094 | 41.094 | 423.0996 | 332.6343 | 135.8308 | 135.879 |  | 67.3522
0 | LS1 | CB03 | BPM | 1359 | LS1_CB03:BPM_D1359 | BPM6 | position monitor, BPMc |  |  |  | 0 | 332.6343 | 423.0996 | 41.094 | 41.094 | 423.0996 | 332.6343 | 135.879 | 135.879 |  | 67.4004
0 | [6 empty] | BPM-box | 40 |  |  | 0.05588 | 332.6064 | 423.0996 | 41.094 | 41.094 | 423.0996 | 332.5784 | 135.907 | 135.9349 |  | 67.4283
0 | [6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 332.5404 | 423.0996 | 41.094 | 41.094 | 423.0996 | 332.5023 | 135.973 | 136.0111 |  | 67.4944
0 | LS1 | CB03 | SOL1 | 1363 | LS1_CB03:SOL1_D1363 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 332.2523 | 423.0996 | 41.094 | 41.094 | 423.0996 | 332.0023 | 136.2611 | 136.5111 |  | 67.7824
0 | [6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 331.9642 | 423.0996 | 41.094 | 41.094 | 423.0996 | 331.9262 | 136.5491 | 136.5872 |  | 68.0705
0 | [6 empty] | bellow |  |  |  | 0.099517 | 331.8764 | 423.0996 | 41.094 | 41.094 | 423.0996 | 331.8266 | 136.6369 | 136.6867 |  | 68.1583
0 | LS1 | CB03 | CAV2 | 1368 | LS1_CB03:CAV2_D1368 |  | b08 resonator | 34 |  |  | 0.29817 | 331.6776 | 423.0996 | 41.094 | 41.094 | 423.0996 | 331.5285 | 136.8358 | 136.9849 |  | 68.3572
0 | [10 empty] | 0 | 331.5285 | 423.0996 | 41.094 | 41.094 | 423.0996 | 331.5285 | 136.9849 | 136.9849 |  | 68.5062
0 | [10 empty] | 0 | 331.5285 | 423.0996 | 41.094 | 41.094 | 423.0996 | 331.5285 | 136.9849 | 136.9849 |  | 68.5062
0 | [6 empty] | bellow |  |  |  | 0.099517 | 331.4787 | 423.0996 | 41.094 | 41.094 | 423.0996 | 331.4289 | 137.0346 | 137.0844 |  | 68.556
0 | LS1 | CB03 | CAV3 | 1372 | LS1_CB03:CAV3_D1372 |  | b08 resonator | 34 |  |  | 0.29817 | 331.2799 | 423.0996 | 41.094 | 41.094 | 423.0996 | 331.1308 | 137.2335 | 137.3826 |  | 68.7548
0 | [10 empty] | 0 | 331.1308 | 423.0996 | 41.094 | 41.094 | 423.0996 | 331.1308 | 137.3826 | 137.3826 |  | 68.9039
0 | [10 empty] | 0 | 331.1308 | 423.0996 | 41.094 | 41.094 | 423.0996 | 331.1308 | 137.3826 | 137.3826 |  | 68.9039
0 | [6 empty] | bellow |  |  |  | 0.099517 | 331.081 | 423.0996 | 41.094 | 41.094 | 423.0996 | 331.0313 | 137.4323 | 137.4821 |  | 68.9537
0 | LS1 | CB03 | CAV4 | 1376 | LS1_CB03:CAV4_D1376 |  | b08 resonator | 34 |  |  | 0.29817 | 330.8822 | 423.0996 | 41.094 | 41.094 | 423.0996 | 330.7331 | 137.6312 | 137.7803 |  | 69.1525
0 | [10 empty] | 0 | 330.7331 | 423.0996 | 41.094 | 41.094 | 423.0996 | 330.7331 | 137.7803 | 137.7803 |  | 69.3016
0 | LS1 | CB03 | BLM | 1378 | LS1_CB03:BLM_D1378 |  | outside loss monitor |  |  |  | 0 | 330.7331 | 423.0996 | 41.094 | 41.094 | 423.0996 | 330.7331 | 137.7803 | 137.7803 |  | 69.3016
0 | [6 empty] | bellow |  |  |  | 0.09651 | 330.6848 | 423.0996 | 41.094 | 41.094 | 423.0996 | 330.6366 | 137.8285 | 137.8768 |  | 69.3499
0 | LS1 | CB03 | BPM | 1379 | LS1_CB03:BPM_D1379 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 330.6366 | 423.0996 | 41.094 | 41.094 | 423.0996 | 330.6366 | 137.8768 | 137.8768 |  | 69.3981
0 | [6 empty] | BPM-box |  |  |  | 0.05588 | 330.6086 | 423.0996 | 41.094 | 41.094 | 423.0996 | 330.5807 | 137.9047 | 137.9326 |  | 69.4261
0 | [6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 330.5426 | 423.0996 | 41.094 | 41.094 | 423.0996 | 330.5046 | 137.9707 | 138.0088 |  | 69.4921
0 | LS1 | CB03 | SOL2 | 1383 | LS1_CB03:SOL2_D1383 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 330.2546 | 423.0996 | 41.094 | 41.094 | 423.0996 | 330.0046 | 138.2588 | 138.5088 |  | 69.7801
0 | [6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 329.9665 | 423.0996 | 41.094 | 41.094 | 423.0996 | 329.9284 | 138.5468 | 138.5849 |  | 70.0682
0 | [6 empty] | bellow |  |  |  | 0.099517 | 329.8787 | 423.0996 | 41.094 | 41.094 | 423.0996 | 329.8289 | 138.6347 | 138.6844 |  | 70.156
0 | LS1 | CB03 | CAV5 | 1388 | LS1_CB03:CAV5_D1388 |  | b08 resonator | 34 |  |  | 0.29817 | 329.6798 | 423.0996 | 41.094 | 41.094 | 423.0996 | 329.5307 | 138.8335 | 138.9826 |  | 70.3549
0 | [10 empty] | 0 | 329.5307 | 423.0996 | 41.094 | 41.094 | 423.0996 | 329.5307 | 138.9826 | 138.9826 |  | 70.504
0 | [10 empty] | 0 | 329.5307 | 423.0996 | 41.094 | 41.094 | 423.0996 | 329.5307 | 138.9826 | 138.9826 |  | 70.504
0 | [6 empty] | bellow |  |  |  | 0.099517 | 329.481 | 423.0996 | 41.094 | 41.094 | 423.0996 | 329.4312 | 139.0324 | 139.0821 |  | 70.5537
0 | LS1 | CB03 | CAV6 | 1392 | LS1_CB03:CAV6_D1392 |  | b08 resonator | 34 |  |  | 0.29817 | 329.2821 | 423.0996 | 41.094 | 41.094 | 423.0996 | 329.1331 | 139.2312 | 139.3803 |  | 70.7526
0 | [10 empty] | 0 | 329.1331 | 423.0996 | 41.094 | 41.094 | 423.0996 | 329.1331 | 139.3803 | 139.3803 |  | 70.9017
0 | [10 empty] | 0 | 329.1331 | 423.0996 | 41.094 | 41.094 | 423.0996 | 329.1331 | 139.3803 | 139.3803 |  | 70.9017
0 | [6 empty] | bellow |  |  |  | 0.099517 | 329.0833 | 423.0996 | 41.094 | 41.094 | 423.0996 | 329.0335 | 139.43 | 139.4798 |  | 70.9514
0 | LS1 | CB03 | CAV7 | 1396 | LS1_CB03:CAV7_D1396 |  | b08 resonator | 34 |  |  | 0.29817 | 328.8844 | 423.0996 | 41.094 | 41.094 | 423.0996 | 328.7354 | 139.6289 | 139.778 |  | 71.1503
0 | [10 empty] | 0 | 328.7354 | 423.0996 | 41.094 | 41.094 | 423.0996 | 328.7354 | 139.778 | 139.778 |  | 71.2993
0 | LS1 | CB03 | BLM | 1398 | LS1_CB03:BLM_D1398 |  | outside loss monitor |  |  |  | 0 | 328.7354 | 423.0996 | 41.094 | 41.094 | 423.0996 | 328.7354 | 139.778 | 139.778 |  | 71.2993
0 | [6 empty] | bellow |  |  |  | 0.09651 | 328.6871 | 423.0996 | 41.094 | 41.094 | 423.0996 | 328.6389 | 139.8262 | 139.8745 |  | 71.3476
0 | LS1 | CB03 | BPM | 1399 | LS1_CB03:BPM_D1399 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 328.6389 | 423.0996 | 41.094 | 41.094 | 423.0996 | 328.6389 | 139.8745 | 139.8745 |  | 71.3959
0 | [6 empty] | BPM-box |  |  |  | 0.05588 | 328.6109 | 423.0996 | 41.094 | 41.094 | 423.0996 | 328.583 | 139.9024 | 139.9304 |  | 71.4238
0 | [6 empty] | solenoid-entry |  |  |  | 0.076136 | 328.5449 | 423.0996 | 41.094 | 41.094 | 423.0996 | 328.5068 | 139.9684 | 140.0065 |  | 71.4898
0 | LS1 | CB03 | SOL3 | 1403 | LS1_CB03:SOL3_D1403 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 328.2568 | 423.0996 | 41.094 | 41.094 | 423.0996 | 328.0068 | 140.2565 | 140.5065 |  | 71.7779
0 | [6 empty] | solenoid-exit |  |  |  | 0.076136 | 327.9688 | 423.0996 | 41.094 | 41.094 | 423.0996 | 327.9307 | 140.5446 | 140.5826 |  | 72.0659
0 | [6 empty] | bellow |  |  |  | 0.099517 | 327.8809 | 423.0996 | 41.094 | 41.094 | 423.0996 | 327.8312 | 140.6324 | 140.6822 |  | 72.1538
0 | LS1 | CB03 | CAV8 | 1408 | LS1_CB03:CAV8_D1408 |  | b08 resonator | 34 |  |  | 0.29817 | 327.6821 | 423.0996 | 41.094 | 41.094 | 423.0996 | 327.533 | 140.8312 | 140.9803 |  | 72.3526
0 | [10 empty] | 0 | 327.533 | 423.0996 | 41.094 | 41.094 | 423.0996 | 327.533 | 140.9803 | 140.9803 |  | 72.5017
0 | [10 empty] | 0 | 327.533 | 423.0996 | 41.094 | 41.094 | 423.0996 | 327.533 | 140.9803 | 140.9803 |  | 72.5017
0 | [6 empty] | bellows |  |  |  | 0.19082 | 327.4376 | 423.0996 | 41.094 | 41.094 | 423.0996 | 327.3422 | 141.0757 | 141.1712 |  | 72.5971
0 | LS1 | CB03 | GV | 1412 | LS1_CB03:GV_D1412 |  | gate valve | 40 |  |  | 0.057404 | 327.3135 | 423.0996 | 41.094 | 41.094 | 423.0996 | 327.2848 | 141.1999 | 141.2286 |  | 72.7212
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 327.233 | 423.0996 | 41.094 | 41.094 | 423.0996 | 327.1813 | 141.2803 | 141.3321 |  | 72.8017
0 | LS1 | WB03 | BPM | 1413 | LS1_WB03:BPM_D1413 | BPM2 | position monitor |  |  |  | 0 | 327.1813 | 423.0996 | 41.094 | 41.094 | 423.0996 | 327.1813 | 141.3321 | 141.3321 |  | 72.8534
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 327.1086 | 423.0996 | 41.094 | 41.094 | 423.0996 | 327.036 | 141.4047 | 141.4773 |  | 72.9261
0 | LS1 | WB03 | HMR | 1415 | LS1_WB03:HMR_D1415 | HR_L | Halo ring, large aperture | 33 |  |  | 0 | 327.036 | 423.0996 | 41.094 | 41.094 | 423.0996 | 327.036 | 141.4773 | 141.4773 |  | 72.9987
0 | [6 empty] | diagnostic box |  |  |  | 0 | 327.036 | 423.0996 | 41.094 | 41.094 | 423.0996 | 327.036 | 141.4773 | 141.4773 |  | 72.9987
0 | LS1 | WB03 | IP | 1415 | LS1_WB03:IP_D1415 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 327.036 | 423.0996 | 41.094 | 41.094 | 423.0996 | 327.036 | 141.4773 | 141.4773 |  | 72.9987
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 326.9704 | 423.0996 | 41.094 | 41.094 | 423.0996 | 326.9048 | 141.5429 | 141.6086 |  | 73.0643
0 | LS1 | CB04 | GV | 1416 | LS1_CB04:GV_D1416 |  | gate valve | 40 |  |  | 0.057404 | 326.8761 | 423.0996 | 41.094 | 41.094 | 423.0996 | 326.8474 | 141.6373 | 141.666 |  | 73.1586
0 | [6 empty] | bellows |  |  |  | 0.19082 | 326.752 | 423.0996 | 41.094 | 41.094 | 423.0996 | 326.6566 | 141.7614 | 141.8568 |  | 73.2827
0 | LS1 | CB04 | CAV1 | 1420 | LS1_CB04:CAV1_D1420 |  | b08 resonator | 34 |  |  | 0.29817 | 326.5075 | 423.0996 | 41.094 | 41.094 | 423.0996 | 326.3584 | 142.0059 | 142.155 |  | 73.5272
0 | [10 empty] | 0 | 326.3584 | 423.0996 | 41.094 | 41.094 | 423.0996 | 326.3584 | 142.155 | 142.155 |  | 73.6763
0 | LS1 | CB04 | BLM | 1422 | LS1_CB04:BLM_D1422 |  | outside loss monitor |  |  |  | 0 | 326.3584 | 423.0996 | 41.094 | 41.094 | 423.0996 | 326.3584 | 142.155 | 142.155 |  | 73.6763
0 | [6 empty] | bellow | 40 |  |  | 0.09651 | 326.3101 | 423.0996 | 41.094 | 41.094 | 423.0996 | 326.2619 | 142.2032 | 142.2515 |  | 73.7246
0 | LS1 | CB04 | BPM | 1423 | LS1_CB04:BPM_D1423 | BPM6 | position monitor, BPMc |  |  |  | 0 | 326.2619 | 423.0996 | 41.094 | 41.094 | 423.0996 | 326.2619 | 142.2515 | 142.2515 |  | 73.7728
0 | [6 empty] | BPM-box | 40 |  |  | 0.05588 | 326.2339 | 423.0996 | 41.094 | 41.094 | 423.0996 | 326.206 | 142.2794 | 142.3073 |  | 73.8008
0 | [6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 326.1679 | 423.0996 | 41.094 | 41.094 | 423.0996 | 326.1299 | 142.3454 | 142.3835 |  | 73.8668
0 | LS1 | CB04 | SOL1 | 1426 | LS1_CB04:SOL1_D1426 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 325.8799 | 423.0996 | 41.094 | 41.094 | 423.0996 | 325.6299 | 142.6335 | 142.8835 |  | 74.1548
0 | [6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 325.5918 | 423.0996 | 41.094 | 41.094 | 423.0996 | 325.5537 | 142.9215 | 142.9596 |  | 74.4429
0 | [6 empty] | bellow |  |  |  | 0.099517 | 325.504 | 423.0996 | 41.094 | 41.094 | 423.0996 | 325.4542 | 143.0094 | 143.0591 |  | 74.5307
0 | LS1 | CB04 | CAV2 | 1432 | LS1_CB04:CAV2_D1432 |  | b08 resonator | 34 |  |  | 0.29817 | 325.3051 | 423.0996 | 41.094 | 41.094 | 423.0996 | 325.156 | 143.2082 | 143.3573 |  | 74.7296
0 | [10 empty] | 0 | 325.156 | 423.0996 | 41.094 | 41.094 | 423.0996 | 325.156 | 143.3573 | 143.3573 |  | 74.8787
0 | [10 empty] | 0 | 325.156 | 423.0996 | 41.094 | 41.094 | 423.0996 | 325.156 | 143.3573 | 143.3573 |  | 74.8787
0 | [6 empty] | bellow |  |  |  | 0.099517 | 325.1063 | 423.0996 | 41.094 | 41.094 | 423.0996 | 325.0565 | 143.4071 | 143.4568 |  | 74.9284
0 | LS1 | CB04 | CAV3 | 1436 | LS1_CB04:CAV3_D1436 |  | b08 resonator | 34 |  |  | 0.29817 | 324.9074 | 423.0996 | 41.094 | 41.094 | 423.0996 | 324.7584 | 143.6059 | 143.755 |  | 75.1273
0 | [10 empty] | 0 | 324.7584 | 423.0996 | 41.094 | 41.094 | 423.0996 | 324.7584 | 143.755 | 143.755 |  | 75.2764
0 | [10 empty] | 0 | 324.7584 | 423.0996 | 41.094 | 41.094 | 423.0996 | 324.7584 | 143.755 | 143.755 |  | 75.2764
0 | [6 empty] | bellow |  |  |  | 0.099517 | 324.7086 | 423.0996 | 41.094 | 41.094 | 423.0996 | 324.6588 | 143.8047 | 143.8545 |  | 75.3261
0 | LS1 | CB04 | CAV4 | 1440 | LS1_CB04:CAV4_D1440 |  | b08 resonator | 34 |  |  | 0.29817 | 324.5097 | 423.0996 | 41.094 | 41.094 | 423.0996 | 324.3607 | 144.0036 | 144.1527 |  | 75.525
0 | [10 empty] | 0 | 324.3607 | 423.0996 | 41.094 | 41.094 | 423.0996 | 324.3607 | 144.1527 | 144.1527 |  | 75.674
0 | LS1 | CB04 | BLM | 1442 | LS1_CB04:BLM_D1442 |  | outside loss monitor |  |  |  | 0 | 324.3607 | 423.0996 | 41.094 | 41.094 | 423.0996 | 324.3607 | 144.1527 | 144.1527 |  | 75.674
0 | [6 empty] | bellow |  |  |  | 0.09651 | 324.3124 | 423.0996 | 41.094 | 41.094 | 423.0996 | 324.2642 | 144.2009 | 144.2492 |  | 75.7223
0 | LS1 | CB04 | BPM | 1442 | LS1_CB04:BPM_D1442 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 324.2642 | 423.0996 | 41.094 | 41.094 | 423.0996 | 324.2642 | 144.2492 | 144.2492 |  | 75.7706
0 | [6 empty] | BPM-box |  |  |  | 0.05588 | 324.2362 | 423.0996 | 41.094 | 41.094 | 423.0996 | 324.2083 | 144.2771 | 144.3051 |  | 75.7985
0 | [6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 324.1702 | 423.0996 | 41.094 | 41.094 | 423.0996 | 324.1321 | 144.3431 | 144.3812 |  | 75.8645
0 | LS1 | CB04 | SOL2 | 1446 | LS1_CB04:SOL2_D1446 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 323.8821 | 423.0996 | 41.094 | 41.094 | 423.0996 | 323.6321 | 144.6312 | 144.8812 |  | 76.1526
0 | [6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 323.5941 | 423.0996 | 41.094 | 41.094 | 423.0996 | 323.556 | 144.9193 | 144.9573 |  | 76.4406
0 | [6 empty] | bellow |  |  |  | 0.099517 | 323.5062 | 423.0996 | 41.094 | 41.094 | 423.0996 | 323.4565 | 145.0071 | 145.0569 |  | 76.5285
0 | LS1 | CB04 | CAV5 | 1452 | LS1_CB04:CAV5_D1452 |  | b08 resonator | 34 |  |  | 0.29817 | 323.3074 | 423.0996 | 41.094 | 41.094 | 423.0996 | 323.1583 | 145.2059 | 145.355 |  | 76.7273
0 | [10 empty] | 0 | 323.1583 | 423.0996 | 41.094 | 41.094 | 423.0996 | 323.1583 | 145.355 | 145.355 |  | 76.8764
0 | [10 empty] | 0 | 323.1583 | 423.0996 | 41.094 | 41.094 | 423.0996 | 323.1583 | 145.355 | 145.355 |  | 76.8764
0 | [6 empty] | bellow |  |  |  | 0.099517 | 323.1086 | 423.0996 | 41.094 | 41.094 | 423.0996 | 323.0588 | 145.4048 | 145.4545 |  | 76.9262
0 | LS1 | CB04 | CAV6 | 1456 | LS1_CB04:CAV6_D1456 |  | b08 resonator | 34 |  |  | 0.29817 | 322.9097 | 423.0996 | 41.094 | 41.094 | 423.0996 | 322.7606 | 145.6036 | 145.7527 |  | 77.125
0 | [10 empty] | 0 | 322.7606 | 423.0996 | 41.094 | 41.094 | 423.0996 | 322.7606 | 145.7527 | 145.7527 |  | 77.2741
0 | [10 empty] | 0 | 322.7606 | 423.0996 | 41.094 | 41.094 | 423.0996 | 322.7606 | 145.7527 | 145.7527 |  | 77.2741
0 | [6 empty] | bellow |  |  |  | 0.099517 | 322.7109 | 423.0996 | 41.094 | 41.094 | 423.0996 | 322.6611 | 145.8025 | 145.8522 |  | 77.3238
0 | LS1 | CB04 | CAV7 | 1460 | LS1_CB04:CAV7_D1460 |  | b08 resonator | 34 |  |  | 0.29817 | 322.512 | 423.0996 | 41.094 | 41.094 | 423.0996 | 322.3629 | 146.0013 | 146.1504 |  | 77.5227
0 | [10 empty] | 0 | 322.3629 | 423.0996 | 41.094 | 41.094 | 423.0996 | 322.3629 | 146.1504 | 146.1504 |  | 77.6718
0 | LS1 | CB04 | BLM | 1462 | LS1_CB04:BLM_D1462 |  | outside loss monitor |  |  |  | 0 | 322.3629 | 423.0996 | 41.094 | 41.094 | 423.0996 | 322.3629 | 146.1504 | 146.1504 |  | 77.6718
0 | [6 empty] | bellow |  |  |  | 0.09651 | 322.3147 | 423.0996 | 41.094 | 41.094 | 423.0996 | 322.2664 | 146.1987 | 146.2469 |  | 77.72
0 | LS1 | CB04 | BPM | 1462 | LS1_CB04:BPM_D1462 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 322.2664 | 423.0996 | 41.094 | 41.094 | 423.0996 | 322.2664 | 146.2469 | 146.2469 |  | 77.7683
0 | [6 empty] | BPM-box |  |  |  | 0.05588 | 322.2385 | 423.0996 | 41.094 | 41.094 | 423.0996 | 322.2105 | 146.2749 | 146.3028 |  | 77.7962
0 | [6 empty] | solenoid-entry |  |  |  | 0.076136 | 322.1725 | 423.0996 | 41.094 | 41.094 | 423.0996 | 322.1344 | 146.3409 | 146.3789 |  | 77.8622
0 | LS1 | CB04 | SOL3 | 1466 | LS1_CB04:SOL3_D1466 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 321.8844 | 423.0996 | 41.094 | 41.094 | 423.0996 | 321.6344 | 146.6289 | 146.8789 |  | 78.1503
0 | [6 empty] | solenoid-exit |  |  |  | 0.076136 | 321.5963 | 423.0996 | 41.094 | 41.094 | 423.0996 | 321.5583 | 146.917 | 146.9551 |  | 78.4384
0 | [6 empty] | bellow |  |  |  | 0.099517 | 321.5085 | 423.0996 | 41.094 | 41.094 | 423.0996 | 321.4588 | 147.0048 | 147.0546 |  | 78.5262
0 | LS1 | CB04 | CAV8 | 1472 | LS1_CB04:CAV8_D1472 |  | b08 resonator | 34 |  |  | 0.29817 | 321.3097 | 423.0996 | 41.094 | 41.094 | 423.0996 | 321.1606 | 147.2037 | 147.3528 |  | 78.725
0 | [10 empty] | 0 | 321.1606 | 423.0996 | 41.094 | 41.094 | 423.0996 | 321.1606 | 147.3528 | 147.3528 |  | 78.8741
0 | [10 empty] | 0 | 321.1606 | 423.0996 | 41.094 | 41.094 | 423.0996 | 321.1606 | 147.3528 | 147.3528 |  | 78.8741
0 | [6 empty] | bellows |  |  |  | 0.19082 | 321.0652 | 423.0996 | 41.094 | 41.094 | 423.0996 | 320.9698 | 147.4482 | 147.5436 |  | 78.9695
0 | LS1 | CB04 | GV | 1476 | LS1_CB04:GV_D1476 |  | gate valve | 40 |  |  | 0.057404 | 320.9411 | 423.0996 | 41.094 | 41.094 | 423.0996 | 320.9124 | 147.5723 | 147.601 |  | 79.0936
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 320.8606 | 423.0996 | 41.094 | 41.094 | 423.0996 | 320.8089 | 147.6527 | 147.7045 |  | 79.1741
0 | LS1 | WB04 | BPM | 1477 | LS1_WB04:BPM_D1477 | BPM2 | position monitor |  |  |  | 0 | 320.8089 | 423.0996 | 41.094 | 41.094 | 423.0996 | 320.8089 | 147.7045 | 147.7045 |  | 79.2259
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 320.7362 | 423.0996 | 41.094 | 41.094 | 423.0996 | 320.6636 | 147.7771 | 147.8498 |  | 79.2985
0 | LS1 | WB04 | HMR | 1478 | LS1_WB04:HMR_D1478 | HR_L | Halo ring, large aperture | 33 |  |  | 0 | 320.6636 | 423.0996 | 41.094 | 41.094 | 423.0996 | 320.6636 | 147.8498 | 147.8498 |  | 79.3711
0 | [6 empty] | diagnostic box |  |  |  | 0 | 320.6636 | 423.0996 | 41.094 | 41.094 | 423.0996 | 320.6636 | 147.8498 | 147.8498 |  | 79.3711
0 | LS1 | WB04 | IP | 1478 | LS1_WB04:IP_D1478 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 320.6636 | 423.0996 | 41.094 | 41.094 | 423.0996 | 320.6636 | 147.8498 | 147.8498 |  | 79.3711
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 320.598 | 423.0996 | 41.094 | 41.094 | 423.0996 | 320.5324 | 147.9154 | 147.981 |  | 79.4367
0 | LS1 | CB05 | GV | 1480 | LS1_CB05:GV_D1480 |  | gate valve | 40 |  |  | 0.057404 | 320.5037 | 423.0996 | 41.094 | 41.094 | 423.0996 | 320.475 | 148.0097 | 148.0384 |  | 79.5311
0 | [6 empty] | bellows |  |  |  | 0.19082 | 320.3795 | 423.0996 | 41.094 | 41.094 | 423.0996 | 320.2841 | 148.1338 | 148.2292 |  | 79.6552
0 | LS1 | CB05 | CAV1 | 1484 | LS1_CB05:CAV1_D1484 |  | b08 resonator | 34 |  |  | 0.29817 | 320.1351 | 423.0996 | 41.094 | 41.094 | 423.0996 | 319.986 | 148.3783 | 148.5274 |  | 79.8997
0 | [10 empty] | 0 | 319.986 | 423.0996 | 41.094 | 41.094 | 423.0996 | 319.986 | 148.5274 | 148.5274 |  | 80.0487
0 | LS1 | CB05 | BLM | 1485 | LS1_CB05:BLM_D1485 |  | outside loss monitor |  |  |  | 0 | 319.986 | 423.0996 | 41.094 | 41.094 | 423.0996 | 319.986 | 148.5274 | 148.5274 |  | 80.0487
0 | [6 empty] | bellow | 40 |  |  | 0.09651 | 319.9377 | 423.0996 | 41.094 | 41.094 | 423.0996 | 319.8895 | 148.5756 | 148.6239 |  | 80.097
0 | LS1 | CB05 | BPM | 1486 | LS1_CB05:BPM_D1486 | BPM6 | position monitor, BPMc |  |  |  | 0 | 319.8895 | 423.0996 | 41.094 | 41.094 | 423.0996 | 319.8895 | 148.6239 | 148.6239 |  | 80.1453
0 | [6 empty] | BPM-box | 40 |  |  | 0.05588 | 319.8615 | 423.0996 | 41.094 | 41.094 | 423.0996 | 319.8336 | 148.6518 | 148.6798 |  | 80.1732
0 | [6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 319.7955 | 423.0996 | 41.094 | 41.094 | 423.0996 | 319.7574 | 148.7178 | 148.7559 |  | 80.2392
0 | LS1 | CB05 | SOL1 | 1490 | LS1_CB05:SOL1_D1490 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 319.5074 | 423.0996 | 41.094 | 41.094 | 423.0996 | 319.2574 | 149.0059 | 149.2559 |  | 80.5273
0 | [6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 319.2194 | 423.0996 | 41.094 | 41.094 | 423.0996 | 319.1813 | 149.294 | 149.332 |  | 80.8153
0 | [6 empty] | bellow |  |  |  | 0.099517 | 319.1315 | 423.0996 | 41.094 | 41.094 | 423.0996 | 319.0818 | 149.3818 | 149.4316 |  | 80.9032
0 | LS1 | CB05 | CAV2 | 1496 | LS1_CB05:CAV2_D1496 |  | b08 resonator | 34 |  |  | 0.29817 | 318.9327 | 423.0996 | 41.094 | 41.094 | 423.0996 | 318.7836 | 149.5806 | 149.7297 |  | 81.102
0 | [10 empty] | 0 | 318.7836 | 423.0996 | 41.094 | 41.094 | 423.0996 | 318.7836 | 149.7297 | 149.7297 |  | 81.2511
0 | [10 empty] | 0 | 318.7836 | 423.0996 | 41.094 | 41.094 | 423.0996 | 318.7836 | 149.7297 | 149.7297 |  | 81.2511
0 | [6 empty] | bellow |  |  |  | 0.099517 | 318.7339 | 423.0996 | 41.094 | 41.094 | 423.0996 | 318.6841 | 149.7795 | 149.8292 |  | 81.3009
0 | LS1 | CB05 | CAV3 | 1500 | LS1_CB05:CAV3_D1500 |  | b08 resonator | 34 |  |  | 0.29817 | 318.535 | 423.0996 | 41.094 | 41.094 | 423.0996 | 318.3859 | 149.9783 | 150.1274 |  | 81.4997
0 | [10 empty] | 0 | 318.3859 | 423.0996 | 41.094 | 41.094 | 423.0996 | 318.3859 | 150.1274 | 150.1274 |  | 81.6488
0 | [10 empty] | 0 | 318.3859 | 423.0996 | 41.094 | 41.094 | 423.0996 | 318.3859 | 150.1274 | 150.1274 |  | 81.6488
0 | [6 empty] | bellow |  |  |  | 0.099517 | 318.3362 | 423.0996 | 41.094 | 41.094 | 423.0996 | 318.2864 | 150.1772 | 150.2269 |  | 81.6985
0 | LS1 | CB05 | CAV4 | 1504 | LS1_CB05:CAV4_D1504 |  | b08 resonator | 34 |  |  | 0.29817 | 318.1373 | 423.0996 | 41.094 | 41.094 | 423.0996 | 317.9882 | 150.376 | 150.5251 |  | 81.8974
0 | [10 empty] | 0 | 317.9882 | 423.0996 | 41.094 | 41.094 | 423.0996 | 317.9882 | 150.5251 | 150.5251 |  | 82.0465
0 | LS1 | CB05 | BLM | 1505 | LS1_CB05:BLM_D1505 |  | outside loss monitor |  |  |  | 0 | 317.9882 | 423.0996 | 41.094 | 41.094 | 423.0996 | 317.9882 | 150.5251 | 150.5251 |  | 82.0465
0 | [6 empty] | bellow |  |  |  | 0.09651 | 317.94 | 423.0996 | 41.094 | 41.094 | 423.0996 | 317.8917 | 150.5734 | 150.6216 |  | 82.0947
0 | LS1 | CB05 | BPM | 1506 | LS1_CB05:BPM_D1506 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 317.8917 | 423.0996 | 41.094 | 41.094 | 423.0996 | 317.8917 | 150.6216 | 150.6216 |  | 82.143
0 | [6 empty] | BPM-box |  |  |  | 0.05588 | 317.8638 | 423.0996 | 41.094 | 41.094 | 423.0996 | 317.8359 | 150.6496 | 150.6775 |  | 82.1709
0 | [6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 317.7978 | 423.0996 | 41.094 | 41.094 | 423.0996 | 317.7597 | 150.7156 | 150.7536 |  | 82.2369
0 | LS1 | CB05 | SOL2 | 1510 | LS1_CB05:SOL2_D1510 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 317.5097 | 423.0996 | 41.094 | 41.094 | 423.0996 | 317.2597 | 151.0036 | 151.2536 |  | 82.525
0 | [6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 317.2216 | 423.0996 | 41.094 | 41.094 | 423.0996 | 317.1836 | 151.2917 | 151.3298 |  | 82.8131
0 | [6 empty] | bellow |  |  |  | 0.099517 | 317.1338 | 423.0996 | 41.094 | 41.094 | 423.0996 | 317.0841 | 151.3795 | 151.4293 |  | 82.9009
0 | LS1 | CB05 | CAV5 | 1516 | LS1_CB05:CAV5_D1516 |  | b08 resonator | 34 |  |  | 0.29817 | 316.935 | 423.0996 | 41.094 | 41.094 | 423.0996 | 316.7859 | 151.5784 | 151.7275 |  | 83.0997
0 | [10 empty] | 0 | 316.7859 | 423.0996 | 41.094 | 41.094 | 423.0996 | 316.7859 | 151.7275 | 151.7275 |  | 83.2488
0 | [10 empty] | 0 | 316.7859 | 423.0996 | 41.094 | 41.094 | 423.0996 | 316.7859 | 151.7275 | 151.7275 |  | 83.2488
0 | [6 empty] | bellow |  |  |  | 0.099517 | 316.7361 | 423.0996 | 41.094 | 41.094 | 423.0996 | 316.6864 | 151.7772 | 151.827 |  | 83.2986
0 | LS1 | CB05 | CAV6 | 1520 | LS1_CB05:CAV6_D1520 |  | b08 resonator | 34 |  |  | 0.29817 | 316.5373 | 423.0996 | 41.094 | 41.094 | 423.0996 | 316.3882 | 151.9761 | 152.1251 |  | 83.4974
0 | [10 empty] | 0 | 316.3882 | 423.0996 | 41.094 | 41.094 | 423.0996 | 316.3882 | 152.1251 | 152.1251 |  | 83.6465
0 | [10 empty] | 0 | 316.3882 | 423.0996 | 41.094 | 41.094 | 423.0996 | 316.3882 | 152.1251 | 152.1251 |  | 83.6465
0 | [6 empty] | bellow |  |  |  | 0.099517 | 316.3384 | 423.0996 | 41.094 | 41.094 | 423.0996 | 316.2887 | 152.1749 | 152.2247 |  | 83.6963
0 | LS1 | CB05 | CAV7 | 1524 | LS1_CB05:CAV7_D1524 |  | b08 resonator | 34 |  |  | 0.29817 | 316.1396 | 423.0996 | 41.094 | 41.094 | 423.0996 | 315.9905 | 152.3737 | 152.5228 |  | 83.8951
0 | [10 empty] | 0 | 315.9905 | 423.0996 | 41.094 | 41.094 | 423.0996 | 315.9905 | 152.5228 | 152.5228 |  | 84.0442
0 | LS1 | CB05 | BLM | 1525 | LS1_CB05:BLM_D1525 |  | outside loss monitor |  |  |  | 0 | 315.9905 | 423.0996 | 41.094 | 41.094 | 423.0996 | 315.9905 | 152.5228 | 152.5228 |  | 84.0442
0 | [6 empty] | bellow |  |  |  | 0.09651 | 315.9423 | 423.0996 | 41.094 | 41.094 | 423.0996 | 315.894 | 152.5711 | 152.6193 |  | 84.0925
0 | LS1 | CB05 | BPM | 1526 | LS1_CB05:BPM_D1526 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 315.894 | 423.0996 | 41.094 | 41.094 | 423.0996 | 315.894 | 152.6193 | 152.6193 |  | 84.1407
0 | [6 empty] | BPM-box |  |  |  | 0.05588 | 315.8661 | 423.0996 | 41.094 | 41.094 | 423.0996 | 315.8381 | 152.6473 | 152.6752 |  | 84.1686
0 | [6 empty] | solenoid-entry |  |  |  | 0.076136 | 315.8001 | 423.0996 | 41.094 | 41.094 | 423.0996 | 315.762 | 152.7133 | 152.7514 |  | 84.2347
0 | LS1 | CB05 | SOL3 | 1530 | LS1_CB05:SOL3_D1530 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 315.512 | 423.0996 | 41.094 | 41.094 | 423.0996 | 315.262 | 153.0014 | 153.2514 |  | 84.5227
0 | [6 empty] | solenoid-exit |  |  |  | 0.076136 | 315.2239 | 423.0996 | 41.094 | 41.094 | 423.0996 | 315.1859 | 153.2894 | 153.3275 |  | 84.8108
0 | [6 empty] | bellow |  |  |  | 0.099517 | 315.1361 | 423.0996 | 41.094 | 41.094 | 423.0996 | 315.0863 | 153.3772 | 153.427 |  | 84.8986
0 | LS1 | CB05 | CAV8 | 1536 | LS1_CB05:CAV8_D1536 |  | b08 resonator | 34 |  |  | 0.29817 | 314.9372 | 423.0996 | 41.094 | 41.094 | 423.0996 | 314.7882 | 153.5761 | 153.7252 |  | 85.0975
0 | [10 empty] | 0 | 314.7882 | 423.0996 | 41.094 | 41.094 | 423.0996 | 314.7882 | 153.7252 | 153.7252 |  | 85.2465
0 | [10 empty] | 0 | 314.7882 | 423.0996 | 41.094 | 41.094 | 423.0996 | 314.7882 | 153.7252 | 153.7252 |  | 85.2465
0 | [6 empty] | bellows |  |  |  | 0.19082 | 314.6928 | 423.0996 | 41.094 | 41.094 | 423.0996 | 314.5973 | 153.8206 | 153.916 |  | 85.342
0 | LS1 | CB05 | GV | 1539 | LS1_CB05:GV_D1539 |  | gate valve | 40 |  |  | 0.057404 | 314.5686 | 423.0996 | 41.094 | 41.094 | 423.0996 | 314.5399 | 153.9447 | 153.9734 |  | 85.4661
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 314.4882 | 423.0996 | 41.094 | 41.094 | 423.0996 | 314.4364 | 154.0252 | 154.0769 |  | 85.5465
0 | LS1 | WB05 | BPM | 1541 | LS1_WB05:BPM_D1541 | BPM2 | position monitor |  |  |  | 0 | 314.4364 | 423.0996 | 41.094 | 41.094 | 423.0996 | 314.4364 | 154.0769 | 154.0769 |  | 85.5983
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 314.3638 | 423.0996 | 41.094 | 41.094 | 423.0996 | 314.2912 | 154.1495 | 154.2222 |  | 85.6709
0 | LS1 | WB05 | HMR | 1542 | LS1_WB05:HMR_D1542 | HR_L | Halo ring, large aperture | 33 |  |  | 0 | 314.2912 | 423.0996 | 41.094 | 41.094 | 423.0996 | 314.2912 | 154.2222 | 154.2222 |  | 85.7436
0 | [6 empty] | diagnostic box |  |  |  | 0 | 314.2912 | 423.0996 | 41.094 | 41.094 | 423.0996 | 314.2912 | 154.2222 | 154.2222 |  | 85.7436
0 | LS1 | WB05 | IP | 1542 | LS1_WB05:IP_D1542 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 314.2912 | 423.0996 | 41.094 | 41.094 | 423.0996 | 314.2912 | 154.2222 | 154.2222 |  | 85.7436
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 314.2255 | 423.0996 | 41.094 | 41.094 | 423.0996 | 314.1599 | 154.2878 | 154.3534 |  | 85.8092
0 | LS1 | CB06 | GV | 1544 | LS1_CB06:GV_D1544 |  | gate valve | 40 |  |  | 0.057404 | 314.1312 | 423.0996 | 41.094 | 41.094 | 423.0996 | 314.1025 | 154.3821 | 154.4108 |  | 85.9035
0 | [6 empty] | bellows |  |  |  | 0.19082 | 314.0071 | 423.0996 | 41.094 | 41.094 | 423.0996 | 313.9117 | 154.5062 | 154.6016 |  | 86.0276
0 | LS1 | CB06 | CAV1 | 1548 | LS1_CB06:CAV1_D1548 |  | b08 resonator | 34 |  |  | 0.29817 | 313.7626 | 423.0996 | 41.094 | 41.094 | 423.0996 | 313.6135 | 154.7507 | 154.8998 |  | 86.2721
0 | [10 empty] | 0 | 313.6135 | 423.0996 | 41.094 | 41.094 | 423.0996 | 313.6135 | 154.8998 | 154.8998 |  | 86.4212
0 | LS1 | CB06 | BLM | 1549 | LS1_CB06:BLM_D1549 |  | outside loss monitor |  |  |  | 0 | 313.6135 | 423.0996 | 41.094 | 41.094 | 423.0996 | 313.6135 | 154.8998 | 154.8998 |  | 86.4212
0 | [6 empty] | bellow | 40 |  |  | 0.09651 | 313.5653 | 423.0996 | 41.094 | 41.094 | 423.0996 | 313.517 | 154.9481 | 154.9963 |  | 86.4694
0 | LS1 | CB06 | BPM | 1550 | LS1_CB06:BPM_D1550 | BPM6 | position monitor, BPMc |  |  |  | 0 | 313.517 | 423.0996 | 41.094 | 41.094 | 423.0996 | 313.517 | 154.9963 | 154.9963 |  | 86.5177
0 | [6 empty] | BPM-box | 40 |  |  | 0.05588 | 313.4891 | 423.0996 | 41.094 | 41.094 | 423.0996 | 313.4612 | 155.0242 | 155.0522 |  | 86.5456
0 | [6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 313.4231 | 423.0996 | 41.094 | 41.094 | 423.0996 | 313.385 | 155.0903 | 155.1283 |  | 86.6116
0 | LS1 | CB06 | SOL1 | 1554 | LS1_CB06:SOL1_D1554 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 313.135 | 423.0996 | 41.094 | 41.094 | 423.0996 | 312.885 | 155.3783 | 155.6283 |  | 86.8997
0 | [6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 312.8469 | 423.0996 | 41.094 | 41.094 | 423.0996 | 312.8089 | 155.6664 | 155.7045 |  | 87.1878
0 | [6 empty] | bellow |  |  |  | 0.099517 | 312.7591 | 423.0996 | 41.094 | 41.094 | 423.0996 | 312.7094 | 155.7542 | 155.804 |  | 87.2756
0 | LS1 | CB06 | CAV2 | 1560 | LS1_CB06:CAV2_D1560 |  | b08 resonator | 34 |  |  | 0.29817 | 312.5603 | 423.0996 | 41.094 | 41.094 | 423.0996 | 312.4112 | 155.9531 | 156.1021 |  | 87.4744
0 | [10 empty] | 0 | 312.4112 | 423.0996 | 41.094 | 41.094 | 423.0996 | 312.4112 | 156.1021 | 156.1021 |  | 87.6235
0 | [10 empty] | 0 | 312.4112 | 423.0996 | 41.094 | 41.094 | 423.0996 | 312.4112 | 156.1021 | 156.1021 |  | 87.6235
0 | [6 empty] | bellow |  |  |  | 0.099517 | 312.3614 | 423.0996 | 41.094 | 41.094 | 423.0996 | 312.3117 | 156.1519 | 156.2017 |  | 87.6733
0 | LS1 | CB06 | CAV3 | 1564 | LS1_CB06:CAV3_D1564 |  | b08 resonator | 34 |  |  | 0.29817 | 312.1626 | 423.0996 | 41.094 | 41.094 | 423.0996 | 312.0135 | 156.3508 | 156.4998 |  | 87.8721
0 | [10 empty] | 0 | 312.0135 | 423.0996 | 41.094 | 41.094 | 423.0996 | 312.0135 | 156.4998 | 156.4998 |  | 88.0212
0 | [10 empty] | 0 | 312.0135 | 423.0996 | 41.094 | 41.094 | 423.0996 | 312.0135 | 156.4998 | 156.4998 |  | 88.0212
0 | [6 empty] | bellow |  |  |  | 0.099517 | 311.9637 | 423.0996 | 41.094 | 41.094 | 423.0996 | 311.914 | 156.5496 | 156.5994 |  | 88.071
0 | LS1 | CB06 | CAV4 | 1567 | LS1_CB06:CAV4_D1567 |  | b08 resonator | 34 |  |  | 0.29817 | 311.7649 | 423.0996 | 41.094 | 41.094 | 423.0996 | 311.6158 | 156.7484 | 156.8975 |  | 88.2698
0 | [10 empty] | 0 | 311.6158 | 423.0996 | 41.094 | 41.094 | 423.0996 | 311.6158 | 156.8975 | 156.8975 |  | 88.4189
0 | LS1 | CB06 | BLM | 1569 | LS1_CB06:BLM_D1569 |  | outside loss monitor |  |  |  | 0 | 311.6158 | 423.0996 | 41.094 | 41.094 | 423.0996 | 311.6158 | 156.8975 | 156.8975 |  | 88.4189
0 | [6 empty] | bellow |  |  |  | 0.09651 | 311.5676 | 423.0996 | 41.094 | 41.094 | 423.0996 | 311.5193 | 156.9458 | 156.994 |  | 88.4672
0 | LS1 | CB06 | BPM | 1570 | LS1_CB06:BPM_D1570 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 311.5193 | 423.0996 | 41.094 | 41.094 | 423.0996 | 311.5193 | 156.994 | 156.994 |  | 88.5154
0 | [6 empty] | BPM-box |  |  |  | 0.05588 | 311.4914 | 423.0996 | 41.094 | 41.094 | 423.0996 | 311.4634 | 157.022 | 157.0499 |  | 88.5433
0 | [6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 311.4254 | 423.0996 | 41.094 | 41.094 | 423.0996 | 311.3873 | 157.088 | 157.1261 |  | 88.6094
0 | LS1 | CB06 | SOL2 | 1574 | LS1_CB06:SOL2_D1574 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 311.1373 | 423.0996 | 41.094 | 41.094 | 423.0996 | 310.8873 | 157.3761 | 157.6261 |  | 88.8974
0 | [6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 310.8492 | 423.0996 | 41.094 | 41.094 | 423.0996 | 310.8112 | 157.6641 | 157.7022 |  | 89.1855
0 | [6 empty] | bellow |  |  |  | 0.099517 | 310.7614 | 423.0996 | 41.094 | 41.094 | 423.0996 | 310.7116 | 157.7519 | 157.8017 |  | 89.2733
0 | LS1 | CB06 | CAV5 | 1580 | LS1_CB06:CAV5_D1580 |  | b08 resonator | 34 |  |  | 0.29817 | 310.5626 | 423.0996 | 41.094 | 41.094 | 423.0996 | 310.4135 | 157.9508 | 158.0999 |  | 89.4722
0 | [10 empty] | 0 | 310.4135 | 423.0996 | 41.094 | 41.094 | 423.0996 | 310.4135 | 158.0999 | 158.0999 |  | 89.6212
0 | [10 empty] | 0 | 310.4135 | 423.0996 | 41.094 | 41.094 | 423.0996 | 310.4135 | 158.0999 | 158.0999 |  | 89.6212
0 | [6 empty] | bellow |  |  |  | 0.099517 | 310.3637 | 423.0996 | 41.094 | 41.094 | 423.0996 | 310.3139 | 158.1496 | 158.1994 |  | 89.671
0 | LS1 | CB06 | CAV6 | 1583 | LS1_CB06:CAV6_D1583 |  | b08 resonator | 34 |  |  | 0.29817 | 310.1649 | 423.0996 | 41.094 | 41.094 | 423.0996 | 310.0158 | 158.3485 | 158.4976 |  | 89.8698
0 | [10 empty] | 0 | 310.0158 | 423.0996 | 41.094 | 41.094 | 423.0996 | 310.0158 | 158.4976 | 158.4976 |  | 90.0189
0 | [10 empty] | 0 | 310.0158 | 423.0996 | 41.094 | 41.094 | 423.0996 | 310.0158 | 158.4976 | 158.4976 |  | 90.0189
0 | [6 empty] | bellow |  |  |  | 0.099517 | 309.966 | 423.0996 | 41.094 | 41.094 | 423.0996 | 309.9163 | 158.5473 | 158.5971 |  | 90.0687
0 | LS1 | CB06 | CAV7 | 1587 | LS1_CB06:CAV7_D1587 |  | b08 resonator | 34 |  |  | 0.29817 | 309.7672 | 423.0996 | 41.094 | 41.094 | 423.0996 | 309.6181 | 158.7462 | 158.8953 |  | 90.2675
0 | [10 empty] | 0 | 309.6181 | 423.0996 | 41.094 | 41.094 | 423.0996 | 309.6181 | 158.8953 | 158.8953 |  | 90.4166
0 | LS1 | CB06 | BLM | 1589 | LS1_CB06:BLM_D1589 |  | outside loss monitor |  |  |  | 0 | 309.6181 | 423.0996 | 41.094 | 41.094 | 423.0996 | 309.6181 | 158.8953 | 158.8953 |  | 90.4166
0 | [6 empty] | bellow |  |  |  | 0.09651 | 309.5698 | 423.0996 | 41.094 | 41.094 | 423.0996 | 309.5216 | 158.9435 | 158.9918 |  | 90.4649
0 | LS1 | CB06 | BPM | 1590 | LS1_CB06:BPM_D1590 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 309.5216 | 423.0996 | 41.094 | 41.094 | 423.0996 | 309.5216 | 158.9918 | 158.9918 |  | 90.5131
0 | [6 empty] | BPM-box |  |  |  | 0.05588 | 309.4936 | 423.0996 | 41.094 | 41.094 | 423.0996 | 309.4657 | 159.0197 | 159.0476 |  | 90.5411
0 | [6 empty] | solenoid-entry |  |  |  | 0.076136 | 309.4276 | 423.0996 | 41.094 | 41.094 | 423.0996 | 309.3896 | 159.0857 | 159.1238 |  | 90.6071
0 | LS1 | CB06 | SOL3 | 1594 | LS1_CB06:SOL3_D1594 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 309.1396 | 423.0996 | 41.094 | 41.094 | 423.0996 | 308.8896 | 159.3738 | 159.6238 |  | 90.8951
0 | [6 empty] | solenoid-exit |  |  |  | 0.076136 | 308.8515 | 423.0996 | 41.094 | 41.094 | 423.0996 | 308.8134 | 159.6618 | 159.6999 |  | 91.1832
0 | [6 empty] | bellow |  |  |  | 0.099517 | 308.7637 | 423.0996 | 41.094 | 41.094 | 423.0996 | 308.7139 | 159.7497 | 159.7994 |  | 91.271
0 | LS1 | CB06 | CAV8 | 1599 | LS1_CB06:CAV8_D1599 |  | b08 resonator | 34 |  |  | 0.29817 | 308.5648 | 423.0996 | 41.094 | 41.094 | 423.0996 | 308.4157 | 159.9485 | 160.0976 |  | 91.4699
0 | [10 empty] | 0 | 308.4157 | 423.0996 | 41.094 | 41.094 | 423.0996 | 308.4157 | 160.0976 | 160.0976 |  | 91.619
0 | [10 empty] | 0 | 308.4157 | 423.0996 | 41.094 | 41.094 | 423.0996 | 308.4157 | 160.0976 | 160.0976 |  | 91.619
0 | [6 empty] | bellows |  |  |  | 0.19082 | 308.3203 | 423.0996 | 41.094 | 41.094 | 423.0996 | 308.2249 | 160.193 | 160.2884 |  | 91.7144
0 | LS1 | CB06 | GV | 1603 | LS1_CB06:GV_D1603 |  | gate valve | 40 |  |  | 0.057404 | 308.1962 | 423.0996 | 41.094 | 41.094 | 423.0996 | 308.1675 | 160.3171 | 160.3458 |  | 91.8385
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 308.1158 | 423.0996 | 41.094 | 41.094 | 423.0996 | 308.064 | 160.3976 | 160.4493 |  | 91.9189
0 | LS1 | WB06 | BPM | 1604 | LS1_WB06:BPM_D1604 | BPM2 | position monitor |  |  |  | 0 | 308.064 | 423.0996 | 41.094 | 41.094 | 423.0996 | 308.064 | 160.4493 | 160.4493 |  | 91.9707
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 307.9914 | 423.0996 | 41.094 | 41.094 | 423.0996 | 307.9187 | 160.522 | 160.5946 |  | 92.0433
0 | LS1 | WB06 | HMR | 1606 | LS1_WB06:HMR_D1606 | HR_L | Halo ring, large aperture | 33 |  |  | 0 | 307.9187 | 423.0996 | 41.094 | 41.094 | 423.0996 | 307.9187 | 160.5946 | 160.5946 |  | 92.116
0 | [6 empty] | diagnostic box |  |  |  | 0 | 307.9187 | 423.0996 | 41.094 | 41.094 | 423.0996 | 307.9187 | 160.5946 | 160.5946 |  | 92.116
0 | LS1 | WB06 | IP | 1606 | LS1_WB06:IP_D1606 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 307.9187 | 423.0996 | 41.094 | 41.094 | 423.0996 | 307.9187 | 160.5946 | 160.5946 |  | 92.116
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 307.8531 | 423.0996 | 41.094 | 41.094 | 423.0996 | 307.7875 | 160.6602 | 160.7258 |  | 92.1816
0 | LS1 | CB07 | GV | 1608 | LS1_CB07:GV_D1608 |  | gate valve | 40 |  |  | 0.057404 | 307.7588 | 423.0996 | 41.094 | 41.094 | 423.0996 | 307.7301 | 160.7545 | 160.7832 |  | 92.2759
0 | [6 empty] | bellows |  |  |  | 0.19082 | 307.6347 | 423.0996 | 41.094 | 41.094 | 423.0996 | 307.5393 | 160.8786 | 160.9741 |  | 92.4
0 | LS1 | CB07 | CAV1 | 1611 | LS1_CB07:CAV1_D1611 |  | b08 resonator | 34 |  |  | 0.29817 | 307.3902 | 423.0996 | 41.094 | 41.094 | 423.0996 | 307.2411 | 161.1231 | 161.2722 |  | 92.6445
0 | [10 empty] | 0 | 307.2411 | 423.0996 | 41.094 | 41.094 | 423.0996 | 307.2411 | 161.2722 | 161.2722 |  | 92.7936
0 | LS1 | CB07 | BLM | 1613 | LS1_CB07:BLM_D1613 |  | outside loss monitor |  |  |  | 0 | 307.2411 | 423.0996 | 41.094 | 41.094 | 423.0996 | 307.2411 | 161.2722 | 161.2722 |  | 92.7936
0 | [6 empty] | bellow | 40 |  |  | 0.09651 | 307.1929 | 423.0996 | 41.094 | 41.094 | 423.0996 | 307.1446 | 161.3205 | 161.3687 |  | 92.8418
0 | LS1 | CB07 | BPM | 1614 | LS1_CB07:BPM_D1614 | BPM6 | position monitor, BPMc |  |  |  | 0 | 307.1446 | 423.0996 | 41.094 | 41.094 | 423.0996 | 307.1446 | 161.3687 | 161.3687 |  | 92.8901
0 | [6 empty] | BPM-box | 40 |  |  | 0.05588 | 307.1167 | 423.0996 | 41.094 | 41.094 | 423.0996 | 307.0887 | 161.3967 | 161.4246 |  | 92.918
0 | [6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 307.0507 | 423.0996 | 41.094 | 41.094 | 423.0996 | 307.0126 | 161.4627 | 161.5008 |  | 92.9841
0 | LS1 | CB07 | SOL1 | 1618 | LS1_CB07:SOL1_D1618 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 306.7626 | 423.0996 | 41.094 | 41.094 | 423.0996 | 306.5126 | 161.7508 | 162.0008 |  | 93.2721
0 | [6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 306.4745 | 423.0996 | 41.094 | 41.094 | 423.0996 | 306.4365 | 162.0388 | 162.0769 |  | 93.5602
0 | [6 empty] | bellow |  |  |  | 0.099517 | 306.3867 | 423.0996 | 41.094 | 41.094 | 423.0996 | 306.3369 | 162.1266 | 162.1764 |  | 93.648
0 | LS1 | CB07 | CAV2 | 1623 | LS1_CB07:CAV2_D1623 |  | b08 resonator | 34 |  |  | 0.29817 | 306.1879 | 423.0996 | 41.094 | 41.094 | 423.0996 | 306.0388 | 162.3255 | 162.4746 |  | 93.8469
0 | [10 empty] | 0 | 306.0388 | 423.0996 | 41.094 | 41.094 | 423.0996 | 306.0388 | 162.4746 | 162.4746 |  | 93.9959
0 | [10 empty] | 0 | 306.0388 | 423.0996 | 41.094 | 41.094 | 423.0996 | 306.0388 | 162.4746 | 162.4746 |  | 93.9959
0 | [6 empty] | bellow |  |  |  | 0.099517 | 305.989 | 423.0996 | 41.094 | 41.094 | 423.0996 | 305.9392 | 162.5243 | 162.5741 |  | 94.0457
0 | LS1 | CB07 | CAV3 | 1627 | LS1_CB07:CAV3_D1627 |  | b08 resonator | 34 |  |  | 0.29817 | 305.7902 | 423.0996 | 41.094 | 41.094 | 423.0996 | 305.6411 | 162.7232 | 162.8723 |  | 94.2445
0 | [10 empty] | 0 | 305.6411 | 423.0996 | 41.094 | 41.094 | 423.0996 | 305.6411 | 162.8723 | 162.8723 |  | 94.3936
0 | [10 empty] | 0 | 305.6411 | 423.0996 | 41.094 | 41.094 | 423.0996 | 305.6411 | 162.8723 | 162.8723 |  | 94.3936
0 | [6 empty] | bellow |  |  |  | 0.099517 | 305.5913 | 423.0996 | 41.094 | 41.094 | 423.0996 | 305.5416 | 162.922 | 162.9718 |  | 94.4434
0 | LS1 | CB07 | CAV4 | 1631 | LS1_CB07:CAV4_D1631 |  | b08 resonator | 34 |  |  | 0.29817 | 305.3925 | 423.0996 | 41.094 | 41.094 | 423.0996 | 305.2434 | 163.1209 | 163.27 |  | 94.6422
0 | [10 empty] | 0 | 305.2434 | 423.0996 | 41.094 | 41.094 | 423.0996 | 305.2434 | 163.27 | 163.27 |  | 94.7913
0 | LS1 | CB07 | BLM | 1633 | LS1_CB07:BLM_D1633 |  | outside loss monitor |  |  |  | 0 | 305.2434 | 423.0996 | 41.094 | 41.094 | 423.0996 | 305.2434 | 163.27 | 163.27 |  | 94.7913
0 | [6 empty] | bellow |  |  |  | 0.09651 | 305.1951 | 423.0996 | 41.094 | 41.094 | 423.0996 | 305.1469 | 163.3182 | 163.3665 |  | 94.8396
0 | LS1 | CB07 | BPM | 1634 | LS1_CB07:BPM_D1634 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 305.1469 | 423.0996 | 41.094 | 41.094 | 423.0996 | 305.1469 | 163.3665 | 163.3665 |  | 94.8878
0 | [6 empty] | BPM-box |  |  |  | 0.05588 | 305.1189 | 423.0996 | 41.094 | 41.094 | 423.0996 | 305.091 | 163.3944 | 163.4223 |  | 94.9158
0 | [6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 305.0529 | 423.0996 | 41.094 | 41.094 | 423.0996 | 305.0149 | 163.4604 | 163.4985 |  | 94.9818
0 | LS1 | CB07 | SOL2 | 1637 | LS1_CB07:SOL2_D1637 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 304.7649 | 423.0996 | 41.094 | 41.094 | 423.0996 | 304.5149 | 163.7485 | 163.9985 |  | 95.2698
0 | [6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 304.4768 | 423.0996 | 41.094 | 41.094 | 423.0996 | 304.4387 | 164.0365 | 164.0746 |  | 95.5579
0 | [6 empty] | bellow |  |  |  | 0.099517 | 304.389 | 423.0996 | 41.094 | 41.094 | 423.0996 | 304.3392 | 164.1244 | 164.1741 |  | 95.6457
0 | LS1 | CB07 | CAV5 | 1643 | LS1_CB07:CAV5_D1643 |  | b08 resonator | 34 |  |  | 0.29817 | 304.1901 | 423.0996 | 41.094 | 41.094 | 423.0996 | 304.041 | 164.3232 | 164.4723 |  | 95.8446
0 | [10 empty] | 0 | 304.041 | 423.0996 | 41.094 | 41.094 | 423.0996 | 304.041 | 164.4723 | 164.4723 |  | 95.9937
0 | [10 empty] | 0 | 304.041 | 423.0996 | 41.094 | 41.094 | 423.0996 | 304.041 | 164.4723 | 164.4723 |  | 95.9937
0 | [6 empty] | bellow |  |  |  | 0.099517 | 303.9913 | 423.0996 | 41.094 | 41.094 | 423.0996 | 303.9415 | 164.5221 | 164.5718 |  | 96.0434
0 | LS1 | CB07 | CAV6 | 1647 | LS1_CB07:CAV6_D1647 |  | b08 resonator | 34 |  |  | 0.29817 | 303.7924 | 423.0996 | 41.094 | 41.094 | 423.0996 | 303.6434 | 164.7209 | 164.87 |  | 96.2423
0 | [10 empty] | 0 | 303.6434 | 423.0996 | 41.094 | 41.094 | 423.0996 | 303.6434 | 164.87 | 164.87 |  | 96.3914
0 | [10 empty] | 0 | 303.6434 | 423.0996 | 41.094 | 41.094 | 423.0996 | 303.6434 | 164.87 | 164.87 |  | 96.3914
0 | [6 empty] | bellow |  |  |  | 0.099517 | 303.5936 | 423.0996 | 41.094 | 41.094 | 423.0996 | 303.5438 | 164.9197 | 164.9695 |  | 96.4411
0 | LS1 | CB07 | CAV7 | 1651 | LS1_CB07:CAV7_D1651 |  | b08 resonator | 34 |  |  | 0.29817 | 303.3948 | 423.0996 | 41.094 | 41.094 | 423.0996 | 303.2457 | 165.1186 | 165.2677 |  | 96.64
0 | [10 empty] | 0 | 303.2457 | 423.0996 | 41.094 | 41.094 | 423.0996 | 303.2457 | 165.2677 | 165.2677 |  | 96.789
0 | LS1 | CB07 | BLM | 1653 | LS1_CB07:BLM_D1653 |  | outside loss monitor |  |  |  | 0 | 303.2457 | 423.0996 | 41.094 | 41.094 | 423.0996 | 303.2457 | 165.2677 | 165.2677 |  | 96.789
0 | [6 empty] | bellow |  |  |  | 0.09651 | 303.1974 | 423.0996 | 41.094 | 41.094 | 423.0996 | 303.1492 | 165.3159 | 165.3642 |  | 96.8373
0 | LS1 | CB07 | BPM | 1654 | LS1_CB07:BPM_D1654 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 303.1492 | 423.0996 | 41.094 | 41.094 | 423.0996 | 303.1492 | 165.3642 | 165.3642 |  | 96.8856
0 | [6 empty] | BPM-box |  |  |  | 0.05588 | 303.1212 | 423.0996 | 41.094 | 41.094 | 423.0996 | 303.0933 | 165.3921 | 165.4201 |  | 96.9135
0 | [6 empty] | solenoid-entry |  |  |  | 0.076136 | 303.0552 | 423.0996 | 41.094 | 41.094 | 423.0996 | 303.0171 | 165.4581 | 165.4962 |  | 96.9795
0 | LS1 | CB07 | SOL3 | 1657 | LS1_CB07:SOL3_D1657 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 302.7671 | 423.0996 | 41.094 | 41.094 | 423.0996 | 302.5171 | 165.7462 | 165.9962 |  | 97.2676
0 | [6 empty] | solenoid-exit |  |  |  | 0.076136 | 302.4791 | 423.0996 | 41.094 | 41.094 | 423.0996 | 302.441 | 166.0343 | 166.0723 |  | 97.5556
0 | [6 empty] | bellow |  |  |  | 0.099517 | 302.3912 | 423.0996 | 41.094 | 41.094 | 423.0996 | 302.3415 | 166.1221 | 166.1719 |  | 97.6435
0 | LS1 | CB07 | CAV8 | 1663 | LS1_CB07:CAV8_D1663 |  | b08 resonator | 34 |  |  | 0.29817 | 302.1924 | 423.0996 | 41.094 | 41.094 | 423.0996 | 302.0433 | 166.3209 | 166.47 |  | 97.8423
0 | [10 empty] | 0 | 302.0433 | 423.0996 | 41.094 | 41.094 | 423.0996 | 302.0433 | 166.47 | 166.47 |  | 97.9914
0 | [10 empty] | 0 | 302.0433 | 423.0996 | 41.094 | 41.094 | 423.0996 | 302.0433 | 166.47 | 166.47 |  | 97.9914
0 | [6 empty] | bellows |  |  |  | 0.19082 | 301.9479 | 423.0996 | 41.094 | 41.094 | 423.0996 | 301.8525 | 166.5654 | 166.6608 |  | 98.0868
0 | LS1 | CB07 | GV | 1667 | LS1_CB07:GV_D1667 |  | gate valve | 40 |  |  | 0.057404 | 301.8238 | 423.0996 | 41.094 | 41.094 | 423.0996 | 301.7951 | 166.6896 | 166.7183 |  | 98.2109
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 301.7433 | 423.0996 | 41.094 | 41.094 | 423.0996 | 301.6916 | 166.77 | 166.8218 |  | 98.2914
0 | LS1 | WB07 | BPM | 1668 | LS1_WB07:BPM_D1668 | BPM2 | position monitor |  |  |  | 0 | 301.6916 | 423.0996 | 41.094 | 41.094 | 423.0996 | 301.6916 | 166.8218 | 166.8218 |  | 98.3431
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 301.6189 | 423.0996 | 41.094 | 41.094 | 423.0996 | 301.5463 | 166.8944 | 166.967 |  | 98.4158
0 | LS1 | WB07 | HMR | 1670 | LS1_WB07:HMR_D1670 | HR_L | Halo ring, large aperture | 33 |  |  | 0 | 301.5463 | 423.0996 | 41.094 | 41.094 | 423.0996 | 301.5463 | 166.967 | 166.967 |  | 98.4884
0 | [6 empty] | diagnostic box |  |  |  | 0 | 301.5463 | 423.0996 | 41.094 | 41.094 | 423.0996 | 301.5463 | 166.967 | 166.967 |  | 98.4884
0 | LS1 | WB07 | IP | 1670 | LS1_WB07:IP_D1670 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 301.5463 | 423.0996 | 41.094 | 41.094 | 423.0996 | 301.5463 | 166.967 | 166.967 |  | 98.4884
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 301.4807 | 423.0996 | 41.094 | 41.094 | 423.0996 | 301.4151 | 167.0326 | 167.0983 |  | 98.554
0 | LS1 | CB08 | GV | 1671 | LS1_CB08:GV_D1671 |  | gate valve | 40 |  |  | 0.057404 | 301.3864 | 423.0996 | 41.094 | 41.094 | 423.0996 | 301.3577 | 167.127 | 167.1557 |  | 98.6483
0 | [6 empty] | bellows |  |  |  | 0.19082 | 301.2623 | 423.0996 | 41.094 | 41.094 | 423.0996 | 301.1669 | 167.2511 | 167.3465 |  | 98.7724
0 | LS1 | CB08 | CAV1 | 1675 | LS1_CB08:CAV1_D1675 |  | b08 resonator | 34 |  |  | 0.29817 | 301.0178 | 423.0996 | 41.094 | 41.094 | 423.0996 | 300.8687 | 167.4956 | 167.6446 |  | 99.0169
0 | [10 empty] | 0 | 300.8687 | 423.0996 | 41.094 | 41.094 | 423.0996 | 300.8687 | 167.6446 | 167.6446 |  | 99.166
0 | LS1 | CB08 | BLM | 1676 | LS1_CB08:BLM_D1676 |  | outside loss monitor |  |  |  | 0 | 300.8687 | 423.0996 | 41.094 | 41.094 | 423.0996 | 300.8687 | 167.6446 | 167.6446 |  | 99.166
0 | [6 empty] | bellow | 40 |  |  | 0.09651 | 300.8204 | 423.0996 | 41.094 | 41.094 | 423.0996 | 300.7722 | 167.6929 | 167.7412 |  | 99.2143
0 | LS1 | CB08 | BPM | 1677 | LS1_CB08:BPM_D1677 | BPM6 | position monitor, BPMc |  |  |  | 0 | 300.7722 | 423.0996 | 41.094 | 41.094 | 423.0996 | 300.7722 | 167.7412 | 167.7412 |  | 99.2625
0 | [6 empty] | BPM-box | 40 |  |  | 0.05588 | 300.7442 | 423.0996 | 41.094 | 41.094 | 423.0996 | 300.7163 | 167.7691 | 167.797 |  | 99.2905
0 | [6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 300.6782 | 423.0996 | 41.094 | 41.094 | 423.0996 | 300.6402 | 167.8351 | 167.8732 |  | 99.3565
0 | LS1 | CB08 | SOL1 | 1681 | LS1_CB08:SOL1_D1681 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 300.3902 | 423.0996 | 41.094 | 41.094 | 423.0996 | 300.1402 | 168.1232 | 168.3732 |  | 99.6445
0 | [6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 300.1021 | 423.0996 | 41.094 | 41.094 | 423.0996 | 300.064 | 168.4112 | 168.4493 |  | 99.9326
0 | [6 empty] | bellow |  |  |  | 0.099517 | 300.0143 | 423.0996 | 41.094 | 41.094 | 423.0996 | 299.9645 | 168.4991 | 168.5488 |  | 100.0204
0 | LS1 | CB08 | CAV2 | 1687 | LS1_CB08:CAV2_D1687 |  | b08 resonator | 34 |  |  | 0.29817 | 299.8154 | 423.0996 | 41.094 | 41.094 | 423.0996 | 299.6663 | 168.6979 | 168.847 |  | 100.2193
0 | [10 empty] | 0 | 299.6663 | 423.0996 | 41.094 | 41.094 | 423.0996 | 299.6663 | 168.847 | 168.847 |  | 100.3684
0 | [10 empty] | 0 | 299.6663 | 423.0996 | 41.094 | 41.094 | 423.0996 | 299.6663 | 168.847 | 168.847 |  | 100.3684
0 | [6 empty] | bellow |  |  |  | 0.099517 | 299.6166 | 423.0996 | 41.094 | 41.094 | 423.0996 | 299.5668 | 168.8968 | 168.9465 |  | 100.4181
0 | LS1 | CB08 | CAV3 | 1691 | LS1_CB08:CAV3_D1691 |  | b08 resonator | 34 |  |  | 0.29817 | 299.4177 | 423.0996 | 41.094 | 41.094 | 423.0996 | 299.2687 | 169.0956 | 169.2447 |  | 100.617
0 | [10 empty] | 0 | 299.2687 | 423.0996 | 41.094 | 41.094 | 423.0996 | 299.2687 | 169.2447 | 169.2447 |  | 100.7661
0 | [10 empty] | 0 | 299.2687 | 423.0996 | 41.094 | 41.094 | 423.0996 | 299.2687 | 169.2447 | 169.2447 |  | 100.7661
0 | [6 empty] | bellow |  |  |  | 0.099517 | 299.2189 | 423.0996 | 41.094 | 41.094 | 423.0996 | 299.1691 | 169.2944 | 169.3442 |  | 100.8158
0 | LS1 | CB08 | CAV4 | 1695 | LS1_CB08:CAV4_D1695 |  | b08 resonator | 34 |  |  | 0.29817 | 299.0201 | 423.0996 | 41.094 | 41.094 | 423.0996 | 298.871 | 169.4933 | 169.6424 |  | 101.0147
0 | [10 empty] | 0 | 298.871 | 423.0996 | 41.094 | 41.094 | 423.0996 | 298.871 | 169.6424 | 169.6424 |  | 101.1637
0 | LS1 | CB08 | BLM | 1696 | LS1_CB08:BLM_D1696 |  | outside loss monitor |  |  |  | 0 | 298.871 | 423.0996 | 41.094 | 41.094 | 423.0996 | 298.871 | 169.6424 | 169.6424 |  | 101.1637
0 | [6 empty] | bellow |  |  |  | 0.09651 | 298.8227 | 423.0996 | 41.094 | 41.094 | 423.0996 | 298.7745 | 169.6906 | 169.7389 |  | 101.212
0 | LS1 | CB08 | BPM | 1697 | LS1_CB08:BPM_D1697 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 298.7745 | 423.0996 | 41.094 | 41.094 | 423.0996 | 298.7745 | 169.7389 | 169.7389 |  | 101.2603
0 | [6 empty] | BPM-box |  |  |  | 0.05588 | 298.7465 | 423.0996 | 41.094 | 41.094 | 423.0996 | 298.7186 | 169.7668 | 169.7948 |  | 101.2882
0 | [6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 298.6805 | 423.0996 | 41.094 | 41.094 | 423.0996 | 298.6424 | 169.8328 | 169.8709 |  | 101.3542
0 | LS1 | CB08 | SOL2 | 1701 | LS1_CB08:SOL2_D1701 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 298.3924 | 423.0996 | 41.094 | 41.094 | 423.0996 | 298.1424 | 170.1209 | 170.3709 |  | 101.6423
0 | [6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 298.1044 | 423.0996 | 41.094 | 41.094 | 423.0996 | 298.0663 | 170.409 | 170.447 |  | 101.9303
0 | [6 empty] | bellow |  |  |  | 0.099517 | 298.0165 | 423.0996 | 41.094 | 41.094 | 423.0996 | 297.9668 | 170.4968 | 170.5466 |  | 102.0182
0 | LS1 | CB08 | CAV5 | 1707 | LS1_CB08:CAV5_D1707 |  | b08 resonator | 34 |  |  | 0.29817 | 297.8177 | 423.0996 | 41.094 | 41.094 | 423.0996 | 297.6686 | 170.6956 | 170.8447 |  | 102.217
0 | [10 empty] | 0 | 297.6686 | 423.0996 | 41.094 | 41.094 | 423.0996 | 297.6686 | 170.8447 | 170.8447 |  | 102.3661
0 | [10 empty] | 0 | 297.6686 | 423.0996 | 41.094 | 41.094 | 423.0996 | 297.6686 | 170.8447 | 170.8447 |  | 102.3661
0 | [6 empty] | bellow |  |  |  | 0.099517 | 297.6189 | 423.0996 | 41.094 | 41.094 | 423.0996 | 297.5691 | 170.8945 | 170.9442 |  | 102.4159
0 | LS1 | CB08 | CAV6 | 1711 | LS1_CB08:CAV6_D1711 |  | b08 resonator | 34 |  |  | 0.29817 | 297.42 | 423.0996 | 41.094 | 41.094 | 423.0996 | 297.2709 | 171.0933 | 171.2424 |  | 102.6147
0 | [10 empty] | 0 | 297.2709 | 423.0996 | 41.094 | 41.094 | 423.0996 | 297.2709 | 171.2424 | 171.2424 |  | 102.7638
0 | [10 empty] | 0 | 297.2709 | 423.0996 | 41.094 | 41.094 | 423.0996 | 297.2709 | 171.2424 | 171.2424 |  | 102.7638
0 | [6 empty] | bellow |  |  |  | 0.099517 | 297.2212 | 423.0996 | 41.094 | 41.094 | 423.0996 | 297.1714 | 171.2922 | 171.3419 |  | 102.8135
0 | LS1 | CB08 | CAV7 | 1715 | LS1_CB08:CAV7_D1715 |  | b08 resonator | 34 |  |  | 0.29817 | 297.0223 | 423.0996 | 41.094 | 41.094 | 423.0996 | 296.8732 | 171.491 | 171.6401 |  | 103.0124
0 | [10 empty] | 0 | 296.8732 | 423.0996 | 41.094 | 41.094 | 423.0996 | 296.8732 | 171.6401 | 171.6401 |  | 103.1615
0 | LS1 | CB08 | BLM | 1716 | LS1_CB08:BLM_D1716 |  | outside loss monitor |  |  |  | 0 | 296.8732 | 423.0996 | 41.094 | 41.094 | 423.0996 | 296.8732 | 171.6401 | 171.6401 |  | 103.1615
0 | [6 empty] | bellow |  |  |  | 0.09651 | 296.825 | 423.0996 | 41.094 | 41.094 | 423.0996 | 296.7767 | 171.6884 | 171.7366 |  | 103.2097
0 | LS1 | CB08 | BPM | 1717 | LS1_CB08:BPM_D1717 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 296.7767 | 423.0996 | 41.094 | 41.094 | 423.0996 | 296.7767 | 171.7366 | 171.7366 |  | 103.258
0 | [6 empty] | BPM-box |  |  |  | 0.05588 | 296.7488 | 423.0996 | 41.094 | 41.094 | 423.0996 | 296.7209 | 171.7646 | 171.7925 |  | 103.2859
0 | [6 empty] | solenoid-entry |  |  |  | 0.076136 | 296.6828 | 423.0996 | 41.094 | 41.094 | 423.0996 | 296.6447 | 171.8306 | 171.8686 |  | 103.3519
0 | LS1 | CB08 | SOL3 | 1721 | LS1_CB08:SOL3_D1721 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 296.3947 | 423.0996 | 41.094 | 41.094 | 423.0996 | 296.1447 | 172.1186 | 172.3686 |  | 103.64
0 | [6 empty] | solenoid-exit |  |  |  | 0.076136 | 296.1066 | 423.0996 | 41.094 | 41.094 | 423.0996 | 296.0686 | 172.4067 | 172.4448 |  | 103.9281
0 | [6 empty] | bellow |  |  |  | 0.099517 | 296.0188 | 423.0996 | 41.094 | 41.094 | 423.0996 | 295.9691 | 172.4945 | 172.5443 |  | 104.0159
0 | LS1 | CB08 | CAV8 | 1727 | LS1_CB08:CAV8_D1727 |  | b08 resonator | 34 |  |  | 0.29817 | 295.82 | 423.0996 | 41.094 | 41.094 | 423.0996 | 295.6709 | 172.6934 | 172.8425 |  | 104.2147
0 | [10 empty] | 0 | 295.6709 | 423.0996 | 41.094 | 41.094 | 423.0996 | 295.6709 | 172.8425 | 172.8425 |  | 104.3638
0 | [10 empty] | 0 | 295.6709 | 423.0996 | 41.094 | 41.094 | 423.0996 | 295.6709 | 172.8425 | 172.8425 |  | 104.3638
0 | [6 empty] | bellows |  |  |  | 0.19082 | 295.5755 | 423.0996 | 41.094 | 41.094 | 423.0996 | 295.4801 | 172.9379 | 173.0333 |  | 104.4592
0 | LS1 | CB08 | GV | 1731 | LS1_CB08:GV_D1731 |  | gate valve | 40 |  |  | 0.057404 | 295.4514 | 423.0996 | 41.094 | 41.094 | 423.0996 | 295.4227 | 173.062 | 173.0907 |  | 104.5833
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 295.3709 | 423.0996 | 41.094 | 41.094 | 423.0996 | 295.3192 | 173.1424 | 173.1942 |  | 104.6638
0 | LS1 | WB08 | BPM | 1732 | LS1_WB08:BPM_D1732 | BPM2 | position monitor |  |  |  | 0 | 295.3192 | 423.0996 | 41.094 | 41.094 | 423.0996 | 295.3192 | 173.1942 | 173.1942 |  | 104.7155
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 295.2465 | 423.0996 | 41.094 | 41.094 | 423.0996 | 295.1739 | 173.2668 | 173.3395 |  | 104.7882
0 | LS1 | WB08 | HMR | 1733 | LS1_WB08:HMR_D1733 | HR_L | Halo ring, large aperture | 33 |  |  | 0 | 295.1739 | 423.0996 | 41.094 | 41.094 | 423.0996 | 295.1739 | 173.3395 | 173.3395 |  | 104.8608
0 | [6 empty] | diagnostic box |  |  |  | 0 | 295.1739 | 423.0996 | 41.094 | 41.094 | 423.0996 | 295.1739 | 173.3395 | 173.3395 |  | 104.8608
0 | LS1 | WB08 | IP | 1733 | LS1_WB08:IP_D1733 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 295.1739 | 423.0996 | 41.094 | 41.094 | 423.0996 | 295.1739 | 173.3395 | 173.3395 |  | 104.8608
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 295.1083 | 423.0996 | 41.094 | 41.094 | 423.0996 | 295.0427 | 173.4051 | 173.4707 |  | 104.9264
0 | LS1 | CB09 | GV | 1735 | LS1_CB09:GV_D1735 |  | gate valve | 40 |  |  | 0.057404 | 295.014 | 423.0996 | 41.094 | 41.094 | 423.0996 | 294.9853 | 173.4994 | 173.5281 |  | 105.0207
0 | [6 empty] | bellows |  |  |  | 0.19082 | 294.8898 | 423.0996 | 41.094 | 41.094 | 423.0996 | 294.7944 | 173.6235 | 173.7189 |  | 105.1449
0 | LS1 | CB09 | CAV1 | 1739 | LS1_CB09:CAV1_D1739 |  | b08 resonator | 34 |  |  | 0.29817 | 294.6454 | 423.0996 | 41.094 | 41.094 | 423.0996 | 294.4963 | 173.868 | 174.0171 |  | 105.3894
0 | [10 empty] | 0 | 294.4963 | 423.0996 | 41.094 | 41.094 | 423.0996 | 294.4963 | 174.0171 | 174.0171 |  | 105.5384
0 | LS1 | CB09 | BLM | 1740 | LS1_CB09:BLM_D1740 |  | outside loss monitor |  |  |  | 0 | 294.4963 | 423.0996 | 41.094 | 41.094 | 423.0996 | 294.4963 | 174.0171 | 174.0171 |  | 105.5384
0 | [6 empty] | bellow | 40 |  |  | 0.09651 | 294.448 | 423.0996 | 41.094 | 41.094 | 423.0996 | 294.3998 | 174.0653 | 174.1136 |  | 105.5867
0 | LS1 | CB09 | BPM | 1741 | LS1_CB09:BPM_D1741 | BPM6 | position monitor, BPMc |  |  |  | 0 | 294.3998 | 423.0996 | 41.094 | 41.094 | 423.0996 | 294.3998 | 174.1136 | 174.1136 |  | 105.635
0 | [6 empty] | BPM-box | 40 |  |  | 0.05588 | 294.3718 | 423.0996 | 41.094 | 41.094 | 423.0996 | 294.3439 | 174.1415 | 174.1695 |  | 105.6629
0 | [6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 294.3058 | 423.0996 | 41.094 | 41.094 | 423.0996 | 294.2677 | 174.2075 | 174.2456 |  | 105.7289
0 | LS1 | CB09 | SOL1 | 1745 | LS1_CB09:SOL1_D1745 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 294.0177 | 423.0996 | 41.094 | 41.094 | 423.0996 | 293.7677 | 174.4956 | 174.7456 |  | 106.017
0 | [6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 293.7297 | 423.0996 | 41.094 | 41.094 | 423.0996 | 293.6916 | 174.7837 | 174.8217 |  | 106.305
0 | [6 empty] | bellow |  |  |  | 0.099517 | 293.6418 | 423.0996 | 41.094 | 41.094 | 423.0996 | 293.5921 | 174.8715 | 174.9213 |  | 106.3929
0 | LS1 | CB09 | CAV2 | 1751 | LS1_CB09:CAV2_D1751 |  | b08 resonator | 34 |  |  | 0.29817 | 293.443 | 423.0996 | 41.094 | 41.094 | 423.0996 | 293.2939 | 175.0703 | 175.2194 |  | 106.5917
0 | [10 empty] | 0 | 293.2939 | 423.0996 | 41.094 | 41.094 | 423.0996 | 293.2939 | 175.2194 | 175.2194 |  | 106.7408
0 | [10 empty] | 0 | 293.2939 | 423.0996 | 41.094 | 41.094 | 423.0996 | 293.2939 | 175.2194 | 175.2194 |  | 106.7408
0 | [6 empty] | bellow |  |  |  | 0.099517 | 293.2442 | 423.0996 | 41.094 | 41.094 | 423.0996 | 293.1944 | 175.2692 | 175.3189 |  | 106.7906
0 | LS1 | CB09 | CAV3 | 1755 | LS1_CB09:CAV3_D1755 |  | b08 resonator | 34 |  |  | 0.29817 | 293.0453 | 423.0996 | 41.094 | 41.094 | 423.0996 | 292.8962 | 175.468 | 175.6171 |  | 106.9894
0 | [10 empty] | 0 | 292.8962 | 423.0996 | 41.094 | 41.094 | 423.0996 | 292.8962 | 175.6171 | 175.6171 |  | 107.1385
0 | [10 empty] | 0 | 292.8962 | 423.0996 | 41.094 | 41.094 | 423.0996 | 292.8962 | 175.6171 | 175.6171 |  | 107.1385
0 | [6 empty] | bellow |  |  |  | 0.099517 | 292.8465 | 423.0996 | 41.094 | 41.094 | 423.0996 | 292.7967 | 175.6669 | 175.7166 |  | 107.1882
0 | LS1 | CB09 | CAV4 | 1759 | LS1_CB09:CAV4_D1759 |  | b08 resonator | 34 |  |  | 0.29817 | 292.6476 | 423.0996 | 41.094 | 41.094 | 423.0996 | 292.4985 | 175.8657 | 176.0148 |  | 107.3871
0 | [10 empty] | 0 | 292.4985 | 423.0996 | 41.094 | 41.094 | 423.0996 | 292.4985 | 176.0148 | 176.0148 |  | 107.5362
0 | LS1 | CB09 | BLM | 1760 | LS1_CB09:BLM_D1760 |  | outside loss monitor |  |  |  | 0 | 292.4985 | 423.0996 | 41.094 | 41.094 | 423.0996 | 292.4985 | 176.0148 | 176.0148 |  | 107.5362
0 | [6 empty] | bellow |  |  |  | 0.09651 | 292.4503 | 423.0996 | 41.094 | 41.094 | 423.0996 | 292.402 | 176.0631 | 176.1113 |  | 107.5844
0 | LS1 | CB09 | BPM | 1761 | LS1_CB09:BPM_D1761 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 292.402 | 423.0996 | 41.094 | 41.094 | 423.0996 | 292.402 | 176.1113 | 176.1113 |  | 107.6327
0 | [6 empty] | BPM-box |  |  |  | 0.05588 | 292.3741 | 423.0996 | 41.094 | 41.094 | 423.0996 | 292.3462 | 176.1392 | 176.1672 |  | 107.6606
0 | [6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 292.3081 | 423.0996 | 41.094 | 41.094 | 423.0996 | 292.27 | 176.2053 | 176.2433 |  | 107.7266
0 | LS1 | CB09 | SOL2 | 1765 | LS1_CB09:SOL2_D1765 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 292.02 | 423.0996 | 41.094 | 41.094 | 423.0996 | 291.77 | 176.4933 | 176.7433 |  | 108.0147
0 | [6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 291.7319 | 423.0996 | 41.094 | 41.094 | 423.0996 | 291.6939 | 176.7814 | 176.8195 |  | 108.3028
0 | [6 empty] | bellow |  |  |  | 0.099517 | 291.6441 | 423.0996 | 41.094 | 41.094 | 423.0996 | 291.5944 | 176.8692 | 176.919 |  | 108.3906
0 | LS1 | CB09 | CAV5 | 1771 | LS1_CB09:CAV5_D1771 |  | b08 resonator | 34 |  |  | 0.29817 | 291.4453 | 423.0996 | 41.094 | 41.094 | 423.0996 | 291.2962 | 177.0681 | 177.2171 |  | 108.5894
0 | [10 empty] | 0 | 291.2962 | 423.0996 | 41.094 | 41.094 | 423.0996 | 291.2962 | 177.2171 | 177.2171 |  | 108.7385
0 | [10 empty] | 0 | 291.2962 | 423.0996 | 41.094 | 41.094 | 423.0996 | 291.2962 | 177.2171 | 177.2171 |  | 108.7385
0 | [6 empty] | bellow |  |  |  | 0.099517 | 291.2464 | 423.0996 | 41.094 | 41.094 | 423.0996 | 291.1967 | 177.2669 | 177.3167 |  | 108.7883
0 | LS1 | CB09 | CAV6 | 1775 | LS1_CB09:CAV6_D1775 |  | b08 resonator | 34 |  |  | 0.29817 | 291.0476 | 423.0996 | 41.094 | 41.094 | 423.0996 | 290.8985 | 177.4658 | 177.6148 |  | 108.9871
0 | [10 empty] | 0 | 290.8985 | 423.0996 | 41.094 | 41.094 | 423.0996 | 290.8985 | 177.6148 | 177.6148 |  | 109.1362
0 | [10 empty] | 0 | 290.8985 | 423.0996 | 41.094 | 41.094 | 423.0996 | 290.8985 | 177.6148 | 177.6148 |  | 109.1362
0 | [6 empty] | bellow |  |  |  | 0.099517 | 290.8487 | 423.0996 | 41.094 | 41.094 | 423.0996 | 290.799 | 177.6646 | 177.7144 |  | 109.186
0 | LS1 | CB09 | CAV7 | 1779 | LS1_CB09:CAV7_D1779 |  | b08 resonator | 34 |  |  | 0.29817 | 290.6499 | 423.0996 | 41.094 | 41.094 | 423.0996 | 290.5008 | 177.8634 | 178.0125 |  | 109.3848
0 | [10 empty] | 0 | 290.5008 | 423.0996 | 41.094 | 41.094 | 423.0996 | 290.5008 | 178.0125 | 178.0125 |  | 109.5339
0 | LS1 | CB09 | BLM | 1780 | LS1_CB09:BLM_D1780 |  | outside loss monitor |  |  |  | 0 | 290.5008 | 423.0996 | 41.094 | 41.094 | 423.0996 | 290.5008 | 178.0125 | 178.0125 |  | 109.5339
0 | [6 empty] | bellow |  |  |  | 0.09651 | 290.4526 | 423.0996 | 41.094 | 41.094 | 423.0996 | 290.4043 | 178.0608 | 178.109 |  | 109.5822
0 | LS1 | CB09 | BPM | 1781 | LS1_CB09:BPM_D1781 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 290.4043 | 423.0996 | 41.094 | 41.094 | 423.0996 | 290.4043 | 178.109 | 178.109 |  | 109.6304
0 | [6 empty] | BPM-box |  |  |  | 0.05588 | 290.3764 | 423.0996 | 41.094 | 41.094 | 423.0996 | 290.3484 | 178.137 | 178.1649 |  | 109.6583
0 | [6 empty] | solenoid-entry |  |  |  | 0.076136 | 290.3104 | 423.0996 | 41.094 | 41.094 | 423.0996 | 290.2723 | 178.203 | 178.2411 |  | 109.7244
0 | LS1 | CB09 | SOL3 | 1785 | LS1_CB09:SOL3_D1785 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 290.0223 | 423.0996 | 41.094 | 41.094 | 423.0996 | 289.7723 | 178.4911 | 178.7411 |  | 110.0124
0 | [6 empty] | solenoid-exit |  |  |  | 0.076136 | 289.7342 | 423.0996 | 41.094 | 41.094 | 423.0996 | 289.6962 | 178.7791 | 178.8172 |  | 110.3005
0 | [6 empty] | bellow |  |  |  | 0.099517 | 289.6464 | 423.0996 | 41.094 | 41.094 | 423.0996 | 289.5966 | 178.8669 | 178.9167 |  | 110.3883
0 | LS1 | CB09 | CAV8 | 1791 | LS1_CB09:CAV8_D1791 |  | b08 resonator | 34 |  |  | 0.29817 | 289.4476 | 423.0996 | 41.094 | 41.094 | 423.0996 | 289.2985 | 179.0658 | 179.2149 |  | 110.5872
0 | [10 empty] | 0 | 289.2985 | 423.0996 | 41.094 | 41.094 | 423.0996 | 289.2985 | 179.2149 | 179.2149 |  | 110.7362
0 | [10 empty] | 0 | 289.2985 | 423.0996 | 41.094 | 41.094 | 423.0996 | 289.2985 | 179.2149 | 179.2149 |  | 110.7362
0 | [6 empty] | bellows |  |  |  | 0.19082 | 289.2031 | 423.0996 | 41.094 | 41.094 | 423.0996 | 289.1076 | 179.3103 | 179.4057 |  | 110.8317
0 | LS1 | CB09 | GV | 1794 | LS1_CB09:GV_D1794 |  | gate valve | 40 |  |  | 0.057404 | 289.0789 | 423.0996 | 41.094 | 41.094 | 423.0996 | 289.0502 | 179.4344 | 179.4631 |  | 110.9558
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 288.9985 | 423.0996 | 41.094 | 41.094 | 423.0996 | 288.9467 | 179.5149 | 179.5666 |  | 111.0362
0 | LS1 | WB09 | BPM | 1796 | LS1_WB09:BPM_D1796 | BPM2 | position monitor |  |  |  | 0 | 288.9467 | 423.0996 | 41.094 | 41.094 | 423.0996 | 288.9467 | 179.5666 | 179.5666 |  | 111.088
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 288.8741 | 423.0996 | 41.094 | 41.094 | 423.0996 | 288.8015 | 179.6392 | 179.7119 |  | 111.1606
0 | LS1 | WB09 | HMR | 1797 | LS1_WB09:HMR_D1797 | HR_L | Halo ring, large aperture | 33 |  |  | 0 | 288.8015 | 423.0996 | 41.094 | 41.094 | 423.0996 | 288.8015 | 179.7119 | 179.7119 |  | 111.2333
0 | [6 empty] | diagnostic box |  |  |  | 0 | 288.8015 | 423.0996 | 41.094 | 41.094 | 423.0996 | 288.8015 | 179.7119 | 179.7119 |  | 111.2333
0 | LS1 | WB09 | IP | 1797 | LS1_WB09:IP_D1797 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 288.8015 | 423.0996 | 41.094 | 41.094 | 423.0996 | 288.8015 | 179.7119 | 179.7119 |  | 111.2333
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 288.7358 | 423.0996 | 41.094 | 41.094 | 423.0996 | 288.6702 | 179.7775 | 179.8431 |  | 111.2989
0 | LS1 | CB10 | GV | 1799 | LS1_CB10:GV_D1799 |  | gate valve | 40 |  |  | 0.057404 | 288.6415 | 423.0996 | 41.094 | 41.094 | 423.0996 | 288.6128 | 179.8718 | 179.9005 |  | 111.3932
0 | [6 empty] | bellows |  |  |  | 0.19082 | 288.5174 | 423.0996 | 41.094 | 41.094 | 423.0996 | 288.422 | 179.9959 | 180.0913 |  | 111.5173
0 | LS1 | CB10 | CAV1 | 1802 | LS1_CB10:CAV1_D1802 |  | b08 resonator | 34 |  |  | 0.29817 | 288.2729 | 423.0996 | 41.094 | 41.094 | 423.0996 | 288.1238 | 180.2404 | 180.3895 |  | 111.7618
0 | [10 empty] | 0 | 288.1238 | 423.0996 | 41.094 | 41.094 | 423.0996 | 288.1238 | 180.3895 | 180.3895 |  | 111.9109
0 | LS1 | CB10 | BLM | 1804 | LS1_CB10:BLM_D1804 |  | outside loss monitor |  |  |  | 0 | 288.1238 | 423.0996 | 41.094 | 41.094 | 423.0996 | 288.1238 | 180.3895 | 180.3895 |  | 111.9109
0 | [6 empty] | bellow | 40 |  |  | 0.09651 | 288.0756 | 423.0996 | 41.094 | 41.094 | 423.0996 | 288.0273 | 180.4378 | 180.486 |  | 111.9591
0 | LS1 | CB10 | BPM | 1805 | LS1_CB10:BPM_D1805 | BPM6 | position monitor, BPMc |  |  |  | 0 | 288.0273 | 423.0996 | 41.094 | 41.094 | 423.0996 | 288.0273 | 180.486 | 180.486 |  | 112.0074
0 | [6 empty] | BPM-box | 40 |  |  | 0.05588 | 287.9994 | 423.0996 | 41.094 | 41.094 | 423.0996 | 287.9715 | 180.5139 | 180.5419 |  | 112.0353
0 | [6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 287.9334 | 423.0996 | 41.094 | 41.094 | 423.0996 | 287.8953 | 180.58 | 180.618 |  | 112.1013
0 | LS1 | CB10 | SOL1 | 1809 | LS1_CB10:SOL1_D1809 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 287.6453 | 423.0996 | 41.094 | 41.094 | 423.0996 | 287.3953 | 180.868 | 181.118 |  | 112.3894
0 | [6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 287.3572 | 423.0996 | 41.094 | 41.094 | 423.0996 | 287.3192 | 181.1561 | 181.1942 |  | 112.6775
0 | [6 empty] | bellow |  |  |  | 0.099517 | 287.2694 | 423.0996 | 41.094 | 41.094 | 423.0996 | 287.2197 | 181.2439 | 181.2937 |  | 112.7653
0 | LS1 | CB10 | CAV2 | 1814 | LS1_CB10:CAV2_D1814 |  | b08 resonator | 34 |  |  | 0.29817 | 287.0706 | 423.0996 | 41.094 | 41.094 | 423.0996 | 286.9215 | 181.4428 | 181.5918 |  | 112.9641
0 | [10 empty] | 0 | 286.9215 | 423.0996 | 41.094 | 41.094 | 423.0996 | 286.9215 | 181.5918 | 181.5918 |  | 113.1132
0 | [10 empty] | 0 | 286.9215 | 423.0996 | 41.094 | 41.094 | 423.0996 | 286.9215 | 181.5918 | 181.5918 |  | 113.1132
0 | [6 empty] | bellow |  |  |  | 0.099517 | 286.8717 | 423.0996 | 41.094 | 41.094 | 423.0996 | 286.822 | 181.6416 | 181.6914 |  | 113.163
0 | LS1 | CB10 | CAV3 | 1818 | LS1_CB10:CAV3_D1818 |  | b08 resonator | 34 |  |  | 0.29817 | 286.6729 | 423.0996 | 41.094 | 41.094 | 423.0996 | 286.5238 | 181.8405 | 181.9895 |  | 113.3618
0 | [10 empty] | 0 | 286.5238 | 423.0996 | 41.094 | 41.094 | 423.0996 | 286.5238 | 181.9895 | 181.9895 |  | 113.5109
0 | [10 empty] | 0 | 286.5238 | 423.0996 | 41.094 | 41.094 | 423.0996 | 286.5238 | 181.9895 | 181.9895 |  | 113.5109
0 | [6 empty] | bellow |  |  |  | 0.099517 | 286.474 | 423.0996 | 41.094 | 41.094 | 423.0996 | 286.4243 | 182.0393 | 182.0891 |  | 113.5607
0 | LS1 | CB10 | CAV4 | 1822 | LS1_CB10:CAV4_D1822 |  | b08 resonator | 34 |  |  | 0.29817 | 286.2752 | 423.0996 | 41.094 | 41.094 | 423.0996 | 286.1261 | 182.2381 | 182.3872 |  | 113.7595
0 | [10 empty] | 0 | 286.1261 | 423.0996 | 41.094 | 41.094 | 423.0996 | 286.1261 | 182.3872 | 182.3872 |  | 113.9086
0 | LS1 | CB10 | BLM | 1824 | LS1_CB10:BLM_D1824 |  | outside loss monitor |  |  |  | 0 | 286.1261 | 423.0996 | 41.094 | 41.094 | 423.0996 | 286.1261 | 182.3872 | 182.3872 |  | 113.9086
0 | [6 empty] | bellow |  |  |  | 0.09651 | 286.0779 | 423.0996 | 41.094 | 41.094 | 423.0996 | 286.0296 | 182.4355 | 182.4837 |  | 113.9568
0 | LS1 | CB10 | BPM | 1825 | LS1_CB10:BPM_D1825 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 286.0296 | 423.0996 | 41.094 | 41.094 | 423.0996 | 286.0296 | 182.4837 | 182.4837 |  | 114.0051
0 | [6 empty] | BPM-box |  |  |  | 0.05588 | 286.0017 | 423.0996 | 41.094 | 41.094 | 423.0996 | 285.9737 | 182.5117 | 182.5396 |  | 114.033
0 | [6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 285.9357 | 423.0996 | 41.094 | 41.094 | 423.0996 | 285.8976 | 182.5777 | 182.6157 |  | 114.0991
0 | LS1 | CB10 | SOL2 | 1829 | LS1_CB10:SOL2_D1829 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 285.6476 | 423.0996 | 41.094 | 41.094 | 423.0996 | 285.3976 | 182.8657 | 183.1157 |  | 114.3871
0 | [6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 285.3595 | 423.0996 | 41.094 | 41.094 | 423.0996 | 285.3215 | 183.1538 | 183.1919 |  | 114.6752
0 | [6 empty] | bellow |  |  |  | 0.099517 | 285.2717 | 423.0996 | 41.094 | 41.094 | 423.0996 | 285.2219 | 183.2416 | 183.2914 |  | 114.763
0 | LS1 | CB10 | CAV5 | 1834 | LS1_CB10:CAV5_D1834 |  | b08 resonator | 34 |  |  | 0.29817 | 285.0729 | 423.0996 | 41.094 | 41.094 | 423.0996 | 284.9238 | 183.4405 | 183.5896 |  | 114.9619
0 | [10 empty] | 0 | 284.9238 | 423.0996 | 41.094 | 41.094 | 423.0996 | 284.9238 | 183.5896 | 183.5896 |  | 115.1109
0 | [10 empty] | 0 | 284.9238 | 423.0996 | 41.094 | 41.094 | 423.0996 | 284.9238 | 183.5896 | 183.5896 |  | 115.1109
0 | [6 empty] | bellow |  |  |  | 0.099517 | 284.874 | 423.0996 | 41.094 | 41.094 | 423.0996 | 284.8243 | 183.6393 | 183.6891 |  | 115.1607
0 | LS1 | CB10 | CAV6 | 1838 | LS1_CB10:CAV6_D1838 |  | b08 resonator | 34 |  |  | 0.29817 | 284.6752 | 423.0996 | 41.094 | 41.094 | 423.0996 | 284.5261 | 183.8382 | 183.9873 |  | 115.3595
0 | [10 empty] | 0 | 284.5261 | 423.0996 | 41.094 | 41.094 | 423.0996 | 284.5261 | 183.9873 | 183.9873 |  | 115.5086
0 | [10 empty] | 0 | 284.5261 | 423.0996 | 41.094 | 41.094 | 423.0996 | 284.5261 | 183.9873 | 183.9873 |  | 115.5086
0 | [6 empty] | bellow |  |  |  | 0.099517 | 284.4763 | 423.0996 | 41.094 | 41.094 | 423.0996 | 284.4266 | 184.037 | 184.0868 |  | 115.5584
0 | LS1 | CB10 | CAV7 | 1842 | LS1_CB10:CAV7_D1842 |  | b08 resonator | 34 |  |  | 0.29817 | 284.2775 | 423.0996 | 41.094 | 41.094 | 423.0996 | 284.1284 | 184.2359 | 184.3849 |  | 115.7572
0 | [10 empty] | 0 | 284.1284 | 423.0996 | 41.094 | 41.094 | 423.0996 | 284.1284 | 184.3849 | 184.3849 |  | 115.9063
0 | LS1 | CB10 | BLM | 1844 | LS1_CB10:BLM_D1844 |  | outside loss monitor |  |  |  | 0 | 284.1284 | 423.0996 | 41.094 | 41.094 | 423.0996 | 284.1284 | 184.3849 | 184.3849 |  | 115.9063
0 | [6 empty] | bellow |  |  |  | 0.09651 | 284.0801 | 423.0996 | 41.094 | 41.094 | 423.0996 | 284.0319 | 184.4332 | 184.4815 |  | 115.9546
0 | LS1 | CB10 | BPM | 1845 | LS1_CB10:BPM_D1845 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 284.0319 | 423.0996 | 41.094 | 41.094 | 423.0996 | 284.0319 | 184.4815 | 184.4815 |  | 116.0028
0 | [6 empty] | BPM-box |  |  |  | 0.05588 | 284.0039 | 423.0996 | 41.094 | 41.094 | 423.0996 | 283.976 | 184.5094 | 184.5373 |  | 116.0308
0 | [6 empty] | solenoid-entry |  |  |  | 0.076136 | 283.9379 | 423.0996 | 41.094 | 41.094 | 423.0996 | 283.8999 | 184.5754 | 184.6135 |  | 116.0968
0 | LS1 | CB10 | SOL3 | 1849 | LS1_CB10:SOL3_D1849 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 283.6499 | 423.0996 | 41.094 | 41.094 | 423.0996 | 283.3999 | 184.8635 | 185.1135 |  | 116.3848
0 | [6 empty] | solenoid-exit |  |  |  | 0.076136 | 283.3618 | 423.0996 | 41.094 | 41.094 | 423.0996 | 283.3237 | 185.1515 | 185.1896 |  | 116.6729
0 | [6 empty] | bellow |  |  |  | 0.099517 | 283.274 | 423.0996 | 41.094 | 41.094 | 423.0996 | 283.2242 | 185.2394 | 185.2891 |  | 116.7607
0 | LS1 | CB10 | CAV8 | 1854 | LS1_CB10:CAV8_D1854 |  | b08 resonator | 34 |  |  | 0.29817 | 283.0751 | 423.0996 | 41.094 | 41.094 | 423.0996 | 282.926 | 185.4382 | 185.5873 |  | 116.9596
0 | [10 empty] | 0 | 282.926 | 423.0996 | 41.094 | 41.094 | 423.0996 | 282.926 | 185.5873 | 185.5873 |  | 117.1087
0 | [10 empty] | 0 | 282.926 | 423.0996 | 41.094 | 41.094 | 423.0996 | 282.926 | 185.5873 | 185.5873 |  | 117.1087
0 | [6 empty] | bellows |  |  |  | 0.19082 | 282.8306 | 423.0996 | 41.094 | 41.094 | 423.0996 | 282.7352 | 185.6827 | 185.7781 |  | 117.2041
0 | LS1 | CB10 | GV | 1858 | LS1_CB10:GV_D1858 |  | gate valve | 40 |  |  | 0.057404 | 282.7065 | 423.0996 | 41.094 | 41.094 | 423.0996 | 282.6778 | 185.8068 | 185.8355 |  | 117.3282
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 282.6261 | 423.0996 | 41.094 | 41.094 | 423.0996 | 282.5743 | 185.8873 | 185.939 |  | 117.4086
0 | LS1 | WB10 | BPM | 1859 | LS1_WB10:BPM_D1859 | BPM2 | position monitor |  |  |  | 0 | 282.5743 | 423.0996 | 41.094 | 41.094 | 423.0996 | 282.5743 | 185.939 | 185.939 |  | 117.4604
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 282.5017 | 423.0996 | 41.094 | 41.094 | 423.0996 | 282.429 | 186.0117 | 186.0843 |  | 117.533
0 | LS1 | WB10 | HMR | 1861 | LS1_WB10:HMR_D1861 | HR_L | Halo ring, large aperture | 33 |  |  | 0 | 282.429 | 423.0996 | 41.094 | 41.094 | 423.0996 | 282.429 | 186.0843 | 186.0843 |  | 117.6057
0 | [6 empty] | diagnostic box |  |  |  | 0 | 282.429 | 423.0996 | 41.094 | 41.094 | 423.0996 | 282.429 | 186.0843 | 186.0843 |  | 117.6057
0 | LS1 | WB10 | IP | 1861 | LS1_WB10:IP_D1861 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 282.429 | 423.0996 | 41.094 | 41.094 | 423.0996 | 282.429 | 186.0843 | 186.0843 |  | 117.6057
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 282.3634 | 423.0996 | 41.094 | 41.094 | 423.0996 | 282.2978 | 186.1499 | 186.2155 |  | 117.6713
0 | LS1 | CB11 | GV | 1862 | LS1_CB11:GV_D1862 |  | gate valve | 40 |  |  | 0.057404 | 282.2691 | 423.0996 | 41.094 | 41.094 | 423.0996 | 282.2404 | 186.2442 | 186.2729 |  | 117.7656
0 | [6 empty] | bellows |  |  |  | 0.19082 | 282.145 | 423.0996 | 41.094 | 41.094 | 423.0996 | 282.0496 | 186.3683 | 186.4638 |  | 117.8897
0 | LS1 | CB11 | CAV1 | 1866 | LS1_CB11:CAV1_D1866 |  | b08 resonator | 34 |  |  | 0.29817 | 281.9005 | 423.0996 | 41.094 | 41.094 | 423.0996 | 281.7514 | 186.6128 | 186.7619 |  | 118.1342
0 | [10 empty] | 0 | 281.7514 | 423.0996 | 41.094 | 41.094 | 423.0996 | 281.7514 | 186.7619 | 186.7619 |  | 118.2833
0 | LS1 | CB11 | BLM | 1868 | LS1_CB11:BLM_D1868 |  | outside loss monitor |  |  |  | 0 | 281.7514 | 423.0996 | 41.094 | 41.094 | 423.0996 | 281.7514 | 186.7619 | 186.7619 |  | 118.2833
0 | [6 empty] | bellow | 40 |  |  | 0.09651 | 281.7032 | 423.0996 | 41.094 | 41.094 | 423.0996 | 281.6549 | 186.8102 | 186.8584 |  | 118.3315
0 | LS1 | CB11 | BPM | 1869 | LS1_CB11:BPM_D1869 | BPM6 | position monitor, BPMc |  |  |  | 0 | 281.6549 | 423.0996 | 41.094 | 41.094 | 423.0996 | 281.6549 | 186.8584 | 186.8584 |  | 118.3798
0 | [6 empty] | BPM-box | 40 |  |  | 0.05588 | 281.627 | 423.0996 | 41.094 | 41.094 | 423.0996 | 281.599 | 186.8864 | 186.9143 |  | 118.4077
0 | [6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 281.561 | 423.0996 | 41.094 | 41.094 | 423.0996 | 281.5229 | 186.9524 | 186.9904 |  | 118.4738
0 | LS1 | CB11 | SOL1 | 1872 | LS1_CB11:SOL1_D1872 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 281.2729 | 423.0996 | 41.094 | 41.094 | 423.0996 | 281.0229 | 187.2404 | 187.4904 |  | 118.7618
0 | [6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 280.9848 | 423.0996 | 41.094 | 41.094 | 423.0996 | 280.9468 | 187.5285 | 187.5666 |  | 119.0499
0 | [6 empty] | bellow |  |  |  | 0.099517 | 280.897 | 423.0996 | 41.094 | 41.094 | 423.0996 | 280.8472 | 187.6163 | 187.6661 |  | 119.1377
0 | LS1 | CB11 | CAV2 | 1878 | LS1_CB11:CAV2_D1878 |  | b08 resonator | 34 |  |  | 0.29817 | 280.6982 | 423.0996 | 41.094 | 41.094 | 423.0996 | 280.5491 | 187.8152 | 187.9643 |  | 119.3366
0 | [10 empty] | 0 | 280.5491 | 423.0996 | 41.094 | 41.094 | 423.0996 | 280.5491 | 187.9643 | 187.9643 |  | 119.4856
0 | [10 empty] | 0 | 280.5491 | 423.0996 | 41.094 | 41.094 | 423.0996 | 280.5491 | 187.9643 | 187.9643 |  | 119.4856
0 | [6 empty] | bellow |  |  |  | 0.099517 | 280.4993 | 423.0996 | 41.094 | 41.094 | 423.0996 | 280.4496 | 188.014 | 188.0638 |  | 119.5354
0 | LS1 | CB11 | CAV3 | 1882 | LS1_CB11:CAV3_D1882 |  | b08 resonator | 34 |  |  | 0.29817 | 280.3005 | 423.0996 | 41.094 | 41.094 | 423.0996 | 280.1514 | 188.2129 | 188.362 |  | 119.7342
0 | [10 empty] | 0 | 280.1514 | 423.0996 | 41.094 | 41.094 | 423.0996 | 280.1514 | 188.362 | 188.362 |  | 119.8833
0 | [10 empty] | 0 | 280.1514 | 423.0996 | 41.094 | 41.094 | 423.0996 | 280.1514 | 188.362 | 188.362 |  | 119.8833
0 | [6 empty] | bellow |  |  |  | 0.099517 | 280.1016 | 423.0996 | 41.094 | 41.094 | 423.0996 | 280.0519 | 188.4117 | 188.4615 |  | 119.9331
0 | LS1 | CB11 | CAV4 | 1886 | LS1_CB11:CAV4_D1886 |  | b08 resonator | 34 |  |  | 0.29817 | 279.9028 | 423.0996 | 41.094 | 41.094 | 423.0996 | 279.7537 | 188.6106 | 188.7596 |  | 120.1319
0 | [10 empty] | 0 | 279.7537 | 423.0996 | 41.094 | 41.094 | 423.0996 | 279.7537 | 188.7596 | 188.7596 |  | 120.281
0 | LS1 | CB11 | BLM | 1888 | LS1_CB11:BLM_D1888 |  | outside loss monitor |  |  |  | 0 | 279.7537 | 423.0996 | 41.094 | 41.094 | 423.0996 | 279.7537 | 188.7596 | 188.7596 |  | 120.281
0 | [6 empty] | bellow |  |  |  | 0.09651 | 279.7054 | 423.0996 | 41.094 | 41.094 | 423.0996 | 279.6572 | 188.8079 | 188.8562 |  | 120.3293
0 | LS1 | CB11 | BPM | 1889 | LS1_CB11:BPM_D1889 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 279.6572 | 423.0996 | 41.094 | 41.094 | 423.0996 | 279.6572 | 188.8562 | 188.8562 |  | 120.3775
0 | [6 empty] | BPM-box |  |  |  | 0.05588 | 279.6292 | 423.0996 | 41.094 | 41.094 | 423.0996 | 279.6013 | 188.8841 | 188.912 |  | 120.4055
0 | [6 empty] | solenoid-entry | 40 |  |  | 0.076136 | 279.5632 | 423.0996 | 41.094 | 41.094 | 423.0996 | 279.5252 | 188.9501 | 188.9882 |  | 120.4715
0 | LS1 | CB11 | SOL2 | 1892 | LS1_CB11:SOL2_D1892 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 279.2752 | 423.0996 | 41.094 | 41.094 | 423.0996 | 279.0252 | 189.2382 | 189.4882 |  | 120.7595
0 | [6 empty] | solenoid-exit | 40 |  |  | 0.076136 | 278.9871 | 423.0996 | 41.094 | 41.094 | 423.0996 | 278.949 | 189.5262 | 189.5643 |  | 121.0476
0 | [6 empty] | bellow |  |  |  | 0.099517 | 278.8993 | 423.0996 | 41.094 | 41.094 | 423.0996 | 278.8495 | 189.6141 | 189.6638 |  | 121.1354
0 | LS1 | CB11 | CAV5 | 1898 | LS1_CB11:CAV5_D1898 |  | b08 resonator | 34 |  |  | 0.29817 | 278.7004 | 423.0996 | 41.094 | 41.094 | 423.0996 | 278.5513 | 189.8129 | 189.962 |  | 121.3343
0 | [10 empty] | 0 | 278.5513 | 423.0996 | 41.094 | 41.094 | 423.0996 | 278.5513 | 189.962 | 189.962 |  | 121.4834
0 | [10 empty] | 0 | 278.5513 | 423.0996 | 41.094 | 41.094 | 423.0996 | 278.5513 | 189.962 | 189.962 |  | 121.4834
0 | [6 empty] | bellow |  |  |  | 0.099517 | 278.5016 | 423.0996 | 41.094 | 41.094 | 423.0996 | 278.4518 | 190.0118 | 190.0615 |  | 121.5331
0 | LS1 | CB11 | CAV6 | 1902 | LS1_CB11:CAV6_D1902 |  | b08 resonator | 34 |  |  | 0.29817 | 278.3027 | 423.0996 | 41.094 | 41.094 | 423.0996 | 278.1537 | 190.2106 | 190.3597 |  | 121.732
0 | [10 empty] | 0 | 278.1537 | 423.0996 | 41.094 | 41.094 | 423.0996 | 278.1537 | 190.3597 | 190.3597 |  | 121.8811
0 | [10 empty] | 0 | 278.1537 | 423.0996 | 41.094 | 41.094 | 423.0996 | 278.1537 | 190.3597 | 190.3597 |  | 121.8811
0 | [6 empty] | bellow |  |  |  | 0.099517 | 278.1039 | 423.0996 | 41.094 | 41.094 | 423.0996 | 278.0541 | 190.4094 | 190.4592 |  | 121.9308
0 | LS1 | CB11 | CAV7 | 1906 | LS1_CB11:CAV7_D1906 |  | b08 resonator | 34 |  |  | 0.29817 | 277.9051 | 423.0996 | 41.094 | 41.094 | 423.0996 | 277.756 | 190.6083 | 190.7574 |  | 122.1297
0 | [10 empty] | 0 | 277.756 | 423.0996 | 41.094 | 41.094 | 423.0996 | 277.756 | 190.7574 | 190.7574 |  | 122.2787
0 | LS1 | CB11 | BLM | 1908 | LS1_CB11:BLM_D1908 |  | outside loss monitor |  |  |  | 0 | 277.756 | 423.0996 | 41.094 | 41.094 | 423.0996 | 277.756 | 190.7574 | 190.7574 |  | 122.2787
0 | [6 empty] | bellow |  |  |  | 0.09651 | 277.7077 | 423.0996 | 41.094 | 41.094 | 423.0996 | 277.6595 | 190.8056 | 190.8539 |  | 122.327
0 | LS1 | CB11 | BPM | 1909 | LS1_CB11:BPM_D1909 | BPM6 | position monitor, BPMc | 40 |  |  | 0 | 277.6595 | 423.0996 | 41.094 | 41.094 | 423.0996 | 277.6595 | 190.8539 | 190.8539 |  | 122.3753
0 | [6 empty] | BPM-box |  |  |  | 0.05588 | 277.6315 | 423.0996 | 41.094 | 41.094 | 423.0996 | 277.6036 | 190.8818 | 190.9098 |  | 122.4032
0 | [6 empty] | solenoid-entry |  |  |  | 0.076136 | 277.5655 | 423.0996 | 41.094 | 41.094 | 423.0996 | 277.5274 | 190.9478 | 190.9859 |  | 122.4692
0 | LS1 | CB11 | SOL3 | 1912 | LS1_CB11:SOL3_D1912 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 277.2774 | 423.0996 | 41.094 | 41.094 | 423.0996 | 277.0274 | 191.2359 | 191.4859 |  | 122.7573
0 | [6 empty] | solenoid-exit |  |  |  | 0.076136 | 276.9894 | 423.0996 | 41.094 | 41.094 | 423.0996 | 276.9513 | 191.524 | 191.562 |  | 123.0453
0 | [6 empty] | bellow |  |  |  | 0.099517 | 276.9015 | 423.0996 | 41.094 | 41.094 | 423.0996 | 276.8518 | 191.6118 | 191.6616 |  | 123.1332
0 | LS1 | CB11 | CAV8 | 1918 | LS1_CB11:CAV8_D1918 |  | b08 resonator | 34 |  |  | 0.29817 | 276.7027 | 423.0996 | 41.094 | 41.094 | 423.0996 | 276.5536 | 191.8106 | 191.9597 |  | 123.332
0 | [10 empty] | 0 | 276.5536 | 423.0996 | 41.094 | 41.094 | 423.0996 | 276.5536 | 191.9597 | 191.9597 |  | 123.4811
0 | [10 empty] | 0 | 276.5536 | 423.0996 | 41.094 | 41.094 | 423.0996 | 276.5536 | 191.9597 | 191.9597 |  | 123.4811
0 | [6 empty] | bellows |  |  |  | 0.19082 | 276.4582 | 423.0996 | 41.094 | 41.094 | 423.0996 | 276.3628 | 192.0551 | 192.1505 |  | 123.5765
0 | LS1 | CB11 | GV | 1922 | LS1_CB11:GV_D1922 |  | gate valve | 40 |  |  | 0.057404 | 276.3341 | 423.0996 | 41.094 | 41.094 | 423.0996 | 276.3054 | 192.1792 | 192.208 |  | 123.7006
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 276.2536 | 423.0996 | 41.094 | 41.094 | 423.0996 | 276.2019 | 192.2597 | 192.3115 |  | 123.7811
0 | LS1 | WB11 | BPM | 1923 | LS1_WB11:BPM_D1923 | BPM2 | position monitor |  |  |  | 0 | 276.2019 | 423.0996 | 41.094 | 41.094 | 423.0996 | 276.2019 | 192.3115 | 192.3115 |  | 123.8328
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 276.1292 | 423.0996 | 41.094 | 41.094 | 423.0996 | 276.0566 | 192.3841 | 192.4567 |  | 123.9055
0 | LS1 | WB11 | HMR | 1925 | LS1_WB11:HMR_D1925 | HR_L | Halo ring, large aperture | 33 |  |  | 0 | 276.0566 | 423.0996 | 41.094 | 41.094 | 423.0996 | 276.0566 | 192.4567 | 192.4567 |  | 123.9781
0 | [6 empty] | diagnostic box |  |  |  | 0 | 276.0566 | 423.0996 | 41.094 | 41.094 | 423.0996 | 276.0566 | 192.4567 | 192.4567 |  | 123.9781
0 | LS1 | WB11 | IP | 1925 | LS1_WB11:IP_D1925 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 276.0566 | 423.0996 | 41.094 | 41.094 | 423.0996 | 276.0566 | 192.4567 | 192.4567 |  | 123.9781
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 275.991 | 423.0996 | 41.094 | 41.094 | 423.0996 | 275.9254 | 192.5223 | 192.588 |  | 124.0437
0 | SEGMENT 1 END | [9 empty] | -9.6e-07 | 275.9254 | 423.0996 | 41.094 | 41.094 | 423.0996 | 275.9254 | 192.5879 | 192.5879 | adjustment for residuals | 124.1093
0 | [10 empty] | 1.05 | 275.4004 | 423.0996 | 41.094 | 41.094 | 423.0996 | 274.8754 | 193.1129 | 193.6379 |  | 124.6343
0 | LS1 | BTS | DCHV | 1937 | LS1_BTS:DCHV_D1937 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 274.7754 | 423.0996 | 41.094 | 41.094 | 423.0996 | 274.6754 | 193.7379 | 193.8379 |  | 125.2593
0 | [10 empty] | 0.25 | 274.5504 | 423.0996 | 41.094 | 41.094 | 423.0996 | 274.4254 | 193.9629 | 194.0879 |  | 125.4843
0 | LS1 | BTS | QH | 1942 | LS1_BTS:QH_D1942 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 274.3004 | 423.0996 | 41.094 | 41.094 | 423.0996 | 274.1754 | 194.2129 | 194.3379 |  | 125.7343
0 | [10 empty] | 0.29595 | 274.0274 | 423.0996 | 41.094 | 41.094 | 423.0996 | 273.8794 | 194.4859 | 194.6339 |  | 126.0073
0 | LS1 | BTS | GV | 1947 | LS1_BTS:GV_D1947 | GV_M | gate valve | 63 | 1.9924e-06 | 273.8419 | 0.075 | 273.8419 | 423.0996 | 41.094 | 41.094 | 423.0996 | 273.8044 | 194.6714 | 194.7089 |  | 126.1928
0 | [10 empty] | 0.12905 | 273.7399 | 423.0996 | 41.094 | 41.094 | 423.0996 | 273.6754 | 194.7734 | 194.8379 |  | 126.2948
0 | LS1 | BTS | QV | 1950 | LS1_BTS:QV_D1950 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 273.5504 | 423.0996 | 41.094 | 41.094 | 423.0996 | 273.4254 | 194.9629 | 195.0879 |  | 126.4843
0 | [10 empty] | 1.198 | 272.8264 | 423.0996 | 41.094 | 41.094 | 423.0996 | 272.2274 | 195.6869 | 196.2859 |  | 127.2083
0 | LS1 | BTS | DCHV | 1964 | LS1_BTS:DCHV_D1964 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 272.1274 | 423.0996 | 41.094 | 41.094 | 423.0996 | 272.0274 | 196.3859 | 196.4859 |  | 127.9073
0 | [10 empty] | 0 | 272.0274 | 423.0996 | 41.094 | 41.094 | 423.0996 | 272.0274 | 196.4859 | 196.4859 |  | 128.0073
0 | [8 empty] | 2.9924e-06 |  | 0.19214 | 271.9313 | 423.0996 | 41.094 | 41.094 | 423.0996 | 271.8352 | 196.582 | 196.6781 |  | 128.1034
0 | LS1 | BTS | BPM | 1967 | LS1_BTS:BPM_D1967 | BPM3 | position monitor | 47.5 | 271.8353 | 271.8352 | 0 | 271.8352 | 423.0996 | 41.094 | 41.094 | 423.0996 | 271.8352 | 196.6781 | 196.6781 | 0.137 | 128.1995
0 | [8 empty] | inconsisent with next one |  | 0.057856 | 271.8063 | 423.0996 | 41.094 | 41.094 | 423.0996 | 271.7774 | 196.707 | 196.7359 |  | 128.2284
0 | LS1 | BTS | QH | 1969 | LS1_BTS:QH_D1969 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 271.6524 | 423.0996 | 41.094 | 41.094 | 423.0996 | 271.5274 | 196.8609 | 196.9859 |  | 128.3823
0 | [10 empty] | 0.5 | 271.2774 | 423.0996 | 41.094 | 41.094 | 423.0996 | 271.0274 | 197.2359 | 197.4859 |  | 128.7573
0 | LS1 | BTS | QV | 1976 | LS1_BTS:QV_D1976 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 270.9024 | 423.0996 | 41.094 | 41.094 | 423.0996 | 270.7774 | 197.6109 | 197.7359 |  | 129.1323
0 | [10 empty] | 0.1325 | 270.7111 | 423.0996 | 41.094 | 41.094 | 423.0996 | 270.6449 | 197.8022 | 197.8685 |  | 129.3236
0 | [6 empty] | bellow |  |  |  | 0.1035 | 270.5931 | 423.0996 | 41.094 | 41.094 | 423.0996 | 270.5414 | 197.9202 | 197.972 |  | 129.4416
0 | LS1 | BTS | BPM | 1980 | LS1_BTS:BPM_D1980 | BPM2 | position monitor | 47.5 | replace by a nipple |  | 0 | 270.5414 | 423.0996 | 41.094 | 41.094 | 423.0996 | 270.5414 | 197.972 | 197.972 | Should be BMP3a, 4/24/2017 | 129.4933
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 270.4687 | 423.0996 | 41.094 | 41.094 | 423.0996 | 270.3961 | 198.0446 | 198.1172 |  | 129.566
0 | LS1 | BTS | PORT | 1981 | LS1_BTS:PORT_D1981 |  | port for future diagnostics |  |  |  | 0 | 270.3961 | 423.0996 | 41.094 | 41.094 | 423.0996 | 270.3961 | 198.1172 | 198.1172 |  | 129.6386
0 | [6 empty] | diagnostic box |  |  |  | 0 | 270.3961 | 423.0996 | 41.094 | 41.094 | 423.0996 | 270.3961 | 198.1172 | 198.1172 |  | 129.6386
0 | LS1 | BTS | IP | 1981 | LS1_BTS:IP_D1981 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 270.3961 | 423.0996 | 41.094 | 41.094 | 423.0996 | 270.3961 | 198.1172 | 198.1172 |  | 129.6386
0 | [6 empty] | bellow |  |  |  | 0.13122 | 270.3305 | 423.0996 | 41.094 | 41.094 | 423.0996 | 270.2649 | 198.1828 | 198.2485 |  | 129.7042
0 | [10 empty] | 0 | 270.2649 | 423.0996 | 41.094 | 41.094 | 423.0996 | 270.2649 | 198.2485 | 198.2485 |  | 129.7698
0 | [10 empty] | 0.059497 | 270.2351 | 423.0996 | 41.094 | 41.094 | 423.0996 | 270.2054 | 198.2782 | 198.3079 |  | 129.7996
0 | [10 empty] | 0.052097 | 270.1793 | 423.0996 | 41.094 | 41.094 | 423.0996 | 270.1533 | 198.334 | 198.36 |  | 129.8554
0 | [10 empty] | 0.32364 | 269.9915 | 423.0996 | 41.094 | 41.094 | 423.0996 | 269.8297 | 198.5219 | 198.6837 |  | 130.0432
0 | [10 empty] | 0 | 269.8297 | 423.0996 | 41.094 | 41.094 | 423.0996 | 269.8297 | 198.6837 | 198.6837 |  | 130.2051
0 | [10 empty] | 0 | 269.8297 | 423.0996 | 41.094 | 41.094 | 423.0996 | 269.8297 | 198.6837 | 198.6837 |  | 130.2051
0 | [10 empty] | 0.099517 | 269.7799 | 423.0996 | 41.094 | 41.094 | 423.0996 | 269.7301 | 198.7334 | 198.7832 |  | 130.2548
0 | [10 empty] | 0.32364 | 269.5683 | 423.0996 | 41.094 | 41.094 | 423.0996 | 269.4065 | 198.945 | 199.1069 |  | 130.4664
0 | [10 empty] | 0 | 269.4065 | 423.0996 | 41.094 | 41.094 | 423.0996 | 269.4065 | 199.1069 | 199.1069 |  | 130.6282
0 | [10 empty] | 0 | 269.4065 | 423.0996 | 41.094 | 41.094 | 423.0996 | 269.4065 | 199.1069 | 199.1069 |  | 130.6282
0 | [10 empty] | 0.052097 | 269.3804 | 423.0996 | 41.094 | 41.094 | 423.0996 | 269.3544 | 199.1329 | 199.1589 |  | 130.6543
0 | [10 empty] | 0.072 | 269.3184 | 423.0996 | 41.094 | 41.094 | 423.0996 | 269.2824 | 199.1949 | 199.2309 |  | 130.7163
0 | [10 empty] | 0 | 269.2824 | 423.0996 | 41.094 | 41.094 | 423.0996 | 269.2824 | 199.2309 | 199.2309 |  | 130.7523
0 | [6 empty] | bellow |  |  |  | 0.094006 | 269.2354 | 423.0996 | 41.094 | 41.094 | 423.0996 | 269.1884 | 199.278 | 199.325 |  | 130.7993
0 | [6 empty] | vacuum box |  |  |  | 0 | 269.1884 | 423.0996 | 41.094 | 41.094 | 423.0996 | 269.1884 | 199.325 | 199.325 |  | 130.8463
0 | LS1 | BTS | PORT | 1993 | LS1_BTS:PORT_D1993 |  | future diagnostic port |  |  |  | 0 | 269.1884 | 423.0996 | 41.094 | 41.094 | 423.0996 | 269.1884 | 199.325 | 199.325 |  | 130.8463
0 | [6 empty] | vacuum box |  | 2.9924e-06 |  | 0 | 269.1884 | 423.0996 | 41.094 | 41.094 | 423.0996 | 269.1884 | 199.325 | 199.325 |  | 130.8463
0 | LS1 | BTS | IP | 1993 | LS1_BTS:IP_D1993 | IP_S | ion pump, 75 l/s |  | 269.1884 | 269.1884 | 0 | 269.1884 | 423.0996 | 41.094 | 41.094 | 423.0996 | 269.1884 | 199.325 | 199.325 |  | 130.8463
0 | [6 empty] | bellow |  |  |  | 0.18999 | 269.0934 | 423.0996 | 41.094 | 41.094 | 423.0996 | 268.9984 | 199.42 | 199.5149 |  | 130.9413
0 | [10 empty] | 0.106 | 268.9454 | 423.0996 | 41.094 | 41.094 | 423.0996 | 268.8924 | 199.5679 | 199.6209 |  | 131.0893
0 | LS1 | BTS | DCHV | 1997 | LS1_BTS:DCHV_D1997 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 268.7924 | 423.0996 | 41.094 | 41.094 | 423.0996 | 268.6924 | 199.7209 | 199.8209 |  | 131.2423
0 | [10 empty] | 0.25 | 268.5674 | 423.0996 | 41.094 | 41.094 | 423.0996 | 268.4424 | 199.9459 | 200.0709 |  | 131.4673
0 | LS1 | BTS | QH | 2002 | LS1_BTS:QH_D2002 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 268.3174 | 423.0996 | 41.094 | 41.094 | 423.0996 | 268.1924 | 200.1959 | 200.3209 |  | 131.7173
0 | [10 empty] | 0.88621 | 267.7493 | 423.0996 | 41.094 | 41.094 | 423.0996 | 267.3062 | 200.7641 | 201.2072 |  | 132.2854
0 | LS1 | BTS | QV | 2013 | LS1_BTS:QV_D2013 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 267.1812 | 423.0996 | 41.094 | 41.094 | 423.0996 | 267.0562 | 201.3322 | 201.4572 |  | 132.8535
0 | [10 empty] | 0.87 | 266.6212 | 423.0996 | 41.094 | 41.094 | 423.0996 | 266.1862 | 201.8922 | 202.3272 |  | 133.4135
0 | LS1 | BTS | DCHV | 2024 | LS1_BTS:DCHV_D2024 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 266.0862 | 423.0996 | 41.094 | 41.094 | 423.0996 | 265.9862 | 202.4272 | 202.5272 |  | 133.9485
0 | [8 empty] | 2.9924e-06 |  | 0.1909 | 265.8907 | 423.0996 | 41.094 | 41.094 | 423.0996 | 265.7953 | 202.6226 | 202.7181 | 0.137 | 134.144
0 | LS1 | BTS | BPM | 2027 | LS1_BTS:BPM_D2027 | BPM3 | position monitor | 47.5 | 265.7953 | 265.7953 | 0 | 265.7953 | 423.0996 | 41.094 | 41.094 | 423.0996 | 265.7953 | 202.7181 | 202.7181 |  | 134.2394
0 | [8 empty] | inconsisent with previous one |  | 0 | 265.7953 | 423.0996 | 41.094 | 41.094 | 423.0996 | 265.7953 | 202.7181 | 202.7181 |  | 134.2394
0 | [10 empty] | 0.059095 | 265.7657 | 423.0996 | 41.094 | 41.094 | 423.0996 | 265.7362 | 202.7476 | 202.7772 |  | 134.269
0 | LS1 | BTS | QH | 2029 | LS1_BTS:QH_D2029 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 265.6112 | 423.0996 | 41.094 | 41.094 | 423.0996 | 265.4862 | 202.9022 | 203.0272 |  | 134.4235
0 | [10 empty] | 1.0862 | 264.9431 | 423.0996 | 41.094 | 41.094 | 423.0996 | 264.4 | 203.5703 | 204.1134 |  | 135.0916
0 | LS1 | BTS | QV | 2042 | LS1_BTS:QV_D2042 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 264.275 | 423.0996 | 41.094 | 41.094 | 423.0996 | 264.15 | 204.2384 | 204.3634 |  | 135.7597
0 | [10 empty] | 0.97033 | 263.6648 | 423.0996 | 41.094 | 41.094 | 423.0996 | 263.1796 | 204.8485 | 205.3337 |  | 136.3699
0 | [6 empty] | bellow |  | 2.9924e-06 |  | 0.1035 | 263.1279 | 423.0996 | 41.094 | 41.094 | 423.0996 | 263.0761 | 205.3855 | 205.4372 |  | 136.9068
0 | LS1 | BTS | BPM | 2054 | LS1_BTS:BPM_D2054 | BPM3 | position monitor | 47.5 | 263.0761 | 263.0761 | 0 | 263.0761 | 423.0996 | 41.094 | 41.094 | 423.0996 | 263.0761 | 205.4372 | 205.4372 |  | 136.9586
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 263.0035 | 423.0996 | 41.094 | 41.094 | 423.0996 | 262.9309 | 205.5098 | 205.5825 |  | 137.0312
0 | LS1 | BTS | PM | 2056 | LS1_BTS:PM_D2056 | PM3 | profile monitor |  |  |  | 0 | 262.9309 | 423.0996 | 41.094 | 41.094 | 423.0996 | 262.9309 | 205.5825 | 205.5825 |  | 137.1039
0 | [6 empty] | diagnostic box |  |  |  | 0 | 262.9309 | 423.0996 | 41.094 | 41.094 | 423.0996 | 262.9309 | 205.5825 | 205.5825 |  | 137.1039
0 | LS1 | BTS | IP | 2056 | LS1_BTS:IP_D2056 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 262.9309 | 423.0996 | 41.094 | 41.094 | 423.0996 | 262.9309 | 205.5825 | 205.5825 |  | 137.1039
0 | [6 empty] | bellow |  |  |  | 0.13122 | 262.8652 | 423.0996 | 41.094 | 41.094 | 423.0996 | 262.7996 | 205.6481 | 205.7137 |  | 137.1695
0 | [10 empty] | 0 | 262.7996 | 423.0996 | 41.094 | 41.094 | 423.0996 | 262.7996 | 205.7137 | 205.7137 |  | 137.2351
0 | [10 empty] | 0.26878 | 262.6652 | 423.0996 | 41.094 | 41.094 | 423.0996 | 262.5309 | 205.8481 | 205.9825 | 2012-03-09 why this odd length ? Make it say 0.78 | 137.3695
0 | LS1 | BTS | DCHV | 2061 | LS1_BTS:DCHV_D2061 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 262.4309 | 423.0996 | 41.094 | 41.094 | 423.0996 | 262.3309 | 206.0825 | 206.1825 |  | 137.6039
0 | [10 empty] | 0.25 | 262.2059 | 423.0996 | 41.094 | 41.094 | 423.0996 | 262.0809 | 206.3075 | 206.4325 |  | 137.8289
0 | LS1 | BTS | QV | 2066 | LS1_BTS:QV_D2066 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 261.9559 | 423.0996 | 41.094 | 41.094 | 423.0996 | 261.8309 | 206.5575 | 206.6825 |  | 138.0789
0 | [10 empty] | 0.5 | 261.5809 | 423.0996 | 41.094 | 41.094 | 423.0996 | 261.3309 | 206.9325 | 207.1825 |  | 138.4539
0 | LS1 | BTS | QH | 2073 | LS1_BTS:QH_D2073 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 261.2059 | 423.0996 | 41.094 | 41.094 | 423.0996 | 261.0809 | 207.3075 | 207.4325 |  | 138.8289
0 | [10 empty] | 0.7 | 260.7309 | 423.0996 | 41.094 | 41.094 | 423.0996 | 260.3809 | 207.7825 | 208.1325 |  | 139.3039
0 | [10 empty] | 0 | 260.3809 | 423.0996 | 41.094 | 41.094 | 423.0996 | 260.3809 | 208.1325 | 208.1325 |  | 139.6539
0 | LS1 | CH01 | GV | 2082 | LS1_CH01:GV_D2082 |  | gate valve | 40 | 0.057404 |  | 0.057404 | 260.3522 | 423.0996 | 41.094 | 41.094 | 423.0996 | 260.3235 | 208.1612 | 208.1899 |  | 139.6826
0 | [6 empty] | bellows |  |  |  | 0.19049 | 260.2282 | 423.0996 | 41.094 | 41.094 | 423.0996 | 260.133 | 208.2851 | 208.3804 |  | 139.8065
0 | LS1 | CH01 | CAV1 | 2085 | LS1_CH01:CAV1_D2085 |  | b08 resonator | 34 | 0.29817 |  | 0.29817 | 259.9839 | 423.0996 | 41.094 | 41.094 | 423.0996 | 259.8348 | 208.5295 | 208.6785 |  | 140.0508
0 | [6 empty] | bellows |  | 0.099517 |  | 0.099517 | 259.785 | 423.0996 | 41.094 | 41.094 | 423.0996 | 259.7353 | 208.7283 | 208.7781 |  | 140.2497
0 | LS1 | CH01 | CAV2 | 2089 | LS1_CH01:CAV2_D2089 |  | b08 resonator | 34 | 0.29817 |  | 0.29817 | 259.5862 | 423.0996 | 41.094 | 41.094 | 423.0996 | 259.4371 | 208.9272 | 209.0762 |  | 140.4485
0 | [6 empty] | bellows |  |  |  | 0.099517 | 259.3873 | 423.0996 | 41.094 | 41.094 | 423.0996 | 259.3376 | 209.126 | 209.1758 |  | 140.6474
0 | [10 empty] | 0.19798 | 259.2386 | 423.0996 | 41.094 | 41.094 | 423.0996 | 259.1396 | 209.2747 | 209.3737 |  | 140.7961
0 | [6 empty] | bellows |  |  |  | 0.099517 | 259.0898 | 423.0996 | 41.094 | 41.094 | 423.0996 | 259.0401 | 209.4235 | 209.4733 |  | 140.9449
0 | LS1 | CH01 | CAV3 | 2096 | LS1_CH01:CAV3_D2096 |  | b08 resonator | 34 | 0.29817 |  | 0.29817 | 258.891 | 423.0996 | 41.094 | 41.094 | 423.0996 | 258.7419 | 209.6223 | 209.7714 |  | 141.1437
0 | [6 empty] | bellows |  | 0.099517 |  | 0.099517 | 258.6922 | 423.0996 | 41.094 | 41.094 | 423.0996 | 258.6424 | 209.8212 | 209.8709 |  | 141.3426
0 | LS1 | CH01 | CAV4 | 2100 | LS1_CH01:CAV4_D2100 |  | b08 resonator | 34 | 0.29817 |  | 0.29817 | 258.4933 | 423.0996 | 41.094 | 41.094 | 423.0996 | 258.3442 | 210.02 | 210.1691 |  | 141.5414
0 | [6 empty] | bellows |  |  |  | 0.19049 | 258.249 | 423.0996 | 41.094 | 41.094 | 423.0996 | 258.1537 | 210.2644 | 210.3596 |  | 141.7857
0 | LS1 | CH01 | GV | 2104 | LS1_CH01:GV_D2104 |  | gate valve | 40 | 0.057404 |  | 0.057404 | 258.125 | 423.0996 | 41.094 | 41.094 | 423.0996 | 258.0963 | 210.3883 | 210.417 |  | 141.9097
0 | [10 empty] | 0.058708 | 258.067 | 423.0996 | 41.094 | 41.094 | 423.0996 | 258.0376 | 210.4464 | 210.4757 |  | 141.9677
0 | LS1 | BTS | GV | 2105 | LS1_BTS:GV_D2105 | GV_M | gate valve | 63 | 1.9924e-06 | 258.0001 | 0.075 | 258.0001 | 423.0996 | 41.094 | 41.094 | 423.0996 | 257.9626 | 210.5132 | 210.5507 |  | 142.0346
0 | [10 empty] | 0.087596 | 257.9188 | 423.0996 | 41.094 | 41.094 | 423.0996 | 257.875 | 210.5945 | 210.6383 |  | 142.1159
0 | LS1 | BTS | IP | 2106 | LS1_BTS:IP_D2106 | IP_S | ion pump, 75 l/s |  | 257.875 | 258.0376 | 0 | 257.875 | 423.0996 | 41.094 | 41.094 | 423.0996 | 257.875 | 210.6383 | 210.6383 |  | 142.1597
0 | [6 empty] | vacuum box |  |  |  | 0.51924 | 257.6154 | 423.0996 | 41.094 | 41.094 | 423.0996 | 257.3558 | 210.8979 | 211.1576 |  | 142.4193
0 | [10 empty] | 0 | 257.3558 | 423.0996 | 41.094 | 41.094 | 423.0996 | 257.3558 | 211.1576 | 211.1576 |  | 142.6789
0 | [10 empty] | 0.1 | 257.3058 | 423.0996 | 41.094 | 41.094 | 423.0996 | 257.2558 | 211.2076 | 211.2576 |  | 142.7289
0 | LS1 | BTS | DCHV | 2114 | LS1_BTS:DCHV_D2114 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 257.1558 | 423.0996 | 41.094 | 41.094 | 423.0996 | 257.0558 | 211.3576 | 211.4576 |  | 142.8789
0 | [10 empty] | 0 | 257.0558 | 423.0996 | 41.094 | 41.094 | 423.0996 | 257.0558 | 211.4576 | 211.4576 |  | 142.9789
0 | [8 empty] | 2.9924e-06 |  | 0.19218 | 256.9597 | 423.0996 | 41.094 | 41.094 | 423.0996 | 256.8636 | 211.5536 | 211.6497 |  | 143.075
0 | LS1 | BTS | BPM | 2116 | LS1_BTS:BPM_D2116 | BPM3 | position monitor | 47.5 | 256.8636 | 256.8636 | 0 | 256.8636 | 423.0996 | 41.094 | 41.094 | 423.0996 | 256.8636 | 211.6497 | 211.6497 | 0.137 | 143.1711
0 | [10 empty] | 0.057824 | 256.8347 | 423.0996 | 41.094 | 41.094 | 423.0996 | 256.8058 | 211.6786 | 211.7076 |  | 143.2
0 | LS1 | BTS | QV | 2118 | LS1_BTS:QV_D2118 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 256.6808 | 423.0996 | 41.094 | 41.094 | 423.0996 | 256.5558 | 211.8326 | 211.9576 |  | 143.3539
0 | [10 empty] | 0.5 | 256.3058 | 423.0996 | 41.094 | 41.094 | 423.0996 | 256.0558 | 212.2076 | 212.4576 |  | 143.7289
0 | LS1 | BTS | QH | 2126 | LS1_BTS:QH_D2126 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 255.9308 | 423.0996 | 41.094 | 41.094 | 423.0996 | 255.8058 | 212.5826 | 212.7076 |  | 144.1039
0 | [10 empty] | 0.16165 | 255.725 | 423.0996 | 41.094 | 41.094 | 423.0996 | 255.6441 | 212.7884 | 212.8692 |  | 144.3097
0 | [6 empty] | bellow |  |  |  | 0.1035 | 255.5924 | 423.0996 | 41.094 | 41.094 | 423.0996 | 255.5406 | 212.921 | 212.9727 |  | 144.4423
0 | LS1 | BTS | BPM | 2130 | LS1_BTS:BPM_D2130 | BPM2 | position monitor | 40 | 1.9924e-06 | 255.5406 | 0 | 255.5406 | 423.0996 | 41.094 | 41.094 | 423.0996 | 255.5406 | 212.9727 | 212.9727 |  | 144.4941
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 255.468 | 423.0996 | 41.094 | 41.094 | 423.0996 | 255.3954 | 213.0453 | 213.118 |  | 144.5667
0 | LS1 | BTS | PM | 2131 | LS1_BTS:PM_D2131 | PM2 | profile monitor | defer installation |  |  | 0 | 255.3954 | 423.0996 | 41.094 | 41.094 | 423.0996 | 255.3954 | 213.118 | 213.118 |  | 144.6394
0 | [6 empty] | diagnostic box |  | 1.9924e-06 | 255.3954 | 0 | 255.3954 | 423.0996 | 41.094 | 41.094 | 423.0996 | 255.3954 | 213.118 | 213.118 |  | 144.6394
0 | LS1 | BTS | IP | 2131 | LS1_BTS:IP_D2131 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 255.3954 | 423.0996 | 41.094 | 41.094 | 423.0996 | 255.3954 | 213.118 | 213.118 |  | 144.6394
0 | [6 empty] | bellow |  | 0.16514 |  | 0.13122 | 255.3297 | 423.0996 | 41.094 | 41.094 | 423.0996 | 255.2641 | 213.1836 | 213.2492 |  | 144.705
0 | [8 empty] | 1.9924e-06 | 255.1755 | 0.088633 | 255.2198 | 423.0996 | 41.094 | 41.094 | 423.0996 | 255.1755 | 213.2935 | 213.3378 |  | 144.8149
0 | FS1 | CH01 | GV | 2134 | FS1_CH02:GV_D2134 |  | gate valve | 40 | 0.057404 | 2.2845 | 0.057404 | 255.1468 | 423.0996 | 41.094 | 41.094 | 423.0996 | 255.1181 | 213.3665 | 213.3952 | b85 buncher moved 0.165123 m upstream | 144.8879
0 | [6 empty] | bellows |  |  |  | 0.19049 | 255.0229 | 423.0996 | 41.094 | 41.094 | 423.0996 | 254.9276 | 213.4905 | 213.5857 |  | 145.0119
0 | FS1 | CH01 | CAV1 | 2137 | FS1_CH02:CAV1_D2137 |  | b08 resonator | 34 | 0.29817 |  | 0.29817 | 254.7785 | 423.0996 | 41.094 | 41.094 | 423.0996 | 254.6294 | 213.7348 | 213.8839 |  | 145.2562
0 | [6 empty] | bellows |  | 0.099517 |  | 0.099517 | 254.5797 | 423.0996 | 41.094 | 41.094 | 423.0996 | 254.5299 | 213.9337 | 213.9834 |  | 145.455
0 | FS1 | CH01 | CAV2 | 2141 | FS1_CH02:CAV2_D2141 |  | b08 resonator | 34 | 0.29817 |  | 0.29817 | 254.3808 | 423.0996 | 41.094 | 41.094 | 423.0996 | 254.2318 | 214.1325 | 214.2816 |  | 145.6539
0 | [6 empty] | bellows |  |  |  | 0.099517 | 254.182 | 423.0996 | 41.094 | 41.094 | 423.0996 | 254.1322 | 214.3313 | 214.3811 |  | 145.8527
0 | [10 empty] | 0.19798 | 254.0332 | 423.0996 | 41.094 | 41.094 | 423.0996 | 253.9343 | 214.4801 | 214.5791 |  | 146.0015
0 | [6 empty] | bellows |  |  |  | 0.099517 | 253.8845 | 423.0996 | 41.094 | 41.094 | 423.0996 | 253.8347 | 214.6288 | 214.6786 |  | 146.1502
0 | FS1 | CH01 | CAV3 | 2148 | FS1_CH02:CAV3_D2148 |  | b08 resonator | 34 | 0.29817 |  | 0.29817 | 253.6856 | 423.0996 | 41.094 | 41.094 | 423.0996 | 253.5366 | 214.8277 | 214.9768 |  | 146.3491
0 | [6 empty] | bellows |  | 0.099517 |  | 0.099517 | 253.4868 | 423.0996 | 41.094 | 41.094 | 423.0996 | 253.437 | 215.0265 | 215.0763 |  | 146.5479
0 | FS1 | CH01 | CAV4 | 2152 | FS1_CH02:CAV4_D2152 |  | b08 resonator | 34 | 0.29817 |  | 0.29817 | 253.288 | 423.0996 | 41.094 | 41.094 | 423.0996 | 253.1389 | 215.2254 | 215.3745 |  | 146.7468
0 | [6 empty] | bellows |  |  |  | 0.19049 | 253.0436 | 423.0996 | 41.094 | 41.094 | 423.0996 | 252.9484 | 215.4697 | 215.565 |  | 146.9911
0 | FS1 | CH01 | GV | 2156 | FS1_CH02:GV_D2156 |  | gate valve | 40 | 0.057404 |  | 0.057404 | 252.9197 | 423.0996 | 41.094 | 41.094 | 423.0996 | 252.891 | 215.5937 | 215.6224 |  | 147.115
0 | [8 empty] | -7.6277e-09 | 252.891 | 0.16514 | 252.8084 | 423.0996 | 41.094 | 41.094 | 423.0996 | 252.7258 | 215.7049 | 215.7875 |  | 147.2263
0 | [6 empty] | bellow |  |  |  | 0.011054 | 252.7203 | 423.0996 | 41.094 | 41.094 | 423.0996 | 252.7148 | 215.793 | 215.7986 |  | 147.3144
0 | [6 empty] | vacuum box |  |  |  | 0 | 252.7148 | 423.0996 | 41.094 | 41.094 | 423.0996 | 252.7148 | 215.7986 | 215.7986 |  | 147.3199
0 | FS1 | CSS | FAVS | 2158 | FS1_CSS:FAVS_D2158 | FVS | fast valve sensor |  | 0.031522 |  | 0 | 252.7148 | 423.0996 | 41.094 | 41.094 | 423.0996 | 252.7148 | 215.7986 | 215.7986 |  | 147.3199
0 | [6 empty] | vacuum box |  | 2.1992e-05 | 252.7148 | 0 | 252.7148 | 423.0996 | 41.094 | 41.094 | 423.0996 | 252.7148 | 215.7986 | 215.7986 |  | 147.3199
0 | FS1 | CSS | IP | 2158 | FS1_CSS:IP_D2158 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 252.7148 | 423.0996 | 41.094 | 41.094 | 423.0996 | 252.7148 | 215.7986 | 215.7986 |  | 147.3199
0 | [6 empty] | bellow |  | 0.027624 | 0.076642 | 0.16868 | 252.6304 | 423.0996 | 41.094 | 41.094 | 423.0996 | 252.5461 | 215.8829 | 215.9672 |  | 147.4043
0 | FS1 | CSS | FAV | 2160 | FS1_CSS:FAV_D2160 | FV | fast valve | 40 | 2.1992e-05 | 252.5461 | 0 | 252.5461 | 423.0996 | 41.094 | 41.094 | 423.0996 | 252.5461 | 215.9672 | 215.9672 | 0.067196 | 147.4886
0 | [10 empty] | 0.20307 | 252.4446 | 423.0996 | 41.094 | 41.094 | 423.0996 | 252.343 | 216.0688 | 216.1703 |  | 147.5901
0 | FS1 | CSS | DH | 2163 | FS1_CSS:DH_D2163 | D1 | dipole-5deg, D1-1, C1 | 40 x 40 |  |  | 0.3 | 252.1932 | 423.0931 | 41.094 | 41.094 | 423.0865 | 252.0434 | 216.3203 | 216.4703 |  | 147.8417
0 | [10 empty] | 0.25 | 251.9189 | 423.0756 | 41.094 | 41.094 | 423.0647 | 251.7944 | 216.5953 | 216.7203 |  | 148.1167
0 | FS1 | CSS | DH | 2169 | FS1_CSS:DH_D2169 | D1 | dipole-5deg, D1-2, C1 | 40 x 40 |  |  | 0.3 | 251.6446 | 423.0582 | 41.094 | 41.094 | 423.0517 | 251.4947 | 216.8703 | 217.0203 |  | 148.3917
0 | [10 empty] | 0.091402 | 251.449 | 423.0517 | 41.094 | 41.094 | 423.0517 | 251.4033 | 217.066 | 217.1117 |  | 148.5874
0 | [10 empty] | 0.067196 | 251.3697 | 423.0517 | 41.094 | 41.094 | 423.0517 | 251.3362 | 217.1453 | 217.1789 | 20170214 moved upstream | 148.6667
0 | [10 empty] | 0.091402 | 251.2905 | 423.0517 | 41.094 | 41.094 | 423.0517 | 251.2447 | 217.2246 | 217.2703 |  | 148.746
0 | FS1 | CSS | DH | 2174 | FS1_CSS:DH_D2174 | D1 | dipole-5deg, D1-3, C1 | 40 x 40 |  |  | 0.3 | 251.0949 | 423.0582 | 41.094 | 41.094 | 423.0647 | 250.9451 | 217.4203 | 217.5703 |  | 148.9417
0 | [10 empty] | 0.25 | 250.8206 | 423.0756 | 41.094 | 41.094 | 423.0865 | 250.6961 | 217.6953 | 217.8203 |  | 149.2167
0 | FS1 | CSS | DH | 2180 | FS1_CSS:DH_D2180 | D1 | dipole-5deg, D1-4, C1 | 40 x 40 |  |  | 0.3 | 250.5463 | 423.0931 | 41.094 | 41.094 | 423.0996 | 250.3965 | 217.9703 | 218.1203 |  | 149.4917
0 | [6 empty] | coil-out (assumed) |  |  |  | 0.06 | 250.3665 | 423.0996 | 41.094 | 41.094 | 423.0996 | 250.3365 | 218.1503 | 218.1803 |  | 149.6717
0 | [8 empty] | 2.2834e-05 |  | 0.11723 | 250.2778 | 423.0996 | 41.094 | 41.094 | 423.0996 | 250.2192 | 218.2389 | 218.2975 |  | 149.7603
0 | FS1 | CSS | BCM | 2183 | FS1_CSS:BCM_D2183 | BCM2 | current monitor |  | 250.1693 | 250.1692 | 0.1 | 250.1692 | 423.0996 | 41.094 | 41.094 | 423.0996 | 250.1192 | 218.3475 | 218.3975 |  | 149.8689
0 | [10 empty] | 0.17244 | 250.033 | 423.0996 | 41.094 | 41.094 | 423.0996 | 249.9468 | 218.4838 | 218.57 |  | 150.0051
0 | FS1 | CSS | TMP | 2186 | FS1_CSS:TMP_D2186 | TMP_L | turbo pump 2050 l/s |  | 2.233e-05 | 249.9468 | 0 | 249.9468 | 423.0996 | 41.094 | 41.094 | 423.0996 | 249.9468 | 218.57 | 218.57 |  | 150.0913
0 | [10 empty] | 0.13933 | 249.8771 | 423.0996 | 41.094 | 41.094 | 423.0996 | 249.8075 | 218.6396 | 218.7093 |  | 150.161
0 | [6 empty] | bellow |  |  |  | 0.1 | 249.7575 | 423.0996 | 41.094 | 41.094 | 423.0996 | 249.7075 | 218.7593 | 218.8093 |  | 150.2807
0 | [6 empty] | coil-out |  |  |  | 0.036 | 249.6895 | 423.0996 | 41.094 | 41.094 | 423.0996 | 249.6715 | 218.8273 | 218.8453 |  | 150.3487
0 | FS1 | CSS | DCHV | 2189 | FS1_CSS:DCHV_D2189 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 249.5715 | 423.0996 | 41.094 | 41.094 | 423.0996 | 249.4715 | 218.9453 | 219.0453 |  | 150.4667
0 | [10 empty] | 0.25 | 249.3465 | 423.0996 | 41.094 | 41.094 | 423.0996 | 249.2215 | 219.1703 | 219.2953 |  | 150.6917
0 | FS1 | CSS | QH | 2194 | FS1_CSS:QH_D2194 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 249.0965 | 423.0996 | 41.094 | 41.094 | 423.0996 | 248.9715 | 219.4203 | 219.5453 |  | 150.9417
0 | [10 empty] | 0.0475 | 248.9477 | 423.0996 | 41.094 | 41.094 | 423.0996 | 248.924 | 219.5691 | 219.5928 |  | 151.0904
0 | [10 empty] | 0 | 248.924 | 423.0996 | 41.094 | 41.094 | 423.0996 | 248.924 | 219.5928 | 219.5928 |  | 151.1142
0 | [6 empty] | diagnostic box |  | 2.2834e-05 |  | 0.19515 | 248.8264 | 423.0996 | 41.094 | 41.094 | 423.0996 | 248.7288 | 219.6904 | 219.788 |  | 151.2117
0 | FS1 | CSS | LPM | 2198 | FS1_CSS:LPM_D2198 | LPMx | bunch length monitor (no costed) | future LPM3 | 248.7288 | 248.7288 | 0 | 248.7288 | 423.0996 | 41.094 | 41.094 | 423.0996 | 248.7288 | 219.788 | 219.788 |  | 151.3093
0 | FS1 | CSS | PM | 2198 | FS1_CSS:PM_D2198 | PM3 | profile monitor | defer installation |  |  | 0 | 248.7288 | 423.0996 | 41.094 | 41.094 | 423.0996 | 248.7288 | 219.788 | 219.788 |  | 151.3093
0 | [10 empty] | 0.12238 | 248.6676 | 423.0996 | 41.094 | 41.094 | 423.0996 | 248.6064 | 219.8491 | 219.9103 |  | 151.3705
0 | FS1 | CSS | GV | 2199 | FS1_CSS:GV_D2199 | GV_M | gate valve | 63 | 2.1834e-05 | 248.569 | 0.075 | 248.5689 | 423.0996 | 41.094 | 41.094 | 423.0996 | 248.5314 | 219.9478 | 219.9853 |  | 151.4692
0 | [10 empty] | 0.097816 | 248.4825 | 423.0996 | 41.094 | 41.094 | 423.0996 | 248.4336 | 220.0342 | 220.0832 |  | 151.5556
0 | FS1 | CSS | QV | 2202 | FS1_CSS:QV_D2202 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 248.3086 | 423.0996 | 41.094 | 41.094 | 423.0996 | 248.1836 | 220.2082 | 220.3332 |  | 151.7295
0 | [10 empty] | 0.13552 | 248.1159 | 423.0996 | 41.094 | 41.094 | 423.0996 | 248.0481 | 220.4009 | 220.4687 |  | 151.9223
0 | [10 empty] | 0.15102 | 247.9726 | 423.0996 | 41.094 | 41.094 | 423.0996 | 247.8971 | 220.5442 | 220.6197 |  | 152.0656
0 | FS1 | CSS | TMP | 2206 | FS1_CSS:TMP_D2206 | TMP_L | turbo pump 2050 l/s |  | 2.3334e-05 |  | 0 | 247.8971 | 423.0996 | 41.094 | 41.094 | 423.0996 | 247.8971 | 220.6197 | 220.6197 |  | 152.1411
0 | FS1 | CSS | FAVS | 2206 | FS1_CSS:FAVS_D2206 | FVS | fast valve sensor |  | 247.8971 | 247.8971 | 0 | 247.8971 | 423.0996 | 41.094 | 41.094 | 423.0996 | 247.8971 | 220.6197 | 220.6197 |  | 152.1411
0 | [10 empty] | 0.26346 | 247.7653 | 423.0996 | 41.094 | 41.094 | 423.0996 | 247.6336 | 220.7514 | 220.8832 |  | 152.2728
0 | FS1 | CSS | DCHV | 2210 | FS1_CSS:DCHV_D2210 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 247.5336 | 423.0996 | 41.094 | 41.094 | 423.0996 | 247.4336 | 220.9832 | 221.0832 |  | 152.5045
0 | [8 empty] | 2.2834e-05 |  | 0.027106 | 247.4201 | 423.0996 | 41.094 | 41.094 | 423.0996 | 247.4065 | 221.0967 | 221.1103 |  | 152.6181
0 | FS1 | CSS | BPM | 2212 | FS1_CSS:BPM_D2212 | BPM3 | position monitor | 47.5 | 247.338 | 247.338 | 0.137 | 247.338 | 423.0996 | 41.094 | 41.094 | 423.0996 | 247.2695 | 221.1788 | 221.2473 |  | 152.7001
0 | [10 empty] | 0.085894 | 247.2266 | 423.0996 | 41.094 | 41.094 | 423.0996 | 247.1836 | 221.2902 | 221.3332 |  | 152.8116
0 | FS1 | CSS | QH | 2215 | FS1_CSS:QH_D2215 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 247.0586 | 423.0996 | 41.094 | 41.094 | 423.0996 | 246.9336 | 221.4582 | 221.5832 |  | 152.9795
0 | [10 empty] | 0.25 | 246.8086 | 423.0996 | 41.094 | 41.094 | 423.0996 | 246.6836 | 221.7082 | 221.8332 |  | 153.2295
0 | FS1 | CSS | QV | 2220 | FS1_CSS:QV_D2220 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 246.5586 | 423.0996 | 41.094 | 41.094 | 423.0996 | 246.4336 | 221.9582 | 222.0832 |  | 153.4795
0 | [10 empty] | 0 | 246.4336 | 423.0996 | 41.094 | 41.094 | 423.0996 | 246.4336 | 222.0832 | 222.0832 |  | 153.6045
0 | [8 empty] | 2.2834e-05 |  | 0.074011 | 246.3966 | 423.0996 | 41.094 | 41.094 | 423.0996 | 246.3596 | 222.1202 | 222.1572 |  | 153.6415
0 | FS1 | CSS | GV | 2222 | FS1_CSS:GV_D2222 | GV_M | gate valve | 63 | 246.3221 | 246.3221 | 0.075 | 246.3221 | 423.0996 | 41.094 | 41.094 | 423.0996 | 246.2846 | 222.1947 | 222.2322 |  | 153.716
0 | [9 empty] | 0.045989 | 0 | 246.2846 | 423.0996 | 41.094 | 41.094 | 423.0996 | 246.2846 | 222.2322 | 222.2322 |  | 153.7535
0 | [6 empty] | bellow |  | 0.054792 |  | 0.094699 | 246.2373 | 423.0996 | 41.094 | 41.094 | 423.0996 | 246.1899 | 222.2795 | 222.3269 | bellow compressed by 0.008803 m, 4/24/2017 | 153.8009
0 | FS1 | CSS | BPM | 2223 | FS1_CSS:BPM_D2223 | BPM3 | position monitor |  | 2.1834e-05 | 246.1899 | 0 | 246.1899 | 423.0996 | 41.094 | 41.094 | 423.0996 | 246.1899 | 222.3269 | 222.3269 |  | 153.8482
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 246.1173 | 423.0996 | 41.094 | 41.094 | 423.0996 | 246.0446 | 222.3995 | 222.4721 |  | 153.9209
0 | FS1 | CSS | PM | 2225 | FS1_CSS:PM_D2225 | PM6 | profile monitor (small) |  |  |  | 0 | 246.0446 | 423.0996 | 41.094 | 41.094 | 423.0996 | 246.0446 | 222.4721 | 222.4721 |  | 153.9935
0 | [6 empty] | diagnostic box |  |  |  | 0 | 246.0446 | 423.0996 | 41.094 | 41.094 | 423.0996 | 246.0446 | 222.4721 | 222.4721 |  | 153.9935
0 | FS1 | CSS | IP | 2225 | FS1_CSS:IP_D2225 | IP_S | ion pump, 75 l/s |  | 2.1834e-05 | 246.0446 | 0 | 246.0446 | 423.0996 | 41.094 | 41.094 | 423.0996 | 246.0446 | 222.4721 | 222.4721 |  | 153.9935
0 | [10 empty] | 0 | 246.0446 | 423.0996 | 41.094 | 41.094 | 423.0996 | 246.0446 | 222.4721 | 222.4721 |  | 153.9935
0 | [6 empty] | bellow |  |  |  | 0.13122 | 245.979 | 423.0996 | 41.094 | 41.094 | 423.0996 | 245.9134 | 222.5378 | 222.6034 |  | 154.0591
0 | [10 empty] | 0.054792 | 245.886 | 423.0996 | 41.094 | 41.094 | 423.0996 | 245.8586 | 222.6308 | 222.6582 |  | 154.1521
0 | FS1 | STRL | GV | 2227 | FS1_STRL:GV_D2227 |  | gate valve |  |  |  | 0.072 | 245.8226 | 423.0996 | 41.094 | 41.094 | 423.0996 | 245.7866 | 222.6942 | 222.7302 | stripper module length: 1.975607 | 154.2155
0 | [6 empty] | stripper module |  |  |  | 0.528 | 245.5226 | 423.0996 | 41.094 | 41.094 | 423.0996 | 245.2586 | 222.9942 | 223.2582 |  | 154.5155
0 | FS1 | STRL | CP | 2233 | FS1_STRL:CP_D2233 | CP_L | cryo pump, 1500 l/s |  |  |  | 0 | 245.2586 | 423.0996 | 41.094 | 41.094 | 423.0996 | 245.2586 | 223.2582 | 223.2582 |  | 154.7795
0 | [10 empty] | 0.4 | 245.0586 | 423.0996 | 41.094 | 41.094 | 423.0996 | 244.8586 | 223.4582 | 223.6582 |  | 154.9795
0 | FS1 | STRL | STRIP | 2237 | FS1_STRL:STRIP_D2237 |  | lithium film stripper |  |  |  | 0 | 244.8586 | 423.0996 | 41.094 | 41.094 | 423.0996 | 244.8586 | 223.6582 | 223.6582 |  | 155.1795
0 | [10 empty] | 0.4 | 244.6586 | 423.0996 | 41.094 | 41.094 | 423.0996 | 244.4586 | 223.8582 | 224.0582 |  | 155.3795
0 | FS1 | STRL | CP | 2241 | FS1_STRL:CP_D2241 | CP_L | cryo pump, 1500 l/s |  |  |  | 0 | 244.4586 | 423.0996 | 41.094 | 41.094 | 423.0996 | 244.4586 | 224.0582 | 224.0582 |  | 155.5795
0 | [6 empty] | stripper module |  |  |  | 0.528 | 244.1946 | 423.0996 | 41.094 | 41.094 | 423.0996 | 243.9306 | 224.3222 | 224.5862 |  | 155.8435
0 | FS1 | STRL | GV | 2246 | FS1_STRL:GV_D2246 |  | gate valve |  |  |  | 0.072 | 243.8946 | 423.0996 | 41.094 | 41.094 | 423.0996 | 243.8586 | 224.6222 | 224.6582 |  | 156.1435
0 | [10 empty] | 0.008836 | 243.8542 | 423.0996 | 41.094 | 41.094 | 423.0996 | 243.8498 | 224.6626 | 224.667 | stripper module increased by 0.008836 m, 4/24/2017 | 156.1839
0 | [6 empty] | bellow |  | -0.008836 |  | 0.1035 | 243.798 | 423.0996 | 41.094 | 41.094 | 423.0996 | 243.7463 | 224.7187 | 224.7705 |  | 156.2401
0 | FS1 | CSS | BPM | 2248 | FS1_CSS:BPM_D2248 | BPM3 | position monitor |  | 2.1834e-05 | 243.7463 | 0 | 243.7463 | 423.0996 | 41.094 | 41.094 | 423.0996 | 243.7463 | 224.7705 | 224.7705 |  | 156.2919
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 243.6736 | 423.0996 | 41.094 | 41.094 | 423.0996 | 243.601 | 224.8431 | 224.9158 |  | 156.3645
0 | FS1 | CSS | PM | 2249 | FS1_CSS:PM_D2249 | PM6 | profile monitor | defer installation |  |  | 0 | 243.601 | 423.0996 | 41.094 | 41.094 | 423.0996 | 243.601 | 224.9158 | 224.9158 |  | 156.4371
0 | [6 empty] | diagnostic box |  |  |  | 0 | 243.601 | 423.0996 | 41.094 | 41.094 | 423.0996 | 243.601 | 224.9158 | 224.9158 |  | 156.4371
0 | FS1 | CSS | IP | 2249 | FS1_CSS:IP_D2249 | IP_S | ion pump, 75 l/s |  | 2.1834e-05 | 243.601 | 0 | 243.601 | 423.0996 | 41.094 | 41.094 | 423.0996 | 243.601 | 224.9158 | 224.9158 |  | 156.4371
0 | [10 empty] | 0 | 243.601 | 423.0996 | 41.094 | 41.094 | 423.0996 | 243.601 | 224.9158 | 224.9158 |  | 156.4371
0 | [6 empty] | bellow |  |  |  | 0.12238 | 243.5398 | 423.0996 | 41.094 | 41.094 | 423.0996 | 243.4786 | 224.977 | 225.0382 | bellow compressed by 0.008836 m, 4/24/2017 | 156.4983
0 | FS1 | CSS | GV | 2251 | FS1_CSS:GV_D2251 | GV_M | gate valve | 63 |  |  | 0.075 | 243.4411 | 423.0996 | 41.094 | 41.094 | 423.0996 | 243.4036 | 225.0757 | 225.1132 |  | 156.597
0 | [10 empty] | 0.12 | 243.3436 | 423.0996 | 41.094 | 41.094 | 423.0996 | 243.2836 | 225.1732 | 225.2332 |  | 156.6945
0 | FS1 | CSS | QV | 2254 | FS1_CSS:QV_D2254 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 243.1586 | 423.0996 | 41.094 | 41.094 | 423.0996 | 243.0336 | 225.3582 | 225.4832 |  | 156.8795
0 | [8 empty] | 2.2834e-05 |  | 0.092506 | 242.9874 | 423.0996 | 41.094 | 41.094 | 423.0996 | 242.9411 | 225.5294 | 225.5757 |  | 157.0508
0 | FS1 | CSS | DCHV | 2257 | FS1_CSS:DCHV_D2257 | C2 | correctors-X,Y | 50 | 242.8411 | 242.8411 | 0.2 | 242.8411 | 423.0996 | 41.094 | 41.094 | 423.0996 | 242.7411 | 225.6757 | 225.7757 |  | 157.197
0 | [10 empty] | 0 | 242.7411 | 423.0996 | 41.094 | 41.094 | 423.0996 | 242.7411 | 225.7757 | 225.7757 |  | 157.297
0 | [10 empty] | 0.10749 | 242.6874 | 423.0996 | 41.094 | 41.094 | 423.0996 | 242.6336 | 225.8294 | 225.8832 |  | 157.3508
0 | [10 empty] | 0 | 242.6336 | 423.0996 | 41.094 | 41.094 | 423.0996 | 242.6336 | 225.8832 | 225.8832 |  | 157.4045
0 | FS1 | CSS | QH | 2260 | FS1_CSS:QH_D2260 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 242.5086 | 423.0996 | 41.094 | 41.094 | 423.0996 | 242.3836 | 226.0082 | 226.1332 |  | 157.5295
0 | [6 empty] | coil out + leads |  |  |  | 0.098 | 242.3346 | 423.0996 | 41.094 | 41.094 | 423.0996 | 242.2856 | 226.1822 | 226.2312 |  | 157.7035
0 | [8 empty] | 2.2834e-05 |  | 0.084251 | 242.2435 | 423.0996 | 41.094 | 41.094 | 423.0996 | 242.2014 | 226.2733 | 226.3154 |  | 157.7946
0 | FS1 | CSS | LPM | 2263 | FS1_CSS:LPM_D2263 | LPM3 | bunch length monitor |  | 242.2014 | 242.2014 | 0 | 242.2014 | 423.0996 | 41.094 | 41.094 | 423.0996 | 242.2014 | 226.3154 | 226.3154 |  | 157.8368
0 | [8 empty] | 2.2834e-05 |  | 0.08375 | 242.1595 | 423.0996 | 41.094 | 41.094 | 423.0996 | 242.1176 | 226.3573 | 226.3992 |  | 157.8787
0 | FS1 | CSS | BCM | 2264 | FS1_CSS:BCM_D2264 | BCM4 | current monitor |  | 242.0676 | 242.0676 | 0.1 | 242.0676 | 423.0996 | 41.094 | 41.094 | 423.0996 | 242.0176 | 226.4492 | 226.4992 |  | 157.9705
0 | [10 empty] | 0.20355 | 241.9158 | 423.0996 | 41.094 | 41.094 | 423.0996 | 241.8141 | 226.6009 | 226.7027 |  | 158.1223
0 | FS1 | CSS | TMP | 2267 | FS1_CSS:TMP_D2267 | TMP_L | turbo pump 2050 l/s |  | 2.2834e-05 |  | 0 | 241.8141 | 423.0996 | 41.094 | 41.094 | 423.0996 | 241.8141 | 226.7027 | 226.7027 |  | 158.2241
0 | FS1 | CSS | FAVS | 2267 | FS1_CSS:FAVS_D2267 | FVS | fast valve sensor |  | 241.8141 | 241.8141 | 0 | 241.8141 | 423.0996 | 41.094 | 41.094 | 423.0996 | 241.8141 | 226.7027 | 226.7027 |  | 158.2241
0 | [7 empty] | 0.69455 |  |  | 0 | 241.8141 | 423.0996 | 41.094 | 41.094 | 423.0996 | 241.8141 | 226.7027 | 226.7027 |  | 158.2241
0 | [6 empty] | collimation flange |  |  |  | 0.28245 | 241.6728 | 423.0996 | 41.094 | 41.094 | 423.0996 | 241.5316 | 226.8439 | 226.9852 |  | 158.3653
0 | [6 empty] | coil out + leads |  |  |  | 0.098 | 241.4826 | 423.0996 | 41.094 | 41.094 | 423.0996 | 241.4336 | 227.0342 | 227.0832 |  | 158.5555
0 | FS1 | CSS | QV | 2272 | FS1_CSS:QV_D2272 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 241.3086 | 423.0996 | 41.094 | 41.094 | 423.0996 | 241.1836 | 227.2082 | 227.3332 |  | 158.7295
0 | [8 empty] | 2.2834e-05 | -0.050806 | 0.13866 | 241.1143 | 423.0996 | 41.094 | 41.094 | 423.0996 | 241.045 | 227.4025 | 227.4718 |  | 158.9239
0 | FS1 | CSS | DCHV | 2276 | FS1_CSS:DCHV_D2276 | C2 | correctors-X,Y | 50 | 240.945 | 240.945 | 0.2 | 240.945 | 423.0996 | 41.094 | 41.094 | 423.0996 | 240.845 | 227.5718 | 227.6718 |  | 159.0932
0 | [8 empty] | 2.1834e-05 | -0.052701 | 0.13939 | 240.7753 | 423.0996 | 41.094 | 41.094 | 423.0996 | 240.7056 | 227.7415 | 227.8112 |  | 159.2629
0 | FS1 | CSS | BPM | 2278 | FS1_CSS:BPM_D2278 | BPM3 | position monitor | 47.5 | 240.7056 | 240.7056 | 0 | 240.7056 | 423.0996 | 41.094 | 41.094 | 423.0996 | 240.7056 | 227.8112 | 227.8112 | 0.137 | 159.3326
0 | [8 empty] | 0.068501 | -0.008701 | 0.0598 | 240.6757 | 423.0996 | 41.094 | 41.094 | 423.0996 | 240.6458 | 227.8411 | 227.871 |  | 159.3625
0 | FS1 | CSS | QH | 2280 | FS1_CSS:QH_D2280 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 240.5208 | 423.0996 | 41.094 | 41.094 | 423.0996 | 240.3958 | 227.996 | 228.121 |  | 159.5174
0 | [10 empty] | 0.28 | 240.2558 | 423.0996 | 41.094 | 41.094 | 423.0996 | 240.1158 | 228.261 | 228.401 |  | 159.7824
0 | [8 empty] | 2.2834e-05 |  | 0.082672 | 240.0744 | 423.0996 | 41.094 | 41.094 | 423.0996 | 240.0331 | 228.4423 | 228.4837 |  | 159.9637
0 | FS1 | CSS | TMP | 2285 | FS1_CSS:TMP_D2285 | TMP_L | turbo pump 2050 l/s |  | 240.0331 | 240.0331 | 0 | 240.0331 | 423.0996 | 41.094 | 41.094 | 423.0996 | 240.0331 | 228.4837 | 228.4837 |  | 160.005
0 | [10 empty] | 0.2199 | 239.9231 | 423.0996 | 41.094 | 41.094 | 423.0996 | 239.8132 | 228.5936 | 228.7036 |  | 160.115
0 | [10 empty] | 0 | 239.8132 | 423.0996 | 41.094 | 41.094 | 423.0996 | 239.8132 | 228.7036 | 228.7036 |  | 160.2249
0 | [8 empty] | -0.017571 |  | 0 | 239.8132 | 423.0996 | 41.094 | 41.094 | 423.0996 | 239.8132 | 228.7036 | 228.7036 |  | 160.2249
0 | [6 empty] | collimation flange |  | 2.1834e-05 | 239.8132 | 0 | 239.8132 | 423.0996 | 41.094 | 41.094 | 423.0996 | 239.8132 | 228.7036 | 228.7036 |  | 160.2249
0 | [10 empty] | 0.18302 | 239.7217 | 423.0996 | 41.094 | 41.094 | 423.0996 | 239.6302 | 228.7951 | 228.8866 |  | 160.3165
0 | [8 empty] | 2.3521e-05 | 0.2294 | 0 | 239.6302 | 423.0996 | 41.094 | 41.094 | 423.0996 | 239.6302 | 228.8866 | 228.8866 | 20170215 D1-1 & D1-2 moved 0.229406 m upstream | 160.408
0 | FS1 | CSS | DH | 2290 | FS1_CSS:DH_D2290 | D1 | dipole-5deg, D1-1, C2 | 40 x 40 | 239.4804 | 239.4804 | 0.3 | 239.4804 | 423.0931 | 41.094 | 41.094 | 423.0865 | 239.3305 | 229.0366 | 229.1866 |  | 160.558
0 | [8 empty] | 2.3568e-05 |  | 0.25 | 239.206 | 423.0756 | 41.094 | 41.094 | 423.0647 | 239.0815 | 229.3116 | 229.4366 |  | 160.833
0 | FS1 | CSS | DH | 2296 | FS1_CSS:DH_D2296 | D1 | dipole-5deg, D1-1, C2 | 40 x 40 | 238.9317 | 238.9317 | 0.3 | 238.9317 | 423.0582 | 41.094 | 41.094 | 423.0517 | 238.7819 | 229.5866 | 229.7366 |  | 161.108
0 | [10 empty] | 0.082804 | 238.7405 | 423.0517 | 41.094 | 41.094 | 423.0517 | 238.6991 | 229.778 | 229.8194 |  | 161.2994
0 | [6 empty] | 2nd chicane not installed initially, replaced by a pipe |  |  |  | 0.1 | 238.6491 | 423.0517 | 41.094 | 41.094 | 423.0517 | 238.5991 | 229.8694 | 229.9194 |  | 161.3908
0 | [10 empty] | 0.067196 | 238.5655 | 423.0517 | 41.094 | 41.094 | 423.0517 | 238.5319 | 229.953 | 229.9866 |  | 161.4744
0 | [8 empty] | 2.3942e-05 |  | 0.1 | 238.4819 | 423.0517 | 41.094 | 41.094 | 423.0517 | 238.4319 | 230.0366 | 230.0866 |  | 161.558
0 | FS1 | CSS | DH | 2302 | FS1_CSS:DH_D2302 | D1 | dipole-5deg, D1-1, C2 | 40 x 40 | 238.2821 | 238.2821 | 0.3 | 238.2821 | 423.0582 | 41.094 | 41.094 | 423.0647 | 238.1323 | 230.2366 | 230.3866 |  | 161.758
0 | [8 empty] | 2.3989e-05 |  | 0.25 | 238.0077 | 423.0756 | 41.094 | 41.094 | 423.0865 | 237.8832 | 230.5116 | 230.6366 | 20170215 added 0.2294 m | 162.033
0 | FS1 | CSS | DH | 2308 | FS1_CSS:DH_D2308 | D1 | dipole-5deg, D1-1, C2 | 40 x 40 | 237.7334 | 237.7334 | 0.3 | 237.7334 | 423.0931 | 41.094 | 41.094 | 423.0996 | 237.5836 | 230.7866 | 230.9366 |  | 162.308
0 | [7 empty] | 237.3497 | 2.3676e-05 | -0.19191 | 0.19645 | 237.4854 | 423.0996 | 41.094 | 41.094 | 423.0996 | 237.3871 | 231.0348 | 231.133 |  | 162.5562
0 | FS1 | CSS | FAV | 2312 | FS1_CSS:FAV_D2312 | FV | fast valve | 40 | 237.3536 | 237.3535 | 0.067196 | 237.3535 | 423.0996 | 41.094 | 41.094 | 423.0996 | 237.32 | 231.1666 | 231.2002 |  | 162.688
0 | [7 empty] | FV still required ? |  |  | 0.003898 | 237.318 | 423.0996 | 41.094 | 41.094 | 423.0996 | 237.3161 | 231.2022 | 231.2041 |  | 162.7236
0 | [6 empty] | bellow | 40 | 2.3676e-05 | 0.0070924 | 0.1035 | 237.2643 | 423.0996 | 41.094 | 41.094 | 423.0996 | 237.2126 | 231.2559 | 231.3076 |  | 162.7773
0 | FS1 | CSS | BPM | 2313 | FS1_CSS:BPM_D2313 | BPM2 | position monitor | 237.3497 | 237.2126 | 237.2126 | 0 | 237.2126 | 423.0996 | 41.094 | 41.094 | 423.0996 | 237.2126 | 231.3076 | 231.3076 |  | 162.829
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 237.1399 | 423.0996 | 41.094 | 41.094 | 423.0996 | 237.0673 | 231.3803 | 231.4529 | 0.39757 | 162.9017
0 | FS1 | CSS | PM | 2315 | FS1_CSS:PM_D2315 | PMx | profile monitor (not costed) | future PM |  |  | 0 | 237.0673 | 423.0996 | 41.094 | 41.094 | 423.0996 | 237.0673 | 231.4529 | 231.4529 | 2.2 | 162.9743
0 | [6 empty] | diagnostic box |  | 2.3676e-05 |  | 0 | 237.0673 | 423.0996 | 41.094 | 41.094 | 423.0996 | 237.0673 | 231.4529 | 231.4529 | 1.4 | 162.9743
0 | FS1 | CSS | IP | 2315 | FS1_CSS:IP_D2315 | IP_S | pump, 75 l/s |  | 237.0673 | 237.0673 | 0 | 237.0673 | 423.0996 | 41.094 | 41.094 | 423.0996 | 237.0673 | 231.4529 | 231.4529 | 0.253 | 162.9743
0 | FS1 | CSS | FAVS | 2315 | FS1_CSS:FAVS_D2315 | FVS | fast valve sensor |  |  |  | 0 | 237.0673 | 423.0996 | 41.094 | 41.094 | 423.0996 | 237.0673 | 231.4529 | 231.4529 |  | 162.9743
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 237.0017 | 423.0996 | 41.094 | 41.094 | 423.0996 | 236.9361 | 231.5185 | 231.5841 |  | 163.0399
0 | [10 empty] | 0.017429 | 236.9273 | 423.0996 | 41.094 | 41.094 | 423.0996 | 236.9186 | 231.5929 | 231.6016 | 20170216 reduced 0.022604 m | 163.1142
0 | [6 empty] | collimation flange |  | 2.2676e-05 | 236.9186 | 0 | 236.9186 | 423.0996 | 41.094 | 41.094 | 423.0996 | 236.9186 | 231.6016 | 231.6016 |  | 163.1229
0 | [10 empty] | 0.096757 | 236.8702 | 423.0996 | 41.094 | 41.094 | 423.0996 | 236.8219 | 231.6499 | 231.6983 |  | 163.1713
0 | [9 empty] | enough or not ? | 0.1 | 236.7719 | 423.0996 | 41.094 | 41.094 | 423.0996 | 236.7219 | 231.7483 | 231.7983 |  | 163.2697
0 | FS1 | MGB01 | GV | 2318 | FS1_MGB01:GV_D2318 |  | gate valve | 40 | -0.34819 | 236.345 | 0.057404 | 236.6932 | 423.0996 | 41.094 | 41.094 | 423.0996 | 236.6645 | 231.827 | 231.8557 |  | 163.3484
0 | [10 empty] | 0 | 236.6645 | 423.0996 | 41.094 | 41.094 | 423.0996 | 236.6645 | 231.8557 | 231.8557 |  | 163.3771
0 | FS1 | MGB01 | CAV | 2320 | FS1_MGB01:CAV_D2320 |  | b19 resonator 1 1/4 | 36 |  |  | 0.31833 | 236.5053 | 423.0996 | 41.094 | 41.094 | 423.0996 | 236.3461 | 232.0149 | 232.1741 |  | 163.5363
0 | FS1 | MGB01 | IP | 2322 | FS1_MGB01:IP_D2322 | IP_L | ion pump, ? l/s |  |  |  | 0 | 236.3461 | 423.0996 | 41.094 | 41.094 | 423.0996 | 236.3461 | 232.1741 | 232.1741 |  | 163.6954
0 | FS1 | MGB01 | CAV | 2323 | FS1_MGB01:CAV_D2323 |  | b19 resonator 1 2/4 | 36 |  |  | 0.31833 | 236.187 | 423.0996 | 41.094 | 41.094 | 423.0996 | 236.0278 | 232.3332 | 232.4924 |  | 163.8546
0 | FS1 | MGB01 | CAV | 2325 | FS1_MGB01:CAV_D2325 |  | rf coupler |  | -0.34819 | 235.6796 | 0 | 236.0278 | 423.0996 | 41.094 | 41.094 | 423.0996 | 236.0278 | 232.4924 | 232.4924 |  | 164.0137
0 | [6 empty] | This buncher will not be installed initially by replacing a pipe |  |  |  | 0 | 236.0278 | 423.0996 | 41.094 | 41.094 | 423.0996 | 236.0278 | 232.4924 | 232.4924 |  | 164.0137
0 | FS1 | MGB01 | CAV | 2327 | FS1_MGB01:CAV_D2327 |  | b19 resonator 1 3/4 | 36 |  |  | 0.31833 | 235.8686 | 423.0996 | 41.094 | 41.094 | 423.0996 | 235.7095 | 232.6515 | 232.8107 |  | 164.1729
0 | FS1 | MGB01 | TMP | 2328 | FS1_MGB01:TMP_D2330 | TMP_L | turbo pump, ? l/s |  |  |  | 0 | 235.7095 | 423.0996 | 41.094 | 41.094 | 423.0996 | 235.7095 | 232.8107 | 232.8107 |  | 164.3321
0 | FS1 | MGB01 | CAV | 2330 | FS1_MGB01:CAV_D2330 |  | b19 resonator 1 4/4 | 36 |  |  | 0.31833 | 235.5503 | 423.0996 | 41.094 | 41.094 | 423.0996 | 235.3912 | 232.9699 | 233.129 |  | 164.4912
0 | [10 empty] | 0 | 235.3912 | 423.0996 | 41.094 | 41.094 | 423.0996 | 235.3912 | 233.129 | 233.129 |  | 164.6504
0 | FS1 | MGB01 | GV | 2332 | FS1_MGB01:GV_D2332 |  | gate valve | 40 | -0.34819 | 235.0143 | 0.057404 | 235.3625 | 423.0996 | 41.094 | 41.094 | 423.0996 | 235.3338 | 233.1577 | 233.1864 |  | 164.6791
0 | [10 empty] | 0.91761 | 234.875 | 423.0996 | 41.094 | 41.094 | 423.0996 | 234.4162 | 233.6452 | 234.104 | diag bx moved down by 0.12119 m, 4/24/2017 | 165.1666
0 | [6 empty] | bellow | 40 |  | 0.13122 | 0.152 | 234.3402 | 423.0996 | 41.094 | 41.094 | 423.0996 | 234.2642 | 234.18 | 234.256 |  | 165.7014
0 | [6 empty] | vacuum box |  |  |  | 0 | 234.2642 | 423.0996 | 41.094 | 41.094 | 423.0996 | 234.2642 | 234.256 | 234.256 |  | 165.7774
0 | FS1 | CSS | FAVS | 2343 | FS1_CSS:FAVS_D2343 | FVS | fast valve sensor |  |  |  | 0 | 234.2642 | 423.0996 | 41.094 | 41.094 | 423.0996 | 234.2642 | 234.256 | 234.256 |  | 165.7774
0 | [6 empty] | vacuum box |  |  |  | 0 | 234.2642 | 423.0996 | 41.094 | 41.094 | 423.0996 | 234.2642 | 234.256 | 234.256 |  | 165.7774
0 | FS1 | CSS | IP | 2343 | FS1_CSS:IP_D2343 | IP_S | ion pump, 75 l/s |  | 2.2676e-05 | 234.2642 | 0 | 234.2642 | 423.0996 | 41.094 | 41.094 | 423.0996 | 234.2642 | 234.256 | 234.256 |  | 165.7774
0 | [6 empty] | bellow | 40 |  |  | 0.132 | 234.1982 | 423.0996 | 41.094 | 41.094 | 423.0996 | 234.1322 | 234.322 | 234.388 |  | 165.8434
0 | [10 empty] | 0.001052 | 234.1316 | 423.0996 | 41.094 | 41.094 | 423.0996 | 234.1311 | 234.3886 | 234.3891 |  | 165.9099
0 | [8 empty] | -0.038013 |  | 0.015594 | 234.1233 | 423.0996 | 41.094 | 41.094 | 423.0996 | 234.1155 | 234.3969 | 234.4047 |  | 165.9183
0 | [6 empty] | collimation flange |  | 2.2676e-05 | 234.1155 | 0 | 234.1155 | 423.0996 | 41.094 | 41.094 | 423.0996 | 234.1155 | 234.4047 | 234.4047 |  | 165.9261
0 | [10 empty] | 0.30958 | 233.9607 | 423.0996 | 41.094 | 41.094 | 423.0996 | 233.8059 | 234.5595 | 234.7143 |  | 166.0808
0 | [10 empty] | 0 | 233.8059 | 423.0996 | 41.094 | 41.094 | 423.0996 | 233.8059 | 234.7143 | 234.7143 |  | 166.2356
0 | [10 empty] | 0 | 233.8059 | 423.0996 | 41.094 | 41.094 | 423.0996 | 233.8059 | 234.7143 | 234.7143 |  | 166.2356
0 | [7 empty] | FV still required ? |  |  | 0.069596 | 233.7711 | 423.0996 | 41.094 | 41.094 | 423.0996 | 233.7363 | 234.7491 | 234.7839 |  | 166.2704
0 | FS1 | CSS | FAV | 2348 | FS1_CSS:FAV_D2348 | FV | fast valve | 40 | 0.067196 |  | 0.067196 | 233.7027 | 423.0996 | 41.094 | 41.094 | 423.0996 | 233.6691 | 234.8175 | 234.8511 |  | 166.3388
0 | [10 empty] | 0.14724 | 233.5955 | 423.0996 | 41.094 | 41.094 | 423.0996 | 233.5219 | 234.9247 | 234.9983 |  | 166.4461
0 | FS1 | CSS | DCHV | 2351 | FS1_CSS:DCHV_D2351 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 233.4219 | 423.0996 | 41.094 | 41.094 | 423.0996 | 233.3219 | 235.0983 | 235.1983 |  | 166.6197
0 | [10 empty] | 0.1067 | 233.2685 | 423.0996 | 41.094 | 41.094 | 423.0996 | 233.2152 | 235.2517 | 235.305 |  | 166.773
0 | FS1 | CSS | BCM | 2353 | FS1_CSS:BCM_D2353 | BCM4 | current monitor |  | 2.2676e-05 | 233.2152 | 0 | 233.2152 | 423.0996 | 41.094 | 41.094 | 423.0996 | 233.2152 | 235.305 | 235.305 |  | 166.8264
0 | [10 empty] | 0.1433 | 233.1435 | 423.0996 | 41.094 | 41.094 | 423.0996 | 233.0719 | 235.3767 | 235.4483 |  | 166.898
0 | FS1 | CSS | QH | 2356 | FS1_CSS:QH_D2356 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 232.9469 | 423.0996 | 41.094 | 41.094 | 423.0996 | 232.8219 | 235.5733 | 235.6983 |  | 167.0947
0 | [6 empty] | coil out |  |  |  | 0.047 | 232.7984 | 423.0996 | 41.094 | 41.094 | 423.0996 | 232.7749 | 235.7218 | 235.7453 |  | 167.2432
0 | [10 empty] | 0.19117 | 232.6793 | 423.0996 | 41.094 | 41.094 | 423.0996 | 232.5837 | 235.8409 | 235.9365 | 20170216 space adjusted | 167.3623
0 | [6 empty] | collimation flange |  | 2.2676e-05 | 232.5837 | 0 | 232.5837 | 423.0996 | 41.094 | 41.094 | 423.0996 | 232.5837 | 235.9365 | 235.9365 |  | 167.4578
0 | [10 empty] | 0.038832 | 232.5643 | 423.0996 | 41.094 | 41.094 | 423.0996 | 232.5449 | 235.9559 | 235.9753 |  | 167.4773
0 | [6 empty] | coil out + leads |  |  |  | 0.098 | 232.4959 | 423.0996 | 41.094 | 41.094 | 423.0996 | 232.4469 | 236.0243 | 236.0733 |  | 167.5457
0 | FS1 | CSS | QH | 2362 | FS1_CSS:QH_D2362 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 232.3219 | 423.0996 | 41.094 | 41.094 | 423.0996 | 232.1969 | 236.1983 | 236.3233 |  | 167.7197
0 | [10 empty] | 0 | 232.1969 | 423.0996 | 41.094 | 41.094 | 423.0996 | 232.1969 | 236.3233 | 236.3233 |  | 167.8447
0 | [10 empty] | 0 | 232.1969 | 423.0996 | 41.094 | 41.094 | 423.0996 | 232.1969 | 236.3233 | 236.3233 |  | 167.8447
0 | [10 empty] | 0.3 | 232.0469 | 423.0996 | 41.094 | 41.094 | 423.0996 | 231.8969 | 236.4733 | 236.6233 |  | 167.9947
0 | FS1 | CSS | DCHV | 2367 | FS1_CSS:DCHV_D2367 | C2 | correctors-X,Y | 50 |  | 0.15635 | 0.2 | 231.7969 | 423.0996 | 41.094 | 41.094 | 423.0996 | 231.6969 | 236.7233 | 236.8233 |  | 168.2447
0 | [8 empty] | 2.3676e-05 | 0.08115 | -0.01215 | 231.703 | 423.0996 | 41.094 | 41.094 | 423.0996 | 231.709 | 236.8172 | 236.8112 |  | 168.3386
0 | FS1 | CSS | BPM | 2369 | FS1_CSS:BPM_D2369 | BPM3 | position monitor | 47.5 | 231.6406 | 231.6405 | 0.137 | 231.6405 | 423.0996 | 41.094 | 41.094 | 423.0996 | 231.572 | 236.8797 | 236.9482 |  | 168.401
0 | [10 empty] | 0.12515 | 231.5095 | 423.0996 | 41.094 | 41.094 | 423.0996 | 231.4469 | 237.0107 | 237.0733 | 20170216 space adjusted | 168.5321
0 | FS1 | CSS | QV | 2372 | FS1_CSS:QV_D2372 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 231.3219 | 423.0996 | 41.094 | 41.094 | 423.0996 | 231.1969 | 237.1983 | 237.3233 |  | 168.7197
0 | [10 empty] | 0.25 | 231.0719 | 423.0996 | 41.094 | 41.094 | 423.0996 | 230.9469 | 237.4483 | 237.5733 |  | 168.9697
0 | FS1 | CSS | QH | 2377 | FS1_CSS:QH_D2377 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 230.8219 | 423.0996 | 41.094 | 41.094 | 423.0996 | 230.6969 | 237.6983 | 237.8233 |  | 169.2197
0 | [10 empty] | 0.145 | 230.6244 | 423.0996 | 41.094 | 41.094 | 423.0996 | 230.5519 | 237.8958 | 237.9683 |  | 169.4172
0 | FS1 | CSS | DCHV | 2381 | FS1_CSS:DCHV_D2381 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 230.4519 | 423.0996 | 41.094 | 41.094 | 423.0996 | 230.3519 | 238.0683 | 238.1683 |  | 169.5897
0 | [10 empty] | 0.059488 | 230.3221 | 423.0996 | 41.094 | 41.094 | 423.0996 | 230.2924 | 238.198 | 238.2278 | 20170216 space adjusted | 169.7194
0 | [6 empty] | bellow | 40 | 2.3676e-05 | 0.033167 | 0.1035 | 230.2407 | 423.0996 | 41.094 | 41.094 | 423.0996 | 230.1889 | 238.2795 | 238.3313 |  | 169.8009
0 | FS1 | CSS | BPM | 2383 | FS1_CSS:BPM_D2383 | BPM3 | position monitor |  | 230.1889 | 230.1889 | 0 | 230.1889 | 423.0996 | 41.094 | 41.094 | 423.0996 | 230.1889 | 238.3313 | 238.3313 |  | 169.8527
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 230.1163 | 423.0996 | 41.094 | 41.094 | 423.0996 | 230.0436 | 238.4039 | 238.4766 |  | 169.9253
0 | FS1 | CSS | PM | 2385 | FS1_CSS:PM_D2385 | PM3 | profile monitor |  |  |  | 0 | 230.0436 | 423.0996 | 41.094 | 41.094 | 423.0996 | 230.0436 | 238.4766 | 238.4766 |  | 169.9979
0 | [6 empty] | diagnostic box |  | 2.4676e-05 |  | 0 | 230.0436 | 423.0996 | 41.094 | 41.094 | 423.0996 | 230.0436 | 238.4766 | 238.4766 |  | 169.9979
0 | FS1 | CSS | FAVS | 2385 | FS1_CSS:FAVS_D2385 | FVS | fast valve sensor |  | 230.0436 | 230.0436 | 0 | 230.0436 | 423.0996 | 41.094 | 41.094 | 423.0996 | 230.0436 | 238.4766 | 238.4766 |  | 169.9979
0 | FS1 | CSS | IP | 2385 | FS1_CSS:IP_D2384 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 230.0436 | 423.0996 | 41.094 | 41.094 | 423.0996 | 230.0436 | 238.4766 | 238.4766 |  | 169.9979
0 | FS1 | CSS | IP | 2385 | FS1_CSS:IP_D2385 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 230.0436 | 423.0996 | 41.094 | 41.094 | 423.0996 | 230.0436 | 238.4766 | 238.4766 |  | 169.9979
0 | [6 empty] | bellow | 40 | 2.3676e-05 | 0.008831 | 0.12238 | 229.9824 | 423.0996 | 41.094 | 41.094 | 423.0996 | 229.9212 | 238.5378 | 238.599 | 20170216 space adjusted | 170.0591
0 | FS1 | CSS | GV | 2386 | FS1_CSS:GV_D2386 | GV_M | gate valve | 63 | 229.8838 | 229.8837 | 0.075 | 229.8837 | 423.0996 | 41.094 | 41.094 | 423.0996 | 229.8462 | 238.6365 | 238.674 |  | 170.1578
0 | First beam dump line starts in "V" column | [9 empty] | 0.18422 | 229.7541 | 423.0996 | 41.094 | 41.094 | 423.0996 | 229.662 | 238.7661 | 238.8582 | adjustment for residuals | 170.2874
0 | FS1 | BBS | DH | 2394 | FS1_BBS:DH_D2394 | D2 | dipole-45deg, D2-1 | 40 x 100 |  |  | 1.0407 | 229.1936 | 422.9056 | 41.094 | 41.094 | 422.7115 | 228.7251 | 239.3785 | 239.8988 |  | 170.8999
0 | [10 empty] | 0.63 | 228.5024 | 422.4888 | 41.094 | 41.094 | 422.266 | 228.2796 | 240.2138 | 240.5288 |  | 171.7352
0 | FS1 | BBS | SLH | 2405 | FS1_BBS:SLH_D2405 |  | charge selection slit high power |  |  |  | 0 | 228.2796 | 422.266 | 41.094 | 41.094 | 422.266 | 228.2796 | 240.5288 | 240.5288 |  | 172.0502
0 | [10 empty] | 0 | 228.2796 | 422.266 | 41.094 | 41.094 | 422.266 | 228.2796 | 240.5288 | 240.5288 |  | 172.0502
0 | FS1 | BBS | TMP | 2405 | FS1_BBS:TMP_D2405 | TMP_L | turbo pump 2050 l/s |  |  |  | 0 | 228.2796 | 422.266 | 41.094 | 41.094 | 422.266 | 228.2796 | 240.5288 | 240.5288 |  | 172.0502
0 | [10 empty] | 0.58 | 228.0746 | 422.061 | 41.094 | 41.094 | 421.8559 | 227.8695 | 240.8188 | 241.1088 |  | 172.3402
0 | FS1 | BBS | DCHV | 2412 | FS1_BBS:DCHV_D2412 | C1 | correctors-X,Y | 140 |  |  | 0.2 | 227.7988 | 421.7852 | 41.094 | 41.094 | 421.7145 | 227.7281 | 241.2088 | 241.3088 |  | 172.7302
0 | [10 empty] | 0.15 | 227.675 | 421.6615 | 41.094 | 41.094 | 421.6084 | 227.622 | 241.3838 | 241.4588 |  | 172.9052
0 | FS1 | BBS | QH | 2416 | FS1_BBS:QH_D2416 | Q6 | quadrupole-X, B1 | 40 x 140 |  |  | 0.25 | 227.5336 | 421.52 | 41.094 | 41.094 | 421.4317 | 227.4452 | 241.5838 | 241.7088 |  | 173.1052
0 | [10 empty] | 0 | 227.4452 | 421.4317 | 41.094 | 41.094 | 421.4317 | 227.4452 | 241.7088 | 241.7088 |  | 173.2302
0 | [10 empty] | 0.14 | 227.3957 | 421.3822 | 41.094 | 41.094 | 421.3327 | 227.3462 | 241.7788 | 241.8488 |  | 173.3002
0 | FS1 | BBS | S | 2419 | FS1_BBS:S_D2419 | H3 | hexapole-X+ | 140 |  |  | 0.15 | 227.2932 | 421.2796 | 41.094 | 41.094 | 421.2266 | 227.2402 | 241.9238 | 241.9988 |  | 173.4452
0 | [8 empty] | -3.5182e-06 |  | 0.10618 | 227.2026 | 421.1891 | 41.094 | 41.094 | 421.1515 | 227.1651 | 242.0519 | 242.105 |  | 173.5733
0 | FS1 | BBS | BPM | 2421 | FS1_BBS:BPM_D2421 | BPM5 | position monitor 142x42 | 133.6 | 227.1353 | 227.1353 | 0.084138 | 227.1353 | 421.1218 | 41.094 | 41.094 | 421.092 | 227.1056 | 242.1471 | 242.1891 | 20170216 move up 0.03175 m | 173.6684
0 | [10 empty] | 0.11968 | 227.0633 | 421.0497 | 41.094 | 41.094 | 421.0074 | 227.021 | 242.249 | 242.3088 |  | 173.7704
0 | FS1 | BBS | QV | 2424 | FS1_BBS:QV_D2424 | Q6 | quadrupole-Y, B1 | 40 x 140 |  |  | 0.25 | 226.9326 | 420.919 | 41.094 | 41.094 | 420.8306 | 226.8442 | 242.4338 | 242.5588 |  | 173.9552
0 | [10 empty] | 0.40322 | 226.7016 | 420.6881 | 41.094 | 41.094 | 420.5455 | 226.5591 | 242.7604 | 242.9621 |  | 174.2818
0 | [10 empty] | 0 | 226.5591 | 420.5455 | 41.094 | 41.094 | 420.5455 | 226.5591 | 242.9621 | 242.9621 |  | 174.4834
0 | [10 empty] | 0 | 226.5591 | 420.5455 | 41.094 | 41.094 | 420.5455 | 226.5591 | 242.9621 | 242.9621 |  | 174.4834
0 | [10 empty] | 0 | 226.5591 | 420.5455 | 41.094 | 41.094 | 420.5455 | 226.5591 | 242.9621 | 242.9621 |  | 174.4834
0 | Second beam dump line starts in "V" column | [9 empty] | 0 | 226.5591 | 420.5455 | 41.094 | 41.094 | 420.5455 | 226.5591 | 242.9621 | 242.9621 |  | 174.4834
0 | FS1 | BBS | DH | 2435 | FS1_BBS:DH_D2435 | D2 | dipole-45deg, D2-2 | 40 x 100 |  |  | 1.0407 | 226.365 | 420.077 | 41.094 | 41.094 | 419.6086 | 226.171 | 243.4824 | 244.0027 |  | 175.0037
0 | [10 empty] | 0 | 226.171 | 419.6086 | 41.094 | 41.094 | 419.6086 | 226.171 | 244.0027 | 244.0027 |  | 175.5241
0 | [10 empty] | 0.38904 | 226.171 | 419.4141 | 41.094 | 41.094 | 419.2195 | 226.171 | 244.1972 | 244.3917 |  | 175.7186
0 | FS1 | BBS | PM | 2444 | FS1_BBS:PM_D2444 | PM4 | profile monitor | 150 | 0.00092926 |  | 0 | 226.171 | 419.2195 | 41.094 | 41.094 | 419.2195 | 226.171 | 244.3917 | 244.3917 | 20170216 moved dwon 0.004971 m | 175.9131
0 | [10 empty] | 0 | 226.171 | 419.2195 | 41.094 | 41.094 | 419.2195 | 226.171 | 244.3917 | 244.3917 |  | 175.9131
0 | [10 empty] | 0 | 226.171 | 419.2195 | 41.094 | 41.094 | 419.2195 | 226.171 | 244.3917 | 244.3917 |  | 175.9131
0 | FS1 | BBS | IP | 2444 | FS1_BBS:IP_D2444 | IP_L | ion pump, 300 l/s |  | 0 | 419.2195 | 0 | 226.171 | 419.2195 | 41.094 | 41.094 | 419.2195 | 226.171 | 244.3917 | 244.3917 |  | 175.9131
0 | FS1 | BBS | FAVS | 2444 | FS1_BBS:FAVS_D2444 | FVS | fast valve sensor |  |  |  | 0 | 226.171 | 419.2195 | 41.094 | 41.094 | 419.2195 | 226.171 | 244.3917 | 244.3917 |  | 175.9131
0 | [10 empty] | 0.38096 | 226.171 | 419.0291 | 41.094 | 41.094 | 418.8386 | 226.171 | 244.5822 | 244.7727 |  | 176.1036
0 | [10 empty] | 0 | 226.171 | 418.8386 | 41.094 | 41.094 | 418.8386 | 226.171 | 244.7727 | 244.7727 |  | 176.2941
0 | FS1 | BBS | DH | 2453 | FS1_BBS:DH_D2453 | D2 | dipole-45deg, D2-3 | 40 x 100 |  |  | 1.0407 | 226.365 | 418.3701 | 41.094 | 41.094 | 417.9017 | 226.5591 | 245.293 | 245.8134 |  | 176.8144
0 | [10 empty] | 0.1873 | 226.6253 | 417.8354 | 41.094 | 41.094 | 417.7692 | 226.6915 | 245.907 | 246.0007 |  | 177.4284
0 | [6 empty] | collimation flange |  | 1.4378e-06 | 226.6915 | 0 | 226.6915 | 417.7692 | 41.094 | 41.094 | 417.7692 | 226.6915 | 246.0007 | 246.0007 |  | 177.522
0 | [8 empty] | 0.358 | 0.34093 | 0.21593 | 226.7679 | 417.6929 | 41.094 | 41.094 | 417.6165 | 226.8442 | 246.1086 | 246.2166 |  | 177.63
0 | FS1 | BBS | QV | 2463 | FS1_BBS:QV_D2463 | Q6 | quadrupole-Y, B1 | 40 x 140 |  | 0.125 | 0.25 | 226.9326 | 417.5282 | 41.094 | 41.094 | 417.4398 | 227.021 | 246.3416 | 246.4666 |  | 177.8629
0 | [10 empty] | 0.11968 | 227.0633 | 417.3975 | 41.094 | 41.094 | 417.3551 | 227.1056 | 246.5264 | 246.5863 |  | 178.0478
0 | FS1 | BBS | BPM | 2466 | FS1_BBS:BPM_D2466 | BPM5 | position monitor R=75mm | 133.6 |  |  | 0.084138 | 227.1353 | 417.3254 | 41.094 | 41.094 | 417.2956 | 227.1651 | 246.6283 | 246.6704 |  | 178.1497
0 | [6 empty] | collimation flange |  |  |  | 0.10618 | 227.2026 | 417.2581 | 41.094 | 41.094 | 417.2206 | 227.2402 | 246.7235 | 246.7766 |  | 178.2449
0 | FS1 | BBS | S | 2469 | FS1_BBS:S_D2469 | H3 | hexapole-X+ | 140 |  |  | 0.15 | 227.2932 | 417.1675 | 41.094 | 41.094 | 417.1145 | 227.3462 | 246.8516 | 246.9266 |  | 178.3729
0 | [10 empty] | 0.14 | 227.3957 | 417.065 | 41.094 | 41.094 | 417.0155 | 227.4452 | 246.9966 | 247.0666 |  | 178.5179
0 | FS1 | BBS | QH | 2472 | FS1_BBS:QH_D2472 | Q6 | quadrupole-Y, B1 | 40 x 140 |  |  | 0.25 | 227.5336 | 416.9271 | 41.094 | 41.094 | 416.8387 | 227.622 | 247.1916 | 247.3166 |  | 178.7129
0 | [10 empty] | 0.15 | 227.675 | 416.7857 | 41.094 | 41.094 | 416.7327 | 227.7281 | 247.3916 | 247.4666 |  | 178.9129
0 | FS1 | BBS | DCHV | 2476 | FS1_BBS:DCHV_D2476 | C1 | correctors-X,Y | 140 |  |  | 0.2 | 227.7988 | 416.6619 | 41.094 | 41.094 | 416.5912 | 227.8695 | 247.5666 | 247.6666 |  | 179.0879
0 | [10 empty] | 0.58 | 228.0746 | 416.3862 | 41.094 | 41.094 | 416.1811 | 228.2796 | 247.9566 | 248.2466 |  | 179.4779
0 | FS1 | BBS | PM | 2482 | FS1_BBS:PM_D2482 | PM4 | profile monitor | 150 |  |  | 0 | 228.2796 | 416.1811 | 41.094 | 41.094 | 416.1811 | 228.2796 | 248.2466 | 248.2466 |  | 179.7679
0 | FS1 | BBS | IP | 2482 | FS1_BBS:IP_D2482 | IP_L | ion pump, 300 l/s |  |  |  | 0 | 228.2796 | 416.1811 | 41.094 | 41.094 | 416.1811 | 228.2796 | 248.2466 | 248.2466 |  | 179.7679
0 | [10 empty] | 0.63 | 228.5024 | 415.9584 | 41.094 | 41.094 | 415.7356 | 228.7251 | 248.5616 | 248.8766 |  | 180.0829
0 | FS1 | BBS | DH | 2494 | FS1_BBS:DH_D2494 | D2 | dipole-45deg, D2-4 | 40 x 100 |  |  | 1.0407 | 229.1936 | 415.5416 | 41.094 | 41.094 | 415.3476 | 229.662 | 249.3969 | 249.9172 |  | 180.9183
0 | [8 empty] | 3.176e-06 | 0.016423 | 0.15866 | 229.7413 | 415.3476 | 41.094 | 41.094 | 415.3476 | 229.8207 | 249.9966 | 250.0759 |  | 181.5179
0 | FS1 | BMS | GV | 2501 | FS1_BMS:GV_D2501 | GV_M | gate valve | 63 | 229.8582 | 229.8582 | 0.075 | 229.8582 | 415.3476 | 41.094 | 41.094 | 415.3476 | 229.8957 | 250.1134 | 250.1509 | 20170216 seems X of GV incorrect | 181.6348
0 | [10 empty] | 0 | 229.8957 | 415.3476 | 41.094 | 41.094 | 415.3476 | 229.8957 | 250.1509 | 250.1509 |  | 181.6723
0 | [6 empty] | bellow | 40 | -0.008801 |  | 0.094701 | 229.943 | 415.3476 | 41.094 | 41.094 | 415.3476 | 229.9904 | 250.1982 | 250.2456 | bellow compressed by 0.008801 m, 4/24/2017 | 181.7196
0 | FS1 | BMS | BPM | 2502 | FS1_BMS:BPM_D2502 | BPM3 | position monitor | 50 | 2.176e-06 | 229.9904 | 0 | 229.9904 | 415.3476 | 41.094 | 41.094 | 415.3476 | 229.9904 | 250.2456 | 250.2456 |  | 181.767
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 230.063 | 415.3476 | 41.094 | 41.094 | 415.3476 | 230.1357 | 250.3182 | 250.3909 |  | 181.8396
0 | FS1 | BMS | PM | 2504 | FS1_BMS:PM_D2504 | PM3 | profile monitor | 50 | defer installation |  | 0 | 230.1357 | 415.3476 | 41.094 | 41.094 | 415.3476 | 230.1357 | 250.3909 | 250.3909 |  | 181.9122
0 | [6 empty] | diagnostic box |  |  |  | 0 | 230.1357 | 415.3476 | 41.094 | 41.094 | 415.3476 | 230.1357 | 250.3909 | 250.3909 |  | 181.9122
0 | FS1 | BMS | FAVS | 2504 | FS1_BMS:FAVS_D2504 | FVS | fast valve sensor |  |  |  | 0 | 230.1357 | 415.3476 | 41.094 | 41.094 | 415.3476 | 230.1357 | 250.3909 | 250.3909 |  | 181.9122
0 | FS1 | BMS | IP | 2504 | FS1_BMS:IP_D2504 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 230.1357 | 415.3476 | 41.094 | 41.094 | 415.3476 | 230.1357 | 250.3909 | 250.3909 |  | 181.9122
0 | [6 empty] | bellow | 40 | 0.085138 |  | 0.13122 | 230.2013 | 415.3476 | 41.094 | 41.094 | 415.3476 | 230.2669 | 250.4565 | 250.5221 |  | 181.9779
0 | [10 empty] | 0.01743 | 230.2756 | 415.3476 | 41.094 | 41.094 | 415.3476 | 230.2843 | 250.5308 | 250.5395 |  | 182.0522
0 | [6 empty] | collimation flange |  | 2.176e-06 | 230.2843 | 0 | 230.2843 | 415.3476 | 41.094 | 41.094 | 415.3476 | 230.2843 | 250.5395 | 250.5395 | 20170216 collimation flange moved from upstream | 182.0609
0 | [10 empty] | 0.067708 | 230.3182 | 415.3476 | 41.094 | 41.094 | 415.3476 | 230.352 | 250.5734 | 250.6072 |  | 182.0947
0 | FS1 | BMS | DCHV | 2507 | FS1_BMS:DCHV_D2507 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 230.452 | 415.3476 | 41.094 | 41.094 | 415.3476 | 230.552 | 250.7072 | 250.8072 |  | 182.2286
0 | [10 empty] | 0.145 | 230.6245 | 415.3476 | 41.094 | 41.094 | 415.3476 | 230.697 | 250.8797 | 250.9522 |  | 182.4011
0 | FS1 | BMS | QV | 2511 | FS1_BMS:QV_D2511 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 230.822 | 415.3476 | 41.094 | 41.094 | 415.3476 | 230.947 | 251.0772 | 251.2022 |  | 182.5986
0 | [6 empty] | collimation flange moved?? |  |  |  | 0.2 | 231.047 | 415.3476 | 41.094 | 41.094 | 415.3476 | 231.147 | 251.3022 | 251.4022 |  | 182.8236
0 | FS1 | BMS | QH | 2515 | FS1_BMS:QH_D2515 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 231.272 | 415.3476 | 41.094 | 41.094 | 415.3476 | 231.397 | 251.5272 | 251.6522 |  | 183.0486
0 | [8 empty] | 3.176e-06 | 0.20683 | 0.23183 | 231.5129 | 415.3476 | 41.094 | 41.094 | 415.3476 | 231.6288 | 251.7681 | 251.8841 | 20170216 space adjusted | 183.2895
0 | FS1 | BMS | BCM | 2519 | FS1_BMS:BCM_D2519 | BCM3 | current monitor |  | 231.6788 | 231.6788 | 0.1 | 231.6788 | 415.3476 | 41.094 | 41.094 | 415.3476 | 231.7288 | 251.9341 | 251.9841 |  | 183.4554
0 | [10 empty] | 1.1932 | 232.3254 | 415.3476 | 41.094 | 41.094 | 415.3476 | 232.922 | 252.5806 | 253.1772 | 20170216 space adjusted | 184.102
0 | [6 empty] | collimation flange moved?? |  |  |  | 0.1 | 232.972 | 415.3476 | 41.094 | 41.094 | 415.3476 | 233.022 | 253.2272 | 253.2772 |  | 184.7486
0 | FS1 | BMS | DCHV | 2534 | FS1_BMS:DCHV_D2534 | C2 | correctors-X,Y | 50 |  | 0.29242 | 0.2 | 233.122 | 415.3476 | 41.094 | 41.094 | 415.3476 | 233.222 | 253.3772 | 253.4772 |  | 184.8986
0 | [8 empty] | 3.176e-06 | 0.054921 | 0.12392 | 233.284 | 415.3476 | 41.094 | 41.094 | 415.3476 | 233.3459 | 253.5392 | 253.6012 | 20170216 space adjusted | 185.0606
0 | FS1 | BMS | BPM | 2537 | FS1_BMS:BPM_D2537 | BPM3 | position monitor | 47.5 | 233.4144 | 233.4144 | 0.137 | 233.4144 | 415.3476 | 41.094 | 41.094 | 415.3476 | 233.4829 | 253.6697 | 253.7382 |  | 185.191
0 | [9 empty] | 0.18258 | -0.010921 | 233.4775 | 415.3476 | 41.094 | 41.094 | 415.3476 | 233.472 | 253.7327 | 253.7272 | 20170216 space adjusted | 185.2541
0 | FS1 | BMS | QV | 2539 | FS1_BMS:QV_D2539 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 233.597 | 415.3476 | 41.094 | 41.094 | 415.3476 | 233.722 | 253.8522 | 253.9772 |  | 185.3736
0 | [10 empty] | 1.2 | 234.322 | 415.3476 | 41.094 | 41.094 | 415.3476 | 234.922 | 254.5772 | 255.1772 |  | 186.0986
0 | FS1 | BMS | IP | 2552 | FS1_BMS:IP_D2552 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 234.922 | 415.3476 | 41.094 | 41.094 | 415.3476 | 234.922 | 255.1772 | 255.1772 |  | 186.6986
0 | [10 empty] | 1 | 235.422 | 415.3476 | 41.094 | 41.094 | 415.3476 | 235.922 | 255.6772 | 256.1772 |  | 187.1986
0 | FS1 | BMS | QH | 2563 | FS1_BMS:QH_D2563 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 236.047 | 415.3476 | 41.094 | 41.094 | 415.3476 | 236.172 | 256.3022 | 256.4272 |  | 187.8236
0 | [10 empty] | 0 | 236.172 | 415.3476 | 41.094 | 41.094 | 415.3476 | 236.172 | 256.4272 | 256.4272 |  | 187.9486
0 | [6 empty] | collimation flange |  |  |  | 0.1 | 236.222 | 415.3476 | 41.094 | 41.094 | 415.3476 | 236.272 | 256.4772 | 256.5272 |  | 187.9986
0 | [8 empty] | 3.176e-06 |  | 0.004193 | 236.2741 | 415.3476 | 41.094 | 41.094 | 415.3476 | 236.2762 | 256.5293 | 256.5314 | 20170216 space adjusted | 188.0507
0 | FS1 | BMS | GV | 2566 | FS1_BMS:GV_D2566 | GV_M | gate valve | 63 | 236.3137 | 236.3137 | 0.075 | 236.3137 | 415.3476 | 41.094 | 41.094 | 415.3476 | 236.3512 | 256.5689 | 256.6064 | 20170216 GV inserted | 188.0903
0 | [8 empty] | 3.076e-06 |  | 1.7184 | 237.2104 | 415.3476 | 41.094 | 41.094 | 415.3476 | 238.0696 | 257.4656 | 258.3248 | 20170216 space adjusted | 188.987
0 | FS1 | BMS | DCHV | 2584 | FS1_BMS:DCHV_D2584 | C2 | correctors-X,Y | 50 | 238.1696 | 238.1696 | 0.2 | 238.1696 | 415.3476 | 41.094 | 41.094 | 415.3476 | 238.2696 | 258.4248 | 258.5248 |  | 189.9462
0 | [8 empty] | 3.076e-06 |  | 0.073805 | 238.3065 | 415.3476 | 41.094 | 41.094 | 415.3476 | 238.3434 | 258.5617 | 258.5986 | 20170216 space adjusted | 190.0831
0 | FS1 | BMS | BPM | 2587 | FS1_BMS:BPM_D2587 | BPM3 | position monitor | 47.5 | 238.4119 | 238.4119 | 0.137 | 238.4119 | 415.3476 | 41.094 | 41.094 | 415.3476 | 238.4804 | 258.6671 | 258.7356 |  | 190.1885
0 | [10 empty] | 0.090002 | 238.5254 | 415.3476 | 41.094 | 41.094 | 415.3476 | 238.5704 | 258.7806 | 258.8256 | 20170216 space adjusted | 190.302
0 | FS1 | BMS | QV | 2590 | FS1_BMS:QV_D2590 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 238.6954 | 415.3476 | 41.094 | 41.094 | 415.3476 | 238.8204 | 258.9506 | 259.0756 |  | 190.472
0 | [8 empty] | 3.076e-06 | -0.25408 | 0.24592 | 238.9434 | 415.3476 | 41.094 | 41.094 | 415.3476 | 239.0663 | 259.1986 | 259.3216 | 20170216 space adjusted | 190.72
0 | FS1 | BMS | FAV | 2594 | FS1_BMS:FAV_D2594 | FV | fast valve | 50 | 239.0999 | 239.0999 | 0.067196 | 239.0999 | 415.3476 | 41.094 | 41.094 | 415.3476 | 239.1335 | 259.3552 | 259.3888 | 20170216 fast valve moved from downstream | 190.8765
0 | [7 empty] | Is it the good place for FV ? |  |  | 0.18688 | 239.227 | 415.3476 | 41.094 | 41.094 | 415.3476 | 239.3204 | 259.4822 | 259.5756 | 20170216 space adjusted | 191.0036
0 | FS1 | BMS | QH | 2597 | FS1_BMS:QH_D2597 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 239.4454 | 415.3476 | 41.094 | 41.094 | 415.3476 | 239.5704 | 259.7006 | 259.8256 |  | 191.222
0 | [10 empty] | 0.10898 | 239.6249 | 415.3476 | 41.094 | 41.094 | 415.3476 | 239.6794 | 259.8801 | 259.9346 | 20170216 space adjusted | 191.4015
0 | [6 empty] | bellow | 50 | 3.076e-06 |  | 0.1035 | 239.7312 | 415.3476 | 41.094 | 41.094 | 415.3476 | 239.7829 | 259.9864 | 260.0381 |  | 191.5077
0 | FS1 | BMS | BPM | 2600 | FS1_BMS:BPM_D2600 | BPM2 | position monitor |  | 239.7829 | 239.7829 | 0 | 239.7829 | 415.3476 | 41.094 | 41.094 | 415.3476 | 239.7829 | 260.0381 | 260.0381 |  | 191.5595
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 239.8555 | 415.3476 | 41.094 | 41.094 | 415.3476 | 239.9282 | 260.1108 | 260.1834 |  | 191.6321
0 | FS1 | BMS | PM | 2602 | FS1_BMS:PM_D2602 | PM2 | profile monitor |  | 3.076e-06 |  | 0 | 239.9282 | 415.3476 | 41.094 | 41.094 | 415.3476 | 239.9282 | 260.1834 | 260.1834 |  | 191.7048
0 | [6 empty] | diagnostic box |  | 239.9282 | 239.9282 | 0 | 239.9282 | 415.3476 | 41.094 | 41.094 | 415.3476 | 239.9282 | 260.1834 | 260.1834 |  | 191.7048
0 | FS1 | BMS | FAVS | 2602 | FS1_BMS:FAVS_D2602 | FVS | fast valve sensor |  |  |  | 0 | 239.9282 | 415.3476 | 41.094 | 41.094 | 415.3476 | 239.9282 | 260.1834 | 260.1834 |  | 191.7048
0 | FS1 | BMS | IP | 2602 | FS1_BMS:IP_D2602 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 239.9282 | 415.3476 | 41.094 | 41.094 | 415.3476 | 239.9282 | 260.1834 | 260.1834 |  | 191.7048
0 | [6 empty] | bellow | 50 | 0 | 0.019014 | 0.13122 | 239.9938 | 415.3476 | 41.094 | 41.094 | 415.3476 | 240.0594 | 260.249 | 260.3146 |  | 191.7704
0 | [8 empty] | 0.5776 | 0.5776 | 0.088621 | 240.1037 | 415.3476 | 41.094 | 41.094 | 415.3476 | 240.148 | 260.3589 | 260.4032 | 20170216 space adjusted | 191.8803
0 | FS1 | MGB02 | GV | 2604 | FS1_MGB02:GV_D2604 |  | gate valve | 40 | -240.1767 |  | 0.057404 | 240.1767 | 415.3476 | 41.094 | 41.094 | 415.3476 | 240.2054 | 260.4319 | 260.4606 |  | 191.9533
0 | [10 empty] | 0 | 240.2054 | 415.3476 | 41.094 | 41.094 | 415.3476 | 240.2054 | 260.4606 | 260.4606 |  | 191.982
0 | FS1 | MGB02 | CAV | 2606 | FS1_MGB02:CAV_D2606 |  | b19 resonator 1 1/4 | 36 |  |  | 0.31833 | 240.3646 | 415.3476 | 41.094 | 41.094 | 415.3476 | 240.5238 | 260.6198 | 260.779 |  | 192.1412
0 | FS1 | MGB02 | IP | 2608 | FS1_MGB02:CAV_D2608 | IP_L | ion pump, ? l/s |  |  |  | 0 | 240.5238 | 415.3476 | 41.094 | 41.094 | 415.3476 | 240.5238 | 260.779 | 260.779 |  | 192.3003
0 | FS1 | MGB02 | CAV | 2609 | FS1_MGB02:CAV_D2609 |  | b19 resonator 1 2/4 | 36 |  |  | 0.31833 | 240.6829 | 415.3476 | 41.094 | 41.094 | 415.3476 | 240.8421 | 260.9381 | 261.0973 |  | 192.4595
0 |  |  |  | 2611 | FS1_MGB02:CAV_D2611 |  | rf coupler |  | -240.8421 |  | 0 | 240.8421 | 415.3476 | 41.094 | 41.094 | 415.3476 | 240.8421 | 261.0973 | 261.0973 |  | 192.6187
0 | [10 empty] | 0 | 240.8421 | 415.3476 | 41.094 | 41.094 | 415.3476 | 240.8421 | 261.0973 | 261.0973 |  | 192.6187
0 | FS1 | MGB02 | CAV | 2613 | FS1_MGB02:CAV_D2613 |  | b19 resonator 1 3/4 | 36 |  |  | 0.31833 | 241.0012 | 415.3476 | 41.094 | 41.094 | 415.3476 | 241.1604 | 261.2565 | 261.4156 |  | 192.7778
0 | FS1 | MGB02 | TMP | 2614 | FS1_MGB02:CAV_D2614 | TMP_L | turbo pump, ? l/s |  |  |  | 0 | 241.1604 | 415.3476 | 41.094 | 41.094 | 415.3476 | 241.1604 | 261.4156 | 261.4156 |  | 192.937
0 | FS1 | MGB02 | CAV | 2616 | FS1_MGB02:CAV_D2616 |  | b19 resonator 1 4/4 | 36 |  |  | 0.31833 | 241.3196 | 415.3476 | 41.094 | 41.094 | 415.3476 | 241.4787 | 261.5748 | 261.7339 |  | 193.0962
0 | [10 empty] | 0 | 241.4787 | 415.3476 | 41.094 | 41.094 | 415.3476 | 241.4787 | 261.7339 | 261.7339 |  | 193.2553
0 | FS1 | MGB02 | GV | 2618 | FS1_MGB02:GV_D2618 |  | gate valve | 40 | -241.5074 |  | 0.057404 | 241.5074 | 415.3476 | 41.094 | 41.094 | 415.3476 | 241.5361 | 261.7626 | 261.7914 |  | 193.284
0 | [10 empty] | 0.89642 | 241.9843 | 415.3476 | 41.094 | 41.094 | 415.3476 | 242.4326 | 262.2396 | 262.6878 |  | 193.7609
0 | [6 empty] | bellow | 40 |  |  | 0 | 242.4326 | 415.3476 | 41.094 | 41.094 | 415.3476 | 242.4326 | 262.6878 | 262.6878 |  | 194.2091
0 | [10 empty] | 0.24879 | 242.5569 | 415.3476 | 41.094 | 41.094 | 415.3476 | 242.6813 | 262.8122 | 262.9366 | 20170216 space adjusted | 194.3335
0 | [6 empty] | vacuum box |  | 3.076e-06 | 0.15479 | 0 | 242.6813 | 415.3476 | 41.094 | 41.094 | 415.3476 | 242.6813 | 262.9366 | 262.9366 |  | 194.4579
0 | FS1 | BMS | FAVS | 2629 | FS1_BMS:FAVS_D2629 | FVS | fast valve sensor |  | 242.6813 | 242.6813 | 0 | 242.6813 | 415.3476 | 41.094 | 41.094 | 415.3476 | 242.6813 | 262.9366 | 262.9366 |  | 194.4579
0 | [6 empty] | vacuum box |  |  |  | 0 | 242.6813 | 415.3476 | 41.094 | 41.094 | 415.3476 | 242.6813 | 262.9366 | 262.9366 |  | 194.4579
0 | FS1 | BMS | IP | 2629 | FS1_BMS:IP_D2629 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 242.6813 | 415.3476 | 41.094 | 41.094 | 415.3476 | 242.6813 | 262.9366 | 262.9366 |  | 194.4579
0 | [6 empty] | bellow | 40 |  |  | 0.035206 | 242.6989 | 415.3476 | 41.094 | 41.094 | 415.3476 | 242.7166 | 262.9542 | 262.9718 | 20170216 space adjusted | 194.4755
0 | [10 empty] | 0.001052 | 242.7171 | 415.3476 | 41.094 | 41.094 | 415.3476 | 242.7176 | 262.9723 | 262.9728 |  | 194.4937
0 | [10 empty] | 0.072 | 242.7536 | 415.3476 | 41.094 | 41.094 | 415.3476 | 242.7896 | 263.0088 | 263.0448 | 20170216 moved from down stream | 194.5302
0 | [10 empty] | 0.37429 | 242.9768 | 415.3476 | 41.094 | 41.094 | 415.3476 | 243.1639 | 263.232 | 263.4191 |  | 194.7533
0 | FS1 | BMS | BD | 2634 | FS1_BMS:BD_D2634 | dump | dump 25W no current reading |  | -2.076e-06 | 243.1639 | 0 | 243.1639 | 415.3476 | 41.094 | 41.094 | 415.3476 | 243.1639 | 263.4191 | 263.4191 |  | 194.9405
0 | [8 empty] | 0.64996 |  | 0.27567 | 243.3017 | 415.3476 | 41.094 | 41.094 | 415.3476 | 243.4396 | 263.5569 | 263.6948 |  | 195.0783
0 | [10 empty] | 0.072 | 243.4756 | 415.3476 | 41.094 | 41.094 | 415.3476 | 243.5116 | 263.7308 | 263.7668 |  | 195.2522
0 | [10 empty] | 0.1775 | 243.6003 | 415.3476 | 41.094 | 41.094 | 415.3476 | 243.6891 | 263.8555 | 263.9443 |  | 195.3769
0 | FS1 | BMS | DCHV | 2640 | FS1_BMS:DCHV_D2640 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 243.7891 | 415.3476 | 41.094 | 41.094 | 415.3476 | 243.8891 | 264.0443 | 264.1443 |  | 195.5657
0 | [10 empty] | 0.25 | 244.0141 | 415.3476 | 41.094 | 41.094 | 415.3476 | 244.1391 | 264.2693 | 264.3943 |  | 195.7907
0 | FS1 | BMS | QH | 2645 | FS1_BMS:QH_D2645 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 244.2641 | 415.3476 | 41.094 | 41.094 | 415.3476 | 244.3891 | 264.5193 | 264.6443 |  | 196.0407
0 | [10 empty] | 0.41601 | 244.5971 | 415.3476 | 41.094 | 41.094 | 415.3476 | 244.8051 | 264.8523 | 265.0603 |  | 196.3737
0 | FS1 | BMS | GV | 2651 | FS1_BMS:GV_D2651 | GV_M | gate valve | 63 | 2.076e-06 | 244.8426 | 0.075 | 244.8426 | 415.3476 | 41.094 | 41.094 | 415.3476 | 244.8801 | 265.0978 | 265.1353 |  | 196.6192
0 | [10 empty] | 0.13139 | 244.9458 | 415.3476 | 41.094 | 41.094 | 415.3476 | 245.0115 | 265.201 | 265.2667 |  | 196.7224
0 | FS1 | BMS | QV | 2654 | FS1_BMS:QV_D2654 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 245.1365 | 415.3476 | 41.094 | 41.094 | 415.3476 | 245.2615 | 265.3917 | 265.5167 |  | 196.9131
0 | [10 empty] | 0.55 | 245.5365 | 415.3476 | 41.094 | 41.094 | 415.3476 | 245.8115 | 265.7917 | 266.0667 |  | 197.3131
0 | FS1 | BMS | DCHV | 2662 | FS1_BMS:DCHV_D2662 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 245.9115 | 415.3476 | 41.094 | 41.094 | 415.3476 | 246.0115 | 266.1667 | 266.2667 |  | 197.6881
0 | [8 empty] | 2.776e-06 | 0.054921 | 0.12392 | 246.0734 | 415.3476 | 41.094 | 41.094 | 415.3476 | 246.1354 | 266.3286 | 266.3906 |  | 197.85
0 | FS1 | BMS | BPM | 2665 | FS1_BMS:BPM_D2665 | BPM3 | position monitor | 47.5 | 246.2039 | 246.2039 | 0.137 | 246.2039 | 415.3476 | 41.094 | 41.094 | 415.3476 | 246.2724 | 266.4591 | 266.5276 |  | 197.9805
0 | [10 empty] | -0.010921 | 246.2669 | 415.3476 | 41.094 | 41.094 | 415.3476 | 246.2615 | 266.5221 | 266.5167 |  | 198.0435
0 | FS1 | BMS | QH | 2666 | FS1_BMS:QH_D2666 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 246.3865 | 415.3476 | 41.094 | 41.094 | 415.3476 | 246.5115 | 266.6417 | 266.7667 |  | 198.1631
0 | [10 empty] | 0.49027 | 246.7566 | 415.3476 | 41.094 | 41.094 | 415.3476 | 247.0017 | 267.0118 | 267.257 | 20170216 space adjusted | 198.5332
0 | FS1 | BMS | LPM | 2673 | FS1_BMS:LPM_D2673 | LPMx | bunch length monitor (not costed) | future LPM2 | 2.976e-06 | -0.034726 | 0 | 247.0017 | 415.3476 | 41.094 | 41.094 | 415.3476 | 247.0017 | 267.257 | 267.257 |  | 198.7783
0 | FS1 | BMS | IP | 2673 | FS1_BMS:IP_D2673 | IP_S | ion pump, 75 l/s |  | 247.0017 | 247.0017 | 0 | 247.0017 | 415.3476 | 41.094 | 41.094 | 415.3476 | 247.0017 | 267.257 | 267.257 |  | 198.7783
0 | [8 empty] | 2.976e-06 |  | 0.14619 | 247.0748 | 415.3476 | 41.094 | 41.094 | 415.3476 | 247.1479 | 267.3301 | 267.4031 | 20170216 space adjusted | 198.8514
0 | FS1 | BMS | BCM | 2675 | FS1_BMS:BCM_D2675 | BCM2 | current monitor |  | 247.2479 | 247.2479 | 0.2 | 247.2479 | 415.3476 | 41.094 | 41.094 | 415.3476 | 247.3479 | 267.5031 | 267.6031 |  | 199.0245
0 | [10 empty] | 0.21354 | 247.4547 | 415.3476 | 41.094 | 41.094 | 415.3476 | 247.5615 | 267.7099 | 267.8167 | 20170216 space adjusted | 199.2313
0 | FS1 | BMS | QH | 2679 | FS1_BMS:QH_D2679 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 247.6865 | 415.3476 | 41.094 | 41.094 | 415.3476 | 247.8115 | 267.9417 | 268.0667 |  | 199.4631
0 | [10 empty] | 0.3495 | 247.9862 | 415.3476 | 41.094 | 41.094 | 415.3476 | 248.161 | 268.2414 | 268.4162 |  | 199.7628
0 | [10 empty] | 0 | 248.161 | 415.3476 | 41.094 | 41.094 | 415.3476 | 248.161 | 268.4162 | 268.4162 |  | 199.9376
0 | [8 empty] | 2.976e-06 | -0.058407 | 0.29109 | 248.3065 | 415.3476 | 41.094 | 41.094 | 415.3476 | 248.4521 | 268.5617 | 268.7073 | 20170216 space adjusted | 200.0831
0 | FS1 | BMS | DCHV | 2688 | FS1_BMS:DCHV_D2688 | C2 | correctors-X,Y | 50 | 248.5521 | 248.5521 | 0.2 | 248.5521 | 415.3476 | 41.094 | 41.094 | 415.3476 | 248.6521 | 268.8073 | 268.9073 |  | 200.3286
0 | [8 empty] | 2.976e-06 | -0.13856 | -0.011153 | 248.6465 | 415.3476 | 41.094 | 41.094 | 415.3476 | 248.6409 | 268.9017 | 268.8961 | 20170216 space adjusted | 200.4231
0 | FS1 | BMS | BPM | 2690 | FS1_BMS:BPM_D2690 | BPM3 | position monitor | 47.5 | 248.7094 | 248.7094 | 0.137 | 248.7094 | 415.3476 | 41.094 | 41.094 | 415.3476 | 248.7779 | 268.9646 | 269.0331 |  | 200.486
0 | [10 empty] | 0.18256 | 248.8692 | 415.3476 | 41.094 | 41.094 | 415.3476 | 248.9605 | 269.1244 | 269.2157 | 20170216 space adjusted | 200.6458
0 | FS1 | BMS | QV | 2693 | FS1_BMS:QV_D2693 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 249.0855 | 415.3476 | 41.094 | 41.094 | 415.3476 | 249.2105 | 269.3407 | 269.4657 |  | 200.8621
0 | [10 empty] | 0.25 | 249.3355 | 415.3476 | 41.094 | 41.094 | 415.3476 | 249.4605 | 269.5907 | 269.7157 |  | 201.1121
0 | FS1 | BMS | QH | 2698 | FS1_BMS:QH_D2698 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 249.5855 | 415.3476 | 41.094 | 41.094 | 415.3476 | 249.7105 | 269.8407 | 269.9657 |  | 201.3621
0 | [10 empty] | 0.12 | 249.7705 | 415.3476 | 41.094 | 41.094 | 415.3476 | 249.8305 | 270.0257 | 270.0857 |  | 201.5471
0 | [6 empty] | bellow | 40 | -0.058788 |  | 0.1035 | 249.8822 | 415.3476 | 41.094 | 41.094 | 415.3476 | 249.934 | 270.1374 | 270.1892 |  | 201.6588
0 | FS1 | BMS | BPM | 2702 | FS1_BMS:BPM_D2702 | BPM2 | position monitor |  | 249.8752 | 249.934 | 0 | 249.934 | 415.3476 | 41.094 | 41.094 | 415.3476 | 249.934 | 270.1892 | 270.1892 |  | 201.7106
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 250.0066 | 415.3476 | 41.094 | 41.094 | 415.3476 | 250.0793 | 270.2618 | 270.3345 |  | 201.7832
0 | FS1 | BMS | PM | 2703 | FS1_BMS:PM_D2703 | PM2 | profile monitor | TO BE CHECKED | defer installation |  | 0 | 250.0793 | 415.3476 | 41.094 | 41.094 | 415.3476 | 250.0793 | 270.3345 | 270.3345 |  | 201.8558
0 | [6 empty] | diagnostic box |  | 250.0205 | 250.0793 | 0 | 250.0793 | 415.3476 | 41.094 | 41.094 | 415.3476 | 250.0793 | 270.3345 | 270.3345 |  | 201.8558
0 | FS1 | BMS | IP | 2703 | FS1_BMS:IP_D2703 | IP_S | ion pump, 75 l/s |  | -0.058788 | 250.0793 | 0 | 250.0793 | 415.3476 | 41.094 | 41.094 | 415.3476 | 250.0793 | 270.3345 | 270.3345 |  | 201.8558
0 | [10 empty] | 0 | 250.0793 | 415.3476 | 41.094 | 41.094 | 415.3476 | 250.0793 | 270.3345 | 270.3345 |  | 201.8558
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 250.1449 | 415.3476 | 41.094 | 41.094 | 415.3476 | 250.2105 | 270.4001 | 270.4657 | adjustment for residuals | 201.9214
0 | SEGMENT 2 START (12 b29 CMs +12 b54 CMs) | [9 empty] | -2e-07 | 250.2105 | 415.3476 | 41.094 | 41.094 | 415.3476 | 250.2105 | 270.4657 | 270.4657 |  | 201.9871
0 | LS2 | CC01 | GV | 2705 | LS2_CC01:GV_D2705 |  | gate valve | 40 |  |  | 0.057404 | 250.2392 | 415.3476 | 41.094 | 41.094 | 415.3476 | 250.2679 | 270.4944 | 270.5231 |  | 202.0158
0 | [6 empty] | bellows |  |  |  | 0.1677 | 250.3517 | 415.3476 | 41.094 | 41.094 | 415.3476 | 250.4356 | 270.6069 | 270.6908 |  | 202.1283
0 | LS2 | CC01 | CAV1 | 2706 | LS2_CC01:CAV1_D2706 |  | b29 resonator | 40 | 0.32364 |  | 0.32364 | 250.5974 | 415.3476 | 41.094 | 41.094 | 415.3476 | 250.7592 | 270.8526 | 271.0144 |  | 202.374
0 | LS2 | CC01 | BLM | 2709 | LS2_CC01:BLM_D2709 |  | outside loss monitor |  |  |  | 0 | 250.7592 | 415.3476 | 41.094 | 41.094 | 415.3476 | 250.7592 | 271.0144 | 271.0144 |  | 202.5358
0 | [10 empty] | 0 | 250.7592 | 415.3476 | 41.094 | 41.094 | 415.3476 | 250.7592 | 271.0144 | 271.0144 |  | 202.5358
0 | [6 empty] | bellow |  |  |  | 0.099517 | 250.809 | 415.3476 | 41.094 | 41.094 | 415.3476 | 250.8587 | 271.0642 | 271.114 |  | 202.5856
0 | LS2 | CC01 | CAV2 | 2711 | LS2_CC01:CAV2_D2711 |  | b29 resonator | 40 |  |  | 0.32364 | 251.0206 | 415.3476 | 41.094 | 41.094 | 415.3476 | 251.1824 | 271.2758 | 271.4376 |  | 202.7971
0 | [10 empty] | 0 | 251.1824 | 415.3476 | 41.094 | 41.094 | 415.3476 | 251.1824 | 271.4376 | 271.4376 |  | 202.959
0 | [10 empty] | 0 | 251.1824 | 415.3476 | 41.094 | 41.094 | 415.3476 | 251.1824 | 271.4376 | 271.4376 |  | 202.959
0 | [6 empty] | bellow |  |  |  | 0.099517 | 251.2321 | 415.3476 | 41.094 | 41.094 | 415.3476 | 251.2819 | 271.4874 | 271.5371 |  | 203.0087
0 | LS2 | CC01 | CAV3 | 2715 | LS2_CC01:CAV3_D2715 |  | b29 resonator | 40 |  |  | 0.32364 | 251.4437 | 415.3476 | 41.094 | 41.094 | 415.3476 | 251.6055 | 271.6989 | 271.8608 |  | 203.2203
0 | LS2 | CC01 | BLM | 2717 | LS2_CC01:BLM_D2717 |  | outside loss monitor |  |  |  | 0 | 251.6055 | 415.3476 | 41.094 | 41.094 | 415.3476 | 251.6055 | 271.8608 | 271.8608 |  | 203.3821
0 | [10 empty] | 0 | 251.6055 | 415.3476 | 41.094 | 41.094 | 415.3476 | 251.6055 | 271.8608 | 271.8608 |  | 203.3821
0 | [6 empty] | bellow+tube |  |  |  | 0.15525 | 251.6832 | 415.3476 | 41.094 | 41.094 | 415.3476 | 251.7608 | 271.9384 | 272.016 |  | 203.4598
0 | LS2 | CC01 | SOL1 | 2723 | LS2_CC01:SOL1_D2723 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 252.0108 | 415.3476 | 41.094 | 41.094 | 415.3476 | 252.2608 | 272.266 | 272.516 |  | 203.7874
0 | [6 empty] | bellow+tube |  |  |  | 0.15525 | 252.3384 | 415.3476 | 41.094 | 41.094 | 415.3476 | 252.416 | 272.5936 | 272.6713 |  | 204.115
0 | LS2 | CC01 | CAV4 | 2726 | LS2_CC01:CAV4_D2726 |  | b29 resonator | 40 |  |  | 0.32364 | 252.5779 | 415.3476 | 41.094 | 41.094 | 415.3476 | 252.7397 | 272.8331 | 272.9949 |  | 204.3545
0 | [10 empty] | 0 | 252.7397 | 415.3476 | 41.094 | 41.094 | 415.3476 | 252.7397 | 272.9949 | 272.9949 |  | 204.5163
0 | [10 empty] | 0 | 252.7397 | 415.3476 | 41.094 | 41.094 | 415.3476 | 252.7397 | 272.9949 | 272.9949 |  | 204.5163
0 | [10 empty] | 0.099517 | 252.7894 | 415.3476 | 41.094 | 41.094 | 415.3476 | 252.8392 | 273.0447 | 273.0944 |  | 204.566
0 | LS2 | CC01 | CAV5 | 2730 | LS2_CC01:CAV5_D2730 |  | b29 resonator | 40 |  |  | 0.32364 | 253.001 | 415.3476 | 41.094 | 41.094 | 415.3476 | 253.1628 | 273.2562 | 273.4181 |  | 204.7776
0 | [10 empty] | 0 | 253.1628 | 415.3476 | 41.094 | 41.094 | 415.3476 | 253.1628 | 273.4181 | 273.4181 |  | 204.9394
0 | [10 empty] | 0 | 253.1628 | 415.3476 | 41.094 | 41.094 | 415.3476 | 253.1628 | 273.4181 | 273.4181 |  | 204.9394
0 | [10 empty] | 0.099517 | 253.2126 | 415.3476 | 41.094 | 41.094 | 415.3476 | 253.2624 | 273.4678 | 273.5176 |  | 204.9892
0 | LS2 | CC01 | CAV6 | 2735 | LS2_CC01:CAV6_D2735 |  | b29 resonator | 40 |  |  | 0.32364 | 253.4242 | 415.3476 | 41.094 | 41.094 | 415.3476 | 253.586 | 273.6794 | 273.8412 |  | 205.2008
0 | LS2 | CC01 | BLM | 2737 | LS2_CC01:BLM_D2737 |  | outside loss monitor |  |  |  | 0 | 253.586 | 415.3476 | 41.094 | 41.094 | 415.3476 | 253.586 | 273.8412 | 273.8412 |  | 205.3626
0 | [10 empty] | 0 | 253.586 | 415.3476 | 41.094 | 41.094 | 415.3476 | 253.586 | 273.8412 | 273.8412 |  | 205.3626
0 | [6 empty] | bellows |  |  |  | 0.1677 | 253.6699 | 415.3476 | 41.094 | 41.094 | 415.3476 | 253.7537 | 273.9251 | 274.0089 |  | 205.4464
0 | LS2 | CC01 | GV | 2740 | LS2_CC01:GV_D2740 |  | gate valve | 40 |  |  | 0.057404 | 253.7824 | 415.3476 | 41.094 | 41.094 | 415.3476 | 253.8111 | 274.0376 | 274.0663 |  | 205.559
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 253.8629 | 415.3476 | 41.094 | 41.094 | 415.3476 | 253.9146 | 274.1181 | 274.1698 |  | 205.6395
0 | LS2 | WC01 | BPM | 2742 | LS2_WC01:BPM_D2742 | BPM2 | position monitor |  |  |  | 0 | 253.9146 | 415.3476 | 41.094 | 41.094 | 415.3476 | 253.9146 | 274.1698 | 274.1698 |  | 205.6912
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 253.9873 | 415.3476 | 41.094 | 41.094 | 415.3476 | 254.0599 | 274.2425 | 274.3151 |  | 205.7638
0 | LS2 | WC01 | PORT | 2743 | LS2_WC01:PORT_D2743 |  | port for future diagnostics |  |  |  | 0 | 254.0599 | 415.3476 | 41.094 | 41.094 | 415.3476 | 254.0599 | 274.3151 | 274.3151 |  | 205.8365
0 | [6 empty] | diagnostic box |  |  |  | 0 | 254.0599 | 415.3476 | 41.094 | 41.094 | 415.3476 | 254.0599 | 274.3151 | 274.3151 |  | 205.8365
0 | LS2 | WC01 | IP | 2743 | LS2_WC01:IP_D2743 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 254.0599 | 415.3476 | 41.094 | 41.094 | 415.3476 | 254.0599 | 274.3151 | 274.3151 |  | 205.8365
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 254.1255 | 415.3476 | 41.094 | 41.094 | 415.3476 | 254.1911 | 274.3807 | 274.4463 |  | 205.9021
0 | LS2 | CC02 | GV | 2745 | LS2_CC02:GV_D2745 |  | gate valve | 40 |  |  | 0.057404 | 254.2198 | 415.3476 | 41.094 | 41.094 | 415.3476 | 254.2485 | 274.475 | 274.5037 |  | 205.9964
0 | [6 empty] | bellows |  |  |  | 0.1677 | 254.3324 | 415.3476 | 41.094 | 41.094 | 415.3476 | 254.4162 | 274.5876 | 274.6714 |  | 206.109
0 | LS2 | CC02 | CAV1 | 2746 | LS2_CC02:CAV1_D2746 |  | b29 resonator | 40 | 0.32364 |  | 0.32364 | 254.578 | 415.3476 | 41.094 | 41.094 | 415.3476 | 254.7399 | 274.8333 | 274.9951 |  | 206.3546
0 | LS2 | CC02 | BLM | 2748 | LS2_CC02:BLM_D2748 |  | outside loss monitor |  |  |  | 0 | 254.7399 | 415.3476 | 41.094 | 41.094 | 415.3476 | 254.7399 | 274.9951 | 274.9951 |  | 206.5165
0 | [10 empty] | 0 | 254.7399 | 415.3476 | 41.094 | 41.094 | 415.3476 | 254.7399 | 274.9951 | 274.9951 |  | 206.5165
0 | [6 empty] | bellow |  |  |  | 0.099517 | 254.7896 | 415.3476 | 41.094 | 41.094 | 415.3476 | 254.8394 | 275.0448 | 275.0946 |  | 206.5662
0 | LS2 | CC02 | CAV2 | 2750 | LS2_CC02:CAV2_D2750 |  | b29 resonator | 40 |  |  | 0.32364 | 255.0012 | 415.3476 | 41.094 | 41.094 | 415.3476 | 255.163 | 275.2564 | 275.4182 |  | 206.7778
0 | [10 empty] | 0 | 255.163 | 415.3476 | 41.094 | 41.094 | 415.3476 | 255.163 | 275.4182 | 275.4182 |  | 206.9396
0 | [10 empty] | 0 | 255.163 | 415.3476 | 41.094 | 41.094 | 415.3476 | 255.163 | 275.4182 | 275.4182 |  | 206.9396
0 | [6 empty] | bellow |  |  |  | 0.099517 | 255.2128 | 415.3476 | 41.094 | 41.094 | 415.3476 | 255.2625 | 275.468 | 275.5178 |  | 206.9894
0 | LS2 | CC02 | CAV3 | 2755 | LS2_CC02:CAV3_D2755 |  | b29 resonator | 40 |  |  | 0.32364 | 255.4244 | 415.3476 | 41.094 | 41.094 | 415.3476 | 255.5862 | 275.6796 | 275.8414 |  | 207.201
0 | LS2 | CC02 | BLM | 2757 | LS2_CC02:BLM_D2757 |  | outside loss monitor |  |  |  | 0 | 255.5862 | 415.3476 | 41.094 | 41.094 | 415.3476 | 255.5862 | 275.8414 | 275.8414 |  | 207.3628
0 | [10 empty] | 0 | 255.5862 | 415.3476 | 41.094 | 41.094 | 415.3476 | 255.5862 | 275.8414 | 275.8414 |  | 207.3628
0 | [6 empty] | bellow+tube |  |  |  | 0.15525 | 255.6638 | 415.3476 | 41.094 | 41.094 | 415.3476 | 255.7414 | 275.919 | 275.9967 |  | 207.4404
0 | LS2 | CC02 | SOL1 | 2762 | LS2_CC02:SOL1_D2762 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 255.9914 | 415.3476 | 41.094 | 41.094 | 415.3476 | 256.2414 | 276.2467 | 276.4967 |  | 207.768
0 | [6 empty] | bellow+tube |  |  |  | 0.15525 | 256.3191 | 415.3476 | 41.094 | 41.094 | 415.3476 | 256.3967 | 276.5743 | 276.6519 |  | 208.0957
0 | LS2 | CC02 | CAV4 | 2766 | LS2_CC02:CAV4_D2766 |  | b29 resonator | 40 |  |  | 0.32364 | 256.5585 | 415.3476 | 41.094 | 41.094 | 415.3476 | 256.7203 | 276.8137 | 276.9756 |  | 208.3351
0 | [10 empty] | 0 | 256.7203 | 415.3476 | 41.094 | 41.094 | 415.3476 | 256.7203 | 276.9756 | 276.9756 |  | 208.4969
0 | [10 empty] | 0 | 256.7203 | 415.3476 | 41.094 | 41.094 | 415.3476 | 256.7203 | 276.9756 | 276.9756 |  | 208.4969
0 | [10 empty] | 0.099517 | 256.7701 | 415.3476 | 41.094 | 41.094 | 415.3476 | 256.8199 | 277.0253 | 277.0751 |  | 208.5467
0 | LS2 | CC02 | CAV5 | 2770 | LS2_CC02:CAV5_D2770 |  | b29 resonator | 40 |  |  | 0.32364 | 256.9817 | 415.3476 | 41.094 | 41.094 | 415.3476 | 257.1435 | 277.2369 | 277.3987 |  | 208.7583
0 | [10 empty] | 0 | 257.1435 | 415.3476 | 41.094 | 41.094 | 415.3476 | 257.1435 | 277.3987 | 277.3987 |  | 208.9201
0 | [10 empty] | 0 | 257.1435 | 415.3476 | 41.094 | 41.094 | 415.3476 | 257.1435 | 277.3987 | 277.3987 |  | 208.9201
0 | [10 empty] | 0.099517 | 257.1933 | 415.3476 | 41.094 | 41.094 | 415.3476 | 257.243 | 277.4485 | 277.4982 |  | 208.9698
0 | LS2 | CC02 | CAV6 | 2774 | LS2_CC02:CAV6_D2774 |  | b29 resonator | 40 |  |  | 0.32364 | 257.4048 | 415.3476 | 41.094 | 41.094 | 415.3476 | 257.5667 | 277.6601 | 277.8219 |  | 209.1814
0 | LS2 | CC02 | BLM | 2777 | LS2_CC02:BLM_D2777 |  | outside loss monitor |  |  |  | 0 | 257.5667 | 415.3476 | 41.094 | 41.094 | 415.3476 | 257.5667 | 277.8219 | 277.8219 |  | 209.3432
0 | [10 empty] | 0 | 257.5667 | 415.3476 | 41.094 | 41.094 | 415.3476 | 257.5667 | 277.8219 | 277.8219 |  | 209.3432
0 | [6 empty] | bellows |  |  |  | 0.1677 | 257.6505 | 415.3476 | 41.094 | 41.094 | 415.3476 | 257.7344 | 277.9057 | 277.9896 |  | 209.4271
0 | LS2 | CC02 | GV | 2780 | LS2_CC02:GV_D2780 |  | gate valve | 40 |  |  | 0.057404 | 257.7631 | 415.3476 | 41.094 | 41.094 | 415.3476 | 257.7918 | 278.0183 | 278.047 |  | 209.5396
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 257.8435 | 415.3476 | 41.094 | 41.094 | 415.3476 | 257.8953 | 278.0987 | 278.1505 |  | 209.6201
0 | LS2 | WC02 | BPM | 2782 | LS2_WC02:BPM_D2782 | BPM2 | position monitor |  |  |  | 0 | 257.8953 | 415.3476 | 41.094 | 41.094 | 415.3476 | 257.8953 | 278.1505 | 278.1505 |  | 209.6719
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 257.9679 | 415.3476 | 41.094 | 41.094 | 415.3476 | 258.0405 | 278.2231 | 278.2958 |  | 209.7445
0 | LS2 | WC02 | PORT | 2783 | LS2_WC02:PORT_D2783 |  | port for future diagnostics |  |  |  | 0 | 258.0405 | 415.3476 | 41.094 | 41.094 | 415.3476 | 258.0405 | 278.2958 | 278.2958 |  | 209.8171
0 | [6 empty] | diagnostic box |  |  |  | 0 | 258.0405 | 415.3476 | 41.094 | 41.094 | 415.3476 | 258.0405 | 278.2958 | 278.2958 |  | 209.8171
0 | LS2 | WC02 | IP | 2783 | LS2_WC02:IP_D2783 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 258.0405 | 415.3476 | 41.094 | 41.094 | 415.3476 | 258.0405 | 278.2958 | 278.2958 |  | 209.8171
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 258.1062 | 415.3476 | 41.094 | 41.094 | 415.3476 | 258.1718 | 278.3614 | 278.427 |  | 209.8827
0 | LS2 | CC03 | GV | 2785 | LS2_CC03:GV_D2785 |  | gate valve | 40 |  |  | 0.057404 | 258.2005 | 415.3476 | 41.094 | 41.094 | 415.3476 | 258.2292 | 278.4557 | 278.4844 |  | 209.9771
0 | [6 empty] | bellows |  |  |  | 0.1677 | 258.313 | 415.3476 | 41.094 | 41.094 | 415.3476 | 258.3969 | 278.5682 | 278.6521 |  | 210.0896
0 | LS2 | CC03 | CAV1 | 2786 | LS2_CC03:CAV1_D2786 |  | b29 resonator | 40 | 0.32364 |  | 0.32364 | 258.5587 | 415.3476 | 41.094 | 41.094 | 415.3476 | 258.7205 | 278.8139 | 278.9757 |  | 210.3353
0 | LS2 | CC03 | BLM | 2788 | LS2_CC03:BLM_D2788 |  | outside loss monitor |  |  |  | 0 | 258.7205 | 415.3476 | 41.094 | 41.094 | 415.3476 | 258.7205 | 278.9757 | 278.9757 |  | 210.4971
0 | [10 empty] | 0 | 258.7205 | 415.3476 | 41.094 | 41.094 | 415.3476 | 258.7205 | 278.9757 | 278.9757 |  | 210.4971
0 | [6 empty] | bellow |  |  |  | 0.099517 | 258.7703 | 415.3476 | 41.094 | 41.094 | 415.3476 | 258.82 | 279.0255 | 279.0752 |  | 210.5469
0 | LS2 | CC03 | CAV2 | 2790 | LS2_CC03:CAV2_D2790 |  | b29 resonator | 40 |  |  | 0.32364 | 258.9819 | 415.3476 | 41.094 | 41.094 | 415.3476 | 259.1437 | 279.2371 | 279.3989 |  | 210.7584
0 | [10 empty] | 0 | 259.1437 | 415.3476 | 41.094 | 41.094 | 415.3476 | 259.1437 | 279.3989 | 279.3989 |  | 210.9203
0 | [10 empty] | 0 | 259.1437 | 415.3476 | 41.094 | 41.094 | 415.3476 | 259.1437 | 279.3989 | 279.3989 |  | 210.9203
0 | [6 empty] | bellow |  |  |  | 0.099517 | 259.1934 | 415.3476 | 41.094 | 41.094 | 415.3476 | 259.2432 | 279.4487 | 279.4984 |  | 210.97
0 | LS2 | CC03 | CAV3 | 2794 | LS2_CC03:CAV3_D2794 |  | b29 resonator | 40 |  |  | 0.32364 | 259.405 | 415.3476 | 41.094 | 41.094 | 415.3476 | 259.5668 | 279.6602 | 279.8221 |  | 211.1816
0 | LS2 | CC03 | BLM | 2797 | LS2_CC03:BLM_D2797 |  | outside loss monitor |  |  |  | 0 | 259.5668 | 415.3476 | 41.094 | 41.094 | 415.3476 | 259.5668 | 279.8221 | 279.8221 |  | 211.3434
0 | [10 empty] | 0 | 259.5668 | 415.3476 | 41.094 | 41.094 | 415.3476 | 259.5668 | 279.8221 | 279.8221 |  | 211.3434
0 | [6 empty] | bellow+tube |  |  |  | 0.15525 | 259.6445 | 415.3476 | 41.094 | 41.094 | 415.3476 | 259.7221 | 279.8997 | 279.9773 |  | 211.4211
0 | LS2 | CC03 | SOL1 | 2802 | LS2_CC03:SOL1_D2802 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 259.9721 | 415.3476 | 41.094 | 41.094 | 415.3476 | 260.2221 | 280.2273 | 280.4773 |  | 211.7487
0 | [6 empty] | bellow+tube |  |  |  | 0.15525 | 260.2997 | 415.3476 | 41.094 | 41.094 | 415.3476 | 260.3773 | 280.5549 | 280.6326 |  | 212.0763
0 | LS2 | CC03 | CAV4 | 2806 | LS2_CC03:CAV4_D2806 |  | b29 resonator | 40 |  |  | 0.32364 | 260.5392 | 415.3476 | 41.094 | 41.094 | 415.3476 | 260.701 | 280.7944 | 280.9562 |  | 212.3157
0 | [10 empty] | 0 | 260.701 | 415.3476 | 41.094 | 41.094 | 415.3476 | 260.701 | 280.9562 | 280.9562 |  | 212.4776
0 | [10 empty] | 0 | 260.701 | 415.3476 | 41.094 | 41.094 | 415.3476 | 260.701 | 280.9562 | 280.9562 |  | 212.4776
0 | [10 empty] | 0.099517 | 260.7507 | 415.3476 | 41.094 | 41.094 | 415.3476 | 260.8005 | 281.006 | 281.0557 |  | 212.5273
0 | LS2 | CC03 | CAV5 | 2810 | LS2_CC03:CAV5_D2810 |  | b29 resonator | 40 |  |  | 0.32364 | 260.9623 | 415.3476 | 41.094 | 41.094 | 415.3476 | 261.1241 | 281.2175 | 281.3794 |  | 212.7389
0 | [10 empty] | 0 | 261.1241 | 415.3476 | 41.094 | 41.094 | 415.3476 | 261.1241 | 281.3794 | 281.3794 |  | 212.9007
0 | [10 empty] | 0 | 261.1241 | 415.3476 | 41.094 | 41.094 | 415.3476 | 261.1241 | 281.3794 | 281.3794 |  | 212.9007
0 | [10 empty] | 0.099517 | 261.1739 | 415.3476 | 41.094 | 41.094 | 415.3476 | 261.2237 | 281.4291 | 281.4789 |  | 212.9505
0 | LS2 | CC03 | CAV6 | 2814 | LS2_CC03:CAV6_D2814 |  | b29 resonator | 40 |  |  | 0.32364 | 261.3855 | 415.3476 | 41.094 | 41.094 | 415.3476 | 261.5473 | 281.6407 | 281.8025 |  | 213.1621
0 | LS2 | CC03 | BLM | 2816 | LS2_CC03:BLM_D2816 |  | outside loss monitor |  |  |  | 0 | 261.5473 | 415.3476 | 41.094 | 41.094 | 415.3476 | 261.5473 | 281.8025 | 281.8025 |  | 213.3239
0 | [10 empty] | 0 | 261.5473 | 415.3476 | 41.094 | 41.094 | 415.3476 | 261.5473 | 281.8025 | 281.8025 |  | 213.3239
0 | [6 empty] | bellows |  |  |  | 0.1677 | 261.6312 | 415.3476 | 41.094 | 41.094 | 415.3476 | 261.715 | 281.8864 | 281.9702 |  | 213.4077
0 | LS2 | CC03 | GV | 2820 | LS2_CC03:GV_D2820 |  | gate valve | 40 |  |  | 0.057404 | 261.7437 | 415.3476 | 41.094 | 41.094 | 415.3476 | 261.7724 | 281.9989 | 282.0276 |  | 213.5203
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 261.8242 | 415.3476 | 41.094 | 41.094 | 415.3476 | 261.8759 | 282.0794 | 282.1311 |  | 213.6008
0 | LS2 | WC03 | BPM | 2821 | LS2_WC03:BPM_D2821 | BPM2 | position monitor |  |  |  | 0 | 261.8759 | 415.3476 | 41.094 | 41.094 | 415.3476 | 261.8759 | 282.1311 | 282.1311 |  | 213.6525
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 261.9486 | 415.3476 | 41.094 | 41.094 | 415.3476 | 262.0212 | 282.2038 | 282.2764 |  | 213.7251
0 | LS2 | WC03 | PORT | 2823 | LS2_WC03:PORT_D2823 |  | port for future diagnostics |  |  |  | 0 | 262.0212 | 415.3476 | 41.094 | 41.094 | 415.3476 | 262.0212 | 282.2764 | 282.2764 |  | 213.7978
0 | [6 empty] | diagnostic box |  |  |  | 0 | 262.0212 | 415.3476 | 41.094 | 41.094 | 415.3476 | 262.0212 | 282.2764 | 282.2764 |  | 213.7978
0 | LS2 | WC03 | IP | 2823 | LS2_WC03:IP_D2823 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 262.0212 | 415.3476 | 41.094 | 41.094 | 415.3476 | 262.0212 | 282.2764 | 282.2764 |  | 213.7978
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 262.0868 | 415.3476 | 41.094 | 41.094 | 415.3476 | 262.1524 | 282.342 | 282.4076 |  | 213.8634
0 | LS2 | CC04 | GV | 2824 | LS2_CC04:GV_D2824 |  | gate valve | 40 |  |  | 0.057404 | 262.1811 | 415.3476 | 41.094 | 41.094 | 415.3476 | 262.2098 | 282.4363 | 282.465 |  | 213.9577
0 | [6 empty] | bellows |  |  |  | 0.1677 | 262.2937 | 415.3476 | 41.094 | 41.094 | 415.3476 | 262.3775 | 282.5489 | 282.6327 |  | 214.0703
0 | LS2 | CC04 | CAV1 | 2825 | LS2_CC04:CAV1_D2825 |  | b29 resonator | 40 | 0.32364 |  | 0.32364 | 262.5393 | 415.3476 | 41.094 | 41.094 | 415.3476 | 262.7012 | 282.7946 | 282.9564 |  | 214.3159
0 | LS2 | CC04 | BLM | 2828 | LS2_CC04:BLM_D2828 |  | outside loss monitor |  |  |  | 0 | 262.7012 | 415.3476 | 41.094 | 41.094 | 415.3476 | 262.7012 | 282.9564 | 282.9564 |  | 214.4778
0 | [10 empty] | 0 | 262.7012 | 415.3476 | 41.094 | 41.094 | 415.3476 | 262.7012 | 282.9564 | 282.9564 |  | 214.4778
0 | [6 empty] | bellow |  |  |  | 0.099517 | 262.7509 | 415.3476 | 41.094 | 41.094 | 415.3476 | 262.8007 | 283.0061 | 283.0559 |  | 214.5275
0 | LS2 | CC04 | CAV2 | 2830 | LS2_CC04:CAV2_D2830 |  | b29 resonator | 40 |  |  | 0.32364 | 262.9625 | 415.3476 | 41.094 | 41.094 | 415.3476 | 263.1243 | 283.2177 | 283.3795 |  | 214.7391
0 | [10 empty] | 0 | 263.1243 | 415.3476 | 41.094 | 41.094 | 415.3476 | 263.1243 | 283.3795 | 283.3795 |  | 214.9009
0 | [10 empty] | 0 | 263.1243 | 415.3476 | 41.094 | 41.094 | 415.3476 | 263.1243 | 283.3795 | 283.3795 |  | 214.9009
0 | [6 empty] | bellow |  |  |  | 0.099517 | 263.1741 | 415.3476 | 41.094 | 41.094 | 415.3476 | 263.2238 | 283.4293 | 283.4791 |  | 214.9507
0 | LS2 | CC04 | CAV3 | 2834 | LS2_CC04:CAV3_D2834 |  | b29 resonator | 40 |  |  | 0.32364 | 263.3857 | 415.3476 | 41.094 | 41.094 | 415.3476 | 263.5475 | 283.6409 | 283.8027 |  | 215.1623
0 | LS2 | CC04 | BLM | 2836 | LS2_CC04:BLM_D2836 |  | outside loss monitor |  |  |  | 0 | 263.5475 | 415.3476 | 41.094 | 41.094 | 415.3476 | 263.5475 | 283.8027 | 283.8027 |  | 215.3241
0 | [10 empty] | 0 | 263.5475 | 415.3476 | 41.094 | 41.094 | 415.3476 | 263.5475 | 283.8027 | 283.8027 |  | 215.3241
0 | [6 empty] | bellow+tube |  |  |  | 0.15525 | 263.6251 | 415.3476 | 41.094 | 41.094 | 415.3476 | 263.7027 | 283.8803 | 283.958 |  | 215.4017
0 | LS2 | CC04 | SOL1 | 2842 | LS2_CC04:SOL1_D2842 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 263.9527 | 415.3476 | 41.094 | 41.094 | 415.3476 | 264.2027 | 284.208 | 284.458 |  | 215.7293
0 | [6 empty] | bellow+tube |  |  |  | 0.15525 | 264.2804 | 415.3476 | 41.094 | 41.094 | 415.3476 | 264.358 | 284.5356 | 284.6132 |  | 216.0569
0 | LS2 | CC04 | CAV4 | 2845 | LS2_CC04:CAV4_D2845 |  | b29 resonator | 40 |  |  | 0.32364 | 264.5198 | 415.3476 | 41.094 | 41.094 | 415.3476 | 264.6816 | 284.775 | 284.9368 |  | 216.2964
0 | [10 empty] | 0 | 264.6816 | 415.3476 | 41.094 | 41.094 | 415.3476 | 264.6816 | 284.9368 | 284.9368 |  | 216.4582
0 | [10 empty] | 0 | 264.6816 | 415.3476 | 41.094 | 41.094 | 415.3476 | 264.6816 | 284.9368 | 284.9368 |  | 216.4582
0 | [10 empty] | 0.099517 | 264.7314 | 415.3476 | 41.094 | 41.094 | 415.3476 | 264.7811 | 284.9866 | 285.0364 |  | 216.508
0 | LS2 | CC04 | CAV5 | 2850 | LS2_CC04:CAV5_D2850 |  | b29 resonator | 40 |  |  | 0.32364 | 264.943 | 415.3476 | 41.094 | 41.094 | 415.3476 | 265.1048 | 285.1982 | 285.36 |  | 216.7196
0 | [10 empty] | 0 | 265.1048 | 415.3476 | 41.094 | 41.094 | 415.3476 | 265.1048 | 285.36 | 285.36 |  | 216.8814
0 | [10 empty] | 0 | 265.1048 | 415.3476 | 41.094 | 41.094 | 415.3476 | 265.1048 | 285.36 | 285.36 |  | 216.8814
0 | [10 empty] | 0.099517 | 265.1546 | 415.3476 | 41.094 | 41.094 | 415.3476 | 265.2043 | 285.4098 | 285.4595 |  | 216.9311
0 | LS2 | CC04 | CAV6 | 2854 | LS2_CC04:CAV6_D2854 |  | b29 resonator | 40 |  |  | 0.32364 | 265.3661 | 415.3476 | 41.094 | 41.094 | 415.3476 | 265.528 | 285.6213 | 285.7832 |  | 217.1427
0 | LS2 | CC04 | BLM | 2856 | LS2_CC04:BLM_D2856 |  | outside loss monitor |  |  |  | 0 | 265.528 | 415.3476 | 41.094 | 41.094 | 415.3476 | 265.528 | 285.7832 | 285.7832 |  | 217.3045
0 | [10 empty] | 0 | 265.528 | 415.3476 | 41.094 | 41.094 | 415.3476 | 265.528 | 285.7832 | 285.7832 |  | 217.3045
0 | [6 empty] | bellows |  |  |  | 0.1677 | 265.6118 | 415.3476 | 41.094 | 41.094 | 415.3476 | 265.6957 | 285.867 | 285.9509 |  | 217.3884
0 | LS2 | CC04 | GV | 2860 | LS2_CC04:GV_D2860 |  | gate valve | 40 |  |  | 0.057404 | 265.7244 | 415.3476 | 41.094 | 41.094 | 415.3476 | 265.7531 | 285.9796 | 286.0083 |  | 217.5009
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 265.8048 | 415.3476 | 41.094 | 41.094 | 415.3476 | 265.8566 | 286.06 | 286.1118 |  | 217.5814
0 | LS2 | WC04 | BPM | 2861 | LS2_WC04:BPM_D2861 | BPM2 | position monitor |  |  |  | 0 | 265.8566 | 415.3476 | 41.094 | 41.094 | 415.3476 | 265.8566 | 286.1118 | 286.1118 |  | 217.6331
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 265.9292 | 415.3476 | 41.094 | 41.094 | 415.3476 | 266.0018 | 286.1844 | 286.2571 |  | 217.7058
0 | LS2 | WC04 | PORT | 2863 | LS2_WC04:PORT_D2863 |  | port for future diagnostics |  |  |  | 0 | 266.0018 | 415.3476 | 41.094 | 41.094 | 415.3476 | 266.0018 | 286.2571 | 286.2571 |  | 217.7784
0 | [6 empty] | diagnostic box |  |  |  | 0 | 266.0018 | 415.3476 | 41.094 | 41.094 | 415.3476 | 266.0018 | 286.2571 | 286.2571 |  | 217.7784
0 | LS2 | WC04 | IP | 2863 | LS2_WC04:IP_D2863 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 266.0018 | 415.3476 | 41.094 | 41.094 | 415.3476 | 266.0018 | 286.2571 | 286.2571 |  | 217.7784
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 266.0675 | 415.3476 | 41.094 | 41.094 | 415.3476 | 266.1331 | 286.3227 | 286.3883 |  | 217.844
0 | LS2 | CC05 | GV | 2864 | LS2_CC05:GV_D2864 |  | gate valve | 40 |  |  | 0.057404 | 266.1618 | 415.3476 | 41.094 | 41.094 | 415.3476 | 266.1905 | 286.417 | 286.4457 |  | 217.9383
0 | [6 empty] | bellows |  |  |  | 0.1677 | 266.2743 | 415.3476 | 41.094 | 41.094 | 415.3476 | 266.3582 | 286.5295 | 286.6134 |  | 218.0509
0 | LS2 | CC05 | CAV1 | 2865 | LS2_CC05:CAV1_D2865 |  | b29 resonator | 40 | 0.32364 |  | 0.32364 | 266.52 | 415.3476 | 41.094 | 41.094 | 415.3476 | 266.6818 | 286.7752 | 286.937 |  | 218.2966
0 | LS2 | CC05 | BLM | 2868 | LS2_CC05:BLM_D2868 |  | outside loss monitor |  |  |  | 0 | 266.6818 | 415.3476 | 41.094 | 41.094 | 415.3476 | 266.6818 | 286.937 | 286.937 |  | 218.4584
0 | [10 empty] | 0 | 266.6818 | 415.3476 | 41.094 | 41.094 | 415.3476 | 266.6818 | 286.937 | 286.937 |  | 218.4584
0 | [6 empty] | bellow |  |  |  | 0.099517 | 266.7316 | 415.3476 | 41.094 | 41.094 | 415.3476 | 266.7813 | 286.9868 | 287.0365 |  | 218.5082
0 | LS2 | CC05 | CAV2 | 2870 | LS2_CC05:CAV2_D2870 |  | b29 resonator | 40 |  |  | 0.32364 | 266.9432 | 415.3476 | 41.094 | 41.094 | 415.3476 | 267.105 | 287.1984 | 287.3602 |  | 218.7197
0 | [10 empty] | 0 | 267.105 | 415.3476 | 41.094 | 41.094 | 415.3476 | 267.105 | 287.3602 | 287.3602 |  | 218.8816
0 | [10 empty] | 0 | 267.105 | 415.3476 | 41.094 | 41.094 | 415.3476 | 267.105 | 287.3602 | 287.3602 |  | 218.8816
0 | [6 empty] | bellow |  |  |  | 0.099517 | 267.1547 | 415.3476 | 41.094 | 41.094 | 415.3476 | 267.2045 | 287.4099 | 287.4597 |  | 218.9313
0 | LS2 | CC05 | CAV3 | 2874 | LS2_CC05:CAV3_D2874 |  | b29 resonator | 40 |  |  | 0.32364 | 267.3663 | 415.3476 | 41.094 | 41.094 | 415.3476 | 267.5281 | 287.6215 | 287.7834 |  | 219.1429
0 | LS2 | CC05 | BLM | 2876 | LS2_CC05:BLM_D2876 |  | outside loss monitor |  |  |  | 0 | 267.5281 | 415.3476 | 41.094 | 41.094 | 415.3476 | 267.5281 | 287.7834 | 287.7834 |  | 219.3047
0 | [10 empty] | 0 | 267.5281 | 415.3476 | 41.094 | 41.094 | 415.3476 | 267.5281 | 287.7834 | 287.7834 |  | 219.3047
0 | [6 empty] | bellow+tube |  |  |  | 0.15525 | 267.6058 | 415.3476 | 41.094 | 41.094 | 415.3476 | 267.6834 | 287.861 | 287.9386 |  | 219.3823
0 | LS2 | CC05 | SOL1 | 2882 | LS2_CC05:SOL1_D2882 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 267.9334 | 415.3476 | 41.094 | 41.094 | 415.3476 | 268.1834 | 288.1886 | 288.4386 |  | 219.71
0 | [6 empty] | bellow+tube |  |  |  | 0.15525 | 268.261 | 415.3476 | 41.094 | 41.094 | 415.3476 | 268.3386 | 288.5162 | 288.5939 |  | 220.0376
0 | LS2 | CC05 | CAV4 | 2885 | LS2_CC05:CAV4_D2885 |  | b29 resonator | 40 |  |  | 0.32364 | 268.5005 | 415.3476 | 41.094 | 41.094 | 415.3476 | 268.6623 | 288.7557 | 288.9175 |  | 220.277
0 | [10 empty] | 0 | 268.6623 | 415.3476 | 41.094 | 41.094 | 415.3476 | 268.6623 | 288.9175 | 288.9175 |  | 220.4389
0 | [10 empty] | 0 | 268.6623 | 415.3476 | 41.094 | 41.094 | 415.3476 | 268.6623 | 288.9175 | 288.9175 |  | 220.4389
0 | [10 empty] | 0.099517 | 268.712 | 415.3476 | 41.094 | 41.094 | 415.3476 | 268.7618 | 288.9673 | 289.017 |  | 220.4886
0 | LS2 | CC05 | CAV5 | 2890 | LS2_CC05:CAV5_D2890 |  | b29 resonator | 40 |  |  | 0.32364 | 268.9236 | 415.3476 | 41.094 | 41.094 | 415.3476 | 269.0854 | 289.1788 | 289.3407 |  | 220.7002
0 | [10 empty] | 0 | 269.0854 | 415.3476 | 41.094 | 41.094 | 415.3476 | 269.0854 | 289.3407 | 289.3407 |  | 220.862
0 | [10 empty] | 0 | 269.0854 | 415.3476 | 41.094 | 41.094 | 415.3476 | 269.0854 | 289.3407 | 289.3407 |  | 220.862
0 | [10 empty] | 0.099517 | 269.1352 | 415.3476 | 41.094 | 41.094 | 415.3476 | 269.185 | 289.3904 | 289.4402 |  | 220.9118
0 | LS2 | CC05 | CAV6 | 2894 | LS2_CC05:CAV6_D2894 |  | b29 resonator | 40 |  |  | 0.32364 | 269.3468 | 415.3476 | 41.094 | 41.094 | 415.3476 | 269.5086 | 289.602 | 289.7638 |  | 221.1234
0 | LS2 | CC05 | BLM | 2896 | LS2_CC05:BLM_D2896 |  | outside loss monitor |  |  |  | 0 | 269.5086 | 415.3476 | 41.094 | 41.094 | 415.3476 | 269.5086 | 289.7638 | 289.7638 |  | 221.2852
0 | [10 empty] | 0 | 269.5086 | 415.3476 | 41.094 | 41.094 | 415.3476 | 269.5086 | 289.7638 | 289.7638 |  | 221.2852
0 | [6 empty] | bellows |  |  |  | 0.1677 | 269.5925 | 415.3476 | 41.094 | 41.094 | 415.3476 | 269.6763 | 289.8477 | 289.9315 |  | 221.369
0 | LS2 | CC05 | GV | 2900 | LS2_CC05:GV_D2900 |  | gate valve | 40 |  |  | 0.057404 | 269.705 | 415.3476 | 41.094 | 41.094 | 415.3476 | 269.7337 | 289.9602 | 289.9889 |  | 221.4816
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 269.7855 | 415.3476 | 41.094 | 41.094 | 415.3476 | 269.8372 | 290.0407 | 290.0924 |  | 221.562
0 | LS2 | WC05 | BPM | 2901 | LS2_WC05:BPM_D2901 | BPM2 | position monitor |  |  |  | 0 | 269.8372 | 415.3476 | 41.094 | 41.094 | 415.3476 | 269.8372 | 290.0924 | 290.0924 |  | 221.6138
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 269.9099 | 415.3476 | 41.094 | 41.094 | 415.3476 | 269.9825 | 290.1651 | 290.2377 |  | 221.6864
0 | LS2 | WC05 | PORT | 2902 | LS2_WC05:PORT_D2902 |  | port for future diagnostics |  |  |  | 0 | 269.9825 | 415.3476 | 41.094 | 41.094 | 415.3476 | 269.9825 | 290.2377 | 290.2377 |  | 221.7591
0 | [6 empty] | diagnostic box |  |  |  | 0 | 269.9825 | 415.3476 | 41.094 | 41.094 | 415.3476 | 269.9825 | 290.2377 | 290.2377 |  | 221.7591
0 | LS2 | WC05 | IP | 2902 | LS2_WC05:IP_D2902 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 269.9825 | 415.3476 | 41.094 | 41.094 | 415.3476 | 269.9825 | 290.2377 | 290.2377 |  | 221.7591
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 270.0481 | 415.3476 | 41.094 | 41.094 | 415.3476 | 270.1137 | 290.3033 | 290.3689 |  | 221.8247
0 | LS2 | CC06 | GV | 2904 | LS2_CC06:GV_D2904 |  | gate valve | 40 |  |  | 0.057404 | 270.1424 | 415.3476 | 41.094 | 41.094 | 415.3476 | 270.1711 | 290.3976 | 290.4263 |  | 221.919
0 | [6 empty] | bellows |  |  |  | 0.1677 | 270.255 | 415.3476 | 41.094 | 41.094 | 415.3476 | 270.3388 | 290.5102 | 290.594 |  | 222.0316
0 | LS2 | CC06 | CAV1 | 2905 | LS2_CC06:CAV1_D2905 |  | b29 resonator | 40 | 0.32364 |  | 0.32364 | 270.5006 | 415.3476 | 41.094 | 41.094 | 415.3476 | 270.6625 | 290.7559 | 290.9177 |  | 222.2772
0 | LS2 | CC06 | BLM | 2908 | LS2_CC06:BLM_D2908 |  | outside loss monitor |  |  |  | 0 | 270.6625 | 415.3476 | 41.094 | 41.094 | 415.3476 | 270.6625 | 290.9177 | 290.9177 |  | 222.439
0 | [10 empty] | 0 | 270.6625 | 415.3476 | 41.094 | 41.094 | 415.3476 | 270.6625 | 290.9177 | 290.9177 |  | 222.439
0 | [6 empty] | bellow |  |  |  | 0.099517 | 270.7122 | 415.3476 | 41.094 | 41.094 | 415.3476 | 270.762 | 290.9674 | 291.0172 |  | 222.4888
0 | LS2 | CC06 | CAV2 | 2910 | LS2_CC06:CAV2_D2910 |  | b29 resonator | 40 |  |  | 0.32364 | 270.9238 | 415.3476 | 41.094 | 41.094 | 415.3476 | 271.0856 | 291.179 | 291.3408 |  | 222.7004
0 | [10 empty] | 0 | 271.0856 | 415.3476 | 41.094 | 41.094 | 415.3476 | 271.0856 | 291.3408 | 291.3408 |  | 222.8622
0 | [10 empty] | 0 | 271.0856 | 415.3476 | 41.094 | 41.094 | 415.3476 | 271.0856 | 291.3408 | 291.3408 |  | 222.8622
0 | [6 empty] | bellow |  |  |  | 0.099517 | 271.1354 | 415.3476 | 41.094 | 41.094 | 415.3476 | 271.1851 | 291.3906 | 291.4404 |  | 222.912
0 | LS2 | CC06 | CAV3 | 2914 | LS2_CC06:CAV3_D2914 |  | b29 resonator | 40 |  |  | 0.32364 | 271.347 | 415.3476 | 41.094 | 41.094 | 415.3476 | 271.5088 | 291.6022 | 291.764 |  | 223.1235
0 | LS2 | CC06 | BLM | 2916 | LS2_CC06:BLM_D2916 |  | outside loss monitor |  |  |  | 0 | 271.5088 | 415.3476 | 41.094 | 41.094 | 415.3476 | 271.5088 | 291.764 | 291.764 |  | 223.2854
0 | [10 empty] | 0 | 271.5088 | 415.3476 | 41.094 | 41.094 | 415.3476 | 271.5088 | 291.764 | 291.764 |  | 223.2854
0 | [6 empty] | bellow+tube |  |  |  | 0.15525 | 271.5864 | 415.3476 | 41.094 | 41.094 | 415.3476 | 271.664 | 291.8416 | 291.9192 |  | 223.363
0 | LS2 | CC06 | SOL1 | 2922 | LS2_CC06:SOL1_D2922 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 271.914 | 415.3476 | 41.094 | 41.094 | 415.3476 | 272.164 | 292.1692 | 292.4192 |  | 223.6906
0 | [6 empty] | bellow+tube |  |  |  | 0.15525 | 272.2417 | 415.3476 | 41.094 | 41.094 | 415.3476 | 272.3193 | 292.4969 | 292.5745 |  | 224.0182
0 | LS2 | CC06 | CAV4 | 2925 | LS2_CC06:CAV4_D2925 |  | b29 resonator | 40 |  |  | 0.32364 | 272.4811 | 415.3476 | 41.094 | 41.094 | 415.3476 | 272.6429 | 292.7363 | 292.8981 |  | 224.2577
0 | [10 empty] | 0 | 272.6429 | 415.3476 | 41.094 | 41.094 | 415.3476 | 272.6429 | 292.8981 | 292.8981 |  | 224.4195
0 | [10 empty] | 0 | 272.6429 | 415.3476 | 41.094 | 41.094 | 415.3476 | 272.6429 | 292.8981 | 292.8981 |  | 224.4195
0 | [10 empty] | 0.099517 | 272.6927 | 415.3476 | 41.094 | 41.094 | 415.3476 | 272.7424 | 292.9479 | 292.9977 |  | 224.4693
0 | LS2 | CC06 | CAV5 | 2929 | LS2_CC06:CAV5_D2929 |  | b29 resonator | 40 |  |  | 0.32364 | 272.9043 | 415.3476 | 41.094 | 41.094 | 415.3476 | 273.0661 | 293.1595 | 293.3213 |  | 224.6809
0 | [10 empty] | 0 | 273.0661 | 415.3476 | 41.094 | 41.094 | 415.3476 | 273.0661 | 293.3213 | 293.3213 |  | 224.8427
0 | [10 empty] | 0 | 273.0661 | 415.3476 | 41.094 | 41.094 | 415.3476 | 273.0661 | 293.3213 | 293.3213 |  | 224.8427
0 | [10 empty] | 0.099517 | 273.1158 | 415.3476 | 41.094 | 41.094 | 415.3476 | 273.1656 | 293.3711 | 293.4208 |  | 224.8924
0 | LS2 | CC06 | CAV6 | 2934 | LS2_CC06:CAV6_D2934 |  | b29 resonator | 40 |  |  | 0.32364 | 273.3274 | 415.3476 | 41.094 | 41.094 | 415.3476 | 273.4892 | 293.5826 | 293.7445 |  | 225.104
0 | LS2 | CC06 | BLM | 2936 | LS2_CC06:BLM_D2936 |  | outside loss monitor |  |  |  | 0 | 273.4892 | 415.3476 | 41.094 | 41.094 | 415.3476 | 273.4892 | 293.7445 | 293.7445 |  | 225.2658
0 | [10 empty] | 0 | 273.4892 | 415.3476 | 41.094 | 41.094 | 415.3476 | 273.4892 | 293.7445 | 293.7445 |  | 225.2658
0 | [6 empty] | bellows |  |  |  | 0.1677 | 273.5731 | 415.3476 | 41.094 | 41.094 | 415.3476 | 273.657 | 293.8283 | 293.9122 |  | 225.3497
0 | LS2 | CC06 | GV | 2939 | LS2_CC06:GV_D2939 |  | gate valve | 40 |  |  | 0.057404 | 273.6857 | 415.3476 | 41.094 | 41.094 | 415.3476 | 273.7144 | 293.9409 | 293.9696 |  | 225.4622
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 273.7661 | 415.3476 | 41.094 | 41.094 | 415.3476 | 273.8179 | 294.0213 | 294.0731 |  | 225.5427
0 | LS2 | WC06 | BPM | 2941 | LS2_WC06:BPM_D2941 | BPM2 | position monitor |  |  |  | 0 | 273.8179 | 415.3476 | 41.094 | 41.094 | 415.3476 | 273.8179 | 294.0731 | 294.0731 |  | 225.5944
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 273.8905 | 415.3476 | 41.094 | 41.094 | 415.3476 | 273.9631 | 294.1457 | 294.2184 |  | 225.6671
0 | LS2 | WC06 | PORT | 2942 | LS2_WC06:PORT_D2942 |  | port for future diagnostics |  |  |  | 0 | 273.9631 | 415.3476 | 41.094 | 41.094 | 415.3476 | 273.9631 | 294.2184 | 294.2184 |  | 225.7397
0 | [6 empty] | diagnostic box |  |  |  | 0 | 273.9631 | 415.3476 | 41.094 | 41.094 | 415.3476 | 273.9631 | 294.2184 | 294.2184 |  | 225.7397
0 | LS2 | WC06 | IP | 2942 | LS2_WC06:IP_D2942 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 273.9631 | 415.3476 | 41.094 | 41.094 | 415.3476 | 273.9631 | 294.2184 | 294.2184 |  | 225.7397
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 274.0287 | 415.3476 | 41.094 | 41.094 | 415.3476 | 274.0944 | 294.284 | 294.3496 |  | 225.8053
0 | LS2 | CC07 | GV | 2944 | LS2_CC07:GV_D2944 |  | gate valve | 40 |  |  | 0.057404 | 274.1231 | 415.3476 | 41.094 | 41.094 | 415.3476 | 274.1518 | 294.3783 | 294.407 |  | 225.8996
0 | [6 empty] | bellows |  |  |  | 0.1677 | 274.2356 | 415.3476 | 41.094 | 41.094 | 415.3476 | 274.3195 | 294.4908 | 294.5747 |  | 226.0122
0 | LS2 | CC07 | CAV1 | 2945 | LS2_CC07:CAV1_D2945 |  | b29 resonator | 40 | 0.32364 |  | 0.32364 | 274.4813 | 415.3476 | 41.094 | 41.094 | 415.3476 | 274.6431 | 294.7365 | 294.8983 |  | 226.2579
0 | LS2 | CC07 | BLM | 2947 | LS2_CC07:BLM_D2947 |  | outside loss monitor |  |  |  | 0 | 274.6431 | 415.3476 | 41.094 | 41.094 | 415.3476 | 274.6431 | 294.8983 | 294.8983 |  | 226.4197
0 | [10 empty] | 0 | 274.6431 | 415.3476 | 41.094 | 41.094 | 415.3476 | 274.6431 | 294.8983 | 294.8983 |  | 226.4197
0 | [6 empty] | bellow |  |  |  | 0.099517 | 274.6929 | 415.3476 | 41.094 | 41.094 | 415.3476 | 274.7426 | 294.9481 | 294.9978 |  | 226.4695
0 | LS2 | CC07 | CAV2 | 2949 | LS2_CC07:CAV2_D2949 |  | b29 resonator | 40 |  |  | 0.32364 | 274.9044 | 415.3476 | 41.094 | 41.094 | 415.3476 | 275.0663 | 295.1597 | 295.3215 |  | 226.681
0 | [10 empty] | 0 | 275.0663 | 415.3476 | 41.094 | 41.094 | 415.3476 | 275.0663 | 295.3215 | 295.3215 |  | 226.8429
0 | [10 empty] | 0 | 275.0663 | 415.3476 | 41.094 | 41.094 | 415.3476 | 275.0663 | 295.3215 | 295.3215 |  | 226.8429
0 | [6 empty] | bellow |  |  |  | 0.099517 | 275.116 | 415.3476 | 41.094 | 41.094 | 415.3476 | 275.1658 | 295.3712 | 295.421 |  | 226.8926
0 | LS2 | CC07 | CAV3 | 2954 | LS2_CC07:CAV3_D2954 |  | b29 resonator | 40 |  |  | 0.32364 | 275.3276 | 415.3476 | 41.094 | 41.094 | 415.3476 | 275.4894 | 295.5828 | 295.7446 |  | 227.1042
0 | LS2 | CC07 | BLM | 2956 | LS2_CC07:BLM_D2956 |  | outside loss monitor |  |  |  | 0 | 275.4894 | 415.3476 | 41.094 | 41.094 | 415.3476 | 275.4894 | 295.7446 | 295.7446 |  | 227.266
0 | [10 empty] | 0 | 275.4894 | 415.3476 | 41.094 | 41.094 | 415.3476 | 275.4894 | 295.7446 | 295.7446 |  | 227.266
0 | [6 empty] | bellow+tube |  |  |  | 0.15525 | 275.5671 | 415.3476 | 41.094 | 41.094 | 415.3476 | 275.6447 | 295.8223 | 295.8999 |  | 227.3436
0 | LS2 | CC07 | SOL1 | 2961 | LS2_CC07:SOL1_D2961 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 275.8947 | 415.3476 | 41.094 | 41.094 | 415.3476 | 276.1447 | 296.1499 | 296.3999 |  | 227.6713
0 | [6 empty] | bellow+tube |  |  |  | 0.15525 | 276.2223 | 415.3476 | 41.094 | 41.094 | 415.3476 | 276.2999 | 296.4775 | 296.5551 |  | 227.9989
0 | LS2 | CC07 | CAV4 | 2965 | LS2_CC07:CAV4_D2965 |  | b29 resonator | 40 |  |  | 0.32364 | 276.4618 | 415.3476 | 41.094 | 41.094 | 415.3476 | 276.6236 | 296.717 | 296.8788 |  | 228.2383
0 | [10 empty] | 0 | 276.6236 | 415.3476 | 41.094 | 41.094 | 415.3476 | 276.6236 | 296.8788 | 296.8788 |  | 228.4002
0 | [10 empty] | 0 | 276.6236 | 415.3476 | 41.094 | 41.094 | 415.3476 | 276.6236 | 296.8788 | 296.8788 |  | 228.4002
0 | [10 empty] | 0.099517 | 276.6733 | 415.3476 | 41.094 | 41.094 | 415.3476 | 276.7231 | 296.9285 | 296.9783 |  | 228.4499
0 | LS2 | CC07 | CAV5 | 2969 | LS2_CC07:CAV5_D2969 |  | b29 resonator | 40 |  |  | 0.32364 | 276.8849 | 415.3476 | 41.094 | 41.094 | 415.3476 | 277.0467 | 297.1401 | 297.302 |  | 228.6615
0 | [10 empty] | 0 | 277.0467 | 415.3476 | 41.094 | 41.094 | 415.3476 | 277.0467 | 297.302 | 297.302 |  | 228.8233
0 | [10 empty] | 0 | 277.0467 | 415.3476 | 41.094 | 41.094 | 415.3476 | 277.0467 | 297.302 | 297.302 |  | 228.8233
0 | [10 empty] | 0.099517 | 277.0965 | 415.3476 | 41.094 | 41.094 | 415.3476 | 277.1463 | 297.3517 | 297.4015 |  | 228.8731
0 | LS2 | CC07 | CAV6 | 2974 | LS2_CC07:CAV6_D2974 |  | b29 resonator | 40 |  |  | 0.32364 | 277.3081 | 415.3476 | 41.094 | 41.094 | 415.3476 | 277.4699 | 297.5633 | 297.7251 |  | 229.0847
0 | LS2 | CC07 | BLM | 2976 | LS2_CC07:BLM_D2976 |  | outside loss monitor |  |  |  | 0 | 277.4699 | 415.3476 | 41.094 | 41.094 | 415.3476 | 277.4699 | 297.7251 | 297.7251 |  | 229.2465
0 | [10 empty] | 0 | 277.4699 | 415.3476 | 41.094 | 41.094 | 415.3476 | 277.4699 | 297.7251 | 297.7251 |  | 229.2465
0 | [6 empty] | bellows |  |  |  | 0.1677 | 277.5537 | 415.3476 | 41.094 | 41.094 | 415.3476 | 277.6376 | 297.809 | 297.8928 |  | 229.3303
0 | LS2 | CC07 | GV | 2979 | LS2_CC07:GV_D2979 |  | gate valve | 40 |  |  | 0.057404 | 277.6663 | 415.3476 | 41.094 | 41.094 | 415.3476 | 277.695 | 297.9215 | 297.9502 |  | 229.4429
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 277.7468 | 415.3476 | 41.094 | 41.094 | 415.3476 | 277.7985 | 298.002 | 298.0537 |  | 229.5233
0 | LS2 | WC07 | BPM | 2981 | LS2_WC07:BPM_D2981 | BPM2 | position monitor |  |  |  | 0 | 277.7985 | 415.3476 | 41.094 | 41.094 | 415.3476 | 277.7985 | 298.0537 | 298.0537 |  | 229.5751
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 277.8711 | 415.3476 | 41.094 | 41.094 | 415.3476 | 277.9438 | 298.1264 | 298.199 |  | 229.6477
0 | LS2 | WC07 | PORT | 2982 | LS2_WC07:PORT_D2982 |  | port for future diagnostics |  |  |  | 0 | 277.9438 | 415.3476 | 41.094 | 41.094 | 415.3476 | 277.9438 | 298.199 | 298.199 |  | 229.7204
0 | [6 empty] | diagnostic box |  |  |  | 0 | 277.9438 | 415.3476 | 41.094 | 41.094 | 415.3476 | 277.9438 | 298.199 | 298.199 |  | 229.7204
0 | LS2 | WC07 | IP | 2982 | LS2_WC07:IP_D2982 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 277.9438 | 415.3476 | 41.094 | 41.094 | 415.3476 | 277.9438 | 298.199 | 298.199 |  | 229.7204
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 278.0094 | 415.3476 | 41.094 | 41.094 | 415.3476 | 278.075 | 298.2646 | 298.3302 |  | 229.786
0 | LS2 | CC08 | GV | 2984 | LS2_CC08:GV_D2984 |  | gate valve | 40 |  |  | 0.057404 | 278.1037 | 415.3476 | 41.094 | 41.094 | 415.3476 | 278.1324 | 298.3589 | 298.3876 |  | 229.8803
0 | [6 empty] | bellows |  |  |  | 0.1677 | 278.2163 | 415.3476 | 41.094 | 41.094 | 415.3476 | 278.3001 | 298.4715 | 298.5553 |  | 229.9928
0 | LS2 | CC08 | CAV1 | 2985 | LS2_CC08:CAV1_D2985 |  | b29 resonator | 40 | 0.32364 |  | 0.32364 | 278.4619 | 415.3476 | 41.094 | 41.094 | 415.3476 | 278.6238 | 298.7172 | 298.879 |  | 230.2385
0 | LS2 | CC08 | BLM | 2987 | LS2_CC08:BLM_D2987 |  | outside loss monitor |  |  |  | 0 | 278.6238 | 415.3476 | 41.094 | 41.094 | 415.3476 | 278.6238 | 298.879 | 298.879 |  | 230.4003
0 | [10 empty] | 0 | 278.6238 | 415.3476 | 41.094 | 41.094 | 415.3476 | 278.6238 | 298.879 | 298.879 |  | 230.4003
0 | [6 empty] | bellow |  |  |  | 0.099517 | 278.6735 | 415.3476 | 41.094 | 41.094 | 415.3476 | 278.7233 | 298.9287 | 298.9785 |  | 230.4501
0 | LS2 | CC08 | CAV2 | 2989 | LS2_CC08:CAV2_D2989 |  | b29 resonator | 40 |  |  | 0.32364 | 278.8851 | 415.3476 | 41.094 | 41.094 | 415.3476 | 279.0469 | 299.1403 | 299.3021 |  | 230.6617
0 | [10 empty] | 0 | 279.0469 | 415.3476 | 41.094 | 41.094 | 415.3476 | 279.0469 | 299.3021 | 299.3021 |  | 230.8235
0 | [10 empty] | 0 | 279.0469 | 415.3476 | 41.094 | 41.094 | 415.3476 | 279.0469 | 299.3021 | 299.3021 |  | 230.8235
0 | [6 empty] | bellow |  |  |  | 0.099517 | 279.0967 | 415.3476 | 41.094 | 41.094 | 415.3476 | 279.1464 | 299.3519 | 299.4017 |  | 230.8733
0 | LS2 | CC08 | CAV3 | 2994 | LS2_CC08:CAV3_D2994 |  | b29 resonator | 40 |  |  | 0.32364 | 279.3083 | 415.3476 | 41.094 | 41.094 | 415.3476 | 279.4701 | 299.5635 | 299.7253 |  | 231.0848
0 | LS2 | CC08 | BLM | 2996 | LS2_CC08:BLM_D2996 |  | outside loss monitor |  |  |  | 0 | 279.4701 | 415.3476 | 41.094 | 41.094 | 415.3476 | 279.4701 | 299.7253 | 299.7253 |  | 231.2467
0 | [10 empty] | 0 | 279.4701 | 415.3476 | 41.094 | 41.094 | 415.3476 | 279.4701 | 299.7253 | 299.7253 |  | 231.2467
0 | [6 empty] | bellow+tube |  |  |  | 0.15525 | 279.5477 | 415.3476 | 41.094 | 41.094 | 415.3476 | 279.6253 | 299.8029 | 299.8805 |  | 231.3243
0 | LS2 | CC08 | SOL1 | 3001 | LS2_CC08:SOL1_D3001 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 279.8753 | 415.3476 | 41.094 | 41.094 | 415.3476 | 280.1253 | 300.1305 | 300.3805 |  | 231.6519
0 | [6 empty] | bellow+tube |  |  |  | 0.15525 | 280.203 | 415.3476 | 41.094 | 41.094 | 415.3476 | 280.2806 | 300.4582 | 300.5358 |  | 231.9795
0 | LS2 | CC08 | CAV4 | 3005 | LS2_CC08:CAV4_D3005 |  | b29 resonator | 40 |  |  | 0.32364 | 280.4424 | 415.3476 | 41.094 | 41.094 | 415.3476 | 280.6042 | 300.6976 | 300.8594 |  | 232.219
0 | [10 empty] | 0 | 280.6042 | 415.3476 | 41.094 | 41.094 | 415.3476 | 280.6042 | 300.8594 | 300.8594 |  | 232.3808
0 | [10 empty] | 0 | 280.6042 | 415.3476 | 41.094 | 41.094 | 415.3476 | 280.6042 | 300.8594 | 300.8594 |  | 232.3808
0 | [10 empty] | 0.099517 | 280.654 | 415.3476 | 41.094 | 41.094 | 415.3476 | 280.7037 | 300.9092 | 300.959 |  | 232.4306
0 | LS2 | CC08 | CAV5 | 3009 | LS2_CC08:CAV5_D3009 |  | b29 resonator | 40 |  |  | 0.32364 | 280.8656 | 415.3476 | 41.094 | 41.094 | 415.3476 | 281.0274 | 301.1208 | 301.2826 |  | 232.6421
0 | [10 empty] | 0 | 281.0274 | 415.3476 | 41.094 | 41.094 | 415.3476 | 281.0274 | 301.2826 | 301.2826 |  | 232.804
0 | [10 empty] | 0 | 281.0274 | 415.3476 | 41.094 | 41.094 | 415.3476 | 281.0274 | 301.2826 | 301.2826 |  | 232.804
0 | [10 empty] | 0.099517 | 281.0771 | 415.3476 | 41.094 | 41.094 | 415.3476 | 281.1269 | 301.3324 | 301.3821 |  | 232.8537
0 | LS2 | CC08 | CAV6 | 3013 | LS2_CC08:CAV6_D3013 |  | b29 resonator | 40 |  |  | 0.32364 | 281.2887 | 415.3476 | 41.094 | 41.094 | 415.3476 | 281.4505 | 301.5439 | 301.7058 |  | 233.0653
0 | LS2 | CC08 | BLM | 3015 | LS2_CC08:BLM_D3015 |  | outside loss monitor |  |  |  | 0 | 281.4505 | 415.3476 | 41.094 | 41.094 | 415.3476 | 281.4505 | 301.7058 | 301.7058 |  | 233.2271
0 | [10 empty] | 0 | 281.4505 | 415.3476 | 41.094 | 41.094 | 415.3476 | 281.4505 | 301.7058 | 301.7058 |  | 233.2271
0 | [6 empty] | bellows |  |  |  | 0.1677 | 281.5344 | 415.3476 | 41.094 | 41.094 | 415.3476 | 281.6182 | 301.7896 | 301.8735 |  | 233.311
0 | LS2 | CC08 | GV | 3019 | LS2_CC08:GV_D3019 |  | gate valve | 40 |  |  | 0.057404 | 281.647 | 415.3476 | 41.094 | 41.094 | 415.3476 | 281.6757 | 301.9022 | 301.9309 |  | 233.4235
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 281.7274 | 415.3476 | 41.094 | 41.094 | 415.3476 | 281.7792 | 301.9826 | 302.0344 |  | 233.504
0 | LS2 | WC08 | BPM | 3020 | LS2_WC08:BPM_D3020 | BPM2 | position monitor |  |  |  | 0 | 281.7792 | 415.3476 | 41.094 | 41.094 | 415.3476 | 281.7792 | 302.0344 | 302.0344 |  | 233.5557
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 281.8518 | 415.3476 | 41.094 | 41.094 | 415.3476 | 281.9244 | 302.107 | 302.1797 |  | 233.6284
0 | LS2 | WC08 | PORT | 3022 | LS2_WC08:PORT_D3022 |  | port for future diagnostics |  |  |  | 0 | 281.9244 | 415.3476 | 41.094 | 41.094 | 415.3476 | 281.9244 | 302.1797 | 302.1797 |  | 233.701
0 | [6 empty] | diagnostic box |  |  |  | 0 | 281.9244 | 415.3476 | 41.094 | 41.094 | 415.3476 | 281.9244 | 302.1797 | 302.1797 |  | 233.701
0 | LS2 | WC08 | IP | 3022 | LS2_WC08:IP_D3022 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 281.9244 | 415.3476 | 41.094 | 41.094 | 415.3476 | 281.9244 | 302.1797 | 302.1797 |  | 233.701
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 281.99 | 415.3476 | 41.094 | 41.094 | 415.3476 | 282.0557 | 302.2453 | 302.3109 |  | 233.7666
0 | LS2 | CC09 | GV | 3023 | LS2_CC09:GV_D3023 |  | gate valve | 40 |  |  | 0.057404 | 282.0844 | 415.3476 | 41.094 | 41.094 | 415.3476 | 282.1131 | 302.3396 | 302.3683 |  | 233.8609
0 | [6 empty] | bellows |  |  |  | 0.1677 | 282.1969 | 415.3476 | 41.094 | 41.094 | 415.3476 | 282.2808 | 302.4521 | 302.536 |  | 233.9735
0 | LS2 | CC09 | CAV1 | 3025 | LS2_CC09:CAV1_D3025 |  | b29 resonator | 40 | 0.32364 |  | 0.32364 | 282.4426 | 415.3476 | 41.094 | 41.094 | 415.3476 | 282.6044 | 302.6978 | 302.8596 |  | 234.2192
0 | LS2 | CC09 | BLM | 3027 | LS2_CC09:BLM_D3027 |  | outside loss monitor |  |  |  | 0 | 282.6044 | 415.3476 | 41.094 | 41.094 | 415.3476 | 282.6044 | 302.8596 | 302.8596 |  | 234.381
0 | [10 empty] | 0 | 282.6044 | 415.3476 | 41.094 | 41.094 | 415.3476 | 282.6044 | 302.8596 | 302.8596 |  | 234.381
0 | [6 empty] | bellow |  |  |  | 0.099517 | 282.6542 | 415.3476 | 41.094 | 41.094 | 415.3476 | 282.7039 | 302.9094 | 302.9591 |  | 234.4307
0 | LS2 | CC09 | CAV2 | 3029 | LS2_CC09:CAV2_D3029 |  | b29 resonator | 40 |  |  | 0.32364 | 282.8657 | 415.3476 | 41.094 | 41.094 | 415.3476 | 283.0276 | 303.121 | 303.2828 |  | 234.6423
0 | [10 empty] | 0 | 283.0276 | 415.3476 | 41.094 | 41.094 | 415.3476 | 283.0276 | 303.2828 | 303.2828 |  | 234.8042
0 | [10 empty] | 0 | 283.0276 | 415.3476 | 41.094 | 41.094 | 415.3476 | 283.0276 | 303.2828 | 303.2828 |  | 234.8042
0 | [6 empty] | bellow |  |  |  | 0.099517 | 283.0773 | 415.3476 | 41.094 | 41.094 | 415.3476 | 283.1271 | 303.3325 | 303.3823 |  | 234.8539
0 | LS2 | CC09 | CAV3 | 3033 | LS2_CC09:CAV3_D3033 |  | b29 resonator | 40 |  |  | 0.32364 | 283.2889 | 415.3476 | 41.094 | 41.094 | 415.3476 | 283.4507 | 303.5441 | 303.7059 |  | 235.0655
0 | LS2 | CC09 | BLM | 3035 | LS2_CC09:BLM_D3035 |  | outside loss monitor |  |  |  | 0 | 283.4507 | 415.3476 | 41.094 | 41.094 | 415.3476 | 283.4507 | 303.7059 | 303.7059 |  | 235.2273
0 | [10 empty] | 0 | 283.4507 | 415.3476 | 41.094 | 41.094 | 415.3476 | 283.4507 | 303.7059 | 303.7059 |  | 235.2273
0 | [6 empty] | bellow+tube |  |  |  | 0.15525 | 283.5284 | 415.3476 | 41.094 | 41.094 | 415.3476 | 283.606 | 303.7836 | 303.8612 |  | 235.3049
0 | LS2 | CC09 | SOL1 | 3041 | LS2_CC09:SOL1_D3041 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 283.856 | 415.3476 | 41.094 | 41.094 | 415.3476 | 284.106 | 304.1112 | 304.3612 |  | 235.6326
0 | [6 empty] | bellow+tube |  |  |  | 0.15525 | 284.1836 | 415.3476 | 41.094 | 41.094 | 415.3476 | 284.2612 | 304.4388 | 304.5164 |  | 235.9602
0 | LS2 | CC09 | CAV4 | 3044 | LS2_CC09:CAV4_D3044 |  | b29 resonator | 40 |  |  | 0.32364 | 284.423 | 415.3476 | 41.094 | 41.094 | 415.3476 | 284.5849 | 304.6783 | 304.8401 |  | 236.1996
0 | [10 empty] | 0 | 284.5849 | 415.3476 | 41.094 | 41.094 | 415.3476 | 284.5849 | 304.8401 | 304.8401 |  | 236.3615
0 | [10 empty] | 0 | 284.5849 | 415.3476 | 41.094 | 41.094 | 415.3476 | 284.5849 | 304.8401 | 304.8401 |  | 236.3615
0 | [10 empty] | 0.099517 | 284.6346 | 415.3476 | 41.094 | 41.094 | 415.3476 | 284.6844 | 304.8898 | 304.9396 |  | 236.4112
0 | LS2 | CC09 | CAV5 | 3049 | LS2_CC09:CAV5_D3049 |  | b29 resonator | 40 |  |  | 0.32364 | 284.8462 | 415.3476 | 41.094 | 41.094 | 415.3476 | 285.008 | 305.1014 | 305.2632 |  | 236.6228
0 | [10 empty] | 0 | 285.008 | 415.3476 | 41.094 | 41.094 | 415.3476 | 285.008 | 305.2632 | 305.2632 |  | 236.7846
0 | [10 empty] | 0 | 285.008 | 415.3476 | 41.094 | 41.094 | 415.3476 | 285.008 | 305.2632 | 305.2632 |  | 236.7846
0 | [10 empty] | 0.099517 | 285.0578 | 415.3476 | 41.094 | 41.094 | 415.3476 | 285.1075 | 305.313 | 305.3628 |  | 236.8344
0 | LS2 | CC09 | CAV6 | 3053 | LS2_CC09:CAV6_D3053 |  | b29 resonator | 40 |  |  | 0.32364 | 285.2694 | 415.3476 | 41.094 | 41.094 | 415.3476 | 285.4312 | 305.5246 | 305.6864 |  | 237.046
0 | LS2 | CC09 | BLM | 3055 | LS2_CC09:BLM_D3055 |  | outside loss monitor |  |  |  | 0 | 285.4312 | 415.3476 | 41.094 | 41.094 | 415.3476 | 285.4312 | 305.6864 | 305.6864 |  | 237.2078
0 | [10 empty] | 0 | 285.4312 | 415.3476 | 41.094 | 41.094 | 415.3476 | 285.4312 | 305.6864 | 305.6864 |  | 237.2078
0 | [6 empty] | bellows |  |  |  | 0.1677 | 285.515 | 415.3476 | 41.094 | 41.094 | 415.3476 | 285.5989 | 305.7703 | 305.8541 |  | 237.2916
0 | LS2 | CC09 | GV | 3059 | LS2_CC09:GV_D3059 |  | gate valve | 40 |  |  | 0.057404 | 285.6276 | 415.3476 | 41.094 | 41.094 | 415.3476 | 285.6563 | 305.8828 | 305.9115 |  | 237.4042
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 285.7081 | 415.3476 | 41.094 | 41.094 | 415.3476 | 285.7598 | 305.9633 | 306.015 |  | 237.4846
0 | LS2 | WC09 | BPM | 3060 | LS2_WC09:BPM_D3060 | BPM2 | position monitor |  |  |  | 0 | 285.7598 | 415.3476 | 41.094 | 41.094 | 415.3476 | 285.7598 | 306.015 | 306.015 |  | 237.5364
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 285.8324 | 415.3476 | 41.094 | 41.094 | 415.3476 | 285.9051 | 306.0877 | 306.1603 |  | 237.609
0 | LS2 | WC09 | PORT | 3062 | LS2_WC09:PORT_D3062 |  | port for future diagnostics |  |  |  | 0 | 285.9051 | 415.3476 | 41.094 | 41.094 | 415.3476 | 285.9051 | 306.1603 | 306.1603 |  | 237.6817
0 | [6 empty] | diagnostic box |  |  |  | 0 | 285.9051 | 415.3476 | 41.094 | 41.094 | 415.3476 | 285.9051 | 306.1603 | 306.1603 |  | 237.6817
0 | LS2 | WC09 | IP | 3062 | LS2_WC09:IP_D3062 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 285.9051 | 415.3476 | 41.094 | 41.094 | 415.3476 | 285.9051 | 306.1603 | 306.1603 |  | 237.6817
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 285.9707 | 415.3476 | 41.094 | 41.094 | 415.3476 | 286.0363 | 306.2259 | 306.2915 |  | 237.7473
0 | LS2 | CC10 | GV | 3063 | LS2_CC10:GV_D3063 |  | gate valve | 40 |  |  | 0.057404 | 286.065 | 415.3476 | 41.094 | 41.094 | 415.3476 | 286.0937 | 306.3202 | 306.3489 |  | 237.8416
0 | [6 empty] | bellows |  |  |  | 0.1677 | 286.1776 | 415.3476 | 41.094 | 41.094 | 415.3476 | 286.2614 | 306.4328 | 306.5166 |  | 237.9541
0 | LS2 | CC10 | CAV1 | 3064 | LS2_CC10:CAV1_D3064 |  | b29 resonator | 40 | 0.32364 |  | 0.32364 | 286.4232 | 415.3476 | 41.094 | 41.094 | 415.3476 | 286.5851 | 306.6784 | 306.8403 |  | 238.1998
0 | LS2 | CC10 | BLM | 3067 | LS2_CC10:BLM_D3067 |  | outside loss monitor |  |  |  | 0 | 286.5851 | 415.3476 | 41.094 | 41.094 | 415.3476 | 286.5851 | 306.8403 | 306.8403 |  | 238.3616
0 | [10 empty] | 0 | 286.5851 | 415.3476 | 41.094 | 41.094 | 415.3476 | 286.5851 | 306.8403 | 306.8403 |  | 238.3616
0 | [6 empty] | bellow |  |  |  | 0.099517 | 286.6348 | 415.3476 | 41.094 | 41.094 | 415.3476 | 286.6846 | 306.89 | 306.9398 |  | 238.4114
0 | LS2 | CC10 | CAV2 | 3069 | LS2_CC10:CAV2_D3069 |  | b29 resonator | 40 |  |  | 0.32364 | 286.8464 | 415.3476 | 41.094 | 41.094 | 415.3476 | 287.0082 | 307.1016 | 307.2634 |  | 238.623
0 | [10 empty] | 0 | 287.0082 | 415.3476 | 41.094 | 41.094 | 415.3476 | 287.0082 | 307.2634 | 307.2634 |  | 238.7848
0 | [10 empty] | 0 | 287.0082 | 415.3476 | 41.094 | 41.094 | 415.3476 | 287.0082 | 307.2634 | 307.2634 |  | 238.7848
0 | [6 empty] | bellow |  |  |  | 0.099517 | 287.058 | 415.3476 | 41.094 | 41.094 | 415.3476 | 287.1077 | 307.3132 | 307.3629 |  | 238.8346
0 | LS2 | CC10 | CAV3 | 3073 | LS2_CC10:CAV3_D3073 |  | b29 resonator | 40 |  |  | 0.32364 | 287.2696 | 415.3476 | 41.094 | 41.094 | 415.3476 | 287.4314 | 307.5248 | 307.6866 |  | 239.0461
0 | LS2 | CC10 | BLM | 3075 | LS2_CC10:BLM_D3075 |  | outside loss monitor |  |  |  | 0 | 287.4314 | 415.3476 | 41.094 | 41.094 | 415.3476 | 287.4314 | 307.6866 | 307.6866 |  | 239.208
0 | [10 empty] | 0 | 287.4314 | 415.3476 | 41.094 | 41.094 | 415.3476 | 287.4314 | 307.6866 | 307.6866 |  | 239.208
0 | [6 empty] | bellow+tube |  |  |  | 0.15525 | 287.509 | 415.3476 | 41.094 | 41.094 | 415.3476 | 287.5866 | 307.7642 | 307.8418 |  | 239.2856
0 | LS2 | CC10 | SOL1 | 3081 | LS2_CC10:SOL1_D3081 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 287.8366 | 415.3476 | 41.094 | 41.094 | 415.3476 | 288.0866 | 308.0918 | 308.3418 |  | 239.6132
0 | [6 empty] | bellow+tube |  |  |  | 0.15525 | 288.1642 | 415.3476 | 41.094 | 41.094 | 415.3476 | 288.2419 | 308.4195 | 308.4971 |  | 239.9408
0 | LS2 | CC10 | CAV4 | 3084 | LS2_CC10:CAV4_D3084 |  | b29 resonator | 40 |  |  | 0.32364 | 288.4037 | 415.3476 | 41.094 | 41.094 | 415.3476 | 288.5655 | 308.6589 | 308.8207 |  | 240.1803
0 | [10 empty] | 0 | 288.5655 | 415.3476 | 41.094 | 41.094 | 415.3476 | 288.5655 | 308.8207 | 308.8207 |  | 240.3421
0 | [10 empty] | 0 | 288.5655 | 415.3476 | 41.094 | 41.094 | 415.3476 | 288.5655 | 308.8207 | 308.8207 |  | 240.3421
0 | [10 empty] | 0.099517 | 288.6153 | 415.3476 | 41.094 | 41.094 | 415.3476 | 288.665 | 308.8705 | 308.9203 |  | 240.3919
0 | LS2 | CC10 | CAV5 | 3089 | LS2_CC10:CAV5_D3089 |  | b29 resonator | 40 |  |  | 0.32364 | 288.8269 | 415.3476 | 41.094 | 41.094 | 415.3476 | 288.9887 | 309.0821 | 309.2439 |  | 240.6034
0 | [10 empty] | 0 | 288.9887 | 415.3476 | 41.094 | 41.094 | 415.3476 | 288.9887 | 309.2439 | 309.2439 |  | 240.7653
0 | [10 empty] | 0 | 288.9887 | 415.3476 | 41.094 | 41.094 | 415.3476 | 288.9887 | 309.2439 | 309.2439 |  | 240.7653
0 | [10 empty] | 0.099517 | 289.0384 | 415.3476 | 41.094 | 41.094 | 415.3476 | 289.0882 | 309.2937 | 309.3434 |  | 240.815
0 | LS2 | CC10 | CAV6 | 3093 | LS2_CC10:CAV6_D3093 |  | b29 resonator | 40 |  |  | 0.32364 | 289.25 | 415.3476 | 41.094 | 41.094 | 415.3476 | 289.4118 | 309.5052 | 309.6671 |  | 241.0266
0 | LS2 | CC10 | BLM | 3095 | LS2_CC10:BLM_D3095 |  | outside loss monitor |  |  |  | 0 | 289.4118 | 415.3476 | 41.094 | 41.094 | 415.3476 | 289.4118 | 309.6671 | 309.6671 |  | 241.1884
0 | [10 empty] | 0 | 289.4118 | 415.3476 | 41.094 | 41.094 | 415.3476 | 289.4118 | 309.6671 | 309.6671 |  | 241.1884
0 | [6 empty] | bellows |  |  |  | 0.1677 | 289.4957 | 415.3476 | 41.094 | 41.094 | 415.3476 | 289.5795 | 309.7509 | 309.8348 |  | 241.2723
0 | LS2 | CC10 | GV | 3099 | LS2_CC10:GV_D3099 |  | gate valve | 40 |  |  | 0.057404 | 289.6082 | 415.3476 | 41.094 | 41.094 | 415.3476 | 289.6369 | 309.8635 | 309.8922 |  | 241.3848
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 289.6887 | 415.3476 | 41.094 | 41.094 | 415.3476 | 289.7405 | 309.9439 | 309.9957 |  | 241.4653
0 | LS2 | WC10 | BPM | 3100 | LS2_WC10:BPM_D3100 | BPM2 | position monitor |  |  |  | 0 | 289.7405 | 415.3476 | 41.094 | 41.094 | 415.3476 | 289.7405 | 309.9957 | 309.9957 |  | 241.517
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 289.8131 | 415.3476 | 41.094 | 41.094 | 415.3476 | 289.8857 | 310.0683 | 310.1409 |  | 241.5897
0 | LS2 | WC10 | PORT | 3101 | LS2_WC10:PORT_D3101 |  | port for future diagnostics |  |  |  | 0 | 289.8857 | 415.3476 | 41.094 | 41.094 | 415.3476 | 289.8857 | 310.1409 | 310.1409 |  | 241.6623
0 | [6 empty] | diagnostic box |  |  |  | 0 | 289.8857 | 415.3476 | 41.094 | 41.094 | 415.3476 | 289.8857 | 310.1409 | 310.1409 |  | 241.6623
0 | LS2 | WC10 | IP | 3101 | LS2_WC10:IP_D3101 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 289.8857 | 415.3476 | 41.094 | 41.094 | 415.3476 | 289.8857 | 310.1409 | 310.1409 |  | 241.6623
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 289.9513 | 415.3476 | 41.094 | 41.094 | 415.3476 | 290.0169 | 310.2066 | 310.2722 |  | 241.7279
0 | LS2 | CC11 | GV | 3103 | LS2_CC11:GV_D3103 |  | gate valve | 40 |  |  | 0.057404 | 290.0457 | 415.3476 | 41.094 | 41.094 | 415.3476 | 290.0744 | 310.3009 | 310.3296 |  | 241.8222
0 | [6 empty] | bellows |  |  |  | 0.1677 | 290.1582 | 415.3476 | 41.094 | 41.094 | 415.3476 | 290.2421 | 310.4134 | 310.4973 |  | 241.9348
0 | LS2 | CC11 | CAV1 | 3104 | LS2_CC11:CAV1_D3104 |  | b29 resonator | 40 | 0.32364 |  | 0.32364 | 290.4039 | 415.3476 | 41.094 | 41.094 | 415.3476 | 290.5657 | 310.6591 | 310.8209 |  | 242.1805
0 | LS2 | CC11 | BLM | 3107 | LS2_CC11:BLM_D3107 |  | outside loss monitor |  |  |  | 0 | 290.5657 | 415.3476 | 41.094 | 41.094 | 415.3476 | 290.5657 | 310.8209 | 310.8209 |  | 242.3423
0 | [10 empty] | 0 | 290.5657 | 415.3476 | 41.094 | 41.094 | 415.3476 | 290.5657 | 310.8209 | 310.8209 |  | 242.3423
0 | [6 empty] | bellow |  |  |  | 0.099517 | 290.6155 | 415.3476 | 41.094 | 41.094 | 415.3476 | 290.6652 | 310.8707 | 310.9204 |  | 242.392
0 | LS2 | CC11 | CAV2 | 3109 | LS2_CC11:CAV2_D3109 |  | b29 resonator | 40 |  |  | 0.32364 | 290.827 | 415.3476 | 41.094 | 41.094 | 415.3476 | 290.9889 | 311.0823 | 311.2441 |  | 242.6036
0 | [10 empty] | 0 | 290.9889 | 415.3476 | 41.094 | 41.094 | 415.3476 | 290.9889 | 311.2441 | 311.2441 |  | 242.7654
0 | [10 empty] | 0 | 290.9889 | 415.3476 | 41.094 | 41.094 | 415.3476 | 290.9889 | 311.2441 | 311.2441 |  | 242.7654
0 | [6 empty] | bellow |  |  |  | 0.099517 | 291.0386 | 415.3476 | 41.094 | 41.094 | 415.3476 | 291.0884 | 311.2938 | 311.3436 |  | 242.8152
0 | LS2 | CC11 | CAV3 | 3113 | LS2_CC11:CAV3_D3113 |  | b29 resonator | 40 |  |  | 0.32364 | 291.2502 | 415.3476 | 41.094 | 41.094 | 415.3476 | 291.412 | 311.5054 | 311.6672 |  | 243.0268
0 | LS2 | CC11 | BLM | 3115 | LS2_CC11:BLM_D3115 |  | outside loss monitor |  |  |  | 0 | 291.412 | 415.3476 | 41.094 | 41.094 | 415.3476 | 291.412 | 311.6672 | 311.6672 |  | 243.1886
0 | [10 empty] | 0 | 291.412 | 415.3476 | 41.094 | 41.094 | 415.3476 | 291.412 | 311.6672 | 311.6672 |  | 243.1886
0 | [6 empty] | bellow+tube |  |  |  | 0.15525 | 291.4896 | 415.3476 | 41.094 | 41.094 | 415.3476 | 291.5673 | 311.7449 | 311.8225 |  | 243.2662
0 | LS2 | CC11 | SOL1 | 3121 | LS2_CC11:SOL1_D3121 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 291.8173 | 415.3476 | 41.094 | 41.094 | 415.3476 | 292.0673 | 312.0725 | 312.3225 |  | 243.5939
0 | [6 empty] | bellow+tube |  |  |  | 0.15525 | 292.1449 | 415.3476 | 41.094 | 41.094 | 415.3476 | 292.2225 | 312.4001 | 312.4777 |  | 243.9215
0 | LS2 | CC11 | CAV4 | 3124 | LS2_CC11:CAV4_D3124 |  | b29 resonator | 40 |  |  | 0.32364 | 292.3843 | 415.3476 | 41.094 | 41.094 | 415.3476 | 292.5462 | 312.6396 | 312.8014 |  | 244.1609
0 | [10 empty] | 0 | 292.5462 | 415.3476 | 41.094 | 41.094 | 415.3476 | 292.5462 | 312.8014 | 312.8014 |  | 244.3228
0 | [10 empty] | 0 | 292.5462 | 415.3476 | 41.094 | 41.094 | 415.3476 | 292.5462 | 312.8014 | 312.8014 |  | 244.3228
0 | [10 empty] | 0.099517 | 292.5959 | 415.3476 | 41.094 | 41.094 | 415.3476 | 292.6457 | 312.8511 | 312.9009 |  | 244.3725
0 | LS2 | CC11 | CAV5 | 3129 | LS2_CC11:CAV5_D3129 |  | b29 resonator | 40 |  |  | 0.32364 | 292.8075 | 415.3476 | 41.094 | 41.094 | 415.3476 | 292.9693 | 313.0627 | 313.2245 |  | 244.5841
0 | [10 empty] | 0 | 292.9693 | 415.3476 | 41.094 | 41.094 | 415.3476 | 292.9693 | 313.2245 | 313.2245 |  | 244.7459
0 | [10 empty] | 0 | 292.9693 | 415.3476 | 41.094 | 41.094 | 415.3476 | 292.9693 | 313.2245 | 313.2245 |  | 244.7459
0 | [10 empty] | 0.099517 | 293.0191 | 415.3476 | 41.094 | 41.094 | 415.3476 | 293.0688 | 313.2743 | 313.3241 |  | 244.7957
0 | LS2 | CC11 | CAV6 | 3133 | LS2_CC11:CAV6_D3133 |  | b29 resonator | 40 |  |  | 0.32364 | 293.2307 | 415.3476 | 41.094 | 41.094 | 415.3476 | 293.3925 | 313.4859 | 313.6477 |  | 245.0073
0 | LS2 | CC11 | BLM | 3135 | LS2_CC11:BLM_D3135 |  | outside loss monitor |  |  |  | 0 | 293.3925 | 415.3476 | 41.094 | 41.094 | 415.3476 | 293.3925 | 313.6477 | 313.6477 |  | 245.1691
0 | [10 empty] | 0 | 293.3925 | 415.3476 | 41.094 | 41.094 | 415.3476 | 293.3925 | 313.6477 | 313.6477 |  | 245.1691
0 | [6 empty] | bellows |  |  |  | 0.1677 | 293.4763 | 415.3476 | 41.094 | 41.094 | 415.3476 | 293.5602 | 313.7316 | 313.8154 |  | 245.2529
0 | LS2 | CC11 | GV | 3138 | LS2_CC11:GV_D3138 |  | gate valve | 40 |  |  | 0.057404 | 293.5889 | 415.3476 | 41.094 | 41.094 | 415.3476 | 293.6176 | 313.8441 | 313.8728 |  | 245.3655
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 293.6693 | 415.3476 | 41.094 | 41.094 | 415.3476 | 293.7211 | 313.9246 | 313.9763 |  | 245.4459
0 | LS2 | WC11 | BPM | 3140 | LS2_WC11:BPM_D3140 | BPM2 | position monitor |  |  |  | 0 | 293.7211 | 415.3476 | 41.094 | 41.094 | 415.3476 | 293.7211 | 313.9763 | 313.9763 |  | 245.4977
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 293.7937 | 415.3476 | 41.094 | 41.094 | 415.3476 | 293.8664 | 314.049 | 314.1216 |  | 245.5703
0 | LS2 | WC11 | PORT | 3141 | LS2_WC11:PORT_D3141 |  | port for future diagnostics |  |  |  | 0 | 293.8664 | 415.3476 | 41.094 | 41.094 | 415.3476 | 293.8664 | 314.1216 | 314.1216 |  | 245.643
0 | [6 empty] | diagnostic box |  |  |  | 0 | 293.8664 | 415.3476 | 41.094 | 41.094 | 415.3476 | 293.8664 | 314.1216 | 314.1216 |  | 245.643
0 | LS2 | WC11 | IP | 3141 | LS2_WC11:IP_D3141 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 293.8664 | 415.3476 | 41.094 | 41.094 | 415.3476 | 293.8664 | 314.1216 | 314.1216 |  | 245.643
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 293.932 | 415.3476 | 41.094 | 41.094 | 415.3476 | 293.9976 | 314.1872 | 314.2528 |  | 245.7086
0 | LS2 | CC12 | GV | 3143 | LS2_CC12:GV_D3143 |  | gate valve | 40 |  |  | 0.057404 | 294.0263 | 415.3476 | 41.094 | 41.094 | 415.3476 | 294.055 | 314.2815 | 314.3102 |  | 245.8029
0 | [6 empty] | bellows |  |  |  | 0.1677 | 294.1389 | 415.3476 | 41.094 | 41.094 | 415.3476 | 294.2227 | 314.3941 | 314.4779 |  | 245.9154
0 | LS2 | CC12 | CAV1 | 3144 | LS2_CC12:CAV1_D3144 |  | b29 resonator | 40 | 0.32364 |  | 0.32364 | 294.3845 | 415.3476 | 41.094 | 41.094 | 415.3476 | 294.5463 | 314.6397 | 314.8016 |  | 246.1611
0 | LS2 | CC12 | BLM | 3146 | LS2_CC12:BLM_D3146 |  | outside loss monitor |  |  |  | 0 | 294.5463 | 415.3476 | 41.094 | 41.094 | 415.3476 | 294.5463 | 314.8016 | 314.8016 |  | 246.3229
0 | [10 empty] | 0 | 294.5463 | 415.3476 | 41.094 | 41.094 | 415.3476 | 294.5463 | 314.8016 | 314.8016 |  | 246.3229
0 | [6 empty] | bellow |  |  |  | 0.099517 | 294.5961 | 415.3476 | 41.094 | 41.094 | 415.3476 | 294.6459 | 314.8513 | 314.9011 |  | 246.3727
0 | LS2 | CC12 | CAV2 | 3149 | LS2_CC12:CAV2_D3149 |  | b29 resonator | 40 |  |  | 0.32364 | 294.8077 | 415.3476 | 41.094 | 41.094 | 415.3476 | 294.9695 | 315.0629 | 315.2247 |  | 246.5843
0 | [10 empty] | 0 | 294.9695 | 415.3476 | 41.094 | 41.094 | 415.3476 | 294.9695 | 315.2247 | 315.2247 |  | 246.7461
0 | [10 empty] | 0 | 294.9695 | 415.3476 | 41.094 | 41.094 | 415.3476 | 294.9695 | 315.2247 | 315.2247 |  | 246.7461
0 | [6 empty] | bellow |  |  |  | 0.099517 | 295.0193 | 415.3476 | 41.094 | 41.094 | 415.3476 | 295.069 | 315.2745 | 315.3242 |  | 246.7959
0 | LS2 | CC12 | CAV3 | 3153 | LS2_CC12:CAV3_D3153 |  | b29 resonator | 40 |  |  | 0.32364 | 295.2308 | 415.3476 | 41.094 | 41.094 | 415.3476 | 295.3927 | 315.4861 | 315.6479 |  | 247.0074
0 | LS2 | CC12 | BLM | 3155 | LS2_CC12:BLM_D3155 |  | outside loss monitor |  |  |  | 0 | 295.3927 | 415.3476 | 41.094 | 41.094 | 415.3476 | 295.3927 | 315.6479 | 315.6479 |  | 247.1693
0 | [10 empty] | 0 | 295.3927 | 415.3476 | 41.094 | 41.094 | 415.3476 | 295.3927 | 315.6479 | 315.6479 |  | 247.1693
0 | [6 empty] | bellow+tube |  |  |  | 0.15525 | 295.4703 | 415.3476 | 41.094 | 41.094 | 415.3476 | 295.5479 | 315.7255 | 315.8031 |  | 247.2469
0 | LS2 | CC12 | SOL1 | 3161 | LS2_CC12:SOL1_D3161 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 295.7979 | 415.3476 | 41.094 | 41.094 | 415.3476 | 296.0479 | 316.0531 | 316.3031 |  | 247.5745
0 | [6 empty] | bellow+tube |  |  |  | 0.15525 | 296.1255 | 415.3476 | 41.094 | 41.094 | 415.3476 | 296.2032 | 316.3808 | 316.4584 |  | 247.9021
0 | LS2 | CC12 | CAV4 | 3164 | LS2_CC12:CAV4_D3164 |  | b29 resonator | 40 |  |  | 0.32364 | 296.365 | 415.3476 | 41.094 | 41.094 | 415.3476 | 296.5268 | 316.6202 | 316.782 |  | 248.1416
0 | [10 empty] | 0 | 296.5268 | 415.3476 | 41.094 | 41.094 | 415.3476 | 296.5268 | 316.782 | 316.782 |  | 248.3034
0 | [10 empty] | 0 | 296.5268 | 415.3476 | 41.094 | 41.094 | 415.3476 | 296.5268 | 316.782 | 316.782 |  | 248.3034
0 | [10 empty] | 0.099517 | 296.5766 | 415.3476 | 41.094 | 41.094 | 415.3476 | 296.6263 | 316.8318 | 316.8815 |  | 248.3532
0 | LS2 | CC12 | CAV5 | 3168 | LS2_CC12:CAV5_D3168 |  | b29 resonator | 40 |  |  | 0.32364 | 296.7882 | 415.3476 | 41.094 | 41.094 | 415.3476 | 296.95 | 317.0434 | 317.2052 |  | 248.5647
0 | [10 empty] | 0 | 296.95 | 415.3476 | 41.094 | 41.094 | 415.3476 | 296.95 | 317.2052 | 317.2052 |  | 248.7266
0 | [10 empty] | 0 | 296.95 | 415.3476 | 41.094 | 41.094 | 415.3476 | 296.95 | 317.2052 | 317.2052 |  | 248.7266
0 | [10 empty] | 0.099517 | 296.9997 | 415.3476 | 41.094 | 41.094 | 415.3476 | 297.0495 | 317.255 | 317.3047 |  | 248.7763
0 | LS2 | CC12 | CAV6 | 3173 | LS2_CC12:CAV6_D3173 |  | b29 resonator | 40 |  |  | 0.32364 | 297.2113 | 415.3476 | 41.094 | 41.094 | 415.3476 | 297.3731 | 317.4665 | 317.6284 |  | 248.9879
0 | LS2 | CC12 | BLM | 3175 | LS2_CC12:BLM_D3175 |  | outside loss monitor |  |  |  | 0 | 297.3731 | 415.3476 | 41.094 | 41.094 | 415.3476 | 297.3731 | 317.6284 | 317.6284 |  | 249.1497
0 | [10 empty] | 0 | 297.3731 | 415.3476 | 41.094 | 41.094 | 415.3476 | 297.3731 | 317.6284 | 317.6284 |  | 249.1497
0 | [6 empty] | bellows |  |  |  | 0.1677 | 297.457 | 415.3476 | 41.094 | 41.094 | 415.3476 | 297.5408 | 317.7122 | 317.7961 |  | 249.2336
0 | LS2 | CC12 | GV | 3178 | LS2_CC12:GV_D3178 |  | gate valve | 40 |  |  | 0.057404 | 297.5695 | 415.3476 | 41.094 | 41.094 | 415.3476 | 297.5982 | 317.8248 | 317.8535 |  | 249.3461
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 297.65 | 415.3476 | 41.094 | 41.094 | 415.3476 | 297.7017 | 317.9052 | 317.957 |  | 249.4266
0 | LS2 | WC12 | BPM | 3180 | LS2_WC12:BPM_D3180 | BPM2 | position monitor |  |  |  | 0 | 297.7017 | 415.3476 | 41.094 | 41.094 | 415.3476 | 297.7017 | 317.957 | 317.957 |  | 249.4783
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 297.7744 | 415.3476 | 41.094 | 41.094 | 415.3476 | 297.847 | 318.0296 | 318.1022 |  | 249.551
0 | LS2 | WC12 | PM | 3181 | LS2_WC12:PM_D3181 | PMx | profile monitor (not costed) | future PM |  |  | 0 | 297.847 | 415.3476 | 41.094 | 41.094 | 415.3476 | 297.847 | 318.1022 | 318.1022 |  | 249.6236
0 | [6 empty] | diagnostic box |  |  |  | 0 | 297.847 | 415.3476 | 41.094 | 41.094 | 415.3476 | 297.847 | 318.1022 | 318.1022 |  | 249.6236
0 | LS2 | WC12 | IP | 3181 | LS2_WC12:IP_D3181 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 297.847 | 415.3476 | 41.094 | 41.094 | 415.3476 | 297.847 | 318.1022 | 318.1022 |  | 249.6236
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 297.9126 | 415.3476 | 41.094 | 41.094 | 415.3476 | 297.9782 | 318.1679 | 318.2335 |  | 249.6892
0 | LS2 | CD01 | GV | 3183 | LS2_CD01:GV_D3183 |  | gate valve | 40 |  |  | 0.057404 | 298.0069 | 415.3476 | 41.094 | 41.094 | 415.3476 | 298.0356 | 318.2622 | 318.2909 |  | 249.7835
0 | [6 empty] | bellows |  |  |  | 0.1677 | 298.1195 | 415.3476 | 41.094 | 41.094 | 415.3476 | 298.2034 | 318.3747 | 318.4586 |  | 249.8961
0 | LS2 | CD01 | CAV1 | 3187 | LS2_CD01:CAV1_D3187 |  | b54 resonator | 40 |  |  | 0.49566 | 298.4512 | 415.3476 | 41.094 | 41.094 | 415.3476 | 298.699 | 318.7064 | 318.9542 |  | 250.2278
0 | LS2 | CD01 | BLM | 3190 | LS2_CD01:BLM_D3190 |  | outside loss monitor |  |  |  | 0 | 298.699 | 415.3476 | 41.094 | 41.094 | 415.3476 | 298.699 | 318.9542 | 318.9542 |  | 250.4756
0 | [10 empty] | 0 | 298.699 | 415.3476 | 41.094 | 41.094 | 415.3476 | 298.699 | 318.9542 | 318.9542 |  | 250.4756
0 | [6 empty] | bellow |  |  |  | 0.099517 | 298.7488 | 415.3476 | 41.094 | 41.094 | 415.3476 | 298.7985 | 319.004 | 319.0537 |  | 250.5254
0 | LS2 | CD01 | CAV2 | 3193 | LS2_CD01:CAV2_D3193 |  | b54 resonator | 40 |  |  | 0.49566 | 299.0464 | 415.3476 | 41.094 | 41.094 | 415.3476 | 299.2942 | 319.3016 | 319.5494 |  | 250.8229
0 | [10 empty] | 0 | 299.2942 | 415.3476 | 41.094 | 41.094 | 415.3476 | 299.2942 | 319.5494 | 319.5494 |  | 251.0708
0 | [10 empty] | 0 | 299.2942 | 415.3476 | 41.094 | 41.094 | 415.3476 | 299.2942 | 319.5494 | 319.5494 |  | 251.0708
0 | [6 empty] | bellow |  |  |  | 0.099517 | 299.3439 | 415.3476 | 41.094 | 41.094 | 415.3476 | 299.3937 | 319.5992 | 319.6489 |  | 251.1205
0 | LS2 | CD01 | CAV3 | 3199 | LS2_CD01:CAV3_D3199 |  | b54 resonator | 40 |  |  | 0.49566 | 299.6415 | 415.3476 | 41.094 | 41.094 | 415.3476 | 299.8894 | 319.8967 | 320.1446 |  | 251.4181
0 | [10 empty] | 0 | 299.8894 | 415.3476 | 41.094 | 41.094 | 415.3476 | 299.8894 | 320.1446 | 320.1446 |  | 251.6659
0 | [10 empty] | 0 | 299.8894 | 415.3476 | 41.094 | 41.094 | 415.3476 | 299.8894 | 320.1446 | 320.1446 |  | 251.6659
0 | [6 empty] | bellow |  |  |  | 0.099517 | 299.9391 | 415.3476 | 41.094 | 41.094 | 415.3476 | 299.9889 | 320.1943 | 320.2441 |  | 251.7157
0 | LS2 | CD01 | CAV4 | 3205 | LS2_CD01:CAV4_D3205 |  | b54 resonator | 40 |  |  | 0.49566 | 300.2367 | 415.3476 | 41.094 | 41.094 | 415.3476 | 300.4845 | 320.4919 | 320.7397 |  | 252.0133
0 | LS2 | CD01 | BLM | 3207 | LS2_CD01:BLM_D3207 |  | outside loss monitor |  |  |  | 0 | 300.4845 | 415.3476 | 41.094 | 41.094 | 415.3476 | 300.4845 | 320.7397 | 320.7397 |  | 252.2611
0 | [10 empty] | 0 | 300.4845 | 415.3476 | 41.094 | 41.094 | 415.3476 | 300.4845 | 320.7397 | 320.7397 |  | 252.2611
0 | [6 empty] | bellow |  |  |  | 0.15525 | 300.5622 | 415.3476 | 41.094 | 41.094 | 415.3476 | 300.6398 | 320.8174 | 320.895 |  | 252.3387
0 | LS2 | CD01 | SOL1 | 3211 | LS2_CD01:SOL1_D3211 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 300.8898 | 415.3476 | 41.094 | 41.094 | 415.3476 | 301.1398 | 321.145 | 321.395 |  | 252.6664
0 | [6 empty] | bellow |  |  |  | 0.15525 | 301.2174 | 415.3476 | 41.094 | 41.094 | 415.3476 | 301.295 | 321.4726 | 321.5502 |  | 252.994
0 | LS2 | CD01 | CAV5 | 3218 | LS2_CD01:CAV5_D3218 |  | b54 resonator | 40 |  |  | 0.49566 | 301.5429 | 415.3476 | 41.094 | 41.094 | 415.3476 | 301.7907 | 321.7981 | 322.0459 |  | 253.3194
0 | [10 empty] | 0 | 301.7907 | 415.3476 | 41.094 | 41.094 | 415.3476 | 301.7907 | 322.0459 | 322.0459 |  | 253.5673
0 | [10 empty] | 0 | 301.7907 | 415.3476 | 41.094 | 41.094 | 415.3476 | 301.7907 | 322.0459 | 322.0459 |  | 253.5673
0 | [6 empty] | bellow |  |  |  | 0.099517 | 301.8404 | 415.3476 | 41.094 | 41.094 | 415.3476 | 301.8902 | 322.0957 | 322.1454 |  | 253.617
0 | LS2 | CD01 | CAV6 | 3224 | LS2_CD01:CAV6_D3224 |  | b54 resonator | 40 |  |  | 0.49566 | 302.138 | 415.3476 | 41.094 | 41.094 | 415.3476 | 302.3859 | 322.3932 | 322.6411 |  | 253.9146
0 | [10 empty] | 0 | 302.3859 | 415.3476 | 41.094 | 41.094 | 415.3476 | 302.3859 | 322.6411 | 322.6411 |  | 254.1624
0 | [10 empty] | 0 | 302.3859 | 415.3476 | 41.094 | 41.094 | 415.3476 | 302.3859 | 322.6411 | 322.6411 |  | 254.1624
0 | [6 empty] | bellow |  |  |  | 0.099517 | 302.4356 | 415.3476 | 41.094 | 41.094 | 415.3476 | 302.4854 | 322.6908 | 322.7406 |  | 254.2122
0 | LS2 | CD01 | CAV7 | 3230 | LS2_CD01:CAV7_D3230 |  | b54 resonator | 40 |  |  | 0.49566 | 302.7332 | 415.3476 | 41.094 | 41.094 | 415.3476 | 302.981 | 322.9884 | 323.2362 |  | 254.5098
0 | [10 empty] | 0 | 302.981 | 415.3476 | 41.094 | 41.094 | 415.3476 | 302.981 | 323.2362 | 323.2362 |  | 254.7576
0 | [10 empty] | 0 | 302.981 | 415.3476 | 41.094 | 41.094 | 415.3476 | 302.981 | 323.2362 | 323.2362 |  | 254.7576
0 | [6 empty] | bellow |  |  |  | 0.099517 | 303.0308 | 415.3476 | 41.094 | 41.094 | 415.3476 | 303.0805 | 323.286 | 323.3358 |  | 254.8074
0 | LS2 | CD01 | CAV8 | 3236 | LS2_CD01:CAV8_D3236 |  | b54 resonator | 40 |  |  | 0.49566 | 303.3284 | 415.3476 | 41.094 | 41.094 | 415.3476 | 303.5762 | 323.5836 | 323.8314 |  | 255.105
0 | LS2 | CD01 | BLM | 3238 | LS2_CD01:BLM_D3238 |  | outside loss monitor |  |  |  | 0 | 303.5762 | 415.3476 | 41.094 | 41.094 | 415.3476 | 303.5762 | 323.8314 | 323.8314 |  | 255.3528
0 | [10 empty] | 0 | 303.5762 | 415.3476 | 41.094 | 41.094 | 415.3476 | 303.5762 | 323.8314 | 323.8314 |  | 255.3528
0 | [10 empty] | 0.1677 | 303.6601 | 415.3476 | 41.094 | 41.094 | 415.3476 | 303.7439 | 323.9153 | 323.9991 |  | 255.4366
0 | LS2 | CD01 | GV | 3240 | LS2_CD01:GV_D3240 |  | gate valve | 40 |  |  | 0.057404 | 303.7726 | 415.3476 | 41.094 | 41.094 | 415.3476 | 303.8013 | 324.0278 | 324.0565 |  | 255.5492
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 303.8531 | 415.3476 | 41.094 | 41.094 | 415.3476 | 303.9048 | 324.1083 | 324.16 |  | 255.6296
0 | LS2 | WD01 | BPM | 3242 | LS2_WD01:BPM_D3242 | BPM2 | position monitor |  |  |  | 0 | 303.9048 | 415.3476 | 41.094 | 41.094 | 415.3476 | 303.9048 | 324.16 | 324.16 |  | 255.6814
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 303.9775 | 415.3476 | 41.094 | 41.094 | 415.3476 | 304.0501 | 324.2327 | 324.3053 |  | 255.754
0 | LS2 | WD01 | PM | 3243 | LS2_WD01:PM_D3243 | PMx | profile monitor (not costed) | future PM |  |  | 0 | 304.0501 | 415.3476 | 41.094 | 41.094 | 415.3476 | 304.0501 | 324.3053 | 324.3053 |  | 255.8267
0 | [6 empty] | diagnostic box |  |  |  | 0 | 304.0501 | 415.3476 | 41.094 | 41.094 | 415.3476 | 304.0501 | 324.3053 | 324.3053 |  | 255.8267
0 | LS2 | WD01 | IP | 3243 | LS2_WD01:IP_D3243 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 304.0501 | 415.3476 | 41.094 | 41.094 | 415.3476 | 304.0501 | 324.3053 | 324.3053 |  | 255.8267
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 304.1157 | 415.3476 | 41.094 | 41.094 | 415.3476 | 304.1813 | 324.3709 | 324.4365 |  | 255.8923
0 | LS2 | CD02 | GV | 3245 | LS2_CD02:GV_D3245 |  | gate valve | 40 |  |  | 0.057404 | 304.21 | 415.3476 | 41.094 | 41.094 | 415.3476 | 304.2387 | 324.4652 | 324.4939 |  | 255.9866
0 | [6 empty] | bellows |  |  |  | 0.1677 | 304.3226 | 415.3476 | 41.094 | 41.094 | 415.3476 | 304.4064 | 324.5778 | 324.6616 |  | 256.0992
0 | LS2 | CD02 | CAV1 | 3249 | LS2_CD02:CAV1_D3249 |  | b54 resonator | 40 |  |  | 0.49566 | 304.6542 | 415.3476 | 41.094 | 41.094 | 415.3476 | 304.9021 | 324.9095 | 325.1573 |  | 256.4308
0 | LS2 | CD02 | BLM | 3252 | LS2_CD02:BLM_D3252 |  | outside loss monitor |  |  |  | 0 | 304.9021 | 415.3476 | 41.094 | 41.094 | 415.3476 | 304.9021 | 325.1573 | 325.1573 |  | 256.6787
0 | [10 empty] | 0 | 304.9021 | 415.3476 | 41.094 | 41.094 | 415.3476 | 304.9021 | 325.1573 | 325.1573 |  | 256.6787
0 | [6 empty] | bellow |  |  |  | 0.099517 | 304.9518 | 415.3476 | 41.094 | 41.094 | 415.3476 | 305.0016 | 325.207 | 325.2568 |  | 256.7284
0 | LS2 | CD02 | CAV2 | 3255 | LS2_CD02:CAV2_D3255 |  | b54 resonator | 40 |  |  | 0.49566 | 305.2494 | 415.3476 | 41.094 | 41.094 | 415.3476 | 305.4972 | 325.5046 | 325.7525 |  | 257.026
0 | [10 empty] | 0 | 305.4972 | 415.3476 | 41.094 | 41.094 | 415.3476 | 305.4972 | 325.7525 | 325.7525 |  | 257.2738
0 | [10 empty] | 0 | 305.4972 | 415.3476 | 41.094 | 41.094 | 415.3476 | 305.4972 | 325.7525 | 325.7525 |  | 257.2738
0 | [6 empty] | bellow |  |  |  | 0.099517 | 305.547 | 415.3476 | 41.094 | 41.094 | 415.3476 | 305.5968 | 325.8022 | 325.852 |  | 257.3236
0 | LS2 | CD02 | CAV3 | 3261 | LS2_CD02:CAV3_D3261 |  | b54 resonator | 40 |  |  | 0.49566 | 305.8446 | 415.3476 | 41.094 | 41.094 | 415.3476 | 306.0924 | 326.0998 | 326.3476 |  | 257.6212
0 | [10 empty] | 0 | 306.0924 | 415.3476 | 41.094 | 41.094 | 415.3476 | 306.0924 | 326.3476 | 326.3476 |  | 257.869
0 | [10 empty] | 0 | 306.0924 | 415.3476 | 41.094 | 41.094 | 415.3476 | 306.0924 | 326.3476 | 326.3476 |  | 257.869
0 | [6 empty] | bellow |  |  |  | 0.099517 | 306.1422 | 415.3476 | 41.094 | 41.094 | 415.3476 | 306.1919 | 326.3974 | 326.4472 |  | 257.9188
0 | LS2 | CD02 | CAV4 | 3267 | LS2_CD02:CAV4_D3267 |  | b54 resonator | 40 |  |  | 0.49566 | 306.4398 | 415.3476 | 41.094 | 41.094 | 415.3476 | 306.6876 | 326.695 | 326.9428 |  | 258.2164
0 | LS2 | CD02 | BLM | 3269 | LS2_CD02:BLM_D3269 |  | outside loss monitor |  |  |  | 0 | 306.6876 | 415.3476 | 41.094 | 41.094 | 415.3476 | 306.6876 | 326.9428 | 326.9428 |  | 258.4642
0 | [10 empty] | 0 | 306.6876 | 415.3476 | 41.094 | 41.094 | 415.3476 | 306.6876 | 326.9428 | 326.9428 |  | 258.4642
0 | [6 empty] | bellow |  |  |  | 0.15525 | 306.7652 | 415.3476 | 41.094 | 41.094 | 415.3476 | 306.8428 | 327.0204 | 327.0981 |  | 258.5418
0 | LS2 | CD02 | SOL1 | 3273 | LS2_CD02:SOL1_D3273 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 307.0928 | 415.3476 | 41.094 | 41.094 | 415.3476 | 307.3428 | 327.3481 | 327.5981 |  | 258.8694
0 | [6 empty] | bellow |  |  |  | 0.15525 | 307.4205 | 415.3476 | 41.094 | 41.094 | 415.3476 | 307.4981 | 327.6757 | 327.7533 |  | 259.1971
0 | LS2 | CD02 | CAV5 | 3280 | LS2_CD02:CAV5_D3280 |  | b54 resonator | 40 |  |  | 0.49566 | 307.7459 | 415.3476 | 41.094 | 41.094 | 415.3476 | 307.9937 | 328.0011 | 328.249 |  | 259.5225
0 | [10 empty] | 0 | 307.9937 | 415.3476 | 41.094 | 41.094 | 415.3476 | 307.9937 | 328.249 | 328.249 |  | 259.7703
0 | [10 empty] | 0 | 307.9937 | 415.3476 | 41.094 | 41.094 | 415.3476 | 307.9937 | 328.249 | 328.249 |  | 259.7703
0 | [6 empty] | bellow |  |  |  | 0.099517 | 308.0435 | 415.3476 | 41.094 | 41.094 | 415.3476 | 308.0933 | 328.2987 | 328.3485 |  | 259.8201
0 | LS2 | CD02 | CAV6 | 3286 | LS2_CD02:CAV6_D3286 |  | b54 resonator | 40 |  |  | 0.49566 | 308.3411 | 415.3476 | 41.094 | 41.094 | 415.3476 | 308.5889 | 328.5963 | 328.8441 |  | 260.1177
0 | [10 empty] | 0 | 308.5889 | 415.3476 | 41.094 | 41.094 | 415.3476 | 308.5889 | 328.8441 | 328.8441 |  | 260.3655
0 | [10 empty] | 0 | 308.5889 | 415.3476 | 41.094 | 41.094 | 415.3476 | 308.5889 | 328.8441 | 328.8441 |  | 260.3655
0 | [6 empty] | bellow |  |  |  | 0.099517 | 308.6387 | 415.3476 | 41.094 | 41.094 | 415.3476 | 308.6884 | 328.8939 | 328.9437 |  | 260.4153
0 | LS2 | CD02 | CAV7 | 3292 | LS2_CD02:CAV7_D3292 |  | b54 resonator | 40 |  |  | 0.49566 | 308.9363 | 415.3476 | 41.094 | 41.094 | 415.3476 | 309.1841 | 329.1915 | 329.4393 |  | 260.7129
0 | [10 empty] | 0 | 309.1841 | 415.3476 | 41.094 | 41.094 | 415.3476 | 309.1841 | 329.4393 | 329.4393 |  | 260.9607
0 | [10 empty] | 0 | 309.1841 | 415.3476 | 41.094 | 41.094 | 415.3476 | 309.1841 | 329.4393 | 329.4393 |  | 260.9607
0 | [6 empty] | bellow |  |  |  | 0.099517 | 309.2339 | 415.3476 | 41.094 | 41.094 | 415.3476 | 309.2836 | 329.4891 | 329.5388 |  | 261.0104
0 | LS2 | CD02 | CAV8 | 3298 | LS2_CD02:CAV8_D3298 |  | b54 resonator | 40 |  |  | 0.49566 | 309.5314 | 415.3476 | 41.094 | 41.094 | 415.3476 | 309.7793 | 329.7867 | 330.0345 |  | 261.308
0 | LS2 | CD02 | BLM | 3300 | LS2_CD02:BLM_D3300 |  | outside loss monitor |  |  |  | 0 | 309.7793 | 415.3476 | 41.094 | 41.094 | 415.3476 | 309.7793 | 330.0345 | 330.0345 |  | 261.5559
0 | [10 empty] | 0 | 309.7793 | 415.3476 | 41.094 | 41.094 | 415.3476 | 309.7793 | 330.0345 | 330.0345 |  | 261.5559
0 | [10 empty] | 0.1677 | 309.8631 | 415.3476 | 41.094 | 41.094 | 415.3476 | 309.947 | 330.1183 | 330.2022 |  | 261.6397
0 | LS2 | CD02 | GV | 3302 | LS2_CD02:GV_D3302 |  | gate valve | 40 |  |  | 0.057404 | 309.9757 | 415.3476 | 41.094 | 41.094 | 415.3476 | 310.0044 | 330.2309 | 330.2596 |  | 261.7523
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 310.0561 | 415.3476 | 41.094 | 41.094 | 415.3476 | 310.1079 | 330.3113 | 330.3631 |  | 261.8327
0 | LS2 | WD02 | BPM | 3304 | LS2_WD02:BPM_D3304 | BPM2 | position monitor |  |  |  | 0 | 310.1079 | 415.3476 | 41.094 | 41.094 | 415.3476 | 310.1079 | 330.3631 | 330.3631 |  | 261.8845
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 310.1805 | 415.3476 | 41.094 | 41.094 | 415.3476 | 310.2532 | 330.4357 | 330.5084 |  | 261.9571
0 | LS2 | WD02 | PM | 3305 | LS2_WD02:PM_D3305 | PMx | profile monitor (not costed) | future PM |  |  | 0 | 310.2532 | 415.3476 | 41.094 | 41.094 | 415.3476 | 310.2532 | 330.5084 | 330.5084 |  | 262.0297
0 | [6 empty] | diagnostic box |  |  |  | 0 | 310.2532 | 415.3476 | 41.094 | 41.094 | 415.3476 | 310.2532 | 330.5084 | 330.5084 |  | 262.0297
0 | LS2 | WD02 | IP | 3305 | LS2_WD02:IP_D3305 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 310.2532 | 415.3476 | 41.094 | 41.094 | 415.3476 | 310.2532 | 330.5084 | 330.5084 |  | 262.0297
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 310.3188 | 415.3476 | 41.094 | 41.094 | 415.3476 | 310.3844 | 330.574 | 330.6396 |  | 262.0954
0 | LS2 | CD03 | GV | 3307 | LS2_CD03:GV_D3307 |  | gate valve | 40 |  |  | 0.057404 | 310.4131 | 415.3476 | 41.094 | 41.094 | 415.3476 | 310.4418 | 330.6683 | 330.697 |  | 262.1897
0 | [6 empty] | bellows |  |  |  | 0.1677 | 310.5256 | 415.3476 | 41.094 | 41.094 | 415.3476 | 310.6095 | 330.7808 | 330.8647 |  | 262.3022
0 | LS2 | CD03 | CAV1 | 3311 | LS2_CD03:CAV1_D3311 |  | b54 resonator | 40 |  |  | 0.49566 | 310.8573 | 415.3476 | 41.094 | 41.094 | 415.3476 | 311.1051 | 331.1125 | 331.3604 |  | 262.6339
0 | LS2 | CD03 | BLM | 3314 | LS2_CD03:BLM_D3314 |  | outside loss monitor |  |  |  | 0 | 311.1051 | 415.3476 | 41.094 | 41.094 | 415.3476 | 311.1051 | 331.3604 | 331.3604 |  | 262.8817
0 | [10 empty] | 0 | 311.1051 | 415.3476 | 41.094 | 41.094 | 415.3476 | 311.1051 | 331.3604 | 331.3604 |  | 262.8817
0 | [6 empty] | bellow |  |  |  | 0.099517 | 311.1549 | 415.3476 | 41.094 | 41.094 | 415.3476 | 311.2047 | 331.4101 | 331.4599 |  | 262.9315
0 | LS2 | CD03 | CAV2 | 3317 | LS2_CD03:CAV2_D3317 |  | b54 resonator | 40 |  |  | 0.49566 | 311.4525 | 415.3476 | 41.094 | 41.094 | 415.3476 | 311.7003 | 331.7077 | 331.9555 |  | 263.2291
0 | [10 empty] | 0 | 311.7003 | 415.3476 | 41.094 | 41.094 | 415.3476 | 311.7003 | 331.9555 | 331.9555 |  | 263.4769
0 | [10 empty] | 0 | 311.7003 | 415.3476 | 41.094 | 41.094 | 415.3476 | 311.7003 | 331.9555 | 331.9555 |  | 263.4769
0 | [6 empty] | bellow |  |  |  | 0.099517 | 311.7501 | 415.3476 | 41.094 | 41.094 | 415.3476 | 311.7998 | 332.0053 | 332.055 |  | 263.5267
0 | LS2 | CD03 | CAV3 | 3323 | LS2_CD03:CAV3_D3323 |  | b54 resonator | 40 |  |  | 0.49566 | 312.0477 | 415.3476 | 41.094 | 41.094 | 415.3476 | 312.2955 | 332.3029 | 332.5507 |  | 263.8242
0 | [10 empty] | 0 | 312.2955 | 415.3476 | 41.094 | 41.094 | 415.3476 | 312.2955 | 332.5507 | 332.5507 |  | 264.0721
0 | [10 empty] | 0 | 312.2955 | 415.3476 | 41.094 | 41.094 | 415.3476 | 312.2955 | 332.5507 | 332.5507 |  | 264.0721
0 | [6 empty] | bellow |  |  |  | 0.099517 | 312.3452 | 415.3476 | 41.094 | 41.094 | 415.3476 | 312.395 | 332.6005 | 332.6502 |  | 264.1218
0 | LS2 | CD03 | CAV4 | 3329 | LS2_CD03:CAV4_D3329 |  | b54 resonator | 40 |  |  | 0.49566 | 312.6428 | 415.3476 | 41.094 | 41.094 | 415.3476 | 312.8907 | 332.898 | 333.1459 |  | 264.4194
0 | LS2 | CD03 | BLM | 3331 | LS2_CD03:BLM_D3331 |  | outside loss monitor |  |  |  | 0 | 312.8907 | 415.3476 | 41.094 | 41.094 | 415.3476 | 312.8907 | 333.1459 | 333.1459 |  | 264.6672
0 | [10 empty] | 0 | 312.8907 | 415.3476 | 41.094 | 41.094 | 415.3476 | 312.8907 | 333.1459 | 333.1459 |  | 264.6672
0 | [6 empty] | bellow |  |  |  | 0.15525 | 312.9683 | 415.3476 | 41.094 | 41.094 | 415.3476 | 313.0459 | 333.2235 | 333.3011 |  | 264.7449
0 | LS2 | CD03 | SOL1 | 3336 | LS2_CD03:SOL1_D3336 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 313.2959 | 415.3476 | 41.094 | 41.094 | 415.3476 | 313.5459 | 333.5511 | 333.8011 |  | 265.0725
0 | [6 empty] | bellow |  |  |  | 0.15525 | 313.6235 | 415.3476 | 41.094 | 41.094 | 415.3476 | 313.7012 | 333.8788 | 333.9564 |  | 265.4001
0 | LS2 | CD03 | CAV5 | 3342 | LS2_CD03:CAV5_D3342 |  | b54 resonator | 40 |  |  | 0.49566 | 313.949 | 415.3476 | 41.094 | 41.094 | 415.3476 | 314.1968 | 334.2042 | 334.452 |  | 265.7256
0 | [10 empty] | 0 | 314.1968 | 415.3476 | 41.094 | 41.094 | 415.3476 | 314.1968 | 334.452 | 334.452 |  | 265.9734
0 | [10 empty] | 0 | 314.1968 | 415.3476 | 41.094 | 41.094 | 415.3476 | 314.1968 | 334.452 | 334.452 |  | 265.9734
0 | [6 empty] | bellow |  |  |  | 0.099517 | 314.2466 | 415.3476 | 41.094 | 41.094 | 415.3476 | 314.2963 | 334.5018 | 334.5515 |  | 266.0232
0 | LS2 | CD03 | CAV6 | 3348 | LS2_CD03:CAV6_D3348 |  | b54 resonator | 40 |  |  | 0.49566 | 314.5442 | 415.3476 | 41.094 | 41.094 | 415.3476 | 314.792 | 334.7994 | 335.0472 |  | 266.3207
0 | [10 empty] | 0 | 314.792 | 415.3476 | 41.094 | 41.094 | 415.3476 | 314.792 | 335.0472 | 335.0472 |  | 266.5686
0 | [10 empty] | 0 | 314.792 | 415.3476 | 41.094 | 41.094 | 415.3476 | 314.792 | 335.0472 | 335.0472 |  | 266.5686
0 | [6 empty] | bellow |  |  |  | 0.099517 | 314.8417 | 415.3476 | 41.094 | 41.094 | 415.3476 | 314.8915 | 335.097 | 335.1467 |  | 266.6183
0 | LS2 | CD03 | CAV7 | 3354 | LS2_CD03:CAV7_D3354 |  | b54 resonator | 40 |  |  | 0.49566 | 315.1393 | 415.3476 | 41.094 | 41.094 | 415.3476 | 315.3872 | 335.3945 | 335.6424 |  | 266.9159
0 | [10 empty] | 0 | 315.3872 | 415.3476 | 41.094 | 41.094 | 415.3476 | 315.3872 | 335.6424 | 335.6424 |  | 267.1637
0 | [10 empty] | 0 | 315.3872 | 415.3476 | 41.094 | 41.094 | 415.3476 | 315.3872 | 335.6424 | 335.6424 |  | 267.1637
0 | [6 empty] | bellow |  |  |  | 0.099517 | 315.4369 | 415.3476 | 41.094 | 41.094 | 415.3476 | 315.4867 | 335.6921 | 335.7419 |  | 267.2135
0 | LS2 | CD03 | CAV8 | 3360 | LS2_CD03:CAV8_D3360 |  | b54 resonator | 40 |  |  | 0.49566 | 315.7345 | 415.3476 | 41.094 | 41.094 | 415.3476 | 315.9823 | 335.9897 | 336.2376 |  | 267.5111
0 | LS2 | CD03 | BLM | 3362 | LS2_CD03:BLM_D3362 |  | outside loss monitor |  |  |  | 0 | 315.9823 | 415.3476 | 41.094 | 41.094 | 415.3476 | 315.9823 | 336.2376 | 336.2376 |  | 267.7589
0 | [10 empty] | 0 | 315.9823 | 415.3476 | 41.094 | 41.094 | 415.3476 | 315.9823 | 336.2376 | 336.2376 |  | 267.7589
0 | [10 empty] | 0.1677 | 316.0662 | 415.3476 | 41.094 | 41.094 | 415.3476 | 316.15 | 336.3214 | 336.4053 |  | 267.8428
0 | LS2 | CD03 | GV | 3364 | LS2_CD03:GV_D3364 |  | gate valve | 40 |  |  | 0.057404 | 316.1787 | 415.3476 | 41.094 | 41.094 | 415.3476 | 316.2074 | 336.434 | 336.4627 |  | 267.9553
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 316.2592 | 415.3476 | 41.094 | 41.094 | 415.3476 | 316.3109 | 336.5144 | 336.5662 |  | 268.0358
0 | LS2 | WD03 | BPM | 3366 | LS2_WD03:BPM_D3366 | BPM2 | position monitor |  |  |  | 0 | 316.3109 | 415.3476 | 41.094 | 41.094 | 415.3476 | 316.3109 | 336.5662 | 336.5662 |  | 268.0875
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 316.3836 | 415.3476 | 41.094 | 41.094 | 415.3476 | 316.4562 | 336.6388 | 336.7114 |  | 268.1602
0 | LS2 | WD03 | PM | 3367 | LS2_WD03:PM_D3367 | PMx | profile monitor (not costed) | future PM |  |  | 0 | 316.4562 | 415.3476 | 41.094 | 41.094 | 415.3476 | 316.4562 | 336.7114 | 336.7114 |  | 268.2328
0 | [6 empty] | diagnostic box |  |  |  | 0 | 316.4562 | 415.3476 | 41.094 | 41.094 | 415.3476 | 316.4562 | 336.7114 | 336.7114 |  | 268.2328
0 | LS2 | WD03 | IP | 3367 | LS2_WD03:IP_D3367 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 316.4562 | 415.3476 | 41.094 | 41.094 | 415.3476 | 316.4562 | 336.7114 | 336.7114 |  | 268.2328
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 316.5218 | 415.3476 | 41.094 | 41.094 | 415.3476 | 316.5874 | 336.7771 | 336.8427 |  | 268.2984
0 | LS2 | CD04 | GV | 3369 | LS2_CD04:GV_D3369 |  | gate valve | 40 |  |  | 0.057404 | 316.6161 | 415.3476 | 41.094 | 41.094 | 415.3476 | 316.6448 | 336.8714 | 336.9001 |  | 268.3927
0 | [6 empty] | bellows |  |  |  | 0.1677 | 316.7287 | 415.3476 | 41.094 | 41.094 | 415.3476 | 316.8126 | 336.9839 | 337.0678 |  | 268.5053
0 | LS2 | CD04 | CAV1 | 3373 | LS2_CD04:CAV1_D3373 |  | b54 resonator | 40 |  |  | 0.49566 | 317.0604 | 415.3476 | 41.094 | 41.094 | 415.3476 | 317.3082 | 337.3156 | 337.5634 |  | 268.837
0 | LS2 | CD04 | BLM | 3376 | LS2_CD04:BLM_D3376 |  | outside loss monitor |  |  |  | 0 | 317.3082 | 415.3476 | 41.094 | 41.094 | 415.3476 | 317.3082 | 337.5634 | 337.5634 |  | 269.0848
0 | [10 empty] | 0 | 317.3082 | 415.3476 | 41.094 | 41.094 | 415.3476 | 317.3082 | 337.5634 | 337.5634 |  | 269.0848
0 | [6 empty] | bellow |  |  |  | 0.099517 | 317.358 | 415.3476 | 41.094 | 41.094 | 415.3476 | 317.4077 | 337.6132 | 337.6629 |  | 269.1346
0 | LS2 | CD04 | CAV2 | 3379 | LS2_CD04:CAV2_D3379 |  | b54 resonator | 40 |  |  | 0.49566 | 317.6556 | 415.3476 | 41.094 | 41.094 | 415.3476 | 317.9034 | 337.9108 | 338.1586 |  | 269.4321
0 | [10 empty] | 0 | 317.9034 | 415.3476 | 41.094 | 41.094 | 415.3476 | 317.9034 | 338.1586 | 338.1586 |  | 269.68
0 | [10 empty] | 0 | 317.9034 | 415.3476 | 41.094 | 41.094 | 415.3476 | 317.9034 | 338.1586 | 338.1586 |  | 269.68
0 | [6 empty] | bellow |  |  |  | 0.099517 | 317.9531 | 415.3476 | 41.094 | 41.094 | 415.3476 | 318.0029 | 338.2084 | 338.2581 |  | 269.7297
0 | LS2 | CD04 | CAV3 | 3385 | LS2_CD04:CAV3_D3385 |  | b54 resonator | 40 |  |  | 0.49566 | 318.2507 | 415.3476 | 41.094 | 41.094 | 415.3476 | 318.4986 | 338.5059 | 338.7538 |  | 270.0273
0 | [10 empty] | 0 | 318.4986 | 415.3476 | 41.094 | 41.094 | 415.3476 | 318.4986 | 338.7538 | 338.7538 |  | 270.2751
0 | [10 empty] | 0 | 318.4986 | 415.3476 | 41.094 | 41.094 | 415.3476 | 318.4986 | 338.7538 | 338.7538 |  | 270.2751
0 | [6 empty] | bellow |  |  |  | 0.099517 | 318.5483 | 415.3476 | 41.094 | 41.094 | 415.3476 | 318.5981 | 338.8035 | 338.8533 |  | 270.3249
0 | LS2 | CD04 | CAV4 | 3391 | LS2_CD04:CAV4_D3391 |  | b54 resonator | 40 |  |  | 0.49566 | 318.8459 | 415.3476 | 41.094 | 41.094 | 415.3476 | 319.0937 | 339.1011 | 339.3489 |  | 270.6225
0 | LS2 | CD04 | BLM | 3393 | LS2_CD04:BLM_D3393 |  | outside loss monitor |  |  |  | 0 | 319.0937 | 415.3476 | 41.094 | 41.094 | 415.3476 | 319.0937 | 339.3489 | 339.3489 |  | 270.8703
0 | [10 empty] | 0 | 319.0937 | 415.3476 | 41.094 | 41.094 | 415.3476 | 319.0937 | 339.3489 | 339.3489 |  | 270.8703
0 | [6 empty] | bellow |  |  |  | 0.15525 | 319.1714 | 415.3476 | 41.094 | 41.094 | 415.3476 | 319.249 | 339.4266 | 339.5042 |  | 270.9479
0 | LS2 | CD04 | SOL1 | 3398 | LS2_CD04:SOL1_D3398 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 319.499 | 415.3476 | 41.094 | 41.094 | 415.3476 | 319.749 | 339.7542 | 340.0042 |  | 271.2756
0 | [6 empty] | bellow |  |  |  | 0.15525 | 319.8266 | 415.3476 | 41.094 | 41.094 | 415.3476 | 319.9042 | 340.0818 | 340.1594 |  | 271.6032
0 | LS2 | CD04 | CAV5 | 3404 | LS2_CD04:CAV5_D3404 |  | b54 resonator | 40 |  |  | 0.49566 | 320.1521 | 415.3476 | 41.094 | 41.094 | 415.3476 | 320.3999 | 340.4073 | 340.6551 |  | 271.9286
0 | [10 empty] | 0 | 320.3999 | 415.3476 | 41.094 | 41.094 | 415.3476 | 320.3999 | 340.6551 | 340.6551 |  | 272.1765
0 | [10 empty] | 0 | 320.3999 | 415.3476 | 41.094 | 41.094 | 415.3476 | 320.3999 | 340.6551 | 340.6551 |  | 272.1765
0 | [6 empty] | bellow |  |  |  | 0.099517 | 320.4496 | 415.3476 | 41.094 | 41.094 | 415.3476 | 320.4994 | 340.7049 | 340.7546 |  | 272.2262
0 | LS2 | CD04 | CAV6 | 3410 | LS2_CD04:CAV6_D3410 |  | b54 resonator | 40 |  |  | 0.49566 | 320.7472 | 415.3476 | 41.094 | 41.094 | 415.3476 | 320.9951 | 341.0024 | 341.2503 |  | 272.5238
0 | [10 empty] | 0 | 320.9951 | 415.3476 | 41.094 | 41.094 | 415.3476 | 320.9951 | 341.2503 | 341.2503 |  | 272.7716
0 | [10 empty] | 0 | 320.9951 | 415.3476 | 41.094 | 41.094 | 415.3476 | 320.9951 | 341.2503 | 341.2503 |  | 272.7716
0 | [6 empty] | bellow |  |  |  | 0.099517 | 321.0448 | 415.3476 | 41.094 | 41.094 | 415.3476 | 321.0946 | 341.3 | 341.3498 |  | 272.8214
0 | LS2 | CD04 | CAV7 | 3416 | LS2_CD04:CAV7_D3416 |  | b54 resonator | 40 |  |  | 0.49566 | 321.3424 | 415.3476 | 41.094 | 41.094 | 415.3476 | 321.5902 | 341.5976 | 341.8454 |  | 273.119
0 | [10 empty] | 0 | 321.5902 | 415.3476 | 41.094 | 41.094 | 415.3476 | 321.5902 | 341.8454 | 341.8454 |  | 273.3668
0 | [10 empty] | 0 | 321.5902 | 415.3476 | 41.094 | 41.094 | 415.3476 | 321.5902 | 341.8454 | 341.8454 |  | 273.3668
0 | [6 empty] | bellow |  |  |  | 0.099517 | 321.64 | 415.3476 | 41.094 | 41.094 | 415.3476 | 321.6897 | 341.8952 | 341.945 |  | 273.4166
0 | LS2 | CD04 | CAV8 | 3422 | LS2_CD04:CAV8_D3422 |  | b54 resonator | 40 |  |  | 0.49566 | 321.9376 | 415.3476 | 41.094 | 41.094 | 415.3476 | 322.1854 | 342.1928 | 342.4406 |  | 273.7142
0 | LS2 | CD04 | BLM | 3424 | LS2_CD04:BLM_D3424 |  | outside loss monitor |  |  |  | 0 | 322.1854 | 415.3476 | 41.094 | 41.094 | 415.3476 | 322.1854 | 342.4406 | 342.4406 |  | 273.962
0 | [10 empty] | 0 | 322.1854 | 415.3476 | 41.094 | 41.094 | 415.3476 | 322.1854 | 342.4406 | 342.4406 |  | 273.962
0 | [10 empty] | 0.1677 | 322.2693 | 415.3476 | 41.094 | 41.094 | 415.3476 | 322.3531 | 342.5245 | 342.6083 |  | 274.0458
0 | LS2 | CD04 | GV | 3426 | LS2_CD04:GV_D3426 |  | gate valve | 40 |  |  | 0.057404 | 322.3818 | 415.3476 | 41.094 | 41.094 | 415.3476 | 322.4105 | 342.637 | 342.6657 |  | 274.1584
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 322.4623 | 415.3476 | 41.094 | 41.094 | 415.3476 | 322.514 | 342.7175 | 342.7692 |  | 274.2388
0 | LS2 | WD04 | BPM | 3428 | LS2_WD04:BPM_D3428 | BPM2 | position monitor |  |  |  | 0 | 322.514 | 415.3476 | 41.094 | 41.094 | 415.3476 | 322.514 | 342.7692 | 342.7692 |  | 274.2906
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 322.5867 | 415.3476 | 41.094 | 41.094 | 415.3476 | 322.6593 | 342.8419 | 342.9145 |  | 274.3632
0 | LS2 | WD04 | PORT | 3429 | LS2_WD04:PORT_D3429 |  | port for future diagnostics |  |  |  | 0 | 322.6593 | 415.3476 | 41.094 | 41.094 | 415.3476 | 322.6593 | 342.9145 | 342.9145 |  | 274.4359
0 | [6 empty] | diagnostic box |  |  |  | 0 | 322.6593 | 415.3476 | 41.094 | 41.094 | 415.3476 | 322.6593 | 342.9145 | 342.9145 |  | 274.4359
0 | LS2 | WD04 | IP | 3429 | LS2_WD04:IP_D3429 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 322.6593 | 415.3476 | 41.094 | 41.094 | 415.3476 | 322.6593 | 342.9145 | 342.9145 |  | 274.4359
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 322.7249 | 415.3476 | 41.094 | 41.094 | 415.3476 | 322.7905 | 342.9801 | 343.0457 |  | 274.5015
0 | LS2 | CD05 | GV | 3431 | LS2_CD05:GV_D3431 |  | gate valve | 40 |  |  | 0.057404 | 322.8192 | 415.3476 | 41.094 | 41.094 | 415.3476 | 322.8479 | 343.0744 | 343.1031 |  | 274.5958
0 | [6 empty] | bellows |  |  |  | 0.1677 | 322.9318 | 415.3476 | 41.094 | 41.094 | 415.3476 | 323.0156 | 343.187 | 343.2708 |  | 274.7084
0 | LS2 | CD05 | CAV1 | 3435 | LS2_CD05:CAV1_D3435 |  | b54 resonator | 40 |  |  | 0.49566 | 323.2634 | 415.3476 | 41.094 | 41.094 | 415.3476 | 323.5113 | 343.5187 | 343.7665 |  | 275.04
0 | LS2 | CD05 | BLM | 3438 | LS2_CD05:BLM_D3438 |  | outside loss monitor |  |  |  | 0 | 323.5113 | 415.3476 | 41.094 | 41.094 | 415.3476 | 323.5113 | 343.7665 | 343.7665 |  | 275.2879
0 | [10 empty] | 0 | 323.5113 | 415.3476 | 41.094 | 41.094 | 415.3476 | 323.5113 | 343.7665 | 343.7665 |  | 275.2879
0 | [6 empty] | bellow |  |  |  | 0.099517 | 323.561 | 415.3476 | 41.094 | 41.094 | 415.3476 | 323.6108 | 343.8162 | 343.866 |  | 275.3376
0 | LS2 | CD05 | CAV2 | 3441 | LS2_CD05:CAV2_D3441 |  | b54 resonator | 40 |  |  | 0.49566 | 323.8586 | 415.3476 | 41.094 | 41.094 | 415.3476 | 324.1064 | 344.1138 | 344.3617 |  | 275.6352
0 | [10 empty] | 0 | 324.1064 | 415.3476 | 41.094 | 41.094 | 415.3476 | 324.1064 | 344.3617 | 344.3617 |  | 275.883
0 | [10 empty] | 0 | 324.1064 | 415.3476 | 41.094 | 41.094 | 415.3476 | 324.1064 | 344.3617 | 344.3617 |  | 275.883
0 | [6 empty] | bellow |  |  |  | 0.099517 | 324.1562 | 415.3476 | 41.094 | 41.094 | 415.3476 | 324.206 | 344.4114 | 344.4612 |  | 275.9328
0 | LS2 | CD05 | CAV3 | 3447 | LS2_CD05:CAV3_D3447 |  | b54 resonator | 40 |  |  | 0.49566 | 324.4538 | 415.3476 | 41.094 | 41.094 | 415.3476 | 324.7016 | 344.709 | 344.9568 |  | 276.2304
0 | [10 empty] | 0 | 324.7016 | 415.3476 | 41.094 | 41.094 | 415.3476 | 324.7016 | 344.9568 | 344.9568 |  | 276.4782
0 | [10 empty] | 0 | 324.7016 | 415.3476 | 41.094 | 41.094 | 415.3476 | 324.7016 | 344.9568 | 344.9568 |  | 276.4782
0 | [6 empty] | bellow |  |  |  | 0.099517 | 324.7514 | 415.3476 | 41.094 | 41.094 | 415.3476 | 324.8011 | 345.0066 | 345.0564 |  | 276.528
0 | LS2 | CD05 | CAV4 | 3453 | LS2_CD05:CAV4_D3453 |  | b54 resonator | 40 |  |  | 0.49566 | 325.049 | 415.3476 | 41.094 | 41.094 | 415.3476 | 325.2968 | 345.3042 | 345.552 |  | 276.8256
0 | LS2 | CD05 | BLM | 3456 | LS2_CD05:BLM_D3456 |  | outside loss monitor |  |  |  | 0 | 325.2968 | 415.3476 | 41.094 | 41.094 | 415.3476 | 325.2968 | 345.552 | 345.552 |  | 277.0734
0 | [10 empty] | 0 | 325.2968 | 415.3476 | 41.094 | 41.094 | 415.3476 | 325.2968 | 345.552 | 345.552 |  | 277.0734
0 | [6 empty] | bellow |  |  |  | 0.15525 | 325.3744 | 415.3476 | 41.094 | 41.094 | 415.3476 | 325.452 | 345.6296 | 345.7073 |  | 277.151
0 | LS2 | CD05 | SOL1 | 3460 | LS2_CD05:SOL1_D3460 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 325.702 | 415.3476 | 41.094 | 41.094 | 415.3476 | 325.952 | 345.9573 | 346.2073 |  | 277.4786
0 | [6 empty] | bellow |  |  |  | 0.15525 | 326.0297 | 415.3476 | 41.094 | 41.094 | 415.3476 | 326.1073 | 346.2849 | 346.3625 |  | 277.8063
0 | LS2 | CD05 | CAV5 | 3466 | LS2_CD05:CAV5_D3466 |  | b54 resonator | 40 |  |  | 0.49566 | 326.3551 | 415.3476 | 41.094 | 41.094 | 415.3476 | 326.6029 | 346.6103 | 346.8582 |  | 278.1317
0 | [10 empty] | 0 | 326.6029 | 415.3476 | 41.094 | 41.094 | 415.3476 | 326.6029 | 346.8582 | 346.8582 |  | 278.3795
0 | [10 empty] | 0 | 326.6029 | 415.3476 | 41.094 | 41.094 | 415.3476 | 326.6029 | 346.8582 | 346.8582 |  | 278.3795
0 | [6 empty] | bellow |  |  |  | 0.099517 | 326.6527 | 415.3476 | 41.094 | 41.094 | 415.3476 | 326.7025 | 346.9079 | 346.9577 |  | 278.4293
0 | LS2 | CD05 | CAV6 | 3472 | LS2_CD05:CAV6_D3472 |  | b54 resonator | 40 |  |  | 0.49566 | 326.9503 | 415.3476 | 41.094 | 41.094 | 415.3476 | 327.1981 | 347.2055 | 347.4533 |  | 278.7269
0 | [10 empty] | 0 | 327.1981 | 415.3476 | 41.094 | 41.094 | 415.3476 | 327.1981 | 347.4533 | 347.4533 |  | 278.9747
0 | [10 empty] | 0 | 327.1981 | 415.3476 | 41.094 | 41.094 | 415.3476 | 327.1981 | 347.4533 | 347.4533 |  | 278.9747
0 | [6 empty] | bellow |  |  |  | 0.099517 | 327.2479 | 415.3476 | 41.094 | 41.094 | 415.3476 | 327.2976 | 347.5031 | 347.5529 |  | 279.0245
0 | LS2 | CD05 | CAV7 | 3478 | LS2_CD05:CAV7_D3478 |  | b54 resonator | 40 |  |  | 0.49566 | 327.5455 | 415.3476 | 41.094 | 41.094 | 415.3476 | 327.7933 | 347.8007 | 348.0485 |  | 279.3221
0 | [10 empty] | 0 | 327.7933 | 415.3476 | 41.094 | 41.094 | 415.3476 | 327.7933 | 348.0485 | 348.0485 |  | 279.5699
0 | [10 empty] | 0 | 327.7933 | 415.3476 | 41.094 | 41.094 | 415.3476 | 327.7933 | 348.0485 | 348.0485 |  | 279.5699
0 | [6 empty] | bellow |  |  |  | 0.099517 | 327.8431 | 415.3476 | 41.094 | 41.094 | 415.3476 | 327.8928 | 348.0983 | 348.148 |  | 279.6196
0 | LS2 | CD05 | CAV8 | 3484 | LS2_CD05:CAV8_D3484 |  | b54 resonator | 40 |  |  | 0.49566 | 328.1406 | 415.3476 | 41.094 | 41.094 | 415.3476 | 328.3885 | 348.3959 | 348.6437 |  | 279.9172
0 | LS2 | CD05 | BLM | 3486 | LS2_CD05:BLM_D3486 |  | outside loss monitor |  |  |  | 0 | 328.3885 | 415.3476 | 41.094 | 41.094 | 415.3476 | 328.3885 | 348.6437 | 348.6437 |  | 280.1651
0 | [10 empty] | 0 | 328.3885 | 415.3476 | 41.094 | 41.094 | 415.3476 | 328.3885 | 348.6437 | 348.6437 |  | 280.1651
0 | [10 empty] | 0.1677 | 328.4723 | 415.3476 | 41.094 | 41.094 | 415.3476 | 328.5562 | 348.7275 | 348.8114 |  | 280.2489
0 | LS2 | CD05 | GV | 3488 | LS2_CD05:GV_D3488 |  | gate valve | 40 |  |  | 0.057404 | 328.5849 | 415.3476 | 41.094 | 41.094 | 415.3476 | 328.6136 | 348.8401 | 348.8688 |  | 280.3615
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 328.6653 | 415.3476 | 41.094 | 41.094 | 415.3476 | 328.7171 | 348.9205 | 348.9723 |  | 280.4419
0 | LS2 | WD05 | BPM | 3490 | LS2_WD05:BPM_D3490 | BPM2 | position monitor |  |  |  | 0 | 328.7171 | 415.3476 | 41.094 | 41.094 | 415.3476 | 328.7171 | 348.9723 | 348.9723 |  | 280.4937
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 328.7897 | 415.3476 | 41.094 | 41.094 | 415.3476 | 328.8624 | 349.0449 | 349.1176 |  | 280.5663
0 | LS2 | WD05 | PORT | 3491 | LS2_WD05:PORT_D3491 |  | port for future diagnostics |  |  |  | 0 | 328.8624 | 415.3476 | 41.094 | 41.094 | 415.3476 | 328.8624 | 349.1176 | 349.1176 |  | 280.6389
0 | [6 empty] | diagnostic box |  |  |  | 0 | 328.8624 | 415.3476 | 41.094 | 41.094 | 415.3476 | 328.8624 | 349.1176 | 349.1176 |  | 280.6389
0 | LS2 | WD05 | IP | 3491 | LS2_WD05:IP_D3491 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 328.8624 | 415.3476 | 41.094 | 41.094 | 415.3476 | 328.8624 | 349.1176 | 349.1176 |  | 280.6389
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 328.928 | 415.3476 | 41.094 | 41.094 | 415.3476 | 328.9936 | 349.1832 | 349.2488 |  | 280.7046
0 | LS2 | CD06 | GV | 3493 | LS2_CD06:GV_D3493 |  | gate valve | 40 |  |  | 0.057404 | 329.0223 | 415.3476 | 41.094 | 41.094 | 415.3476 | 329.051 | 349.2775 | 349.3062 |  | 280.7989
0 | [6 empty] | bellows |  |  |  | 0.1677 | 329.1348 | 415.3476 | 41.094 | 41.094 | 415.3476 | 329.2187 | 349.39 | 349.4739 |  | 280.9114
0 | LS2 | CD06 | CAV1 | 3497 | LS2_CD06:CAV1_D3497 |  | b54 resonator | 40 |  |  | 0.49566 | 329.4665 | 415.3476 | 41.094 | 41.094 | 415.3476 | 329.7143 | 349.7217 | 349.9696 |  | 281.2431
0 | LS2 | CD06 | BLM | 3500 | LS2_CD06:BLM_D3500 |  | outside loss monitor |  |  |  | 0 | 329.7143 | 415.3476 | 41.094 | 41.094 | 415.3476 | 329.7143 | 349.9696 | 349.9696 |  | 281.4909
0 | [10 empty] | 0 | 329.7143 | 415.3476 | 41.094 | 41.094 | 415.3476 | 329.7143 | 349.9696 | 349.9696 |  | 281.4909
0 | [6 empty] | bellow |  |  |  | 0.099517 | 329.7641 | 415.3476 | 41.094 | 41.094 | 415.3476 | 329.8139 | 350.0193 | 350.0691 |  | 281.5407
0 | LS2 | CD06 | CAV2 | 3503 | LS2_CD06:CAV2_D3503 |  | b54 resonator | 40 |  |  | 0.49566 | 330.0617 | 415.3476 | 41.094 | 41.094 | 415.3476 | 330.3095 | 350.3169 | 350.5647 |  | 281.8383
0 | [10 empty] | 0 | 330.3095 | 415.3476 | 41.094 | 41.094 | 415.3476 | 330.3095 | 350.5647 | 350.5647 |  | 282.0861
0 | [10 empty] | 0 | 330.3095 | 415.3476 | 41.094 | 41.094 | 415.3476 | 330.3095 | 350.5647 | 350.5647 |  | 282.0861
0 | [6 empty] | bellow |  |  |  | 0.099517 | 330.3593 | 415.3476 | 41.094 | 41.094 | 415.3476 | 330.409 | 350.6145 | 350.6642 |  | 282.1359
0 | LS2 | CD06 | CAV3 | 3509 | LS2_CD06:CAV3_D3509 |  | b54 resonator | 40 |  |  | 0.49566 | 330.6569 | 415.3476 | 41.094 | 41.094 | 415.3476 | 330.9047 | 350.9121 | 351.1599 |  | 282.4334
0 | [10 empty] | 0 | 330.9047 | 415.3476 | 41.094 | 41.094 | 415.3476 | 330.9047 | 351.1599 | 351.1599 |  | 282.6813
0 | [10 empty] | 0 | 330.9047 | 415.3476 | 41.094 | 41.094 | 415.3476 | 330.9047 | 351.1599 | 351.1599 |  | 282.6813
0 | [6 empty] | bellow |  |  |  | 0.099517 | 330.9544 | 415.3476 | 41.094 | 41.094 | 415.3476 | 331.0042 | 351.2097 | 351.2594 |  | 282.731
0 | LS2 | CD06 | CAV4 | 3515 | LS2_CD06:CAV4_D3515 |  | b54 resonator | 40 |  |  | 0.49566 | 331.252 | 415.3476 | 41.094 | 41.094 | 415.3476 | 331.4999 | 351.5072 | 351.7551 |  | 283.0286
0 | LS2 | CD06 | BLM | 3518 | LS2_CD06:BLM_D3518 |  | outside loss monitor |  |  |  | 0 | 331.4999 | 415.3476 | 41.094 | 41.094 | 415.3476 | 331.4999 | 351.7551 | 351.7551 |  | 283.2764
0 | [10 empty] | 0 | 331.4999 | 415.3476 | 41.094 | 41.094 | 415.3476 | 331.4999 | 351.7551 | 351.7551 |  | 283.2764
0 | [6 empty] | bellow |  |  |  | 0.15525 | 331.5775 | 415.3476 | 41.094 | 41.094 | 415.3476 | 331.6551 | 351.8327 | 351.9103 |  | 283.3541
0 | LS2 | CD06 | SOL1 | 3522 | LS2_CD06:SOL1_D3522 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 331.9051 | 415.3476 | 41.094 | 41.094 | 415.3476 | 332.1551 | 352.1603 | 352.4103 |  | 283.6817
0 | [6 empty] | bellow |  |  |  | 0.15525 | 332.2327 | 415.3476 | 41.094 | 41.094 | 415.3476 | 332.3104 | 352.4879 | 352.5656 |  | 284.0093
0 | LS2 | CD06 | CAV5 | 3528 | LS2_CD06:CAV5_D3528 |  | b54 resonator | 40 |  |  | 0.49566 | 332.5582 | 415.3476 | 41.094 | 41.094 | 415.3476 | 332.806 | 352.8134 | 353.0612 |  | 284.3348
0 | [10 empty] | 0 | 332.806 | 415.3476 | 41.094 | 41.094 | 415.3476 | 332.806 | 353.0612 | 353.0612 |  | 284.5826
0 | [10 empty] | 0 | 332.806 | 415.3476 | 41.094 | 41.094 | 415.3476 | 332.806 | 353.0612 | 353.0612 |  | 284.5826
0 | [6 empty] | bellow |  |  |  | 0.099517 | 332.8558 | 415.3476 | 41.094 | 41.094 | 415.3476 | 332.9055 | 353.111 | 353.1607 |  | 284.6324
0 | LS2 | CD06 | CAV6 | 3534 | LS2_CD06:CAV6_D3534 |  | b54 resonator | 40 |  |  | 0.49566 | 333.1534 | 415.3476 | 41.094 | 41.094 | 415.3476 | 333.4012 | 353.4086 | 353.6564 |  | 284.9299
0 | [10 empty] | 0 | 333.4012 | 415.3476 | 41.094 | 41.094 | 415.3476 | 333.4012 | 353.6564 | 353.6564 |  | 285.1778
0 | [10 empty] | 0 | 333.4012 | 415.3476 | 41.094 | 41.094 | 415.3476 | 333.4012 | 353.6564 | 353.6564 |  | 285.1778
0 | [6 empty] | bellow |  |  |  | 0.099517 | 333.4509 | 415.3476 | 41.094 | 41.094 | 415.3476 | 333.5007 | 353.7062 | 353.7559 |  | 285.2275
0 | LS2 | CD06 | CAV7 | 3540 | LS2_CD06:CAV7_D3540 |  | b54 resonator | 40 |  |  | 0.49566 | 333.7485 | 415.3476 | 41.094 | 41.094 | 415.3476 | 333.9964 | 354.0037 | 354.2516 |  | 285.5251
0 | [10 empty] | 0 | 333.9964 | 415.3476 | 41.094 | 41.094 | 415.3476 | 333.9964 | 354.2516 | 354.2516 |  | 285.7729
0 | [10 empty] | 0 | 333.9964 | 415.3476 | 41.094 | 41.094 | 415.3476 | 333.9964 | 354.2516 | 354.2516 |  | 285.7729
0 | [6 empty] | bellow |  |  |  | 0.099517 | 334.0461 | 415.3476 | 41.094 | 41.094 | 415.3476 | 334.0959 | 354.3013 | 354.3511 |  | 285.8227
0 | LS2 | CD06 | CAV8 | 3546 | LS2_CD06:CAV8_D3546 |  | b54 resonator | 40 |  |  | 0.49566 | 334.3437 | 415.3476 | 41.094 | 41.094 | 415.3476 | 334.5915 | 354.5989 | 354.8467 |  | 286.1203
0 | LS2 | CD06 | BLM | 3548 | LS2_CD06:BLM_D3548 |  | outside loss monitor |  |  |  | 0 | 334.5915 | 415.3476 | 41.094 | 41.094 | 415.3476 | 334.5915 | 354.8467 | 354.8467 |  | 286.3681
0 | [10 empty] | 0 | 334.5915 | 415.3476 | 41.094 | 41.094 | 415.3476 | 334.5915 | 354.8467 | 354.8467 |  | 286.3681
0 | [10 empty] | 0.1677 | 334.6754 | 415.3476 | 41.094 | 41.094 | 415.3476 | 334.7592 | 354.9306 | 355.0145 |  | 286.452
0 | LS2 | CD06 | GV | 3550 | LS2_CD06:GV_D3550 |  | gate valve | 40 |  |  | 0.057404 | 334.7879 | 415.3476 | 41.094 | 41.094 | 415.3476 | 334.8166 | 355.0432 | 355.0719 |  | 286.5645
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 334.8684 | 415.3476 | 41.094 | 41.094 | 415.3476 | 334.9201 | 355.1236 | 355.1754 |  | 286.645
0 | LS2 | WD06 | BPM | 3552 | LS2_WD06:BPM_D3552 | BPM2 | position monitor |  |  |  | 0 | 334.9201 | 415.3476 | 41.094 | 41.094 | 415.3476 | 334.9201 | 355.1754 | 355.1754 |  | 286.6967
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 334.9928 | 415.3476 | 41.094 | 41.094 | 415.3476 | 335.0654 | 355.248 | 355.3206 |  | 286.7694
0 | LS2 | WD06 | PORT | 3553 | LS2_WD06:PORT_D3553 |  | port for future diagnostics |  |  |  | 0 | 335.0654 | 415.3476 | 41.094 | 41.094 | 415.3476 | 335.0654 | 355.3206 | 355.3206 |  | 286.842
0 | [6 empty] | diagnostic box |  |  |  | 0 | 335.0654 | 415.3476 | 41.094 | 41.094 | 415.3476 | 335.0654 | 355.3206 | 355.3206 |  | 286.842
0 | LS2 | WD06 | IP | 3553 | LS2_WD06:IP_D3553 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 335.0654 | 415.3476 | 41.094 | 41.094 | 415.3476 | 335.0654 | 355.3206 | 355.3206 |  | 286.842
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 335.131 | 415.3476 | 41.094 | 41.094 | 415.3476 | 335.1966 | 355.3862 | 355.4519 |  | 286.9076
0 | LS2 | CD07 | GV | 3555 | LS2_CD07:GV_D3555 |  | gate valve | 40 |  |  | 0.057404 | 335.2253 | 415.3476 | 41.094 | 41.094 | 415.3476 | 335.254 | 355.4806 | 355.5093 |  | 287.0019
0 | [6 empty] | bellows |  |  |  | 0.1677 | 335.3379 | 415.3476 | 41.094 | 41.094 | 415.3476 | 335.4217 | 355.5931 | 355.677 |  | 287.1145
0 | LS2 | CD07 | CAV1 | 3559 | LS2_CD07:CAV1_D3559 |  | b54 resonator | 40 |  |  | 0.49566 | 335.6696 | 415.3476 | 41.094 | 41.094 | 415.3476 | 335.9174 | 355.9248 | 356.1726 |  | 287.4462
0 | LS2 | CD07 | BLM | 3562 | LS2_CD07:BLM_D3562 |  | outside loss monitor |  |  |  | 0 | 335.9174 | 415.3476 | 41.094 | 41.094 | 415.3476 | 335.9174 | 356.1726 | 356.1726 |  | 287.694
0 | [10 empty] | 0 | 335.9174 | 415.3476 | 41.094 | 41.094 | 415.3476 | 335.9174 | 356.1726 | 356.1726 |  | 287.694
0 | [6 empty] | bellow |  |  |  | 0.099517 | 335.9672 | 415.3476 | 41.094 | 41.094 | 415.3476 | 336.0169 | 356.2224 | 356.2721 |  | 287.7437
0 | LS2 | CD07 | CAV2 | 3565 | LS2_CD07:CAV2_D3565 |  | b54 resonator | 40 |  |  | 0.49566 | 336.2647 | 415.3476 | 41.094 | 41.094 | 415.3476 | 336.5126 | 356.52 | 356.7678 |  | 288.0413
0 | [10 empty] | 0 | 336.5126 | 415.3476 | 41.094 | 41.094 | 415.3476 | 336.5126 | 356.7678 | 356.7678 |  | 288.2892
0 | [10 empty] | 0 | 336.5126 | 415.3476 | 41.094 | 41.094 | 415.3476 | 336.5126 | 356.7678 | 356.7678 |  | 288.2892
0 | [6 empty] | bellow |  |  |  | 0.099517 | 336.5623 | 415.3476 | 41.094 | 41.094 | 415.3476 | 336.6121 | 356.8176 | 356.8673 |  | 288.3389
0 | LS2 | CD07 | CAV3 | 3571 | LS2_CD07:CAV3_D3571 |  | b54 resonator | 40 |  |  | 0.49566 | 336.8599 | 415.3476 | 41.094 | 41.094 | 415.3476 | 337.1078 | 357.1151 | 357.363 |  | 288.6365
0 | [10 empty] | 0 | 337.1078 | 415.3476 | 41.094 | 41.094 | 415.3476 | 337.1078 | 357.363 | 357.363 |  | 288.8843
0 | [10 empty] | 0 | 337.1078 | 415.3476 | 41.094 | 41.094 | 415.3476 | 337.1078 | 357.363 | 357.363 |  | 288.8843
0 | [6 empty] | bellow |  |  |  | 0.099517 | 337.1575 | 415.3476 | 41.094 | 41.094 | 415.3476 | 337.2073 | 357.4127 | 357.4625 |  | 288.9341
0 | LS2 | CD07 | CAV4 | 3577 | LS2_CD07:CAV4_D3577 |  | b54 resonator | 40 |  |  | 0.49566 | 337.4551 | 415.3476 | 41.094 | 41.094 | 415.3476 | 337.7029 | 357.7103 | 357.9581 |  | 289.2317
0 | LS2 | CD07 | BLM | 3580 | LS2_CD07:BLM_D3580 |  | outside loss monitor |  |  |  | 0 | 337.7029 | 415.3476 | 41.094 | 41.094 | 415.3476 | 337.7029 | 357.9581 | 357.9581 |  | 289.4795
0 | [10 empty] | 0 | 337.7029 | 415.3476 | 41.094 | 41.094 | 415.3476 | 337.7029 | 357.9581 | 357.9581 |  | 289.4795
0 | [6 empty] | bellow |  |  |  | 0.15525 | 337.7805 | 415.3476 | 41.094 | 41.094 | 415.3476 | 337.8582 | 358.0358 | 358.1134 |  | 289.5571
0 | LS2 | CD07 | SOL1 | 3584 | LS2_CD07:SOL1_D3584 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 338.1082 | 415.3476 | 41.094 | 41.094 | 415.3476 | 338.3582 | 358.3634 | 358.6134 |  | 289.8848
0 | [6 empty] | bellow |  |  |  | 0.15525 | 338.4358 | 415.3476 | 41.094 | 41.094 | 415.3476 | 338.5134 | 358.691 | 358.7686 |  | 290.2124
0 | LS2 | CD07 | CAV5 | 3590 | LS2_CD07:CAV5_D3590 |  | b54 resonator | 40 |  |  | 0.49566 | 338.7613 | 415.3476 | 41.094 | 41.094 | 415.3476 | 339.0091 | 359.0165 | 359.2643 |  | 290.5378
0 | [10 empty] | 0 | 339.0091 | 415.3476 | 41.094 | 41.094 | 415.3476 | 339.0091 | 359.2643 | 359.2643 |  | 290.7857
0 | [10 empty] | 0 | 339.0091 | 415.3476 | 41.094 | 41.094 | 415.3476 | 339.0091 | 359.2643 | 359.2643 |  | 290.7857
0 | [6 empty] | bellow |  |  |  | 0.099517 | 339.0588 | 415.3476 | 41.094 | 41.094 | 415.3476 | 339.1086 | 359.3141 | 359.3638 |  | 290.8354
0 | LS2 | CD07 | CAV6 | 3596 | LS2_CD07:CAV6_D3596 |  | b54 resonator | 40 |  |  | 0.49566 | 339.3564 | 415.3476 | 41.094 | 41.094 | 415.3476 | 339.6043 | 359.6116 | 359.8595 |  | 291.133
0 | [10 empty] | 0 | 339.6043 | 415.3476 | 41.094 | 41.094 | 415.3476 | 339.6043 | 359.8595 | 359.8595 |  | 291.3808
0 | [10 empty] | 0 | 339.6043 | 415.3476 | 41.094 | 41.094 | 415.3476 | 339.6043 | 359.8595 | 359.8595 |  | 291.3808
0 | [6 empty] | bellow |  |  |  | 0.099517 | 339.654 | 415.3476 | 41.094 | 41.094 | 415.3476 | 339.7038 | 359.9092 | 359.959 |  | 291.4306
0 | LS2 | CD07 | CAV7 | 3602 | LS2_CD07:CAV7_D3602 |  | b54 resonator | 40 |  |  | 0.49566 | 339.9516 | 415.3476 | 41.094 | 41.094 | 415.3476 | 340.1994 | 360.2068 | 360.4546 |  | 291.7282
0 | [10 empty] | 0 | 340.1994 | 415.3476 | 41.094 | 41.094 | 415.3476 | 340.1994 | 360.4546 | 360.4546 |  | 291.976
0 | [10 empty] | 0 | 340.1994 | 415.3476 | 41.094 | 41.094 | 415.3476 | 340.1994 | 360.4546 | 360.4546 |  | 291.976
0 | [6 empty] | bellow |  |  |  | 0.099517 | 340.2492 | 415.3476 | 41.094 | 41.094 | 415.3476 | 340.2989 | 360.5044 | 360.5542 |  | 292.0258
0 | LS2 | CD07 | CAV8 | 3608 | LS2_CD07:CAV8_D3608 |  | b54 resonator | 40 |  |  | 0.49566 | 340.5468 | 415.3476 | 41.094 | 41.094 | 415.3476 | 340.7946 | 360.802 | 361.0498 |  | 292.3234
0 | LS2 | CD07 | BLM | 3610 | LS2_CD07:BLM_D3610 |  | outside loss monitor |  |  |  | 0 | 340.7946 | 415.3476 | 41.094 | 41.094 | 415.3476 | 340.7946 | 361.0498 | 361.0498 |  | 292.5712
0 | [10 empty] | 0 | 340.7946 | 415.3476 | 41.094 | 41.094 | 415.3476 | 340.7946 | 361.0498 | 361.0498 |  | 292.5712
0 | [10 empty] | 0.1677 | 340.8785 | 415.3476 | 41.094 | 41.094 | 415.3476 | 340.9623 | 361.1337 | 361.2175 |  | 292.655
0 | LS2 | CD07 | GV | 3612 | LS2_CD07:GV_D3612 |  | gate valve | 40 |  |  | 0.057404 | 340.991 | 415.3476 | 41.094 | 41.094 | 415.3476 | 341.0197 | 361.2462 | 361.2749 |  | 292.7676
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 341.0715 | 415.3476 | 41.094 | 41.094 | 415.3476 | 341.1232 | 361.3267 | 361.3784 |  | 292.848
0 | LS2 | WD07 | BPM | 3614 | LS2_WD07:BPM_D3614 | BPM2 | position monitor |  |  |  | 0 | 341.1232 | 415.3476 | 41.094 | 41.094 | 415.3476 | 341.1232 | 361.3784 | 361.3784 |  | 292.8998
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 341.1958 | 415.3476 | 41.094 | 41.094 | 415.3476 | 341.2685 | 361.4511 | 361.5237 |  | 292.9724
0 | LS2 | WD07 | PORT | 3615 | LS2_WD07:PORT_D3615 |  | port for future diagnostics |  |  |  | 0 | 341.2685 | 415.3476 | 41.094 | 41.094 | 415.3476 | 341.2685 | 361.5237 | 361.5237 |  | 293.0451
0 | [6 empty] | diagnostic box |  |  |  | 0 | 341.2685 | 415.3476 | 41.094 | 41.094 | 415.3476 | 341.2685 | 361.5237 | 361.5237 |  | 293.0451
0 | LS2 | WD07 | IP | 3615 | LS2_WD07:IP_D3615 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 341.2685 | 415.3476 | 41.094 | 41.094 | 415.3476 | 341.2685 | 361.5237 | 361.5237 |  | 293.0451
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 341.3341 | 415.3476 | 41.094 | 41.094 | 415.3476 | 341.3997 | 361.5893 | 361.6549 |  | 293.1107
0 | LS2 | CD08 | GV | 3617 | LS2_CD08:GV_D3617 |  | gate valve | 40 |  |  | 0.057404 | 341.4284 | 415.3476 | 41.094 | 41.094 | 415.3476 | 341.4571 | 361.6836 | 361.7123 |  | 293.205
0 | [6 empty] | bellows |  |  |  | 0.1677 | 341.541 | 415.3476 | 41.094 | 41.094 | 415.3476 | 341.6248 | 361.7962 | 361.88 |  | 293.3175
0 | LS2 | CD08 | CAV1 | 3621 | LS2_CD08:CAV1_D3621 |  | b54 resonator | 40 |  |  | 0.49566 | 341.8726 | 415.3476 | 41.094 | 41.094 | 415.3476 | 342.1205 | 362.1279 | 362.3757 |  | 293.6492
0 | LS2 | CD08 | BLM | 3624 | LS2_CD08:BLM_D3624 |  | outside loss monitor |  |  |  | 0 | 342.1205 | 415.3476 | 41.094 | 41.094 | 415.3476 | 342.1205 | 362.3757 | 362.3757 |  | 293.8971
0 | [10 empty] | 0 | 342.1205 | 415.3476 | 41.094 | 41.094 | 415.3476 | 342.1205 | 362.3757 | 362.3757 |  | 293.8971
0 | [6 empty] | bellow |  |  |  | 0.099517 | 342.1702 | 415.3476 | 41.094 | 41.094 | 415.3476 | 342.22 | 362.4254 | 362.4752 |  | 293.9468
0 | LS2 | CD08 | CAV2 | 3627 | LS2_CD08:CAV2_D3627 |  | b54 resonator | 40 |  |  | 0.49566 | 342.4678 | 415.3476 | 41.094 | 41.094 | 415.3476 | 342.7156 | 362.723 | 362.9709 |  | 294.2444
0 | [10 empty] | 0 | 342.7156 | 415.3476 | 41.094 | 41.094 | 415.3476 | 342.7156 | 362.9709 | 362.9709 |  | 294.4922
0 | [10 empty] | 0 | 342.7156 | 415.3476 | 41.094 | 41.094 | 415.3476 | 342.7156 | 362.9709 | 362.9709 |  | 294.4922
0 | [6 empty] | bellow |  |  |  | 0.099517 | 342.7654 | 415.3476 | 41.094 | 41.094 | 415.3476 | 342.8152 | 363.0206 | 363.0704 |  | 294.542
0 | LS2 | CD08 | CAV3 | 3633 | LS2_CD08:CAV3_D3633 |  | b54 resonator | 40 |  |  | 0.49566 | 343.063 | 415.3476 | 41.094 | 41.094 | 415.3476 | 343.3108 | 363.3182 | 363.566 |  | 294.8396
0 | [10 empty] | 0 | 343.3108 | 415.3476 | 41.094 | 41.094 | 415.3476 | 343.3108 | 363.566 | 363.566 |  | 295.0874
0 | [10 empty] | 0 | 343.3108 | 415.3476 | 41.094 | 41.094 | 415.3476 | 343.3108 | 363.566 | 363.566 |  | 295.0874
0 | [6 empty] | bellow |  |  |  | 0.099517 | 343.3606 | 415.3476 | 41.094 | 41.094 | 415.3476 | 343.4103 | 363.6158 | 363.6655 |  | 295.1372
0 | LS2 | CD08 | CAV4 | 3639 | LS2_CD08:CAV4_D3639 |  | b54 resonator | 40 |  |  | 0.49566 | 343.6582 | 415.3476 | 41.094 | 41.094 | 415.3476 | 343.906 | 363.9134 | 364.1612 |  | 295.4347
0 | LS2 | CD08 | BLM | 3642 | LS2_CD08:BLM_D3642 |  | outside loss monitor |  |  |  | 0 | 343.906 | 415.3476 | 41.094 | 41.094 | 415.3476 | 343.906 | 364.1612 | 364.1612 |  | 295.6826
0 | [10 empty] | 0 | 343.906 | 415.3476 | 41.094 | 41.094 | 415.3476 | 343.906 | 364.1612 | 364.1612 |  | 295.6826
0 | [6 empty] | bellow |  |  |  | 0.15525 | 343.9836 | 415.3476 | 41.094 | 41.094 | 415.3476 | 344.0612 | 364.2388 | 364.3165 |  | 295.7602
0 | LS2 | CD08 | SOL1 | 3646 | LS2_CD08:SOL1_D3646 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 344.3112 | 415.3476 | 41.094 | 41.094 | 415.3476 | 344.5612 | 364.5665 | 364.8165 |  | 296.0878
0 | [6 empty] | bellow |  |  |  | 0.15525 | 344.6389 | 415.3476 | 41.094 | 41.094 | 415.3476 | 344.7165 | 364.8941 | 364.9717 |  | 296.4155
0 | LS2 | CD08 | CAV5 | 3652 | LS2_CD08:CAV5_D3652 |  | b54 resonator | 40 |  |  | 0.49566 | 344.9643 | 415.3476 | 41.094 | 41.094 | 415.3476 | 345.2121 | 365.2195 | 365.4674 |  | 296.7409
0 | [10 empty] | 0 | 345.2121 | 415.3476 | 41.094 | 41.094 | 415.3476 | 345.2121 | 365.4674 | 365.4674 |  | 296.9887
0 | [10 empty] | 0 | 345.2121 | 415.3476 | 41.094 | 41.094 | 415.3476 | 345.2121 | 365.4674 | 365.4674 |  | 296.9887
0 | [6 empty] | bellow |  |  |  | 0.099517 | 345.2619 | 415.3476 | 41.094 | 41.094 | 415.3476 | 345.3117 | 365.5171 | 365.5669 |  | 297.0385
0 | LS2 | CD08 | CAV6 | 3658 | LS2_CD08:CAV6_D3658 |  | b54 resonator | 40 |  |  | 0.49566 | 345.5595 | 415.3476 | 41.094 | 41.094 | 415.3476 | 345.8073 | 365.8147 | 366.0625 |  | 297.3361
0 | [10 empty] | 0 | 345.8073 | 415.3476 | 41.094 | 41.094 | 415.3476 | 345.8073 | 366.0625 | 366.0625 |  | 297.5839
0 | [10 empty] | 0 | 345.8073 | 415.3476 | 41.094 | 41.094 | 415.3476 | 345.8073 | 366.0625 | 366.0625 |  | 297.5839
0 | [6 empty] | bellow |  |  |  | 0.099517 | 345.8571 | 415.3476 | 41.094 | 41.094 | 415.3476 | 345.9068 | 366.1123 | 366.1621 |  | 297.6337
0 | LS2 | CD08 | CAV7 | 3664 | LS2_CD08:CAV7_D3664 |  | b54 resonator | 40 |  |  | 0.49566 | 346.1547 | 415.3476 | 41.094 | 41.094 | 415.3476 | 346.4025 | 366.4099 | 366.6577 |  | 297.9313
0 | [10 empty] | 0 | 346.4025 | 415.3476 | 41.094 | 41.094 | 415.3476 | 346.4025 | 366.6577 | 366.6577 |  | 298.1791
0 | [10 empty] | 0 | 346.4025 | 415.3476 | 41.094 | 41.094 | 415.3476 | 346.4025 | 366.6577 | 366.6577 |  | 298.1791
0 | [6 empty] | bellow |  |  |  | 0.099517 | 346.4523 | 415.3476 | 41.094 | 41.094 | 415.3476 | 346.502 | 366.7075 | 366.7572 |  | 298.2288
0 | LS2 | CD08 | CAV8 | 3670 | LS2_CD08:CAV8_D3670 |  | b54 resonator | 40 |  |  | 0.49566 | 346.7498 | 415.3476 | 41.094 | 41.094 | 415.3476 | 346.9977 | 367.0051 | 367.2529 |  | 298.5264
0 | LS2 | CD08 | BLM | 3673 | LS2_CD08:BLM_D3673 |  | outside loss monitor |  |  |  | 0 | 346.9977 | 415.3476 | 41.094 | 41.094 | 415.3476 | 346.9977 | 367.2529 | 367.2529 |  | 298.7743
0 | [10 empty] | 0 | 346.9977 | 415.3476 | 41.094 | 41.094 | 415.3476 | 346.9977 | 367.2529 | 367.2529 |  | 298.7743
0 | [10 empty] | 0.1677 | 347.0815 | 415.3476 | 41.094 | 41.094 | 415.3476 | 347.1654 | 367.3367 | 367.4206 |  | 298.8581
0 | LS2 | CD08 | GV | 3674 | LS2_CD08:GV_D3674 |  | gate valve | 40 |  |  | 0.057404 | 347.1941 | 415.3476 | 41.094 | 41.094 | 415.3476 | 347.2228 | 367.4493 | 367.478 |  | 298.9707
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 347.2745 | 415.3476 | 41.094 | 41.094 | 415.3476 | 347.3263 | 367.5297 | 367.5815 |  | 299.0511
0 | LS2 | WD08 | BPM | 3676 | LS2_WD08:BPM_D3676 | BPM2 | position monitor |  |  |  | 0 | 347.3263 | 415.3476 | 41.094 | 41.094 | 415.3476 | 347.3263 | 367.5815 | 367.5815 |  | 299.1029
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 347.3989 | 415.3476 | 41.094 | 41.094 | 415.3476 | 347.4716 | 367.6541 | 367.7268 |  | 299.1755
0 | LS2 | WD08 | PORT | 3677 | LS2_WD08:PORT_D3677 |  | port for future diagnostics |  |  |  | 0 | 347.4716 | 415.3476 | 41.094 | 41.094 | 415.3476 | 347.4716 | 367.7268 | 367.7268 |  | 299.2481
0 | [6 empty] | diagnostic box |  |  |  | 0 | 347.4716 | 415.3476 | 41.094 | 41.094 | 415.3476 | 347.4716 | 367.7268 | 367.7268 |  | 299.2481
0 | LS2 | WD08 | IP | 3677 | LS2_WD08:IP_D3677 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 347.4716 | 415.3476 | 41.094 | 41.094 | 415.3476 | 347.4716 | 367.7268 | 367.7268 |  | 299.2481
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 347.5372 | 415.3476 | 41.094 | 41.094 | 415.3476 | 347.6028 | 367.7924 | 367.858 |  | 299.3138
0 | LS2 | CD09 | GV | 3679 | LS2_CD09:GV_D3679 |  | gate valve | 40 |  |  | 0.057404 | 347.6315 | 415.3476 | 41.094 | 41.094 | 415.3476 | 347.6602 | 367.8867 | 367.9154 |  | 299.4081
0 | [6 empty] | bellows |  |  |  | 0.1677 | 347.744 | 415.3476 | 41.094 | 41.094 | 415.3476 | 347.8279 | 367.9992 | 368.0831 |  | 299.5206
0 | LS2 | CD09 | CAV1 | 3683 | LS2_CD09:CAV1_D3683 |  | b54 resonator | 40 |  |  | 0.49566 | 348.0757 | 415.3476 | 41.094 | 41.094 | 415.3476 | 348.3235 | 368.3309 | 368.5788 |  | 299.8523
0 | LS2 | CD09 | BLM | 3686 | LS2_CD09:BLM_D3686 |  | outside loss monitor |  |  |  | 0 | 348.3235 | 415.3476 | 41.094 | 41.094 | 415.3476 | 348.3235 | 368.5788 | 368.5788 |  | 300.1001
0 | [10 empty] | 0 | 348.3235 | 415.3476 | 41.094 | 41.094 | 415.3476 | 348.3235 | 368.5788 | 368.5788 |  | 300.1001
0 | [6 empty] | bellow |  |  |  | 0.099517 | 348.3733 | 415.3476 | 41.094 | 41.094 | 415.3476 | 348.4231 | 368.6285 | 368.6783 |  | 300.1499
0 | LS2 | CD09 | CAV2 | 3689 | LS2_CD09:CAV2_D3689 |  | b54 resonator | 40 |  |  | 0.49566 | 348.6709 | 415.3476 | 41.094 | 41.094 | 415.3476 | 348.9187 | 368.9261 | 369.1739 |  | 300.4475
0 | [10 empty] | 0 | 348.9187 | 415.3476 | 41.094 | 41.094 | 415.3476 | 348.9187 | 369.1739 | 369.1739 |  | 300.6953
0 | [10 empty] | 0 | 348.9187 | 415.3476 | 41.094 | 41.094 | 415.3476 | 348.9187 | 369.1739 | 369.1739 |  | 300.6953
0 | [6 empty] | bellow |  |  |  | 0.099517 | 348.9685 | 415.3476 | 41.094 | 41.094 | 415.3476 | 349.0182 | 369.2237 | 369.2734 |  | 300.7451
0 | LS2 | CD09 | CAV3 | 3695 | LS2_CD09:CAV3_D3695 |  | b54 resonator | 40 |  |  | 0.49566 | 349.2661 | 415.3476 | 41.094 | 41.094 | 415.3476 | 349.5139 | 369.5213 | 369.7691 |  | 301.0426
0 | [10 empty] | 0 | 349.5139 | 415.3476 | 41.094 | 41.094 | 415.3476 | 349.5139 | 369.7691 | 369.7691 |  | 301.2905
0 | [10 empty] | 0 | 349.5139 | 415.3476 | 41.094 | 41.094 | 415.3476 | 349.5139 | 369.7691 | 369.7691 |  | 301.2905
0 | [6 empty] | bellow |  |  |  | 0.099517 | 349.5636 | 415.3476 | 41.094 | 41.094 | 415.3476 | 349.6134 | 369.8189 | 369.8686 |  | 301.3402
0 | LS2 | CD09 | CAV4 | 3701 | LS2_CD09:CAV4_D3701 |  | b54 resonator | 40 |  |  | 0.49566 | 349.8612 | 415.3476 | 41.094 | 41.094 | 415.3476 | 350.1091 | 370.1164 | 370.3643 |  | 301.6378
0 | LS2 | CD09 | BLM | 3704 | LS2_CD09:BLM_D3704 |  | outside loss monitor |  |  |  | 0 | 350.1091 | 415.3476 | 41.094 | 41.094 | 415.3476 | 350.1091 | 370.3643 | 370.3643 |  | 301.8856
0 | [10 empty] | 0 | 350.1091 | 415.3476 | 41.094 | 41.094 | 415.3476 | 350.1091 | 370.3643 | 370.3643 |  | 301.8856
0 | [6 empty] | bellow |  |  |  | 0.15525 | 350.1867 | 415.3476 | 41.094 | 41.094 | 415.3476 | 350.2643 | 370.4419 | 370.5195 |  | 301.9633
0 | LS2 | CD09 | SOL1 | 3708 | LS2_CD09:SOL1_D3708 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 350.5143 | 415.3476 | 41.094 | 41.094 | 415.3476 | 350.7643 | 370.7695 | 371.0195 |  | 302.2909
0 | [6 empty] | bellow |  |  |  | 0.15525 | 350.8419 | 415.3476 | 41.094 | 41.094 | 415.3476 | 350.9196 | 371.0971 | 371.1748 |  | 302.6185
0 | LS2 | CD09 | CAV5 | 3714 | LS2_CD09:CAV5_D3714 |  | b54 resonator | 40 |  |  | 0.49566 | 351.1674 | 415.3476 | 41.094 | 41.094 | 415.3476 | 351.4152 | 371.4226 | 371.6704 |  | 302.944
0 | [10 empty] | 0 | 351.4152 | 415.3476 | 41.094 | 41.094 | 415.3476 | 351.4152 | 371.6704 | 371.6704 |  | 303.1918
0 | [10 empty] | 0 | 351.4152 | 415.3476 | 41.094 | 41.094 | 415.3476 | 351.4152 | 371.6704 | 371.6704 |  | 303.1918
0 | [6 empty] | bellow |  |  |  | 0.099517 | 351.465 | 415.3476 | 41.094 | 41.094 | 415.3476 | 351.5147 | 371.7202 | 371.7699 |  | 303.2416
0 | LS2 | CD09 | CAV6 | 3720 | LS2_CD09:CAV6_D3720 |  | b54 resonator | 40 |  |  | 0.49566 | 351.7626 | 415.3476 | 41.094 | 41.094 | 415.3476 | 352.0104 | 372.0178 | 372.2656 |  | 303.5391
0 | [10 empty] | 0 | 352.0104 | 415.3476 | 41.094 | 41.094 | 415.3476 | 352.0104 | 372.2656 | 372.2656 |  | 303.787
0 | [10 empty] | 0 | 352.0104 | 415.3476 | 41.094 | 41.094 | 415.3476 | 352.0104 | 372.2656 | 372.2656 |  | 303.787
0 | [6 empty] | bellow |  |  |  | 0.099517 | 352.0601 | 415.3476 | 41.094 | 41.094 | 415.3476 | 352.1099 | 372.3154 | 372.3651 |  | 303.8367
0 | LS2 | CD09 | CAV7 | 3726 | LS2_CD09:CAV7_D3726 |  | b54 resonator | 40 |  |  | 0.49566 | 352.3577 | 415.3476 | 41.094 | 41.094 | 415.3476 | 352.6056 | 372.6129 | 372.8608 |  | 304.1343
0 | [10 empty] | 0 | 352.6056 | 415.3476 | 41.094 | 41.094 | 415.3476 | 352.6056 | 372.8608 | 372.8608 |  | 304.3821
0 | [10 empty] | 0 | 352.6056 | 415.3476 | 41.094 | 41.094 | 415.3476 | 352.6056 | 372.8608 | 372.8608 |  | 304.3821
0 | [6 empty] | bellow |  |  |  | 0.099517 | 352.6553 | 415.3476 | 41.094 | 41.094 | 415.3476 | 352.7051 | 372.9105 | 372.9603 |  | 304.4319
0 | LS2 | CD09 | CAV8 | 3732 | LS2_CD09:CAV8_D3732 |  | b54 resonator | 40 |  |  | 0.49566 | 352.9529 | 415.3476 | 41.094 | 41.094 | 415.3476 | 353.2007 | 373.2081 | 373.4559 |  | 304.7295
0 | LS2 | CD09 | BLM | 3735 | LS2_CD09:BLM_D3735 |  | outside loss monitor |  |  |  | 0 | 353.2007 | 415.3476 | 41.094 | 41.094 | 415.3476 | 353.2007 | 373.4559 | 373.4559 |  | 304.9773
0 | [10 empty] | 0 | 353.2007 | 415.3476 | 41.094 | 41.094 | 415.3476 | 353.2007 | 373.4559 | 373.4559 |  | 304.9773
0 | [10 empty] | 0.1677 | 353.2846 | 415.3476 | 41.094 | 41.094 | 415.3476 | 353.3684 | 373.5398 | 373.6237 |  | 305.0612
0 | LS2 | CD09 | GV | 3737 | LS2_CD09:GV_D3737 |  | gate valve | 40 |  |  | 0.057404 | 353.3971 | 415.3476 | 41.094 | 41.094 | 415.3476 | 353.4258 | 373.6524 | 373.6811 |  | 305.1737
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 353.4776 | 415.3476 | 41.094 | 41.094 | 415.3476 | 353.5293 | 373.7328 | 373.7846 |  | 305.2542
0 | LS2 | WD09 | BPM | 3738 | LS2_WD09:BPM_D3738 | BPM2 | position monitor |  |  |  | 0 | 353.5293 | 415.3476 | 41.094 | 41.094 | 415.3476 | 353.5293 | 373.7846 | 373.7846 |  | 305.3059
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 353.602 | 415.3476 | 41.094 | 41.094 | 415.3476 | 353.6746 | 373.8572 | 373.9298 |  | 305.3786
0 | LS2 | WD09 | PORT | 3739 | LS2_WD09:PORT_D3739 |  | port for future diagnostics |  |  |  | 0 | 353.6746 | 415.3476 | 41.094 | 41.094 | 415.3476 | 353.6746 | 373.9298 | 373.9298 |  | 305.4512
0 | [6 empty] | diagnostic box |  |  |  | 0 | 353.6746 | 415.3476 | 41.094 | 41.094 | 415.3476 | 353.6746 | 373.9298 | 373.9298 |  | 305.4512
0 | LS2 | WD09 | IP | 3739 | LS2_WD09:IP_D3739 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 353.6746 | 415.3476 | 41.094 | 41.094 | 415.3476 | 353.6746 | 373.9298 | 373.9298 |  | 305.4512
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 353.7402 | 415.3476 | 41.094 | 41.094 | 415.3476 | 353.8058 | 373.9954 | 374.0611 |  | 305.5168
0 | LS2 | CD10 | GV | 3741 | LS2_CD10:GV_D3741 |  | gate valve | 40 |  |  | 0.057404 | 353.8345 | 415.3476 | 41.094 | 41.094 | 415.3476 | 353.8632 | 374.0898 | 374.1185 |  | 305.6111
0 | [6 empty] | bellows |  |  |  | 0.1677 | 353.9471 | 415.3476 | 41.094 | 41.094 | 415.3476 | 354.0309 | 374.2023 | 374.2862 |  | 305.7237
0 | LS2 | CD10 | CAV1 | 3745 | LS2_CD10:CAV1_D3745 |  | b54 resonator | 40 |  |  | 0.49566 | 354.2788 | 415.3476 | 41.094 | 41.094 | 415.3476 | 354.5266 | 374.534 | 374.7818 |  | 306.0554
0 | LS2 | CD10 | BLM | 3748 | LS2_CD10:BLM_D3748 |  | outside loss monitor |  |  |  | 0 | 354.5266 | 415.3476 | 41.094 | 41.094 | 415.3476 | 354.5266 | 374.7818 | 374.7818 |  | 306.3032
0 | [10 empty] | 0 | 354.5266 | 415.3476 | 41.094 | 41.094 | 415.3476 | 354.5266 | 374.7818 | 374.7818 |  | 306.3032
0 | [6 empty] | bellow |  |  |  | 0.099517 | 354.5764 | 415.3476 | 41.094 | 41.094 | 415.3476 | 354.6261 | 374.8316 | 374.8813 |  | 306.3529
0 | LS2 | CD10 | CAV2 | 3751 | LS2_CD10:CAV2_D3751 |  | b54 resonator | 40 |  |  | 0.49566 | 354.8739 | 415.3476 | 41.094 | 41.094 | 415.3476 | 355.1218 | 375.1292 | 375.377 |  | 306.6505
0 | [10 empty] | 0 | 355.1218 | 415.3476 | 41.094 | 41.094 | 415.3476 | 355.1218 | 375.377 | 375.377 |  | 306.8984
0 | [10 empty] | 0 | 355.1218 | 415.3476 | 41.094 | 41.094 | 415.3476 | 355.1218 | 375.377 | 375.377 |  | 306.8984
0 | [6 empty] | bellow |  |  |  | 0.099517 | 355.1715 | 415.3476 | 41.094 | 41.094 | 415.3476 | 355.2213 | 375.4268 | 375.4765 |  | 306.9481
0 | LS2 | CD10 | CAV3 | 3757 | LS2_CD10:CAV3_D3757 |  | b54 resonator | 40 |  |  | 0.49566 | 355.4691 | 415.3476 | 41.094 | 41.094 | 415.3476 | 355.7169 | 375.7243 | 375.9722 |  | 307.2457
0 | [10 empty] | 0 | 355.7169 | 415.3476 | 41.094 | 41.094 | 415.3476 | 355.7169 | 375.9722 | 375.9722 |  | 307.4935
0 | [10 empty] | 0 | 355.7169 | 415.3476 | 41.094 | 41.094 | 415.3476 | 355.7169 | 375.9722 | 375.9722 |  | 307.4935
0 | [6 empty] | bellow |  |  |  | 0.099517 | 355.7667 | 415.3476 | 41.094 | 41.094 | 415.3476 | 355.8165 | 376.0219 | 376.0717 |  | 307.5433
0 | LS2 | CD10 | CAV4 | 3763 | LS2_CD10:CAV4_D3763 |  | b54 resonator | 40 |  |  | 0.49566 | 356.0643 | 415.3476 | 41.094 | 41.094 | 415.3476 | 356.3121 | 376.3195 | 376.5673 |  | 307.8409
0 | LS2 | CD10 | BLM | 3766 | LS2_CD10:BLM_D3766 |  | outside loss monitor |  |  |  | 0 | 356.3121 | 415.3476 | 41.094 | 41.094 | 415.3476 | 356.3121 | 376.5673 | 376.5673 |  | 308.0887
0 | [10 empty] | 0 | 356.3121 | 415.3476 | 41.094 | 41.094 | 415.3476 | 356.3121 | 376.5673 | 376.5673 |  | 308.0887
0 | [6 empty] | bellow |  |  |  | 0.15525 | 356.3897 | 415.3476 | 41.094 | 41.094 | 415.3476 | 356.4674 | 376.645 | 376.7226 |  | 308.1663
0 | LS2 | CD10 | SOL1 | 3770 | LS2_CD10:SOL1_D3770 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 356.7174 | 415.3476 | 41.094 | 41.094 | 415.3476 | 356.9674 | 376.9726 | 377.2226 |  | 308.494
0 | [6 empty] | bellow |  |  |  | 0.15525 | 357.045 | 415.3476 | 41.094 | 41.094 | 415.3476 | 357.1226 | 377.3002 | 377.3778 |  | 308.8216
0 | LS2 | CD10 | CAV5 | 3776 | LS2_CD10:CAV5_D3776 |  | b54 resonator | 40 |  |  | 0.49566 | 357.3704 | 415.3476 | 41.094 | 41.094 | 415.3476 | 357.6183 | 377.6257 | 377.8735 |  | 309.147
0 | [10 empty] | 0 | 357.6183 | 415.3476 | 41.094 | 41.094 | 415.3476 | 357.6183 | 377.8735 | 377.8735 |  | 309.3949
0 | [10 empty] | 0 | 357.6183 | 415.3476 | 41.094 | 41.094 | 415.3476 | 357.6183 | 377.8735 | 377.8735 |  | 309.3949
0 | [6 empty] | bellow |  |  |  | 0.099517 | 357.668 | 415.3476 | 41.094 | 41.094 | 415.3476 | 357.7178 | 377.9233 | 377.973 |  | 309.4446
0 | LS2 | CD10 | CAV6 | 3782 | LS2_CD10:CAV6_D3782 |  | b54 resonator | 40 |  |  | 0.49566 | 357.9656 | 415.3476 | 41.094 | 41.094 | 415.3476 | 358.2135 | 378.2208 | 378.4687 |  | 309.7422
0 | [10 empty] | 0 | 358.2135 | 415.3476 | 41.094 | 41.094 | 415.3476 | 358.2135 | 378.4687 | 378.4687 |  | 309.99
0 | [10 empty] | 0 | 358.2135 | 415.3476 | 41.094 | 41.094 | 415.3476 | 358.2135 | 378.4687 | 378.4687 |  | 309.99
0 | [6 empty] | bellow |  |  |  | 0.099517 | 358.2632 | 415.3476 | 41.094 | 41.094 | 415.3476 | 358.313 | 378.5184 | 378.5682 |  | 310.0398
0 | LS2 | CD10 | CAV7 | 3788 | LS2_CD10:CAV7_D3788 |  | b54 resonator | 40 |  |  | 0.49566 | 358.5608 | 415.3476 | 41.094 | 41.094 | 415.3476 | 358.8086 | 378.816 | 379.0638 |  | 310.3374
0 | [10 empty] | 0 | 358.8086 | 415.3476 | 41.094 | 41.094 | 415.3476 | 358.8086 | 379.0638 | 379.0638 |  | 310.5852
0 | [10 empty] | 0 | 358.8086 | 415.3476 | 41.094 | 41.094 | 415.3476 | 358.8086 | 379.0638 | 379.0638 |  | 310.5852
0 | [6 empty] | bellow |  |  |  | 0.099517 | 358.8584 | 415.3476 | 41.094 | 41.094 | 415.3476 | 358.9081 | 379.1136 | 379.1634 |  | 310.635
0 | LS2 | CD10 | CAV8 | 3794 | LS2_CD10:CAV8_D3794 |  | b54 resonator | 40 |  |  | 0.49566 | 359.156 | 415.3476 | 41.094 | 41.094 | 415.3476 | 359.4038 | 379.4112 | 379.659 |  | 310.9326
0 | LS2 | CD10 | BLM | 3797 | LS2_CD10:BLM_D3797 |  | outside loss monitor |  |  |  | 0 | 359.4038 | 415.3476 | 41.094 | 41.094 | 415.3476 | 359.4038 | 379.659 | 379.659 |  | 311.1804
0 | [10 empty] | 0 | 359.4038 | 415.3476 | 41.094 | 41.094 | 415.3476 | 359.4038 | 379.659 | 379.659 |  | 311.1804
0 | [10 empty] | 0.1677 | 359.4876 | 415.3476 | 41.094 | 41.094 | 415.3476 | 359.5715 | 379.7429 | 379.8267 |  | 311.2642
0 | LS2 | CD10 | GV | 3799 | LS2_CD10:GV_D3799 |  | gate valve | 40 |  |  | 0.057404 | 359.6002 | 415.3476 | 41.094 | 41.094 | 415.3476 | 359.6289 | 379.8554 | 379.8841 |  | 311.3768
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 359.6807 | 415.3476 | 41.094 | 41.094 | 415.3476 | 359.7324 | 379.9359 | 379.9876 |  | 311.4572
0 | LS2 | WD10 | BPM | 3800 | LS2_WD10:BPM_D3800 | BPM2 | position monitor |  |  |  | 0 | 359.7324 | 415.3476 | 41.094 | 41.094 | 415.3476 | 359.7324 | 379.9876 | 379.9876 |  | 311.509
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 359.805 | 415.3476 | 41.094 | 41.094 | 415.3476 | 359.8777 | 380.0603 | 380.1329 |  | 311.5816
0 | LS2 | WD10 | PORT | 3801 | LS2_WD10:PORT_D3801 |  | port for future diagnostics |  |  |  | 0 | 359.8777 | 415.3476 | 41.094 | 41.094 | 415.3476 | 359.8777 | 380.1329 | 380.1329 |  | 311.6543
0 | [6 empty] | diagnostic box |  |  |  | 0 | 359.8777 | 415.3476 | 41.094 | 41.094 | 415.3476 | 359.8777 | 380.1329 | 380.1329 |  | 311.6543
0 | LS2 | WD10 | IP | 3801 | LS2_WD10:IP_D3801 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 359.8777 | 415.3476 | 41.094 | 41.094 | 415.3476 | 359.8777 | 380.1329 | 380.1329 |  | 311.6543
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 359.9433 | 415.3476 | 41.094 | 41.094 | 415.3476 | 360.0089 | 380.1985 | 380.2641 |  | 311.7199
0 | LS2 | CD11 | GV | 3803 | LS2_CD11:GV_D3803 |  | gate valve | 40 |  |  | 0.057404 | 360.0376 | 415.3476 | 41.094 | 41.094 | 415.3476 | 360.0663 | 380.2928 | 380.3215 |  | 311.8142
0 | [6 empty] | bellows |  |  |  | 0.1677 | 360.1502 | 415.3476 | 41.094 | 41.094 | 415.3476 | 360.234 | 380.4054 | 380.4892 |  | 311.9267
0 | LS2 | CD11 | CAV1 | 3807 | LS2_CD11:CAV1_D3807 |  | b54 resonator | 40 |  |  | 0.49566 | 360.4818 | 415.3476 | 41.094 | 41.094 | 415.3476 | 360.7297 | 380.7371 | 380.9849 |  | 312.2584
0 | LS2 | CD11 | BLM | 3810 | LS2_CD11:BLM_D3810 |  | outside loss monitor |  |  |  | 0 | 360.7297 | 415.3476 | 41.094 | 41.094 | 415.3476 | 360.7297 | 380.9849 | 380.9849 |  | 312.5063
0 | [10 empty] | 0 | 360.7297 | 415.3476 | 41.094 | 41.094 | 415.3476 | 360.7297 | 380.9849 | 380.9849 |  | 312.5063
0 | [6 empty] | bellow |  |  |  | 0.099517 | 360.7794 | 415.3476 | 41.094 | 41.094 | 415.3476 | 360.8292 | 381.0346 | 381.0844 |  | 312.556
0 | LS2 | CD11 | CAV2 | 3813 | LS2_CD11:CAV2_D3813 |  | b54 resonator | 40 |  |  | 0.49566 | 361.077 | 415.3476 | 41.094 | 41.094 | 415.3476 | 361.3248 | 381.3322 | 381.5801 |  | 312.8536
0 | [10 empty] | 0 | 361.3248 | 415.3476 | 41.094 | 41.094 | 415.3476 | 361.3248 | 381.5801 | 381.5801 |  | 313.1014
0 | [10 empty] | 0 | 361.3248 | 415.3476 | 41.094 | 41.094 | 415.3476 | 361.3248 | 381.5801 | 381.5801 |  | 313.1014
0 | [6 empty] | bellow |  |  |  | 0.099517 | 361.3746 | 415.3476 | 41.094 | 41.094 | 415.3476 | 361.4244 | 381.6298 | 381.6796 |  | 313.1512
0 | LS2 | CD11 | CAV3 | 3819 | LS2_CD11:CAV3_D3819 |  | b54 resonator | 40 |  |  | 0.49566 | 361.6722 | 415.3476 | 41.094 | 41.094 | 415.3476 | 361.92 | 381.9274 | 382.1752 |  | 313.4488
0 | [10 empty] | 0 | 361.92 | 415.3476 | 41.094 | 41.094 | 415.3476 | 361.92 | 382.1752 | 382.1752 |  | 313.6966
0 | [10 empty] | 0 | 361.92 | 415.3476 | 41.094 | 41.094 | 415.3476 | 361.92 | 382.1752 | 382.1752 |  | 313.6966
0 | [6 empty] | bellow |  |  |  | 0.099517 | 361.9698 | 415.3476 | 41.094 | 41.094 | 415.3476 | 362.0195 | 382.225 | 382.2747 |  | 313.7464
0 | LS2 | CD11 | CAV4 | 3825 | LS2_CD11:CAV4_D3825 |  | b54 resonator | 40 |  |  | 0.49566 | 362.2674 | 415.3476 | 41.094 | 41.094 | 415.3476 | 362.5152 | 382.5226 | 382.7704 |  | 314.0439
0 | LS2 | CD11 | BLM | 3828 | LS2_CD11:BLM_D3828 |  | outside loss monitor |  |  |  | 0 | 362.5152 | 415.3476 | 41.094 | 41.094 | 415.3476 | 362.5152 | 382.7704 | 382.7704 |  | 314.2918
0 | [10 empty] | 0 | 362.5152 | 415.3476 | 41.094 | 41.094 | 415.3476 | 362.5152 | 382.7704 | 382.7704 |  | 314.2918
0 | [6 empty] | bellow |  |  |  | 0.15525 | 362.5928 | 415.3476 | 41.094 | 41.094 | 415.3476 | 362.6704 | 382.848 | 382.9257 |  | 314.3694
0 | LS2 | CD11 | SOL1 | 3832 | LS2_CD11:SOL1_D3832 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 362.9204 | 415.3476 | 41.094 | 41.094 | 415.3476 | 363.1704 | 383.1757 | 383.4257 |  | 314.697
0 | [6 empty] | bellow |  |  |  | 0.15525 | 363.2481 | 415.3476 | 41.094 | 41.094 | 415.3476 | 363.3257 | 383.5033 | 383.5809 |  | 315.0246
0 | LS2 | CD11 | CAV5 | 3838 | LS2_CD11:CAV5_D3838 |  | b54 resonator | 40 |  |  | 0.49566 | 363.5735 | 415.3476 | 41.094 | 41.094 | 415.3476 | 363.8213 | 383.8287 | 384.0766 |  | 315.3501
0 | [10 empty] | 0 | 363.8213 | 415.3476 | 41.094 | 41.094 | 415.3476 | 363.8213 | 384.0766 | 384.0766 |  | 315.5979
0 | [10 empty] | 0 | 363.8213 | 415.3476 | 41.094 | 41.094 | 415.3476 | 363.8213 | 384.0766 | 384.0766 |  | 315.5979
0 | [6 empty] | bellow |  |  |  | 0.099517 | 363.8711 | 415.3476 | 41.094 | 41.094 | 415.3476 | 363.9209 | 384.1263 | 384.1761 |  | 315.6477
0 | LS2 | CD11 | CAV6 | 3844 | LS2_CD11:CAV6_D3844 |  | b54 resonator | 40 |  |  | 0.49566 | 364.1687 | 415.3476 | 41.094 | 41.094 | 415.3476 | 364.4165 | 384.4239 | 384.6717 |  | 315.9453
0 | [10 empty] | 0 | 364.4165 | 415.3476 | 41.094 | 41.094 | 415.3476 | 364.4165 | 384.6717 | 384.6717 |  | 316.1931
0 | [10 empty] | 0 | 364.4165 | 415.3476 | 41.094 | 41.094 | 415.3476 | 364.4165 | 384.6717 | 384.6717 |  | 316.1931
0 | [6 empty] | bellow |  |  |  | 0.099517 | 364.4663 | 415.3476 | 41.094 | 41.094 | 415.3476 | 364.516 | 384.7215 | 384.7712 |  | 316.2429
0 | LS2 | CD11 | CAV7 | 3850 | LS2_CD11:CAV7_D3850 |  | b54 resonator | 40 |  |  | 0.49566 | 364.7639 | 415.3476 | 41.094 | 41.094 | 415.3476 | 365.0117 | 385.0191 | 385.2669 |  | 316.5404
0 | [10 empty] | 0 | 365.0117 | 415.3476 | 41.094 | 41.094 | 415.3476 | 365.0117 | 385.2669 | 385.2669 |  | 316.7883
0 | [10 empty] | 0 | 365.0117 | 415.3476 | 41.094 | 41.094 | 415.3476 | 365.0117 | 385.2669 | 385.2669 |  | 316.7883
0 | [6 empty] | bellow |  |  |  | 0.099517 | 365.0614 | 415.3476 | 41.094 | 41.094 | 415.3476 | 365.1112 | 385.3167 | 385.3664 |  | 316.838
0 | LS2 | CD11 | CAV8 | 3856 | LS2_CD11:CAV8_D3856 |  | b54 resonator | 40 |  |  | 0.49566 | 365.359 | 415.3476 | 41.094 | 41.094 | 415.3476 | 365.6069 | 385.6143 | 385.8621 |  | 317.1356
0 | LS2 | CD11 | BLM | 3859 | LS2_CD11:BLM_D3859 |  | outside loss monitor |  |  |  | 0 | 365.6069 | 415.3476 | 41.094 | 41.094 | 415.3476 | 365.6069 | 385.8621 | 385.8621 |  | 317.3834
0 | [10 empty] | 0 | 365.6069 | 415.3476 | 41.094 | 41.094 | 415.3476 | 365.6069 | 385.8621 | 385.8621 |  | 317.3834
0 | [10 empty] | 0.1677 | 365.6907 | 415.3476 | 41.094 | 41.094 | 415.3476 | 365.7746 | 385.9459 | 386.0298 |  | 317.4673
0 | LS2 | CD11 | GV | 3861 | LS2_CD11:GV_D3861 |  | gate valve | 40 |  |  | 0.057404 | 365.8033 | 415.3476 | 41.094 | 41.094 | 415.3476 | 365.832 | 386.0585 | 386.0872 |  | 317.5799
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 365.8837 | 415.3476 | 41.094 | 41.094 | 415.3476 | 365.9355 | 386.1389 | 386.1907 |  | 317.6603
0 | LS2 | WD11 | BPM | 3862 | LS2_WD11:BPM_D3862 | BPM2 | position monitor |  |  |  | 0 | 365.9355 | 415.3476 | 41.094 | 41.094 | 415.3476 | 365.9355 | 386.1907 | 386.1907 |  | 317.7121
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 366.0081 | 415.3476 | 41.094 | 41.094 | 415.3476 | 366.0808 | 386.2633 | 386.336 |  | 317.7847
0 | LS2 | WD11 | PORT | 3863 | LS2_WD11:PORT_D3863 |  | port for future diagnostics |  |  |  | 0 | 366.0808 | 415.3476 | 41.094 | 41.094 | 415.3476 | 366.0808 | 386.336 | 386.336 |  | 317.8573
0 | [6 empty] | diagnostic box |  |  |  | 0 | 366.0808 | 415.3476 | 41.094 | 41.094 | 415.3476 | 366.0808 | 386.336 | 386.336 |  | 317.8573
0 | LS2 | WD11 | IP | 3863 | LS2_WD11:IP_D3863 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 366.0808 | 415.3476 | 41.094 | 41.094 | 415.3476 | 366.0808 | 386.336 | 386.336 |  | 317.8573
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 366.1464 | 415.3476 | 41.094 | 41.094 | 415.3476 | 366.212 | 386.4016 | 386.4672 |  | 317.9229
0 | LS2 | CD12 | GV | 3865 | LS2_CD12:GV_D3865 |  | gate valve | 40 |  |  | 0.057404 | 366.2407 | 415.3476 | 41.094 | 41.094 | 415.3476 | 366.2694 | 386.4959 | 386.5246 |  | 318.0173
0 | [6 empty] | bellows |  |  |  | 0.1677 | 366.3532 | 415.3476 | 41.094 | 41.094 | 415.3476 | 366.4371 | 386.6084 | 386.6923 |  | 318.1298
0 | LS2 | CD12 | CAV1 | 3869 | LS2_CD12:CAV1_D3869 |  | b54 resonator | 40 |  |  | 0.49566 | 366.6849 | 415.3476 | 41.094 | 41.094 | 415.3476 | 366.9327 | 386.9401 | 387.188 |  | 318.4615
0 | LS2 | CD12 | BLM | 3872 | LS2_CD12:BLM_D3872 |  | outside loss monitor |  |  |  | 0 | 366.9327 | 415.3476 | 41.094 | 41.094 | 415.3476 | 366.9327 | 387.188 | 387.188 |  | 318.7093
0 | [10 empty] | 0 | 366.9327 | 415.3476 | 41.094 | 41.094 | 415.3476 | 366.9327 | 387.188 | 387.188 |  | 318.7093
0 | [6 empty] | bellow |  |  |  | 0.099517 | 366.9825 | 415.3476 | 41.094 | 41.094 | 415.3476 | 367.0323 | 387.2377 | 387.2875 |  | 318.7591
0 | LS2 | CD12 | CAV2 | 3875 | LS2_CD12:CAV2_D3875 |  | b54 resonator | 40 |  |  | 0.49566 | 367.2801 | 415.3476 | 41.094 | 41.094 | 415.3476 | 367.5279 | 387.5353 | 387.7831 |  | 319.0567
0 | [10 empty] | 0 | 367.5279 | 415.3476 | 41.094 | 41.094 | 415.3476 | 367.5279 | 387.7831 | 387.7831 |  | 319.3045
0 | [10 empty] | 0 | 367.5279 | 415.3476 | 41.094 | 41.094 | 415.3476 | 367.5279 | 387.7831 | 387.7831 |  | 319.3045
0 | [6 empty] | bellow |  |  |  | 0.099517 | 367.5777 | 415.3476 | 41.094 | 41.094 | 415.3476 | 367.6274 | 387.8329 | 387.8826 |  | 319.3543
0 | LS2 | CD12 | CAV3 | 3881 | LS2_CD12:CAV3_D3881 |  | b54 resonator | 40 |  |  | 0.49566 | 367.8753 | 415.3476 | 41.094 | 41.094 | 415.3476 | 368.1231 | 388.1305 | 388.3783 |  | 319.6518
0 | [10 empty] | 0 | 368.1231 | 415.3476 | 41.094 | 41.094 | 415.3476 | 368.1231 | 388.3783 | 388.3783 |  | 319.8997
0 | [10 empty] | 0 | 368.1231 | 415.3476 | 41.094 | 41.094 | 415.3476 | 368.1231 | 388.3783 | 388.3783 |  | 319.8997
0 | [6 empty] | bellow |  |  |  | 0.099517 | 368.1728 | 415.3476 | 41.094 | 41.094 | 415.3476 | 368.2226 | 388.4281 | 388.4778 |  | 319.9494
0 | LS2 | CD12 | CAV4 | 3887 | LS2_CD12:CAV4_D3887 |  | b54 resonator | 40 |  |  | 0.49566 | 368.4704 | 415.3476 | 41.094 | 41.094 | 415.3476 | 368.7183 | 388.7256 | 388.9735 |  | 320.247
0 | LS2 | CD12 | BLM | 3890 | LS2_CD12:BLM_D3890 |  | outside loss monitor |  |  |  | 0 | 368.7183 | 415.3476 | 41.094 | 41.094 | 415.3476 | 368.7183 | 388.9735 | 388.9735 |  | 320.4948
0 | [10 empty] | 0 | 368.7183 | 415.3476 | 41.094 | 41.094 | 415.3476 | 368.7183 | 388.9735 | 388.9735 |  | 320.4948
0 | [6 empty] | bellow |  |  |  | 0.15525 | 368.7959 | 415.3476 | 41.094 | 41.094 | 415.3476 | 368.8735 | 389.0511 | 389.1287 |  | 320.5725
0 | LS2 | CD12 | SOL1 | 3894 | LS2_CD12:SOL1_D3894 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 369.1235 | 415.3476 | 41.094 | 41.094 | 415.3476 | 369.3735 | 389.3787 | 389.6287 |  | 320.9001
0 | [6 empty] | bellow |  |  |  | 0.15525 | 369.4511 | 415.3476 | 41.094 | 41.094 | 415.3476 | 369.5288 | 389.7063 | 389.784 |  | 321.2277
0 | LS2 | CD12 | CAV5 | 3900 | LS2_CD12:CAV5_D3900 |  | b54 resonator | 40 |  |  | 0.49566 | 369.7766 | 415.3476 | 41.094 | 41.094 | 415.3476 | 370.0244 | 390.0318 | 390.2796 |  | 321.5532
0 | [10 empty] | 0 | 370.0244 | 415.3476 | 41.094 | 41.094 | 415.3476 | 370.0244 | 390.2796 | 390.2796 |  | 321.801
0 | [10 empty] | 0 | 370.0244 | 415.3476 | 41.094 | 41.094 | 415.3476 | 370.0244 | 390.2796 | 390.2796 |  | 321.801
0 | [6 empty] | bellow |  |  |  | 0.099517 | 370.0742 | 415.3476 | 41.094 | 41.094 | 415.3476 | 370.1239 | 390.3294 | 390.3791 |  | 321.8508
0 | LS2 | CD12 | CAV6 | 3906 | LS2_CD12:CAV6_D3906 |  | b54 resonator | 40 |  |  | 0.49566 | 370.3718 | 415.3476 | 41.094 | 41.094 | 415.3476 | 370.6196 | 390.627 | 390.8748 |  | 322.1483
0 | [10 empty] | 0 | 370.6196 | 415.3476 | 41.094 | 41.094 | 415.3476 | 370.6196 | 390.8748 | 390.8748 |  | 322.3962
0 | [10 empty] | 0 | 370.6196 | 415.3476 | 41.094 | 41.094 | 415.3476 | 370.6196 | 390.8748 | 390.8748 |  | 322.3962
0 | [6 empty] | bellow |  |  |  | 0.099517 | 370.6693 | 415.3476 | 41.094 | 41.094 | 415.3476 | 370.7191 | 390.9246 | 390.9743 |  | 322.4459
0 | LS2 | CD12 | CAV7 | 3912 | LS2_CD12:CAV7_D3912 |  | b54 resonator | 40 |  |  | 0.49566 | 370.9669 | 415.3476 | 41.094 | 41.094 | 415.3476 | 371.2148 | 391.2221 | 391.47 |  | 322.7435
0 | [10 empty] | 0 | 371.2148 | 415.3476 | 41.094 | 41.094 | 415.3476 | 371.2148 | 391.47 | 391.47 |  | 322.9913
0 | [10 empty] | 0 | 371.2148 | 415.3476 | 41.094 | 41.094 | 415.3476 | 371.2148 | 391.47 | 391.47 |  | 322.9913
0 | [6 empty] | bellow |  |  |  | 0.099517 | 371.2645 | 415.3476 | 41.094 | 41.094 | 415.3476 | 371.3143 | 391.5197 | 391.5695 |  | 323.0411
0 | LS2 | CD12 | CAV8 | 3918 | LS2_CD12:CAV8_D3918 |  | b54 resonator | 40 |  |  | 0.49566 | 371.5621 | 415.3476 | 41.094 | 41.094 | 415.3476 | 371.8099 | 391.8173 | 392.0651 |  | 323.3387
0 | LS2 | CD12 | BLM | 3921 | LS2_CD12:BLM_D3921 |  | outside loss monitor |  |  |  | 0 | 371.8099 | 415.3476 | 41.094 | 41.094 | 415.3476 | 371.8099 | 392.0651 | 392.0651 |  | 323.5865
0 | [10 empty] | 0 | 371.8099 | 415.3476 | 41.094 | 41.094 | 415.3476 | 371.8099 | 392.0651 | 392.0651 |  | 323.5865
0 | [10 empty] | 0.1677 | 371.8938 | 415.3476 | 41.094 | 41.094 | 415.3476 | 371.9776 | 392.149 | 392.2328 |  | 323.6704
0 | LS2 | CD12 | GV | 3923 | LS2_CD12:GV_D3923 |  | gate valve | 40 |  |  | 0.057404 | 372.0063 | 415.3476 | 41.094 | 41.094 | 415.3476 | 372.035 | 392.2616 | 392.2903 |  | 323.7829
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 372.0868 | 415.3476 | 41.094 | 41.094 | 415.3476 | 372.1385 | 392.342 | 392.3938 |  | 323.8634
0 | LS2 | WD12 | BPM | 3924 | LS2_WD12:BPM_D3924 | BPM2 | position monitor |  |  |  | 0 | 372.1385 | 415.3476 | 41.094 | 41.094 | 415.3476 | 372.1385 | 392.3938 | 392.3938 |  | 323.9151
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 372.2112 | 415.3476 | 41.094 | 41.094 | 415.3476 | 372.2838 | 392.4664 | 392.539 |  | 323.9878
0 | LS2 | WD12 | PORT | 3925 | LS2_WD12:PORT_D3925 |  | port for future diagnostics |  |  |  | 0 | 372.2838 | 415.3476 | 41.094 | 41.094 | 415.3476 | 372.2838 | 392.539 | 392.539 |  | 324.0604
0 | [6 empty] | diagnostic box |  |  |  | 0 | 372.2838 | 415.3476 | 41.094 | 41.094 | 415.3476 | 372.2838 | 392.539 | 392.539 |  | 324.0604
0 | LS2 | WD12 | IP | 3925 | LS2_WD12:IP_D3925 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 372.2838 | 415.3476 | 41.094 | 41.094 | 415.3476 | 372.2838 | 392.539 | 392.539 |  | 324.0604
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 372.3494 | 415.3476 | 41.094 | 41.094 | 415.3476 | 372.415 | 392.6046 | 392.6703 |  | 324.126
0 | SEGMENT 2 END (12 b29 CMs +12 b54 CMs) | [9 empty] | 0 | 372.415 | 415.3476 | 41.094 | 41.094 | 415.3476 | 372.415 | 392.6703 | 392.6703 |  | 324.1916
0 | [8 empty] | 0.003903 |  | 0.037404 | 372.4337 | 415.3476 | 41.094 | 41.094 | 415.3476 | 372.4524 | 392.689 | 392.7077 |  | 324.2103
0 | FS2 | BTS | FAV | 3927 | FS2_BTS:FAV_D3927 | FV | fast valve | 40 | 2.176e-06 | 372.4524 | 0 | 372.4524 | 415.3476 | 41.094 | 41.094 | 415.3476 | 372.4524 | 392.7077 | 392.7077 | 0.067196 | 324.229
0 | [10 empty] | 0.1626 | 372.5337 | 415.3476 | 41.094 | 41.094 | 415.3476 | 372.615 | 392.789 | 392.8703 |  | 324.3103
0 | FS2 | BTS | DCHV | 3930 | FS2_BTS:DCHV_D3930 | C2 | correctors-X,Y | 50 |  | 372.715 | 0.2 | 372.715 | 415.3476 | 41.094 | 41.094 | 415.3476 | 372.815 | 392.9703 | 393.0703 |  | 324.4916
0 | [10 empty] | 0.2 | 372.915 | 415.3476 | 41.094 | 41.094 | 415.3476 | 373.015 | 393.1703 | 393.2703 |  | 324.6916
0 | FS2 | BTS | QH | 3934 | FS2_BTS:QH_D3934 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 373.14 | 415.3476 | 41.094 | 41.094 | 415.3476 | 373.265 | 393.3953 | 393.5203 |  | 324.9166
0 | [6 empty] | coil out + leads |  |  |  | 0.098 | 373.314 | 415.3476 | 41.094 | 41.094 | 415.3476 | 373.363 | 393.5693 | 393.6183 |  | 325.0906
0 | [8 empty] | 3.176e-06 |  | 0.001485 | 373.3638 | 415.3476 | 41.094 | 41.094 | 415.3476 | 373.3645 | 393.619 | 393.6197 |  | 325.1404
0 | FS2 | BTS | BCM | 3936 | FS2_BTS:BCM_D3936 | BCM3 | current monitor |  | 373.3645 | 373.3645 | 0 | 373.3645 | 415.3476 | 41.094 | 41.094 | 415.3476 | 373.3645 | 393.6197 | 393.6197 | 0.1 | 325.1411
0 | [10 empty] | 0.17852 | 373.4538 | 415.3476 | 41.094 | 41.094 | 415.3476 | 373.543 | 393.709 | 393.7983 |  | 325.2304
0 | [6 empty] | coil out |  |  |  | 0.047 | 373.5665 | 415.3476 | 41.094 | 41.094 | 415.3476 | 373.59 | 393.8218 | 393.8453 |  | 325.3431
0 | FS2 | BTS | QV | 3940 | FS2_BTS:QV_D3940 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 373.715 | 415.3476 | 41.094 | 41.094 | 415.3476 | 373.84 | 393.9703 | 394.0953 |  | 325.4916
0 | [8 empty] | 3.176e-06 |  | 0.21871 | 373.9494 | 415.3476 | 41.094 | 41.094 | 415.3476 | 374.0587 | 394.2046 | 394.314 |  | 325.726
0 | FS2 | BTS | BPM | 3943 | FS2_BTS:BPM_D3943 | BPM3 | position monitor | 47.5 | 374.0587 | 374.0587 | 0 | 374.0587 | 415.3476 | 41.094 | 41.094 | 415.3476 | 374.0587 | 394.314 | 394.314 | 0.137 | 325.8353
0 | [10 empty] | 0.081294 | 374.0994 | 415.3476 | 41.094 | 41.094 | 415.3476 | 374.14 | 394.3546 | 394.3953 |  | 325.876
0 | FS2 | BTS | DCHV | 3945 | FS2_BTS:DCHV_D3945 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 374.24 | 415.3476 | 41.094 | 41.094 | 415.3476 | 374.34 | 394.4953 | 394.5953 |  | 326.0166
0 | [10 empty] | 0 | 374.34 | 415.3476 | 41.094 | 41.094 | 415.3476 | 374.34 | 394.5953 | 394.5953 |  | 326.1166
0 | FS2 | BTS | BPM | 3946 | [6 empty] | 0 | 374.34 | 415.3476 | 41.094 | 41.094 | 415.3476 | 374.34 | 394.5953 | 394.5953 | 20170214 BPM moved upstream | 326.1166
0 | [10 empty] | 0.25 | 374.465 | 415.3476 | 41.094 | 41.094 | 415.3476 | 374.59 | 394.7203 | 394.8453 |  | 326.2416
0 | FS2 | BTS | QV | 3950 | FS2_BTS:QV_D3950 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 374.715 | 415.3476 | 41.094 | 41.094 | 415.3476 | 374.84 | 394.9703 | 395.0953 |  | 326.4916
0 | [10 empty] | 0 | 374.84 | 415.3476 | 41.094 | 41.094 | 415.3476 | 374.84 | 395.0953 | 395.0953 |  | 326.6166
0 | [10 empty] | 0.25 | 374.965 | 415.3476 | 41.094 | 41.094 | 415.3476 | 375.09 | 395.2203 | 395.3453 |  | 326.7416
0 | [10 empty] | 0 | 375.09 | 415.3476 | 41.094 | 41.094 | 415.3476 | 375.09 | 395.3453 | 395.3453 |  | 326.8666
0 | FS2 | BTS | QH | 3955 | FS2_BTS:QH_D3955 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 375.215 | 415.3476 | 41.094 | 41.094 | 415.3476 | 375.34 | 395.4703 | 395.5953 |  | 326.9916
0 | [10 empty] | 0.062353 | 375.3712 | 415.3476 | 41.094 | 41.094 | 415.3476 | 375.4024 | 395.6264 | 395.6576 |  | 327.1478
0 | [6 empty] | bellow |  | -0.0016469 |  | 0.1035 | 375.4541 | 415.3476 | 41.094 | 41.094 | 415.3476 | 375.5059 | 395.7094 | 395.7611 | 20170427 adjusted based on ME | 327.2307
0 | FS2 | BTS | BPM | 3958 | FS2_BTS:BPM_D3958 | BPM3 | position monitor |  | 2.03e-06 | 375.5059 | 0 | 375.5059 | 415.3476 | 41.094 | 41.094 | 415.3476 | 375.5059 | 395.7611 | 395.7611 |  | 327.2825
0 | [6 empty] | diagnostic box |  |  |  | 0.1718 | 375.5918 | 415.3476 | 41.094 | 41.094 | 415.3476 | 375.6777 | 395.847 | 395.9329 |  | 327.3684
0 | FS2 | BTS | PM | 3959 | FS2_BTS:PM_D3959 | PM3 | profile monitor |  |  |  | 0 | 375.6777 | 415.3476 | 41.094 | 41.094 | 415.3476 | 375.6777 | 395.9329 | 395.9329 |  | 327.4543
0 | [6 empty] | diagnostic box |  | 0.026517 |  | 0 | 375.6777 | 415.3476 | 41.094 | 41.094 | 415.3476 | 375.6777 | 395.9329 | 395.9329 |  | 327.4543
0 | FS2 | BTS | FAVS | 3959 | FS2_BTS:FAVS_D3959 | FVS | fast valve sensor |  | 2.03e-06 | 375.6777 | 0 | 375.6777 | 415.3476 | 41.094 | 41.094 | 415.3476 | 375.6777 | 395.9329 | 395.9329 |  | 327.4543
0 | FS2 | BTS | IP | 3959 | FS2_BTS:IP_D3959 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 375.6777 | 415.3476 | 41.094 | 41.094 | 415.3476 | 375.6777 | 395.9329 | 395.9329 |  | 327.4543
0 | [6 empty] | bellow |  |  |  | 0.13122 | 375.7433 | 415.3476 | 41.094 | 41.094 | 415.3476 | 375.8089 | 395.9985 | 396.0641 |  | 327.5199
0 | [8 empty] | 3.176e-06 |  | 0.069773 | 375.8438 | 415.3476 | 41.094 | 41.094 | 415.3476 | 375.8787 | 396.099 | 396.1339 | 20170427 adjusted based on ME | 327.6204
0 | FS2 | BTS | DCHV | 3962 | FS2_BTS:DCHV_D3962 | C2 | correctors-X,Y | 50 | 375.9787 | 375.9787 | 0.2 | 375.9787 | 415.3476 | 41.094 | 41.094 | 415.3476 | 376.0787 | 396.2339 | 396.3339 |  | 327.7553
0 | [8 empty] | 0.039879 |  | 0.11309 | 376.1352 | 415.3476 | 41.094 | 41.094 | 415.3476 | 376.1918 | 396.3904 | 396.447 |  | 327.9118
0 | FS2 | BTS | GV | 3965 | FS2_BTS:GV_D3965 | GV_M | gate valve | 63 | 2.176e-06 | 376.2293 | 0.075 | 376.2293 | 415.3476 | 41.094 | 41.094 | 415.3476 | 376.2668 | 396.4845 | 396.522 | 0.075 | 328.0059
0 | Third beam dump line starts in "V" column | [9 empty] | 0.20827 | 376.3709 | 415.3476 | 41.094 | 41.094 | 415.3476 | 376.475 | 396.6261 | 396.7303 |  | 328.1475
0 | FS2 | BBS | DH | 3979 | FS2_BBS:DH_D3979 | D3 | dipole-45deg, D3-1 | 35 x 100 |  |  | 2.3375 | 377.5273 | 415.7834 | 41.094 | 41.094 | 416.2193 | 378.5795 | 397.899 | 399.0678 |  | 329.4204
0 | [10 empty] | 0.4 | 378.7209 | 416.3607 | 41.094 | 41.094 | 416.5021 | 378.8624 | 399.2678 | 399.4678 |  | 330.7891
0 | FS2 | BBS | QH | 3996 | FS2_BBS:QH_D3996 | Q4 | quadrupole-X,33cm,B2 | 100 |  |  | 0.33 | 378.979 | 416.6188 | 41.094 | 41.094 | 416.7354 | 379.0957 | 399.6328 | 399.7978 |  | 331.1541
0 | [10 empty] | 0.12 | 379.1381 | 416.7779 | 41.094 | 41.094 | 416.8203 | 379.1806 | 399.8578 | 399.9178 |  | 331.3791
0 | FS2 | BBS | S | 4000 | FS2_BBS:S_D4000 | H1 | hexapole-X+ | 100 |  |  | 0.1 | 379.2159 | 416.8557 | 41.094 | 41.094 | 416.891 | 379.2513 | 399.9678 | 400.0178 |  | 331.4891
0 | [10 empty] | 0.15 | 379.3043 | 416.944 | 41.094 | 41.094 | 416.9971 | 379.3573 | 400.0928 | 400.1678 |  | 331.6141
0 | FS2 | BBS | QV | 4004 | FS2_BBS:QV_D4004 | Q5 | quadrupole-Y,40cm,B2 | 100 | 0 | 0.2 | 0.4 | 379.4988 | 417.1385 | 41.094 | 41.094 | 417.2799 | 379.6402 | 400.3678 | 400.5678 |  | 331.8891
0 | [8 empty] | 0.36323 | 0.11323 | 0.11323 | 379.6802 | 417.32 | 41.094 | 41.094 | 417.36 | 379.7203 | 400.6244 | 400.681 | 20170216 .13 -> .11323 | 332.1457
0 | FS2 | BBS | S | 4007 | FS2_BBS:S_D4007 | H1 | hexapole-X- | 100 |  | 0.05 | 0.1 | 379.7556 | 417.3953 | 41.094 | 41.094 | 417.4307 | 379.791 | 400.731 | 400.781 |  | 332.2524
0 | [8 empty] | 1e-06 |  | 0.095715 | 379.8248 | 417.4645 | 41.094 | 41.094 | 417.4984 | 379.8586 | 400.8288 | 400.8767 | 20170216 .07 -> .09572 | 332.3502
0 | [8 empty] | 0.24572 | 0.095716 | 0 | 379.8586 | 417.4984 | 41.094 | 41.094 | 417.4984 | 379.8586 | 400.8767 | 400.8767 |  | 332.3981
0 | FS2 | BBS | DCHV | 4010 | FS2_BBS:DCHV_D4010 | C3 | correctors-X,Y | 100 | 0 | 0.1 | 0.2 | 379.9294 | 417.5691 | 41.094 | 41.094 | 417.6398 | 380.0001 | 400.9767 | 401.0767 |  | 332.4981
0 | [8 empty] | 0.37605 | 0.11105 | 0.11105 | 380.0393 | 417.6791 | 41.094 | 41.094 | 417.7183 | 380.0786 | 401.1322 | 401.1878 | 20170216 .12 -> .111055 | 332.6536
0 | FS2 | BBS | QH | 4014 | FS2_BBS:QH_D4014 | Q4 | quadrupole-X,33cm,B2 | 100 | 0 | 0.165 | 0.33 | 380.1953 | 417.835 | 41.094 | 41.094 | 417.9517 | 380.3119 | 401.3528 | 401.5178 |  | 332.8741
0 | [8 empty] | 0.40344 | 0.23844 | 0.23844 | 380.3962 | 418.036 | 41.094 | 41.094 | 418.1203 | 380.4805 | 401.637 | 401.7562 | 20170216 .157931 -> .23844 | 333.1583
0 | FS2 | BBS | IP | 4018 | FS2_BBS:IP_D4018 | IP_L | ion pump 300 l/s |  | 2.5e-07 | 0 | 0 | 380.4805 | 418.1203 | 41.094 | 41.094 | 418.1203 | 380.4805 | 401.7562 | 401.7562 |  | 333.2776
0 | [8 empty] | 0.19262 | 0.15055 | 0.15055 | 380.5338 | 418.1735 | 41.094 | 41.094 | 418.2267 | 380.587 | 401.8315 | 401.9067 | 20170216 .0 -> .150549 | 333.3528
0 | FS2 | BBS | BPM | 4019 | FS2_BBS:BPM_D4019 | BPM4 | position monitor R=50mm | 100 |  | 0.042069 | 0.084138 | 380.6167 | 418.2565 | 41.094 | 41.094 | 418.2862 | 380.6465 | 401.9488 | 401.9909 | 20170216 BPM moved after box | 333.4702
0 | [10 empty] | 0.2548 | 380.7366 | 418.3763 | 41.094 | 41.094 | 418.4664 | 380.8267 | 402.1183 | 402.2457 | 20170216 .485861 -> .254803 | 333.6397
0 | FS2 | BBS | DH | 4034 | FS2_BBS:DH_D4034 | D3 | dipole-45deg, D3-2 | 35 x 100 |  |  | 2.3375 | 381.2625 | 419.5186 | 41.094 | 41.094 | 420.5709 | 381.6984 | 403.4144 | 404.5832 |  | 334.9358
0 | [10 empty] | 0.50759 | 381.6984 | 420.8247 | 41.094 | 41.094 | 421.0785 | 381.6984 | 404.837 | 405.0908 | 20170216 space adjusted | 336.3583
0 | FS2 | BBS | PM | 4051 | FS2_BBS:PM_D4051 | PM5 | profile monitor | 100 |  |  | 0 | 381.6984 | 421.0785 | 41.094 | 41.094 | 421.0785 | 381.6984 | 405.0908 | 405.0908 |  | 336.6121
0 | FS2 | BBS | IP | 4051 | FS2_BBS:IP_D4051 | IP_L | ion pump 300 l/s |  |  |  | 0 | 381.6984 | 421.0785 | 41.094 | 41.094 | 421.0785 | 381.6984 | 405.0908 | 405.0908 |  | 336.6121
0 | FS2 | BBS | FAVS | 4051 | FS2_BBS:FAVS_D4051 | FVS | fast valve sensor |  |  |  | 0 | 381.6984 | 421.0785 | 41.094 | 41.094 | 421.0785 | 381.6984 | 405.0908 | 405.0908 |  | 336.6121
0 | [10 empty] | 0.21854 | 381.6984 | 421.1877 | 41.094 | 41.094 | 421.297 | 381.6984 | 405.2 | 405.3093 | 20170216 space adjusted | 336.7214
0 | FS2 | BBS | BPM | 4054 | FS2_BBS:BPM_D4054 | BPM4 | position monitor R=50mm | 100 |  |  | 0.084138 | 381.6984 | 421.3391 | 41.094 | 41.094 | 421.3812 | 381.6984 | 405.3514 | 405.3934 | 20170216 BPM moved after box | 336.8728
0 | [10 empty] | 0.021733 | 381.6984 | 421.392 | 41.094 | 41.094 | 421.4029 | 381.6984 | 405.4043 | 405.4152 | 20170216 space adjusted | 336.9257
0 | FS2 | BBS | DCHV | 4055 | FS2_BBS:DCHV_D4055 | C3 | correctors-X,Y | 100 |  |  | 0.2 | 381.6984 | 421.5029 | 41.094 | 41.094 | 421.6029 | 381.6984 | 405.5152 | 405.6152 |  | 337.0366
0 | [10 empty] | 0.368 | 381.6984 | 421.7869 | 41.094 | 41.094 | 421.9709 | 381.6984 | 405.7992 | 405.9832 |  | 337.3206
0 | FS2 | BBS | DH | 4072 | FS2_BBS:DH_D4072 | D3 | dipole-45deg, D3-3 | 35 x 100 |  |  | 2.3375 | 381.2625 | 423.0231 | 41.094 | 41.094 | 424.0754 | 380.8267 | 407.1519 | 408.3207 |  | 338.6733
0 | [10 empty] | 0.32554 | 380.7116 | 424.1905 | 41.094 | 41.094 | 424.3056 | 380.5965 | 408.4834 | 408.6462 | 20170216 this space NOT symmetric | 340.0048
0 | FS2 | BBS | BPM | 4087 | FS2_BBS:BPM_D4087 | BPM4 | position monitor R=50mm | 100 |  |  | 0.084138 | 380.5667 | 424.3353 | 41.094 | 41.094 | 424.3651 | 380.537 | 408.6883 | 408.7304 |  | 340.2097
0 | [10 empty] | 0.15055 | 380.4838 | 424.4183 | 41.094 | 41.094 | 424.4715 | 380.4305 | 408.8056 | 408.8809 |  | 340.327
0 | FS2 | BBS | IP | 4089 | FS2_BBS:IP_D4089 | IP_L | ion pump 300 l/s |  | 0 | 0 | 0 | 380.4305 | 424.4715 | 41.094 | 41.094 | 424.4715 | 380.4305 | 408.8809 | 408.8809 |  | 340.4023
0 | [8 empty] | 0.33271 | 0.16771 | 0.16771 | 380.3712 | 424.5308 | 41.094 | 41.094 | 424.5901 | 380.3119 | 408.9648 | 409.0486 | 20170216 this space NOT symmetric | 340.4861
0 | FS2 | BBS | QH | 4092 | FS2_BBS:QH_D4092 | Q4 | quadrupole-X,33cm,B2 | 100 |  | 0.165 | 0.33 | 380.1953 | 424.7068 | 41.094 | 41.094 | 424.8234 | 380.0786 | 409.2136 | 409.3786 |  | 340.735
0 | [10 empty] | 0.11105 | 380.0393 | 424.8627 | 41.094 | 41.094 | 424.902 | 380.0001 | 409.4341 | 409.4897 |  | 340.9555
0 | FS2 | BBS | DCHV | 4096 | FS2_BBS:DCHV_D4096 | C3 | correctors-X,Y | 100 |  |  | 0.2 | 379.9294 | 424.9727 | 41.094 | 41.094 | 425.0434 | 379.8586 | 409.5897 | 409.6897 |  | 341.111
0 | [10 empty] | 0 | 379.8586 | 425.0434 | 41.094 | 41.094 | 425.0434 | 379.8586 | 409.6897 | 409.6897 |  | 341.211
0 | [10 empty] | 0.095715 | 379.8248 | 425.0772 | 41.094 | 41.094 | 425.1111 | 379.791 | 409.7375 | 409.7854 |  | 341.2589
0 | FS2 | BBS | S | 4098 | FS2_BBS:S_D4098 | H1 | hexapole-X- | 100 |  |  | 0.1 | 379.7556 | 425.1464 | 41.094 | 41.094 | 425.1818 | 379.7203 | 409.8354 | 409.8854 |  | 341.3568
0 | [10 empty] | 0.11323 | 379.6802 | 425.2218 | 41.094 | 41.094 | 425.2619 | 379.6402 | 409.942 | 409.9986 |  | 341.4634
0 | FS2 | BBS | QV | 4102 | FS2_BBS:QV_D4102 | Q5 | quadrupole-Y,40cm,B2 | 100 |  |  | 0.4 | 379.4988 | 425.4033 | 41.094 | 41.094 | 425.5447 | 379.3573 | 410.1986 | 410.3986 |  | 341.72
0 | [10 empty] | 0.15 | 379.3043 | 425.5977 | 41.094 | 41.094 | 425.6508 | 379.2513 | 410.4736 | 410.5486 |  | 341.995
0 | FS2 | BBS | S | 4106 | FS2_BBS:S_D4106 | H1 | hexapole-X+ | 100 |  |  | 0.1 | 379.2159 | 425.6861 | 41.094 | 41.094 | 425.7215 | 379.1806 | 410.5986 | 410.6486 |  | 342.12
0 | [10 empty] | 0.12 | 379.1381 | 425.7639 | 41.094 | 41.094 | 425.8063 | 379.0957 | 410.7086 | 410.7686 |  | 342.23
0 | FS2 | BBS | QH | 4109 | FS2_BBS:QH_D4109 | Q4 | quadrupole-X,33cm,B2 | 100 |  |  | 0.33 | 378.979 | 425.923 | 41.094 | 41.094 | 426.0397 | 378.8624 | 410.9336 | 411.0986 |  | 342.455
0 | [10 empty] | 0.4 | 378.7209 | 426.1811 | 41.094 | 41.094 | 426.3225 | 378.5795 | 411.2986 | 411.4986 |  | 342.82
0 | FS2 | BBS | DH | 4127 | FS2_BBS:DH_D4127 | D3 | dipole-45deg, D3-4 | 35 x 100 |  |  | 2.3375 | 377.5273 | 426.7584 | 41.094 | 41.094 | 427.1942 | 376.475 | 412.6674 | 413.8361 |  | 344.1887
0 | [7 empty] | -0.17121 | 2.176e-06 |  | 0.20827 | 376.3709 | 427.1942 | 41.094 | 41.094 | 427.1942 | 376.2668 | 413.9402 | 414.0444 | 20170214 more space needed | 345.4616
0 | FS2 | BMS | GV | 4141 | FS2_BMS:GV_D4141 | GV_M | gate valve | 63 | 376.2293 | 376.2293 | 0.075 | 376.2293 | 427.1942 | 41.094 | 41.094 | 427.1942 | 376.1918 | 414.0819 | 414.1194 |  | 345.6032
0 | [6 empty] | bellow | 40 | 2.176e-06 |  | 0.094701 | 376.1444 | 427.1942 | 41.094 | 41.094 | 427.1942 | 376.0971 | 414.1667 | 414.2141 |  | 345.6881
0 | FS2 | BMS | BPM | 4142 | FS2_BMS:BPM_D4142 | BPM3 | position monitor |  | 376.0971 | 376.0971 | 0 | 376.0971 | 427.1942 | 41.094 | 41.094 | 427.1942 | 376.0971 | 414.2141 | 414.2141 |  | 345.7354
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 376.0244 | 427.1942 | 41.094 | 41.094 | 427.1942 | 375.9518 | 414.2867 | 414.3594 |  | 345.8081
0 | FS2 | BMS | PM | 4144 | FS2_BMS:PM_D4144 | PM3 | profile monitor |  | 2.176e-06 |  | 0 | 375.9518 | 427.1942 | 41.094 | 41.094 | 427.1942 | 375.9518 | 414.3594 | 414.3594 |  | 345.8807
0 | [6 empty] | diagnostic box |  | 375.9518 | 375.9518 | 0 | 375.9518 | 427.1942 | 41.094 | 41.094 | 427.1942 | 375.9518 | 414.3594 | 414.3594 |  | 345.8807
0 | FS2 | BMS | FAVS | 4144 | FS2_BMS:FAVS_D4144 | FVS | fast valve sensor |  |  |  | 0 | 375.9518 | 427.1942 | 41.094 | 41.094 | 427.1942 | 375.9518 | 414.3594 | 414.3594 |  | 345.8807
0 | FS2 | BMS | IP | 4144 | FS2_BMS:IP_D4144 | IP_S | ion pump, 75 l/s |  | -8.8818e-15 |  | 0 | 375.9518 | 427.1942 | 41.094 | 41.094 | 427.1942 | 375.9518 | 414.3594 | 414.3594 |  | 345.8807
0 | [6 empty] | bellow | 40 | 1.135 |  | 0.099 | 375.9023 | 427.1942 | 41.094 | 41.094 | 427.1942 | 375.8528 | 414.4089 | 414.4584 |  | 345.9302
0 | [8 empty] | 1.135 |  | 0.087751 | 375.8089 | 427.1942 | 41.094 | 41.094 | 427.1942 | 375.765 | 414.5022 | 414.5461 |  | 346.0236
0 | FS2 | BMS | DCHV | 4146 | FS2_BMS:DCHV_D4146 | C2 | correctors-X,Y | 50 |  | 375.665 | 0.2 | 375.665 | 427.1942 | 41.094 | 41.094 | 427.1942 | 375.565 | 414.6461 | 414.7461 |  | 346.1675
0 | [10 empty] | 0.225 | 375.4525 | 427.1942 | 41.094 | 41.094 | 427.1942 | 375.34 | 414.8586 | 414.9711 | 20170214 0.2 --> 0.225 per ME | 346.38
0 | FS2 | BMS | QH | 4151 | FS2_BMS:QH_D4151 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 375.215 | 427.1942 | 41.094 | 41.094 | 427.1942 | 375.09 | 415.0961 | 415.2211 |  | 346.6175
0 | [10 empty] | 0.25 | 374.965 | 427.1942 | 41.094 | 41.094 | 427.1942 | 374.84 | 415.3461 | 415.4711 |  | 346.8675
0 | FS2 | BMS | QV | 4156 | FS2_BMS:QV_D4156 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 374.715 | 427.1942 | 41.094 | 41.094 | 427.1942 | 374.59 | 415.5961 | 415.7211 |  | 347.1175
0 | [10 empty] | 0.3 | 374.44 | 427.1942 | 41.094 | 41.094 | 427.1942 | 374.29 | 415.8711 | 416.0211 |  | 347.3925
0 | FS2 | BMS | DCHV | 4161 | FS2_BMS:DCHV_D4161 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 374.19 | 427.1942 | 41.094 | 41.094 | 427.1942 | 374.09 | 416.1211 | 416.2211 |  | 347.6425
0 | [8 empty] | 3.176e-06 |  | 0.19316 | 373.9935 | 427.1942 | 41.094 | 41.094 | 427.1942 | 373.8969 | 416.3177 | 416.4143 |  | 347.8391
0 | FS2 | BMS | BPM | 4164 | FS2_BMS:BPM_D4164 | BPM3 | position monitor | 47.5 | 373.8969 | 373.8969 | 0 | 373.8969 | 427.1942 | 41.094 | 41.094 | 427.1942 | 373.8969 | 416.4143 | 416.4143 | 0.137 | 347.9356
0 | [10 empty] | 0.056842 | 373.8685 | 427.1942 | 41.094 | 41.094 | 427.1942 | 373.84 | 416.4427 | 416.4711 |  | 347.9641
0 | FS2 | BMS | QV | 4166 | FS2_BMS:QV_D4166 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 373.715 | 427.1942 | 41.094 | 41.094 | 427.1942 | 373.59 | 416.5961 | 416.7211 |  | 348.1175
0 | [8 empty] | 3.176e-06 |  | 0.18331 | 373.4984 | 427.1942 | 41.094 | 41.094 | 427.1942 | 373.4067 | 416.8128 | 416.9044 |  | 348.3341
0 | FS2 | BBS | BCM | 4169 | FS2_BBS:BCM_D4169 | BCM3 | current monitor |  | 373.4067 | 373.4067 | 0 | 373.4067 | 427.1942 | 41.094 | 41.094 | 427.1942 | 373.4067 | 416.9044 | 416.9044 | 0.1 | 348.4258
0 | [10 empty] | 0.14169 | 373.3359 | 427.1942 | 41.094 | 41.094 | 427.1942 | 373.265 | 416.9753 | 417.0461 |  | 348.4966
0 | FS2 | BMS | QH | 4172 | FS2_BMS:QH_D4172 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 373.14 | 427.1942 | 41.094 | 41.094 | 427.1942 | 373.015 | 417.1711 | 417.2961 |  | 348.6925
0 | [10 empty] | 0.13938 | 372.9453 | 427.1942 | 41.094 | 41.094 | 427.1942 | 372.8757 | 417.3658 | 417.4355 |  | 348.8872
0 | [10 empty] | 0.1 | 372.8257 | 427.1942 | 41.094 | 41.094 | 427.1942 | 372.7757 | 417.4855 | 417.5355 |  | 349.0069
0 | [8 empty] | 3.176e-06 |  | 0 | 372.7757 | 427.1942 | 41.094 | 41.094 | 427.1942 | 372.7757 | 417.5355 | 417.5355 |  | 349.0569
0 | FS2 | BMS | FAV | 4175 | FS2_BMS:FAV_D4175 | FV | fast valve | 40 | 372.7757 | 372.7757 | 0 | 372.7757 | 427.1942 | 41.094 | 41.094 | 427.1942 | 372.7757 | 417.5355 | 417.5355 | 0.067196 | 349.0569
0 | [10 empty] | 0.037495 | 372.7569 | 427.1942 | 41.094 | 41.094 | 427.1942 | 372.7382 | 417.5542 | 417.573 |  | 349.0756
0 | [6 empty] | bellow | 40 | -0.0044369 |  | 0.1035 | 372.6864 | 427.1942 | 41.094 | 41.094 | 427.1942 | 372.6347 | 417.6247 | 417.6765 |  | 349.1461
0 | FS2 | BMS | BPM | 4177 | FS2_BMS:BPM_D4177 | BPM2 | position monitor |  | 2.03e-06 | 372.6347 | 0 | 372.6347 | 427.1942 | 41.094 | 41.094 | 427.1942 | 372.6347 | 417.6765 | 417.6765 |  | 349.1979
0 | [6 empty] | diagnostic box | 0.11862 |  |  | 0.14528 | 372.562 | 427.1942 | 41.094 | 41.094 | 427.1942 | 372.4894 | 417.7491 | 417.8218 |  | 349.2705
0 | FS2 | BMS | PM | 4178 | FS2_BMS:PM_D4178 | PMx | profile monitor (not costed) | future PM |  |  | 0 | 372.4894 | 427.1942 | 41.094 | 41.094 | 427.1942 | 372.4894 | 417.8218 | 417.8218 |  | 349.3431
0 | [6 empty] | diagnostic box |  | 2.03e-06 | 372.4894 | 0 | 372.4894 | 427.1942 | 41.094 | 41.094 | 427.1942 | 372.4894 | 417.8218 | 417.8218 |  | 349.3431
0 | FS2 | BMS | IP | 4178 | FS2_BMS:IP_D4178 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 372.4894 | 427.1942 | 41.094 | 41.094 | 427.1942 | 372.4894 | 417.8218 | 417.8218 |  | 349.3431
0 | [10 empty] | 0 | 372.4894 | 427.1942 | 41.094 | 41.094 | 427.1942 | 372.4894 | 417.8218 | 417.8218 |  | 349.3431
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 372.4238 | 427.1942 | 41.094 | 41.094 | 427.1942 | 372.3582 | 417.8874 | 417.953 |  | 349.4087
0 | [9 empty] | ??? | 0.0086001 | 372.3539 | 427.1942 | 41.094 | 41.094 | 427.1942 | 372.3496 | 417.9573 | 417.9616 | 20120312: why this odd number? | 349.4787
0 | FS2 | CG01 | GV | 4180 | FS2_CG01:GV_D4180 |  | gate valve | 40 |  |  | 0.057404 | 372.3209 | 427.1942 | 41.094 | 41.094 | 427.1942 | 372.2922 | 417.9903 | 418.019 |  | 349.5117
0 | [6 empty] | bellows |  |  |  | 0.17589 | 372.2042 | 427.1942 | 41.094 | 41.094 | 427.1942 | 372.1163 | 418.1069 | 418.1949 |  | 349.6283
0 | FS2 | CG01 | CAV1 | 4184 | FS2_CG01:CAV1_D4184 |  | b54 resonator | 40 |  |  | 0.49566 | 371.8684 | 427.1942 | 41.094 | 41.094 | 427.1942 | 371.6206 | 418.4427 | 418.6905 |  | 349.9641
0 | [6 empty] | bellow |  |  |  | 0.099517 | 371.5709 | 427.1942 | 41.094 | 41.094 | 427.1942 | 371.5211 | 418.7403 | 418.7901 |  | 350.2617
0 | FS2 | CG01 | CAV2 | 4190 | FS2_CG01:CAV2_D4190 |  | b54 resonator | 40 |  |  | 0.49566 | 371.2733 | 427.1942 | 41.094 | 41.094 | 427.1942 | 371.0254 | 419.0379 | 419.2857 |  | 350.5593
0 | [6 empty] | bellow |  |  |  | 0.099517 | 370.9757 | 427.1942 | 41.094 | 41.094 | 427.1942 | 370.9259 | 419.3355 | 419.3852 |  | 350.8568
0 | [10 empty] | 0.20048 | 370.8257 | 427.1942 | 41.094 | 41.094 | 427.1942 | 370.7254 | 419.4855 | 419.5857 |  | 351.0068
0 | [6 empty] | bellow |  |  |  | 0.099517 | 370.6757 | 427.1942 | 41.094 | 41.094 | 427.1942 | 370.6259 | 419.6355 | 419.6852 |  | 351.1568
0 | FS2 | CG01 | CAV3 | 4199 | FS2_CG01:CAV3_D4199 |  | b54 resonator | 40 |  |  | 0.49566 | 370.3781 | 427.1942 | 41.094 | 41.094 | 427.1942 | 370.1303 | 419.9331 | 420.1809 |  | 351.4544
0 | [6 empty] | bellow |  |  |  | 0.099517 | 370.0805 | 427.1942 | 41.094 | 41.094 | 427.1942 | 370.0307 | 420.2306 | 420.2804 |  | 351.752
0 | FS2 | CG01 | CAV4 | 4205 | FS2_CG01:CAV4_D4205 |  | b54 resonator | 40 |  |  | 0.49566 | 369.7829 | 427.1942 | 41.094 | 41.094 | 427.1942 | 369.5351 | 420.5282 | 420.7761 |  | 352.0496
0 | [6 empty] | bellows |  | 3.176e-06 |  | 0.17589 | 369.4471 | 427.1942 | 41.094 | 41.094 | 427.1942 | 369.3592 | 420.864 | 420.952 |  | 352.3854
0 | FS2 | CG01 | GV | 4210 | FS2_CG01:GV_D4210 |  | gate valve | 40 | 369.3018 | 369.3018 | 0.057404 | 369.3305 | 427.1942 | 41.094 | 41.094 | 427.1942 | 369.3018 | 420.9807 | 421.0094 |  | 352.502
0 | [7 empty] | 40 |  |  | 0 | 369.3018 | 427.1942 | 41.094 | 41.094 | 427.1942 | 369.3018 | 421.0094 | 421.0094 |  | 352.5307
0 | [10 empty] | 0 | 369.3018 | 427.1942 | 41.094 | 41.094 | 427.1942 | 369.3018 | 421.0094 | 421.0094 |  | 352.5307
0 | [10 empty] | 0 | 369.3018 | 427.1942 | 41.094 | 41.094 | 427.1942 | 369.3018 | 421.0094 | 421.0094 |  | 352.5307
0 | [10 empty] | 0 | 369.3018 | 427.1942 | 41.094 | 41.094 | 427.1942 | 369.3018 | 421.0094 | 421.0094 |  | 352.5307
0 | [10 empty] | 0 | 369.3018 | 427.1942 | 41.094 | 41.094 | 427.1942 | 369.3018 | 421.0094 | 421.0094 |  | 352.5307
0 | [8 empty] | 0 |  | 0.18837 | 369.2076 | 427.1942 | 41.094 | 41.094 | 427.1942 | 369.1134 | 421.1035 | 421.1977 | 20170214 box moved downstream | 352.6249
0 | [8 empty] | 3.9382 | 3.9382 | 0.036554 | 369.0951 | 427.1942 | 41.094 | 41.094 | 427.1942 | 369.0769 | 421.216 | 421.2343 |  | 352.7374
0 | FS2 | BMS | DCHV | 4213 | FS2_BMS:DCHV_D4213 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 368.9769 | 427.1942 | 41.094 | 41.094 | 427.1942 | 368.8769 | 421.3343 | 421.4343 |  | 352.8557
0 | [8 empty] | 3.176e-06 |  | 0.19472 | 368.7795 | 427.1942 | 41.094 | 41.094 | 427.1942 | 368.6821 | 421.5316 | 421.629 |  | 353.053
0 | FS2 | BMS | BPM | 4216 | FS2_BMS:BPM_D4216 | BPM3 | position monitor | 47.5 | 368.6821 | 368.6821 | 0 | 368.6821 | 427.1942 | 41.094 | 41.094 | 427.1942 | 368.6821 | 421.629 | 421.629 |  | 353.1504
0 | [10 empty] | 0.05528 | 368.6545 | 427.1942 | 41.094 | 41.094 | 427.1942 | 368.6269 | 421.6566 | 421.6843 |  | 353.178
0 | FS2 | BMS | QH | 4218 | FS2_BMS:QH_D4218 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 368.5019 | 427.1942 | 41.094 | 41.094 | 427.1942 | 368.3769 | 421.8093 | 421.9343 |  | 353.3307
0 | [10 empty] | 0 | 368.3769 | 427.1942 | 41.094 | 41.094 | 427.1942 | 368.3769 | 421.9343 | 421.9343 | 20170214 split between the bellow and nipple | 353.4557
0 | [6 empty] | bellow | 40 |  |  | 0.28777 | 368.233 | 427.1942 | 41.094 | 41.094 | 427.1942 | 368.0891 | 422.0782 | 422.2221 |  | 353.5995
0 | [6 empty] | vacuum box |  | 3.176e-06 |  | 0 | 368.0891 | 427.1942 | 41.094 | 41.094 | 427.1942 | 368.0891 | 422.2221 | 422.2221 |  | 353.7434
0 | FS2 | BMS | FAVS | 3537 | FS2_BMS:FAVS_D3537 | FVS | fast valve sensor |  | 368.0891 | 368.0891 | 0 | 368.0891 | 427.1942 | 41.094 | 41.094 | 427.1942 | 368.0891 | 422.2221 | 422.2221 |  | 353.7434
0 | [6 empty] | vacuum box |  |  |  | 0 | 368.0891 | 427.1942 | 41.094 | 41.094 | 427.1942 | 368.0891 | 422.2221 | 422.2221 |  | 353.7434
0 | FS2 | BMS | IP | 3537 | FS2_BMS:IP_D3537 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 368.0891 | 427.1942 | 41.094 | 41.094 | 427.1942 | 368.0891 | 422.2221 | 422.2221 |  | 353.7434
0 | [6 empty] | bellow | 40 |  |  | 0.21223 | 367.983 | 427.1942 | 41.094 | 41.094 | 427.1942 | 367.8769 | 422.3282 | 422.4343 |  | 353.8495
0 | FS2 | BMS | QV | 4226 | FS2_BMS:QV_D4226 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 367.7519 | 427.1942 | 41.094 | 41.094 | 427.1942 | 367.6269 | 422.5593 | 422.6843 |  | 354.0807
0 | [10 empty] | 2.5923 | 366.3307 | 427.1942 | 41.094 | 41.094 | 427.1942 | 365.0346 | 423.9804 | 425.2766 |  | 355.5018
0 | FS2 | BMS | BD | 4253 | FS2_BMS:BD_D4253 | dump | dump 25W no current reading |  | 2.176e-06 | 365.0346 | 0 | 365.0346 | 427.1942 | 41.094 | 41.094 | 427.1942 | 365.0346 | 425.2766 | 425.2766 | 20170214 check position with ME | 356.7979
0 | [10 empty] | 2.3777 | 363.8457 | 427.1942 | 41.094 | 41.094 | 427.1942 | 362.6569 | 426.4654 | 427.6543 |  | 357.9868
0 | FS2 | BMS | LPM | 4277 | FS2_BMS:LPM_D4277 | LPMx | bunch length monitor (not costed) | future LPM2 |  |  | 0 | 362.6569 | 427.1942 | 41.094 | 41.094 | 427.1942 | 362.6569 | 427.6543 | 427.6543 |  | 359.1757
0 | FS2 | BMS | IP | 4277 | FS2_BMS:IP_D4277 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 362.6569 | 427.1942 | 41.094 | 41.094 | 427.1942 | 362.6569 | 427.6543 | 427.6543 |  | 359.1757
0 | [10 empty] | 0.32183 | 362.496 | 427.1942 | 41.094 | 41.094 | 427.1942 | 362.335 | 427.8152 | 427.9761 |  | 359.3366
0 | FS2 | BMS | DCHV | 4281 | FS2_BMS:DCHV_D4281 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 362.235 | 427.1942 | 41.094 | 41.094 | 427.1942 | 362.135 | 428.0761 | 428.1761 |  | 359.5975
0 | [8 empty] | 3.176e-06 |  | 0.12715 | 362.0715 | 427.1942 | 41.094 | 41.094 | 427.1942 | 362.0079 | 428.2397 | 428.3033 |  | 359.7611
0 | FS2 | BMS | BPM | 4283 | FS2_BMS:BPM_D4283 | BPM3 | position monitor | 47.5 | 362.0079 | 362.0079 | 0 | 362.0079 | 427.1942 | 41.094 | 41.094 | 427.1942 | 362.0079 | 428.3033 | 428.3033 | 0.137 | 359.8246
0 | [10 empty] | 0.12285 | 361.9465 | 427.1942 | 41.094 | 41.094 | 427.1942 | 361.885 | 428.3647 | 428.4261 |  | 359.8861
0 | FS2 | BMS | QH | 4286 | FS2_BMS:QH_D4286 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 361.76 | 427.1942 | 41.094 | 41.094 | 427.1942 | 361.635 | 428.5511 | 428.6761 |  | 360.0725
0 | [10 empty] | 0.75 | 361.26 | 427.1942 | 41.094 | 41.094 | 427.1942 | 360.885 | 429.0511 | 429.4261 |  | 360.5725
0 | FS2 | BMS | QV | 4296 | FS2_BMS:QV_D4296 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 360.76 | 427.1942 | 41.094 | 41.094 | 427.1942 | 360.635 | 429.5511 | 429.6761 |  | 361.0725
0 | [10 empty] | 0.87 | 360.2 | 427.1942 | 41.094 | 41.094 | 427.1942 | 359.765 | 430.1111 | 430.5461 |  | 361.6325
0 | [10 empty] | 0 | 359.765 | 427.1942 | 41.094 | 41.094 | 427.1942 | 359.765 | 430.5461 | 430.5461 | dump moved upstream | 362.0675
0 | [10 empty] | 2 | 358.765 | 427.1942 | 41.094 | 41.094 | 427.1942 | 357.765 | 431.5461 | 432.5461 |  | 363.0675
0 | [6 empty] | bellow | 40 | -3.3824e-05 |  | 0.1035 | 357.7133 | 427.1942 | 41.094 | 41.094 | 427.1942 | 357.6615 | 432.5979 | 432.6496 |  | 364.1192
0 | FS2 | BMS | BPM | 4326 | FS2_BMS:BPM_D4326 | BPM3 | position monitor |  | 357.6615 | 357.6615 | 0 | 357.6615 | 427.1942 | 41.094 | 41.094 | 427.1942 | 357.6615 | 432.6496 | 432.6496 |  | 364.171
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 357.5889 | 427.1942 | 41.094 | 41.094 | 427.1942 | 357.5163 | 432.7223 | 432.7949 |  | 364.2436
0 | FS2 | BMS | PM | 4328 | FS2_BMS:PM_D4328 | PMx | profile monitor (not costed) | future PM |  |  | 0 | 357.5163 | 427.1942 | 41.094 | 41.094 | 427.1942 | 357.5163 | 432.7949 | 432.7949 |  | 364.3163
0 | [6 empty] | diagnostic box |  | -3.3824e-05 |  | 0 | 357.5163 | 427.1942 | 41.094 | 41.094 | 427.1942 | 357.5163 | 432.7949 | 432.7949 |  | 364.3163
0 | FS2 | BMS | IP | 4328 | FS2_BMS:IP_D4328 | IP_S | ion pump, 75 l/s |  | 357.5162 | 357.5163 | 0 | 357.5163 | 427.1942 | 41.094 | 41.094 | 427.1942 | 357.5163 | 432.7949 | 432.7949 |  | 364.3163
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 357.4506 | 427.1942 | 41.094 | 41.094 | 427.1942 | 357.385 | 432.8605 | 432.9261 |  | 364.3819
0 | [10 empty] | 0 | 357.385 | 427.1942 | 41.094 | 41.094 | 427.1942 | 357.385 | 432.9261 | 432.9261 |  | 364.4475
0 | LS3 | CD01 | GV | 4330 | LS3_CD01:GV_D4330 |  | gate valve | 40 |  |  | 0.057404 | 357.3563 | 427.1942 | 41.094 | 41.094 | 427.1942 | 357.3276 | 432.9548 | 432.9835 |  | 364.4762
0 | [6 empty] | bellows |  |  |  | 0.1677 | 357.2438 | 427.1942 | 41.094 | 41.094 | 427.1942 | 357.1599 | 433.0674 | 433.1512 |  | 364.5887
0 | LS3 | CD01 | CAV1 | 4334 | LS3_CD01:CAV1_D4334 |  | b54 resonator | 40 |  |  | 0.49566 | 356.9121 | 427.1942 | 41.094 | 41.094 | 427.1942 | 356.6643 | 433.399 | 433.6469 |  | 364.9204
0 | LS3 | CD01 | BLM | 4336 | LS3_CD01:BLM_D4336 |  | outside loss monitor |  |  |  | 0 | 356.6643 | 427.1942 | 41.094 | 41.094 | 427.1942 | 356.6643 | 433.6469 | 433.6469 |  | 365.1682
0 | [10 empty] | 0 | 356.6643 | 427.1942 | 41.094 | 41.094 | 427.1942 | 356.6643 | 433.6469 | 433.6469 |  | 365.1682
0 | [6 empty] | bellow |  |  |  | 0.099517 | 356.6145 | 427.1942 | 41.094 | 41.094 | 427.1942 | 356.5648 | 433.6966 | 433.7464 |  | 365.218
0 | LS3 | CD01 | CAV2 | 4340 | LS3_CD01:CAV2_D4340 |  | b54 resonator | 40 |  |  | 0.49566 | 356.3169 | 427.1942 | 41.094 | 41.094 | 427.1942 | 356.0691 | 433.9942 | 434.242 |  | 365.5156
0 | [10 empty] | 0 | 356.0691 | 427.1942 | 41.094 | 41.094 | 427.1942 | 356.0691 | 434.242 | 434.242 |  | 365.7634
0 | [10 empty] | 0 | 356.0691 | 427.1942 | 41.094 | 41.094 | 427.1942 | 356.0691 | 434.242 | 434.242 |  | 365.7634
0 | [6 empty] | bellow |  |  |  | 0.099517 | 356.0193 | 427.1942 | 41.094 | 41.094 | 427.1942 | 355.9696 | 434.2918 | 434.3416 |  | 365.8132
0 | LS3 | CD01 | CAV3 | 4346 | LS3_CD01:CAV3_D4346 |  | b54 resonator | 40 |  |  | 0.49566 | 355.7218 | 427.1942 | 41.094 | 41.094 | 427.1942 | 355.4739 | 434.5894 | 434.8372 |  | 366.1108
0 | [10 empty] | 0 | 355.4739 | 427.1942 | 41.094 | 41.094 | 427.1942 | 355.4739 | 434.8372 | 434.8372 |  | 366.3586
0 | [10 empty] | 0 | 355.4739 | 427.1942 | 41.094 | 41.094 | 427.1942 | 355.4739 | 434.8372 | 434.8372 |  | 366.3586
0 | [6 empty] | bellow |  |  |  | 0.099517 | 355.4242 | 427.1942 | 41.094 | 41.094 | 427.1942 | 355.3744 | 434.887 | 434.9367 |  | 366.4084
0 | LS3 | CD01 | CAV4 | 4352 | LS3_CD01:CAV4_D4352 |  | b54 resonator | 40 |  |  | 0.49566 | 355.1266 | 427.1942 | 41.094 | 41.094 | 427.1942 | 354.8788 | 435.1846 | 435.4324 |  | 366.7059
0 | LS3 | CD01 | BLM | 4354 | LS3_CD01:BLM_D4354 |  | outside loss monitor |  |  |  | 0 | 354.8788 | 427.1942 | 41.094 | 41.094 | 427.1942 | 354.8788 | 435.4324 | 435.4324 |  | 366.9538
0 | [10 empty] | 0 | 354.8788 | 427.1942 | 41.094 | 41.094 | 427.1942 | 354.8788 | 435.4324 | 435.4324 |  | 366.9538
0 | [6 empty] | bellow |  |  |  | 0.15525 | 354.8011 | 427.1942 | 41.094 | 41.094 | 427.1942 | 354.7235 | 435.51 | 435.5876 |  | 367.0314
0 | LS3 | CD01 | SOL1 | 4358 | LS3_CD01:SOL1_D4358 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 354.4735 | 427.1942 | 41.094 | 41.094 | 427.1942 | 354.2235 | 435.8376 | 436.0876 |  | 367.359
0 | [6 empty] | bellow |  |  |  | 0.15525 | 354.1459 | 427.1942 | 41.094 | 41.094 | 427.1942 | 354.0683 | 436.1653 | 436.2429 |  | 367.6866
0 | LS3 | CD01 | CAV5 | 4365 | LS3_CD01:CAV5_D4365 |  | b54 resonator | 40 |  |  | 0.49566 | 353.8204 | 427.1942 | 41.094 | 41.094 | 427.1942 | 353.5726 | 436.4907 | 436.7386 |  | 368.0121
0 | [10 empty] | 0 | 353.5726 | 427.1942 | 41.094 | 41.094 | 427.1942 | 353.5726 | 436.7386 | 436.7386 |  | 368.2599
0 | [10 empty] | 0 | 353.5726 | 427.1942 | 41.094 | 41.094 | 427.1942 | 353.5726 | 436.7386 | 436.7386 |  | 368.2599
0 | [6 empty] | bellow |  |  |  | 0.099517 | 353.5228 | 427.1942 | 41.094 | 41.094 | 427.1942 | 353.4731 | 436.7883 | 436.8381 |  | 368.3097
0 | LS3 | CD01 | CAV6 | 4371 | LS3_CD01:CAV6_D4371 |  | b54 resonator | 40 |  |  | 0.49566 | 353.2253 | 427.1942 | 41.094 | 41.094 | 427.1942 | 352.9774 | 437.0859 | 437.3337 |  | 368.6073
0 | [10 empty] | 0 | 352.9774 | 427.1942 | 41.094 | 41.094 | 427.1942 | 352.9774 | 437.3337 | 437.3337 |  | 368.8551
0 | [10 empty] | 0 | 352.9774 | 427.1942 | 41.094 | 41.094 | 427.1942 | 352.9774 | 437.3337 | 437.3337 |  | 368.8551
0 | [6 empty] | bellow |  |  |  | 0.099517 | 352.9277 | 427.1942 | 41.094 | 41.094 | 427.1942 | 352.8779 | 437.3835 | 437.4332 |  | 368.9049
0 | LS3 | CD01 | CAV7 | 4377 | LS3_CD01:CAV7_D4377 |  | b54 resonator | 40 |  |  | 0.49566 | 352.6301 | 427.1942 | 41.094 | 41.094 | 427.1942 | 352.3823 | 437.6811 | 437.9289 |  | 369.2024
0 | [10 empty] | 0 | 352.3823 | 427.1942 | 41.094 | 41.094 | 427.1942 | 352.3823 | 437.9289 | 437.9289 |  | 369.4503
0 | [10 empty] | 0 | 352.3823 | 427.1942 | 41.094 | 41.094 | 427.1942 | 352.3823 | 437.9289 | 437.9289 |  | 369.4503
0 | [6 empty] | bellow |  |  |  | 0.099517 | 352.3325 | 427.1942 | 41.094 | 41.094 | 427.1942 | 352.2827 | 437.9787 | 438.0284 |  | 369.5
0 | LS3 | CD01 | CAV8 | 4383 | LS3_CD01:CAV8_D4383 |  | b54 resonator | 40 |  |  | 0.49566 | 352.0349 | 427.1942 | 41.094 | 41.094 | 427.1942 | 351.7871 | 438.2762 | 438.5241 |  | 369.7976
0 | LS3 | CD01 | BLM | 4385 | LS3_CD01:BLM_D4385 |  | outside loss monitor |  |  |  | 0 | 351.7871 | 427.1942 | 41.094 | 41.094 | 427.1942 | 351.7871 | 438.5241 | 438.5241 |  | 370.0454
0 | [10 empty] | 0 | 351.7871 | 427.1942 | 41.094 | 41.094 | 427.1942 | 351.7871 | 438.5241 | 438.5241 |  | 370.0454
0 | [10 empty] | 0.1677 | 351.7032 | 427.1942 | 41.094 | 41.094 | 427.1942 | 351.6194 | 438.6079 | 438.6918 |  | 370.1293
0 | LS3 | CD01 | GV | 4387 | LS3_CD01:GV_D4387 |  | gate valve | 40 |  |  | 0.057404 | 351.5907 | 427.1942 | 41.094 | 41.094 | 427.1942 | 351.562 | 438.7205 | 438.7492 |  | 370.2418
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 351.5102 | 427.1942 | 41.094 | 41.094 | 427.1942 | 351.4585 | 438.8009 | 438.8527 |  | 370.3223
0 | LS3 | WD01 | BPM | 4389 | LS3_WD01:BPM_D4389 | BPM2 | position monitor |  |  |  | 0 | 351.4585 | 427.1942 | 41.094 | 41.094 | 427.1942 | 351.4585 | 438.8527 | 438.8527 |  | 370.374
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 351.3858 | 427.1942 | 41.094 | 41.094 | 427.1942 | 351.3132 | 438.9253 | 438.998 |  | 370.4467
0 | LS3 | WD01 | PORT | 4390 | LS3_WD01:PORT_D4390 |  | port for future diagnostics |  |  |  | 0 | 351.3132 | 427.1942 | 41.094 | 41.094 | 427.1942 | 351.3132 | 438.998 | 438.998 |  | 370.5193
0 | [6 empty] | diagnostic box |  |  |  | 0 | 351.3132 | 427.1942 | 41.094 | 41.094 | 427.1942 | 351.3132 | 438.998 | 438.998 |  | 370.5193
0 | LS3 | WD01 | IP | 4390 | LS3_WD01:IP_D4390 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 351.3132 | 427.1942 | 41.094 | 41.094 | 427.1942 | 351.3132 | 438.998 | 438.998 |  | 370.5193
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 351.2476 | 427.1942 | 41.094 | 41.094 | 427.1942 | 351.182 | 439.0636 | 439.1292 |  | 370.5849
0 | LS3 | CD02 | GV | 4392 | LS3_CD02:GV_D4392 |  | gate valve | 40 |  |  | 0.057404 | 351.1533 | 427.1942 | 41.094 | 41.094 | 427.1942 | 351.1246 | 439.1579 | 439.1866 |  | 370.6792
0 | [6 empty] | bellows |  |  |  | 0.1677 | 351.0407 | 427.1942 | 41.094 | 41.094 | 427.1942 | 350.9569 | 439.2704 | 439.3543 |  | 370.7918
0 | LS3 | CD02 | CAV1 | 4396 | LS3_CD02:CAV1_D4396 |  | b54 resonator | 40 |  |  | 0.49566 | 350.709 | 427.1942 | 41.094 | 41.094 | 427.1942 | 350.4612 | 439.6021 | 439.8499 |  | 371.1235
0 | LS3 | CD02 | BLM | 4398 | LS3_CD02:BLM_D4398 |  | outside loss monitor |  |  |  | 0 | 350.4612 | 427.1942 | 41.094 | 41.094 | 427.1942 | 350.4612 | 439.8499 | 439.8499 |  | 371.3713
0 | [10 empty] | 0 | 350.4612 | 427.1942 | 41.094 | 41.094 | 427.1942 | 350.4612 | 439.8499 | 439.8499 |  | 371.3713
0 | [6 empty] | bellow |  |  |  | 0.099517 | 350.4114 | 427.1942 | 41.094 | 41.094 | 427.1942 | 350.3617 | 439.8997 | 439.9495 |  | 371.4211
0 | LS3 | CD02 | CAV2 | 4402 | LS3_CD02:CAV2_D4402 |  | b54 resonator | 40 |  |  | 0.49566 | 350.1139 | 427.1942 | 41.094 | 41.094 | 427.1942 | 349.866 | 440.1973 | 440.4451 |  | 371.7187
0 | [10 empty] | 0 | 349.866 | 427.1942 | 41.094 | 41.094 | 427.1942 | 349.866 | 440.4451 | 440.4451 |  | 371.9665
0 | [10 empty] | 0 | 349.866 | 427.1942 | 41.094 | 41.094 | 427.1942 | 349.866 | 440.4451 | 440.4451 |  | 371.9665
0 | [6 empty] | bellow |  |  |  | 0.099517 | 349.8163 | 427.1942 | 41.094 | 41.094 | 427.1942 | 349.7665 | 440.4949 | 440.5446 |  | 372.0162
0 | LS3 | CD02 | CAV3 | 4408 | LS3_CD02:CAV3_D4408 |  | b54 resonator | 40 |  |  | 0.49566 | 349.5187 | 427.1942 | 41.094 | 41.094 | 427.1942 | 349.2709 | 440.7925 | 441.0403 |  | 372.3138
0 | [10 empty] | 0 | 349.2709 | 427.1942 | 41.094 | 41.094 | 427.1942 | 349.2709 | 441.0403 | 441.0403 |  | 372.5617
0 | [10 empty] | 0 | 349.2709 | 427.1942 | 41.094 | 41.094 | 427.1942 | 349.2709 | 441.0403 | 441.0403 |  | 372.5617
0 | [6 empty] | bellow |  |  |  | 0.099517 | 349.2211 | 427.1942 | 41.094 | 41.094 | 427.1942 | 349.1713 | 441.09 | 441.1398 |  | 372.6114
0 | LS3 | CD02 | CAV4 | 4414 | LS3_CD02:CAV4_D4414 |  | b54 resonator | 40 |  |  | 0.49566 | 348.9235 | 427.1942 | 41.094 | 41.094 | 427.1942 | 348.6757 | 441.3876 | 441.6355 |  | 372.909
0 | LS3 | CD02 | BLM | 4416 | LS3_CD02:BLM_D4416 |  | outside loss monitor |  |  |  | 0 | 348.6757 | 427.1942 | 41.094 | 41.094 | 427.1942 | 348.6757 | 441.6355 | 441.6355 |  | 373.1568
0 | [10 empty] | 0 | 348.6757 | 427.1942 | 41.094 | 41.094 | 427.1942 | 348.6757 | 441.6355 | 441.6355 |  | 373.1568
0 | [6 empty] | bellow |  |  |  | 0.15525 | 348.5981 | 427.1942 | 41.094 | 41.094 | 427.1942 | 348.5204 | 441.7131 | 441.7907 |  | 373.2345
0 | LS3 | CD02 | SOL1 | 4420 | LS3_CD02:SOL1_D4420 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 348.2704 | 427.1942 | 41.094 | 41.094 | 427.1942 | 348.0204 | 442.0407 | 442.2907 |  | 373.5621
0 | [6 empty] | bellow |  |  |  | 0.15525 | 347.9428 | 427.1942 | 41.094 | 41.094 | 427.1942 | 347.8652 | 442.3683 | 442.446 |  | 373.8897
0 | LS3 | CD02 | CAV5 | 4427 | LS3_CD02:CAV5_D4427 |  | b54 resonator | 40 |  |  | 0.49566 | 347.6174 | 427.1942 | 41.094 | 41.094 | 427.1942 | 347.3695 | 442.6938 | 442.9416 |  | 374.2152
0 | [10 empty] | 0 | 347.3695 | 427.1942 | 41.094 | 41.094 | 427.1942 | 347.3695 | 442.9416 | 442.9416 |  | 374.463
0 | [10 empty] | 0 | 347.3695 | 427.1942 | 41.094 | 41.094 | 427.1942 | 347.3695 | 442.9416 | 442.9416 |  | 374.463
0 | [6 empty] | bellow |  |  |  | 0.099517 | 347.3198 | 427.1942 | 41.094 | 41.094 | 427.1942 | 347.27 | 442.9914 | 443.0411 |  | 374.5127
0 | LS3 | CD02 | CAV6 | 4433 | LS3_CD02:CAV6_D4433 |  | b54 resonator | 40 |  |  | 0.49566 | 347.0222 | 427.1942 | 41.094 | 41.094 | 427.1942 | 346.7744 | 443.289 | 443.5368 |  | 374.8103
0 | [10 empty] | 0 | 346.7744 | 427.1942 | 41.094 | 41.094 | 427.1942 | 346.7744 | 443.5368 | 443.5368 |  | 375.0582
0 | [10 empty] | 0 | 346.7744 | 427.1942 | 41.094 | 41.094 | 427.1942 | 346.7744 | 443.5368 | 443.5368 |  | 375.0582
0 | [6 empty] | bellow |  |  |  | 0.099517 | 346.7246 | 427.1942 | 41.094 | 41.094 | 427.1942 | 346.6748 | 443.5865 | 443.6363 |  | 375.1079
0 | LS3 | CD02 | CAV7 | 4439 | LS3_CD02:CAV7_D4439 |  | b54 resonator | 40 |  |  | 0.49566 | 346.427 | 427.1942 | 41.094 | 41.094 | 427.1942 | 346.1792 | 443.8841 | 444.132 |  | 375.4055
0 | [10 empty] | 0 | 346.1792 | 427.1942 | 41.094 | 41.094 | 427.1942 | 346.1792 | 444.132 | 444.132 |  | 375.6533
0 | [10 empty] | 0 | 346.1792 | 427.1942 | 41.094 | 41.094 | 427.1942 | 346.1792 | 444.132 | 444.132 |  | 375.6533
0 | [6 empty] | bellow |  |  |  | 0.099517 | 346.1294 | 427.1942 | 41.094 | 41.094 | 427.1942 | 346.0797 | 444.1817 | 444.2315 |  | 375.7031
0 | LS3 | CD02 | CAV8 | 4445 | LS3_CD02:CAV8_D4445 |  | b54 resonator | 40 |  |  | 0.49566 | 345.8318 | 427.1942 | 41.094 | 41.094 | 427.1942 | 345.584 | 444.4793 | 444.7271 |  | 376.0007
0 | LS3 | CD02 | BLM | 4447 | LS3_CD02:BLM_D4447 |  | outside loss monitor |  |  |  | 0 | 345.584 | 427.1942 | 41.094 | 41.094 | 427.1942 | 345.584 | 444.7271 | 444.7271 |  | 376.2485
0 | [10 empty] | 0 | 345.584 | 427.1942 | 41.094 | 41.094 | 427.1942 | 345.584 | 444.7271 | 444.7271 |  | 376.2485
0 | [10 empty] | 0.1677 | 345.5002 | 427.1942 | 41.094 | 41.094 | 427.1942 | 345.4163 | 444.811 | 444.8948 |  | 376.3324
0 | LS3 | CD02 | GV | 4449 | LS3_CD02:GV_D4449 |  | gate valve | 40 |  |  | 0.057404 | 345.3876 | 427.1942 | 41.094 | 41.094 | 427.1942 | 345.3589 | 444.9235 | 444.9522 |  | 376.4449
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 345.3072 | 427.1942 | 41.094 | 41.094 | 427.1942 | 345.2554 | 445.004 | 445.0557 |  | 376.5254
0 | LS3 | WD02 | BPM | 4451 | LS3_WD02:BPM_D4451 | BPM2 | position monitor |  |  |  | 0 | 345.2554 | 427.1942 | 41.094 | 41.094 | 427.1942 | 345.2554 | 445.0557 | 445.0557 |  | 376.5771
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 345.1828 | 427.1942 | 41.094 | 41.094 | 427.1942 | 345.1101 | 445.1284 | 445.201 |  | 376.6498
0 | LS3 | WD02 | PORT | 4452 | LS3_WD02:PORT_D4452 |  | port for future diagnostics |  |  |  | 0 | 345.1101 | 427.1942 | 41.094 | 41.094 | 427.1942 | 345.1101 | 445.201 | 445.201 |  | 376.7224
0 | [6 empty] | diagnostic box |  |  |  | 0 | 345.1101 | 427.1942 | 41.094 | 41.094 | 427.1942 | 345.1101 | 445.201 | 445.201 |  | 376.7224
0 | LS3 | WD02 | IP | 4452 | LS3_WD02:IP_D4452 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 345.1101 | 427.1942 | 41.094 | 41.094 | 427.1942 | 345.1101 | 445.201 | 445.201 |  | 376.7224
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 345.0445 | 427.1942 | 41.094 | 41.094 | 427.1942 | 344.9789 | 445.2666 | 445.3322 |  | 376.788
0 | LS3 | CD03 | GV | 4454 | LS3_CD03:GV_D4454 |  | gate valve | 40 |  |  | 0.057404 | 344.9502 | 427.1942 | 41.094 | 41.094 | 427.1942 | 344.9215 | 445.3609 | 445.3896 |  | 376.8823
0 | [6 empty] | bellows |  |  |  | 0.1677 | 344.8376 | 427.1942 | 41.094 | 41.094 | 427.1942 | 344.7538 | 445.4735 | 445.5574 |  | 376.9949
0 | LS3 | CD03 | CAV1 | 4458 | LS3_CD03:CAV1_D4458 |  | b54 resonator | 40 |  |  | 0.49566 | 344.506 | 427.1942 | 41.094 | 41.094 | 427.1942 | 344.2581 | 445.8052 | 446.053 |  | 377.3265
0 | LS3 | CD03 | BLM | 4461 | LS3_CD03:BLM_D4461 |  | outside loss monitor |  |  |  | 0 | 344.2581 | 427.1942 | 41.094 | 41.094 | 427.1942 | 344.2581 | 446.053 | 446.053 |  | 377.5744
0 | [10 empty] | 0 | 344.2581 | 427.1942 | 41.094 | 41.094 | 427.1942 | 344.2581 | 446.053 | 446.053 |  | 377.5744
0 | [6 empty] | bellow |  |  |  | 0.099517 | 344.2084 | 427.1942 | 41.094 | 41.094 | 427.1942 | 344.1586 | 446.1028 | 446.1525 |  | 377.6241
0 | LS3 | CD03 | CAV2 | 4464 | LS3_CD03:CAV2_D4464 |  | b54 resonator | 40 |  |  | 0.49566 | 343.9108 | 427.1942 | 41.094 | 41.094 | 427.1942 | 343.663 | 446.4004 | 446.6482 |  | 377.9217
0 | [10 empty] | 0 | 343.663 | 427.1942 | 41.094 | 41.094 | 427.1942 | 343.663 | 446.6482 | 446.6482 |  | 378.1696
0 | [10 empty] | 0 | 343.663 | 427.1942 | 41.094 | 41.094 | 427.1942 | 343.663 | 446.6482 | 446.6482 |  | 378.1696
0 | [6 empty] | bellow |  |  |  | 0.099517 | 343.6132 | 427.1942 | 41.094 | 41.094 | 427.1942 | 343.5635 | 446.6979 | 446.7477 |  | 378.2193
0 | LS3 | CD03 | CAV3 | 4470 | LS3_CD03:CAV3_D4470 |  | b54 resonator | 40 |  |  | 0.49566 | 343.3156 | 427.1942 | 41.094 | 41.094 | 427.1942 | 343.0678 | 446.9955 | 447.2434 |  | 378.5169
0 | [10 empty] | 0 | 343.0678 | 427.1942 | 41.094 | 41.094 | 427.1942 | 343.0678 | 447.2434 | 447.2434 |  | 378.7647
0 | [10 empty] | 0 | 343.0678 | 427.1942 | 41.094 | 41.094 | 427.1942 | 343.0678 | 447.2434 | 447.2434 |  | 378.7647
0 | [6 empty] | bellow |  |  |  | 0.099517 | 343.018 | 427.1942 | 41.094 | 41.094 | 427.1942 | 342.9683 | 447.2931 | 447.3429 |  | 378.8145
0 | LS3 | CD03 | CAV4 | 4476 | LS3_CD03:CAV4_D4476 |  | b54 resonator | 40 |  |  | 0.49566 | 342.7204 | 427.1942 | 41.094 | 41.094 | 427.1942 | 342.4726 | 447.5907 | 447.8385 |  | 379.1121
0 | LS3 | CD03 | BLM | 4478 | LS3_CD03:BLM_D4478 |  | outside loss monitor |  |  |  | 0 | 342.4726 | 427.1942 | 41.094 | 41.094 | 427.1942 | 342.4726 | 447.8385 | 447.8385 |  | 379.3599
0 | [10 empty] | 0 | 342.4726 | 427.1942 | 41.094 | 41.094 | 427.1942 | 342.4726 | 447.8385 | 447.8385 |  | 379.3599
0 | [6 empty] | bellow |  |  |  | 0.15525 | 342.395 | 427.1942 | 41.094 | 41.094 | 427.1942 | 342.3174 | 447.9162 | 447.9938 |  | 379.4375
0 | LS3 | CD03 | SOL1 | 4482 | LS3_CD03:SOL1_D4482 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 342.0674 | 427.1942 | 41.094 | 41.094 | 427.1942 | 341.8174 | 448.2438 | 448.4938 |  | 379.7651
0 | [6 empty] | bellow |  |  |  | 0.15525 | 341.7397 | 427.1942 | 41.094 | 41.094 | 427.1942 | 341.6621 | 448.5714 | 448.649 |  | 380.0928
0 | LS3 | CD03 | CAV5 | 4489 | LS3_CD03:CAV5_D4489 |  | b54 resonator | 40 |  |  | 0.49566 | 341.4143 | 427.1942 | 41.094 | 41.094 | 427.1942 | 341.1665 | 448.8969 | 449.1447 |  | 380.4182
0 | [10 empty] | 0 | 341.1665 | 427.1942 | 41.094 | 41.094 | 427.1942 | 341.1665 | 449.1447 | 449.1447 |  | 380.6661
0 | [10 empty] | 0 | 341.1665 | 427.1942 | 41.094 | 41.094 | 427.1942 | 341.1665 | 449.1447 | 449.1447 |  | 380.6661
0 | [6 empty] | bellow |  |  |  | 0.099517 | 341.1167 | 427.1942 | 41.094 | 41.094 | 427.1942 | 341.0669 | 449.1944 | 449.2442 |  | 380.7158
0 | LS3 | CD03 | CAV6 | 4495 | LS3_CD03:CAV6_D4495 |  | b54 resonator | 40 |  |  | 0.49566 | 340.8191 | 427.1942 | 41.094 | 41.094 | 427.1942 | 340.5713 | 449.492 | 449.7399 |  | 381.0134
0 | [10 empty] | 0 | 340.5713 | 427.1942 | 41.094 | 41.094 | 427.1942 | 340.5713 | 449.7399 | 449.7399 |  | 381.2612
0 | [10 empty] | 0 | 340.5713 | 427.1942 | 41.094 | 41.094 | 427.1942 | 340.5713 | 449.7399 | 449.7399 |  | 381.2612
0 | [6 empty] | bellow |  |  |  | 0.099517 | 340.5215 | 427.1942 | 41.094 | 41.094 | 427.1942 | 340.4718 | 449.7896 | 449.8394 |  | 381.311
0 | LS3 | CD03 | CAV7 | 4501 | LS3_CD03:CAV7_D4501 |  | b54 resonator | 40 |  |  | 0.49566 | 340.2239 | 427.1942 | 41.094 | 41.094 | 427.1942 | 339.9761 | 450.0872 | 450.335 |  | 381.6086
0 | [10 empty] | 0 | 339.9761 | 427.1942 | 41.094 | 41.094 | 427.1942 | 339.9761 | 450.335 | 450.335 |  | 381.8564
0 | [10 empty] | 0 | 339.9761 | 427.1942 | 41.094 | 41.094 | 427.1942 | 339.9761 | 450.335 | 450.335 |  | 381.8564
0 | [6 empty] | bellow |  |  |  | 0.099517 | 339.9264 | 427.1942 | 41.094 | 41.094 | 427.1942 | 339.8766 | 450.3848 | 450.4345 |  | 381.9062
0 | LS3 | CD03 | CAV8 | 4507 | LS3_CD03:CAV8_D4507 |  | b54 resonator | 40 |  |  | 0.49566 | 339.6288 | 427.1942 | 41.094 | 41.094 | 427.1942 | 339.3809 | 450.6824 | 450.9302 |  | 382.2037
0 | LS3 | CD03 | BLM | 4509 | LS3_CD03:BLM_D4509 |  | outside loss monitor |  |  |  | 0 | 339.3809 | 427.1942 | 41.094 | 41.094 | 427.1942 | 339.3809 | 450.9302 | 450.9302 |  | 382.4516
0 | [10 empty] | 0 | 339.3809 | 427.1942 | 41.094 | 41.094 | 427.1942 | 339.3809 | 450.9302 | 450.9302 |  | 382.4516
0 | [10 empty] | 0.1677 | 339.2971 | 427.1942 | 41.094 | 41.094 | 427.1942 | 339.2132 | 451.0141 | 451.0979 |  | 382.5354
0 | LS3 | CD03 | GV | 4511 | LS3_CD03:GV_D4511 |  | gate valve | 40 |  |  | 0.057404 | 339.1845 | 427.1942 | 41.094 | 41.094 | 427.1942 | 339.1558 | 451.1266 | 451.1553 |  | 382.648
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 339.1041 | 427.1942 | 41.094 | 41.094 | 427.1942 | 339.0523 | 451.2071 | 451.2588 |  | 382.7284
0 | LS3 | WD03 | BPM | 4513 | LS3_WD03:BPM_D4513 | BPM2 | position monitor |  |  |  | 0 | 339.0523 | 427.1942 | 41.094 | 41.094 | 427.1942 | 339.0523 | 451.2588 | 451.2588 |  | 382.7802
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 338.9797 | 427.1942 | 41.094 | 41.094 | 427.1942 | 338.9071 | 451.3315 | 451.4041 |  | 382.8528
0 | LS3 | WD03 | PORT | 4514 | LS3_WD03:PORT_D4514 |  | port for future diagnostics |  |  |  | 0 | 338.9071 | 427.1942 | 41.094 | 41.094 | 427.1942 | 338.9071 | 451.4041 | 451.4041 |  | 382.9255
0 | [6 empty] | diagnostic box |  |  |  | 0 | 338.9071 | 427.1942 | 41.094 | 41.094 | 427.1942 | 338.9071 | 451.4041 | 451.4041 |  | 382.9255
0 | LS3 | WD03 | IP | 4514 | LS3_WD03:IP_D4514 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 338.9071 | 427.1942 | 41.094 | 41.094 | 427.1942 | 338.9071 | 451.4041 | 451.4041 |  | 382.9255
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 338.8414 | 427.1942 | 41.094 | 41.094 | 427.1942 | 338.7758 | 451.4697 | 451.5353 |  | 382.9911
0 | LS3 | CD04 | GV | 4516 | LS3_CD04:GV_D4516 |  | gate valve | 40 |  |  | 0.057404 | 338.7471 | 427.1942 | 41.094 | 41.094 | 427.1942 | 338.7184 | 451.564 | 451.5927 |  | 383.0854
0 | [6 empty] | bellows |  |  |  | 0.1677 | 338.6346 | 427.1942 | 41.094 | 41.094 | 427.1942 | 338.5507 | 451.6766 | 451.7604 |  | 383.1979
0 | LS3 | CD04 | CAV1 | 4520 | LS3_CD04:CAV1_D4520 |  | b54 resonator | 40 |  |  | 0.49566 | 338.3029 | 427.1942 | 41.094 | 41.094 | 427.1942 | 338.0551 | 452.0082 | 452.2561 |  | 383.5296
0 | LS3 | CD04 | BLM | 4523 | LS3_CD04:BLM_D4523 |  | outside loss monitor |  |  |  | 0 | 338.0551 | 427.1942 | 41.094 | 41.094 | 427.1942 | 338.0551 | 452.2561 | 452.2561 |  | 383.7774
0 | [10 empty] | 0 | 338.0551 | 427.1942 | 41.094 | 41.094 | 427.1942 | 338.0551 | 452.2561 | 452.2561 |  | 383.7774
0 | [6 empty] | bellow |  |  |  | 0.099517 | 338.0053 | 427.1942 | 41.094 | 41.094 | 427.1942 | 337.9556 | 452.3058 | 452.3556 |  | 383.8272
0 | LS3 | CD04 | CAV2 | 4526 | LS3_CD04:CAV2_D4526 |  | b54 resonator | 40 |  |  | 0.49566 | 337.7077 | 427.1942 | 41.094 | 41.094 | 427.1942 | 337.4599 | 452.6034 | 452.8512 |  | 384.1248
0 | [10 empty] | 0 | 337.4599 | 427.1942 | 41.094 | 41.094 | 427.1942 | 337.4599 | 452.8512 | 452.8512 |  | 384.3726
0 | [10 empty] | 0 | 337.4599 | 427.1942 | 41.094 | 41.094 | 427.1942 | 337.4599 | 452.8512 | 452.8512 |  | 384.3726
0 | [6 empty] | bellow |  |  |  | 0.099517 | 337.4101 | 427.1942 | 41.094 | 41.094 | 427.1942 | 337.3604 | 452.901 | 452.9508 |  | 384.4224
0 | LS3 | CD04 | CAV3 | 4532 | LS3_CD04:CAV3_D4532 |  | b54 resonator | 40 |  |  | 0.49566 | 337.1126 | 427.1942 | 41.094 | 41.094 | 427.1942 | 336.8647 | 453.1986 | 453.4464 |  | 384.72
0 | [10 empty] | 0 | 336.8647 | 427.1942 | 41.094 | 41.094 | 427.1942 | 336.8647 | 453.4464 | 453.4464 |  | 384.9678
0 | [10 empty] | 0 | 336.8647 | 427.1942 | 41.094 | 41.094 | 427.1942 | 336.8647 | 453.4464 | 453.4464 |  | 384.9678
0 | [6 empty] | bellow |  |  |  | 0.099517 | 336.815 | 427.1942 | 41.094 | 41.094 | 427.1942 | 336.7652 | 453.4962 | 453.5459 |  | 385.0175
0 | LS3 | CD04 | CAV4 | 4538 | LS3_CD04:CAV4_D4538 |  | b54 resonator | 40 |  |  | 0.49566 | 336.5174 | 427.1942 | 41.094 | 41.094 | 427.1942 | 336.2696 | 453.7938 | 454.0416 |  | 385.3151
0 | LS3 | CD04 | BLM | 4540 | LS3_CD04:BLM_D4540 |  | outside loss monitor |  |  |  | 0 | 336.2696 | 427.1942 | 41.094 | 41.094 | 427.1942 | 336.2696 | 454.0416 | 454.0416 |  | 385.563
0 | [10 empty] | 0 | 336.2696 | 427.1942 | 41.094 | 41.094 | 427.1942 | 336.2696 | 454.0416 | 454.0416 |  | 385.563
0 | [6 empty] | bellow |  |  |  | 0.15525 | 336.1919 | 427.1942 | 41.094 | 41.094 | 427.1942 | 336.1143 | 454.1192 | 454.1968 |  | 385.6406
0 | LS3 | CD04 | SOL1 | 4544 | LS3_CD04:SOL1_D4544 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 335.8643 | 427.1942 | 41.094 | 41.094 | 427.1942 | 335.6143 | 454.4468 | 454.6968 |  | 385.9682
0 | [6 empty] | bellow |  |  |  | 0.15525 | 335.5367 | 427.1942 | 41.094 | 41.094 | 427.1942 | 335.4591 | 454.7745 | 454.8521 |  | 386.2958
0 | LS3 | CD04 | CAV5 | 4551 | LS3_CD04:CAV5_D4551 |  | b54 resonator | 40 |  |  | 0.49566 | 335.2112 | 427.1942 | 41.094 | 41.094 | 427.1942 | 334.9634 | 455.0999 | 455.3477 |  | 386.6213
0 | [10 empty] | 0 | 334.9634 | 427.1942 | 41.094 | 41.094 | 427.1942 | 334.9634 | 455.3477 | 455.3477 |  | 386.8691
0 | [10 empty] | 0 | 334.9634 | 427.1942 | 41.094 | 41.094 | 427.1942 | 334.9634 | 455.3477 | 455.3477 |  | 386.8691
0 | [6 empty] | bellow |  |  |  | 0.099517 | 334.9136 | 427.1942 | 41.094 | 41.094 | 427.1942 | 334.8639 | 455.3975 | 455.4473 |  | 386.9189
0 | LS3 | CD04 | CAV6 | 4557 | LS3_CD04:CAV6_D4557 |  | b54 resonator | 40 |  |  | 0.49566 | 334.6161 | 427.1942 | 41.094 | 41.094 | 427.1942 | 334.3682 | 455.6951 | 455.9429 |  | 387.2165
0 | [10 empty] | 0 | 334.3682 | 427.1942 | 41.094 | 41.094 | 427.1942 | 334.3682 | 455.9429 | 455.9429 |  | 387.4643
0 | [10 empty] | 0 | 334.3682 | 427.1942 | 41.094 | 41.094 | 427.1942 | 334.3682 | 455.9429 | 455.9429 |  | 387.4643
0 | [6 empty] | bellow |  |  |  | 0.099517 | 334.3185 | 427.1942 | 41.094 | 41.094 | 427.1942 | 334.2687 | 455.9927 | 456.0424 |  | 387.5141
0 | LS3 | CD04 | CAV7 | 4563 | LS3_CD04:CAV7_D4563 |  | b54 resonator | 40 |  |  | 0.49566 | 334.0209 | 427.1942 | 41.094 | 41.094 | 427.1942 | 333.7731 | 456.2903 | 456.5381 |  | 387.8116
0 | [10 empty] | 0 | 333.7731 | 427.1942 | 41.094 | 41.094 | 427.1942 | 333.7731 | 456.5381 | 456.5381 |  | 388.0595
0 | [10 empty] | 0 | 333.7731 | 427.1942 | 41.094 | 41.094 | 427.1942 | 333.7731 | 456.5381 | 456.5381 |  | 388.0595
0 | [6 empty] | bellow |  |  |  | 0.099517 | 333.7233 | 427.1942 | 41.094 | 41.094 | 427.1942 | 333.6735 | 456.5879 | 456.6376 |  | 388.1092
0 | LS3 | CD04 | CAV8 | 4569 | LS3_CD04:CAV8_D4569 |  | b54 resonator | 40 |  |  | 0.49566 | 333.4257 | 427.1942 | 41.094 | 41.094 | 427.1942 | 333.1779 | 456.8854 | 457.1333 |  | 388.4068
0 | LS3 | CD04 | BLM | 4571 | LS3_CD04:BLM_D4571 |  | outside loss monitor |  |  |  | 0 | 333.1779 | 427.1942 | 41.094 | 41.094 | 427.1942 | 333.1779 | 457.1333 | 457.1333 |  | 388.6546
0 | [10 empty] | 0 | 333.1779 | 427.1942 | 41.094 | 41.094 | 427.1942 | 333.1779 | 457.1333 | 457.1333 |  | 388.6546
0 | [10 empty] | 0.1677 | 333.094 | 427.1942 | 41.094 | 41.094 | 427.1942 | 333.0102 | 457.2171 | 457.301 |  | 388.7385
0 | LS3 | CD04 | GV | 4573 | LS3_CD04:GV_D4573 |  | gate valve | 40 |  |  | 0.057404 | 332.9815 | 427.1942 | 41.094 | 41.094 | 427.1942 | 332.9528 | 457.3297 | 457.3584 |  | 388.851
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 332.901 | 427.1942 | 41.094 | 41.094 | 427.1942 | 332.8493 | 457.4101 | 457.4619 |  | 388.9315
0 | LS3 | WD04 | BPM | 4575 | LS3_WD04:BPM_D4575 | BPM2 | position monitor |  |  |  | 0 | 332.8493 | 427.1942 | 41.094 | 41.094 | 427.1942 | 332.8493 | 457.4619 | 457.4619 |  | 388.9832
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 332.7766 | 427.1942 | 41.094 | 41.094 | 427.1942 | 332.704 | 457.5345 | 457.6072 |  | 389.0559
0 | LS3 | WD04 | PORT | 4576 | LS3_WD04:PORT_D4576 |  | port for future diagnostics |  |  |  | 0 | 332.704 | 427.1942 | 41.094 | 41.094 | 427.1942 | 332.704 | 457.6072 | 457.6072 |  | 389.1285
0 | [6 empty] | diagnostic box |  |  |  | 0 | 332.704 | 427.1942 | 41.094 | 41.094 | 427.1942 | 332.704 | 457.6072 | 457.6072 |  | 389.1285
0 | LS3 | WD04 | IP | 4576 | LS3_WD04:IP_D4576 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 332.704 | 427.1942 | 41.094 | 41.094 | 427.1942 | 332.704 | 457.6072 | 457.6072 |  | 389.1285
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 332.6384 | 427.1942 | 41.094 | 41.094 | 427.1942 | 332.5728 | 457.6728 | 457.7384 |  | 389.1941
0 | LS3 | CD05 | GV | 4578 | LS3_CD05:GV_D4578 |  | gate valve | 40 |  |  | 0.057404 | 332.5441 | 427.1942 | 41.094 | 41.094 | 427.1942 | 332.5154 | 457.7671 | 457.7958 |  | 389.2884
0 | [6 empty] | bellows |  |  |  | 0.1677 | 332.4315 | 427.1942 | 41.094 | 41.094 | 427.1942 | 332.3477 | 457.8796 | 457.9635 |  | 389.401
0 | LS3 | CD05 | CAV1 | 4582 | LS3_CD05:CAV1_D4582 |  | b54 resonator | 40 |  |  | 0.49566 | 332.0998 | 427.1942 | 41.094 | 41.094 | 427.1942 | 331.852 | 458.2113 | 458.4591 |  | 389.7327
0 | LS3 | CD05 | BLM | 4585 | LS3_CD05:BLM_D4585 |  | outside loss monitor |  |  |  | 0 | 331.852 | 427.1942 | 41.094 | 41.094 | 427.1942 | 331.852 | 458.4591 | 458.4591 |  | 389.9805
0 | [10 empty] | 0 | 331.852 | 427.1942 | 41.094 | 41.094 | 427.1942 | 331.852 | 458.4591 | 458.4591 |  | 389.9805
0 | [6 empty] | bellow |  |  |  | 0.099517 | 331.8023 | 427.1942 | 41.094 | 41.094 | 427.1942 | 331.7525 | 458.5089 | 458.5587 |  | 390.0303
0 | LS3 | CD05 | CAV2 | 4588 | LS3_CD05:CAV2_D4588 |  | b54 resonator | 40 |  |  | 0.49566 | 331.5047 | 427.1942 | 41.094 | 41.094 | 427.1942 | 331.2568 | 458.8065 | 459.0543 |  | 390.3279
0 | [10 empty] | 0 | 331.2568 | 427.1942 | 41.094 | 41.094 | 427.1942 | 331.2568 | 459.0543 | 459.0543 |  | 390.5757
0 | [10 empty] | 0 | 331.2568 | 427.1942 | 41.094 | 41.094 | 427.1942 | 331.2568 | 459.0543 | 459.0543 |  | 390.5757
0 | [6 empty] | bellow |  |  |  | 0.099517 | 331.2071 | 427.1942 | 41.094 | 41.094 | 427.1942 | 331.1573 | 459.1041 | 459.1538 |  | 390.6254
0 | LS3 | CD05 | CAV3 | 4594 | LS3_CD05:CAV3_D4594 |  | b54 resonator | 40 |  |  | 0.49566 | 330.9095 | 427.1942 | 41.094 | 41.094 | 427.1942 | 330.6617 | 459.4017 | 459.6495 |  | 390.923
0 | [10 empty] | 0 | 330.6617 | 427.1942 | 41.094 | 41.094 | 427.1942 | 330.6617 | 459.6495 | 459.6495 |  | 391.1709
0 | [10 empty] | 0 | 330.6617 | 427.1942 | 41.094 | 41.094 | 427.1942 | 330.6617 | 459.6495 | 459.6495 |  | 391.1709
0 | [6 empty] | bellow |  |  |  | 0.099517 | 330.6119 | 427.1942 | 41.094 | 41.094 | 427.1942 | 330.5621 | 459.6992 | 459.749 |  | 391.2206
0 | LS3 | CD05 | CAV4 | 4600 | LS3_CD05:CAV4_D4600 |  | b54 resonator | 40 |  |  | 0.49566 | 330.3143 | 427.1942 | 41.094 | 41.094 | 427.1942 | 330.0665 | 459.9968 | 460.2447 |  | 391.5182
0 | LS3 | CD05 | BLM | 4602 | LS3_CD05:BLM_D4602 |  | outside loss monitor |  |  |  | 0 | 330.0665 | 427.1942 | 41.094 | 41.094 | 427.1942 | 330.0665 | 460.2447 | 460.2447 |  | 391.766
0 | [10 empty] | 0 | 330.0665 | 427.1942 | 41.094 | 41.094 | 427.1942 | 330.0665 | 460.2447 | 460.2447 |  | 391.766
0 | [6 empty] | bellow |  |  |  | 0.15525 | 329.9889 | 427.1942 | 41.094 | 41.094 | 427.1942 | 329.9112 | 460.3223 | 460.3999 |  | 391.8437
0 | LS3 | CD05 | SOL1 | 4606 | LS3_CD05:SOL1_D4606 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 329.6612 | 427.1942 | 41.094 | 41.094 | 427.1942 | 329.4112 | 460.6499 | 460.8999 |  | 392.1713
0 | [6 empty] | bellow |  |  |  | 0.15525 | 329.3336 | 427.1942 | 41.094 | 41.094 | 427.1942 | 329.256 | 460.9775 | 461.0552 |  | 392.4989
0 | LS3 | CD05 | CAV5 | 4613 | LS3_CD05:CAV5_D4613 |  | b54 resonator | 40 |  |  | 0.49566 | 329.0082 | 427.1942 | 41.094 | 41.094 | 427.1942 | 328.7603 | 461.303 | 461.5508 |  | 392.8244
0 | [10 empty] | 0 | 328.7603 | 427.1942 | 41.094 | 41.094 | 427.1942 | 328.7603 | 461.5508 | 461.5508 |  | 393.0722
0 | [10 empty] | 0 | 328.7603 | 427.1942 | 41.094 | 41.094 | 427.1942 | 328.7603 | 461.5508 | 461.5508 |  | 393.0722
0 | [6 empty] | bellow |  |  |  | 0.099517 | 328.7106 | 427.1942 | 41.094 | 41.094 | 427.1942 | 328.6608 | 461.6006 | 461.6503 |  | 393.1219
0 | LS3 | CD05 | CAV6 | 4619 | LS3_CD05:CAV6_D4619 |  | b54 resonator | 40 |  |  | 0.49566 | 328.413 | 427.1942 | 41.094 | 41.094 | 427.1942 | 328.1652 | 461.8982 | 462.146 |  | 393.4195
0 | [10 empty] | 0 | 328.1652 | 427.1942 | 41.094 | 41.094 | 427.1942 | 328.1652 | 462.146 | 462.146 |  | 393.6674
0 | [10 empty] | 0 | 328.1652 | 427.1942 | 41.094 | 41.094 | 427.1942 | 328.1652 | 462.146 | 462.146 |  | 393.6674
0 | [6 empty] | bellow |  |  |  | 0.099517 | 328.1154 | 427.1942 | 41.094 | 41.094 | 427.1942 | 328.0656 | 462.1957 | 462.2455 |  | 393.7171
0 | LS3 | CD05 | CAV7 | 4625 | LS3_CD05:CAV7_D4625 |  | b54 resonator | 40 |  |  | 0.49566 | 327.8178 | 427.1942 | 41.094 | 41.094 | 427.1942 | 327.57 | 462.4933 | 462.7412 |  | 394.0147
0 | [10 empty] | 0 | 327.57 | 427.1942 | 41.094 | 41.094 | 427.1942 | 327.57 | 462.7412 | 462.7412 |  | 394.2625
0 | [10 empty] | 0 | 327.57 | 427.1942 | 41.094 | 41.094 | 427.1942 | 327.57 | 462.7412 | 462.7412 |  | 394.2625
0 | [6 empty] | bellow |  |  |  | 0.099517 | 327.5202 | 427.1942 | 41.094 | 41.094 | 427.1942 | 327.4705 | 462.7909 | 462.8407 |  | 394.3123
0 | LS3 | CD05 | CAV8 | 4631 | LS3_CD05:CAV8_D4631 |  | b54 resonator | 40 |  |  | 0.49566 | 327.2226 | 427.1942 | 41.094 | 41.094 | 427.1942 | 326.9748 | 463.0885 | 463.3363 |  | 394.6099
0 | LS3 | CD05 | BLM | 4633 | LS3_CD05:BLM_D4633 |  | outside loss monitor |  |  |  | 0 | 326.9748 | 427.1942 | 41.094 | 41.094 | 427.1942 | 326.9748 | 463.3363 | 463.3363 |  | 394.8577
0 | [10 empty] | 0 | 326.9748 | 427.1942 | 41.094 | 41.094 | 427.1942 | 326.9748 | 463.3363 | 463.3363 |  | 394.8577
0 | [10 empty] | 0.1677 | 326.891 | 427.1942 | 41.094 | 41.094 | 427.1942 | 326.8071 | 463.4202 | 463.504 |  | 394.9416
0 | LS3 | CD05 | GV | 4635 | LS3_CD05:GV_D4635 |  | gate valve | 40 |  |  | 0.057404 | 326.7784 | 427.1942 | 41.094 | 41.094 | 427.1942 | 326.7497 | 463.5327 | 463.5614 |  | 395.0541
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 326.698 | 427.1942 | 41.094 | 41.094 | 427.1942 | 326.6462 | 463.6132 | 463.6649 |  | 395.1346
0 | LS3 | WD05 | BPM | 4637 | LS3_WD05:BPM_D4637 | BPM2 | position monitor |  |  |  | 0 | 326.6462 | 427.1942 | 41.094 | 41.094 | 427.1942 | 326.6462 | 463.6649 | 463.6649 |  | 395.1863
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 326.5736 | 427.1942 | 41.094 | 41.094 | 427.1942 | 326.5009 | 463.7376 | 463.8102 |  | 395.259
0 | LS3 | WD05 | PORT | 4638 | LS3_WD05:PORT_D4638 |  | port for future diagnostics |  |  |  | 0 | 326.5009 | 427.1942 | 41.094 | 41.094 | 427.1942 | 326.5009 | 463.8102 | 463.8102 |  | 395.3316
0 | [6 empty] | diagnostic box |  |  |  | 0 | 326.5009 | 427.1942 | 41.094 | 41.094 | 427.1942 | 326.5009 | 463.8102 | 463.8102 |  | 395.3316
0 | LS3 | WD05 | IP | 4638 | LS3_WD05:IP_D4638 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 326.5009 | 427.1942 | 41.094 | 41.094 | 427.1942 | 326.5009 | 463.8102 | 463.8102 |  | 395.3316
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 326.4353 | 427.1942 | 41.094 | 41.094 | 427.1942 | 326.3697 | 463.8758 | 463.9414 |  | 395.3972
0 | LS3 | CD06 | GV | 4640 | LS3_CD06:GV_D4640 |  | gate valve | 40 |  |  | 0.057404 | 326.341 | 427.1942 | 41.094 | 41.094 | 427.1942 | 326.3123 | 463.9701 | 463.9988 |  | 395.4915
0 | [6 empty] | bellows |  |  |  | 0.1677 | 326.2285 | 427.1942 | 41.094 | 41.094 | 427.1942 | 326.1446 | 464.0827 | 464.1665 |  | 395.6041
0 | LS3 | CD06 | CAV1 | 4644 | LS3_CD06:CAV1_D4644 |  | b54 resonator | 40 |  |  | 0.49566 | 325.8968 | 427.1942 | 41.094 | 41.094 | 427.1942 | 325.6489 | 464.4144 | 464.6622 |  | 395.9357
0 | LS3 | CD06 | BLM | 4647 | LS3_CD06:BLM_D4647 |  | outside loss monitor |  |  |  | 0 | 325.6489 | 427.1942 | 41.094 | 41.094 | 427.1942 | 325.6489 | 464.6622 | 464.6622 |  | 396.1836
0 | [10 empty] | 0 | 325.6489 | 427.1942 | 41.094 | 41.094 | 427.1942 | 325.6489 | 464.6622 | 464.6622 |  | 396.1836
0 | [6 empty] | bellow |  |  |  | 0.099517 | 325.5992 | 427.1942 | 41.094 | 41.094 | 427.1942 | 325.5494 | 464.712 | 464.7617 |  | 396.2333
0 | LS3 | CD06 | CAV2 | 4650 | LS3_CD06:CAV2_D4650 |  | b54 resonator | 40 |  |  | 0.49566 | 325.3016 | 427.1942 | 41.094 | 41.094 | 427.1942 | 325.0538 | 465.0096 | 465.2574 |  | 396.5309
0 | [10 empty] | 0 | 325.0538 | 427.1942 | 41.094 | 41.094 | 427.1942 | 325.0538 | 465.2574 | 465.2574 |  | 396.7787
0 | [10 empty] | 0 | 325.0538 | 427.1942 | 41.094 | 41.094 | 427.1942 | 325.0538 | 465.2574 | 465.2574 |  | 396.7787
0 | [6 empty] | bellow |  |  |  | 0.099517 | 325.004 | 427.1942 | 41.094 | 41.094 | 427.1942 | 324.9543 | 465.3071 | 465.3569 |  | 396.8285
0 | LS3 | CD06 | CAV3 | 4656 | LS3_CD06:CAV3_D4656 |  | b54 resonator | 40 |  |  | 0.49566 | 324.7064 | 427.1942 | 41.094 | 41.094 | 427.1942 | 324.4586 | 465.6047 | 465.8526 |  | 397.1261
0 | [10 empty] | 0 | 324.4586 | 427.1942 | 41.094 | 41.094 | 427.1942 | 324.4586 | 465.8526 | 465.8526 |  | 397.3739
0 | [10 empty] | 0 | 324.4586 | 427.1942 | 41.094 | 41.094 | 427.1942 | 324.4586 | 465.8526 | 465.8526 |  | 397.3739
0 | [6 empty] | bellow |  |  |  | 0.099517 | 324.4088 | 427.1942 | 41.094 | 41.094 | 427.1942 | 324.3591 | 465.9023 | 465.9521 |  | 397.4237
0 | LS3 | CD06 | CAV4 | 4662 | LS3_CD06:CAV4_D4662 |  | b54 resonator | 40 |  |  | 0.49566 | 324.1113 | 427.1942 | 41.094 | 41.094 | 427.1942 | 323.8634 | 466.1999 | 466.4477 |  | 397.7213
0 | LS3 | CD06 | BLM | 4664 | LS3_CD06:BLM_D4664 |  | outside loss monitor |  |  |  | 0 | 323.8634 | 427.1942 | 41.094 | 41.094 | 427.1942 | 323.8634 | 466.4477 | 466.4477 |  | 397.9691
0 | [10 empty] | 0 | 323.8634 | 427.1942 | 41.094 | 41.094 | 427.1942 | 323.8634 | 466.4477 | 466.4477 |  | 397.9691
0 | [6 empty] | bellow |  |  |  | 0.15525 | 323.7858 | 427.1942 | 41.094 | 41.094 | 427.1942 | 323.7082 | 466.5253 | 466.603 |  | 398.0467
0 | LS3 | CD06 | SOL1 | 4669 | LS3_CD06:SOL1_D4669 | S2 | solenoid+correctors | 40 |  |  | 0.5 | 323.4582 | 427.1942 | 41.094 | 41.094 | 427.1942 | 323.2082 | 466.853 | 467.103 |  | 398.3743
0 | [6 empty] | bellow |  |  |  | 0.15525 | 323.1305 | 427.1942 | 41.094 | 41.094 | 427.1942 | 323.0529 | 467.1806 | 467.2582 |  | 398.702
0 | LS3 | CD06 | CAV5 | 4675 | LS3_CD06:CAV5_D4675 |  | b54 resonator | 40 |  |  | 0.49566 | 322.8051 | 427.1942 | 41.094 | 41.094 | 427.1942 | 322.5573 | 467.5061 | 467.7539 |  | 399.0274
0 | [10 empty] | 0 | 322.5573 | 427.1942 | 41.094 | 41.094 | 427.1942 | 322.5573 | 467.7539 | 467.7539 |  | 399.2753
0 | [10 empty] | 0 | 322.5573 | 427.1942 | 41.094 | 41.094 | 427.1942 | 322.5573 | 467.7539 | 467.7539 |  | 399.2753
0 | [6 empty] | bellow |  |  |  | 0.099517 | 322.5075 | 427.1942 | 41.094 | 41.094 | 427.1942 | 322.4578 | 467.8036 | 467.8534 |  | 399.325
0 | LS3 | CD06 | CAV6 | 4681 | LS3_CD06:CAV6_D4681 |  | b54 resonator | 40 |  |  | 0.49566 | 322.2099 | 427.1942 | 41.094 | 41.094 | 427.1942 | 321.9621 | 468.1012 | 468.3491 |  | 399.6226
0 | [10 empty] | 0 | 321.9621 | 427.1942 | 41.094 | 41.094 | 427.1942 | 321.9621 | 468.3491 | 468.3491 |  | 399.8704
0 | [10 empty] | 0 | 321.9621 | 427.1942 | 41.094 | 41.094 | 427.1942 | 321.9621 | 468.3491 | 468.3491 |  | 399.8704
0 | [6 empty] | bellow |  |  |  | 0.099517 | 321.9123 | 427.1942 | 41.094 | 41.094 | 427.1942 | 321.8626 | 468.3988 | 468.4486 |  | 399.9202
0 | LS3 | CD06 | CAV7 | 4687 | LS3_CD06:CAV7_D4687 |  | b54 resonator | 40 |  |  | 0.49566 | 321.6147 | 427.1942 | 41.094 | 41.094 | 427.1942 | 321.3669 | 468.6964 | 468.9442 |  | 400.2178
0 | [10 empty] | 0 | 321.3669 | 427.1942 | 41.094 | 41.094 | 427.1942 | 321.3669 | 468.9442 | 468.9442 |  | 400.4656
0 | [10 empty] | 0 | 321.3669 | 427.1942 | 41.094 | 41.094 | 427.1942 | 321.3669 | 468.9442 | 468.9442 |  | 400.4656
0 | [6 empty] | bellow |  |  |  | 0.099517 | 321.3172 | 427.1942 | 41.094 | 41.094 | 427.1942 | 321.2674 | 468.994 | 469.0437 |  | 400.5154
0 | LS3 | CD06 | CAV8 | 4693 | LS3_CD06:CAV8_D4693 |  | b54 resonator | 40 |  |  | 0.49566 | 321.0196 | 427.1942 | 41.094 | 41.094 | 427.1942 | 320.7717 | 469.2916 | 469.5394 |  | 400.8129
0 | LS3 | CD06 | BLM | 4695 | LS3_CD06:BLM_D4695 |  | outside loss monitor |  |  |  | 0 | 320.7717 | 427.1942 | 41.094 | 41.094 | 427.1942 | 320.7717 | 469.5394 | 469.5394 |  | 401.0608
0 | [10 empty] | 0 | 320.7717 | 427.1942 | 41.094 | 41.094 | 427.1942 | 320.7717 | 469.5394 | 469.5394 |  | 401.0608
0 | [10 empty] | 0.1677 | 320.6879 | 427.1942 | 41.094 | 41.094 | 427.1942 | 320.604 | 469.6233 | 469.7071 |  | 401.1446
0 | LS3 | CD06 | GV | 4697 | LS3_CD06:GV_D4697 |  | gate valve | 40 |  |  | 0.057404 | 320.5753 | 427.1942 | 41.094 | 41.094 | 427.1942 | 320.5466 | 469.7358 | 469.7645 |  | 401.2572
0 | [6 empty] | bellow | 40 |  |  | 0.1035 | 320.4949 | 427.1942 | 41.094 | 41.094 | 427.1942 | 320.4431 | 469.8163 | 469.868 |  | 401.3376
0 | LS3 | WD06 | BPM | 4699 | LS3_WD06:BPM_D4699 | BPM2 | position monitor |  |  |  | 0 | 320.4431 | 427.1942 | 41.094 | 41.094 | 427.1942 | 320.4431 | 469.868 | 469.868 |  | 401.3894
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 320.3705 | 427.1942 | 41.094 | 41.094 | 427.1942 | 320.2979 | 469.9407 | 470.0133 |  | 401.462
0 | LS3 | WD06 | PORT | 4700 | LS3_WD06:PORT_D4700 |  | port for future diagnostics |  |  |  | 0 | 320.2979 | 427.1942 | 41.094 | 41.094 | 427.1942 | 320.2979 | 470.0133 | 470.0133 |  | 401.5347
0 | [6 empty] | diagnostic box |  |  |  | 0 | 320.2979 | 427.1942 | 41.094 | 41.094 | 427.1942 | 320.2979 | 470.0133 | 470.0133 |  | 401.5347
0 | LS3 | WD06 | IP | 4700 | LS3_WD06:IP_D4700 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 320.2979 | 427.1942 | 41.094 | 41.094 | 427.1942 | 320.2979 | 470.0133 | 470.0133 |  | 401.5347
0 | [6 empty] | bellow | 40 |  |  | 0.13122 | 320.2322 | 427.1942 | 41.094 | 41.094 | 427.1942 | 320.1666 | 470.0789 | 470.1445 |  | 401.6003
0 | LINAC SEGMENT 3 CM END | [9 empty] | 0 | 320.1666 | 427.1942 | 41.094 | 41.094 | 427.1942 | 320.1666 | 470.1445 | 470.1445 |  | 401.6659
0 | [8 empty] | -0.00046882 |  | 0 | 320.1666 | 427.1942 | 41.094 | 41.094 | 427.1942 | 320.1666 | 470.1445 | 470.1445 |  | 401.6659
0 | LS3 | BTS | FAV | 4702 | LS3_BTS:FAV_D4702 | FV | fast valve | 40 | 320.1326 | 320.133 | 0.067196 | 320.133 | 427.1942 | 41.094 | 41.094 | 427.1942 | 320.0994 | 470.1781 | 470.2117 |  | 401.6995
0 | [10 empty] | 0.5558 | 319.8215 | 427.1942 | 41.094 | 41.094 | 427.1942 | 319.5436 | 470.4896 | 470.7675 |  | 402.011
0 | LS3 | BTS | DCHV | 4709 | LS3_BTS:DCHV_D4709 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 319.4436 | 427.1942 | 41.094 | 41.094 | 427.1942 | 319.3436 | 470.8675 | 470.9675 |  | 402.3889
0 | [10 empty] | 0.2 | 319.2436 | 427.1942 | 41.094 | 41.094 | 427.1942 | 319.1436 | 471.0675 | 471.1675 |  | 402.5889
0 | LS3 | BTS | QV | 4713 | LS3_BTS:QV_D4713 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 319.0186 | 427.1942 | 41.094 | 41.094 | 427.1942 | 318.8936 | 471.2925 | 471.4175 |  | 402.8139
0 | [10 empty] | 0.25 | 318.7686 | 427.1942 | 41.094 | 41.094 | 427.1942 | 318.6436 | 471.5425 | 471.6675 |  | 403.0639
0 | LS3 | BTS | QH | 4718 | LS3_BTS:QH_D4718 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 318.5186 | 427.1942 | 41.094 | 41.094 | 427.1942 | 318.3936 | 471.7925 | 471.9175 |  | 403.3139
0 | [10 empty] | 0.66963 | 318.0588 | 427.1942 | 41.094 | 41.094 | 427.1942 | 317.724 | 472.2523 | 472.5871 |  | 403.7737
0 | LS3 | BTS | GV | 4726 | LS3_BTS:GV_D4726 | GV_M | gate valve | 63 | 2.176e-06 | 317.6865 | 0.075 | 317.6865 | 427.1942 | 41.094 | 41.094 | 427.1942 | 317.649 | 472.6246 | 472.6621 |  | 404.146
0 | [10 empty] | 2.2843 | 316.5068 | 427.1942 | 41.094 | 41.094 | 427.1942 | 315.3647 | 473.8043 | 474.9465 | 20170217 space adjusted | 405.3257
0 | LS3 | BTS | DCHV | 4750 | LS3_BTS:DCHV_D4750 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 315.2647 | 427.1942 | 41.094 | 41.094 | 427.1942 | 315.1647 | 475.0465 | 475.1465 |  | 406.5678
0 | [10 empty] | 0.13321 | 315.0981 | 427.1942 | 41.094 | 41.094 | 427.1942 | 315.0315 | 475.2131 | 475.2797 | 20170217 space adjusted | 406.7345
0 | LS3 | BTS | BPM | 4753 | LS3_BTS:BPM_D4753 | BPM3 | position monitor | 47.5 |  |  | 0.02 | 315.0215 | 427.1942 | 41.094 | 41.094 | 427.1942 | 315.0115 | 475.2897 | 475.2997 | 20170217 changed from .137 to .02 m (button diameter) | 406.8111
0 | [10 empty] | 0.04782 | 314.9876 | 427.1942 | 41.094 | 41.094 | 427.1942 | 314.9636 | 475.3236 | 475.3475 | 20170217 space adjusted | 406.845
0 | LS3 | BTS | QV | 4755 | LS3_BTS:QV_D4755 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 314.8386 | 427.1942 | 41.094 | 41.094 | 427.1942 | 314.7136 | 475.4725 | 475.5975 |  | 406.9939
0 | [10 empty] | 0.25 | 314.5886 | 427.1942 | 41.094 | 41.094 | 427.1942 | 314.4636 | 475.7225 | 475.8475 |  | 407.2439
0 | LS3 | BTS | QH | 4760 | LS3_BTS:QH_D4760 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 314.3386 | 427.1942 | 41.094 | 41.094 | 427.1942 | 314.2136 | 475.9725 | 476.0975 |  | 407.4939
0 | [10 empty] | 0.72812 | 313.8496 | 427.1942 | 41.094 | 41.094 | 427.1942 | 313.4855 | 476.4616 | 476.8256 | 20170217 space adjusted | 407.9829
0 | [6 empty] | bellow |  | 3.176e-06 | -0.067115 | 0.1035 | 313.4338 | 427.1942 | 41.094 | 41.094 | 427.1942 | 313.382 | 476.8774 | 476.9291 |  | 408.3987
0 | LS3 | BTS | BPM | 4769 | LS3_BTS:BPM_D4769 | BPM3 | position monitor |  | 313.382 | 313.382 | 0 | 313.382 | 427.1942 | 41.094 | 41.094 | 427.1942 | 313.382 | 476.9291 | 476.9291 |  | 408.4505
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 313.3094 | 427.1942 | 41.094 | 41.094 | 427.1942 | 313.2367 | 477.0018 | 477.0744 |  | 408.5231
0 | LS3 | BTS | PM | 4771 | LS3_BTS:PM_D4771 | PM3 | profile monitor |  | 4.176e-06 |  | 0 | 313.2367 | 427.1942 | 41.094 | 41.094 | 427.1942 | 313.2367 | 477.0744 | 477.0744 |  | 408.5958
0 | [6 empty] | diagnostic box |  | 313.2367 | 313.2367 | 0 | 313.2367 | 427.1942 | 41.094 | 41.094 | 427.1942 | 313.2367 | 477.0744 | 477.0744 |  | 408.5958
0 | LS3 | BTS | IP | 4771 | LS3_BTS:IP_D4771 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 313.2367 | 427.1942 | 41.094 | 41.094 | 427.1942 | 313.2367 | 477.0744 | 477.0744 |  | 408.5958
0 | LS3 | BTS | FAVS | 4771 | LS3_BTS:FAVS_D4771 | FVS | fast valve sensor |  |  |  | 0 | 313.2367 | 427.1942 | 41.094 | 41.094 | 427.1942 | 313.2367 | 477.0744 | 477.0744 |  | 408.5958
0 | [6 empty] | bellow |  |  |  | 0.13122 | 313.1711 | 427.1942 | 41.094 | 41.094 | 427.1942 | 313.1055 | 477.14 | 477.2056 |  | 408.6614
0 | [10 empty] | 0.59388 | 312.8086 | 427.1942 | 41.094 | 41.094 | 427.1942 | 312.5116 | 477.5026 | 477.7995 | 20170217 space adjusted | 409.0239
0 | LS3 | BTS | DCHV | 4779 | LS3_BTS:DCHV_D4779 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 312.4116 | 427.1942 | 41.094 | 41.094 | 427.1942 | 312.3116 | 477.8995 | 477.9995 |  | 409.4209
0 | [10 empty] | 0.2 | 312.2116 | 427.1942 | 41.094 | 41.094 | 427.1942 | 312.1116 | 478.0995 | 478.1995 |  | 409.6209
0 | LS3 | BTS | QH | 4783 | LS3_BTS:QH_D4783 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 311.9866 | 427.1942 | 41.094 | 41.094 | 427.1942 | 311.8616 | 478.3245 | 478.4495 |  | 409.8459
0 | [8 empty] | 3.176e-06 | -0.20458 | 1.2046 | 311.2594 | 427.1942 | 41.094 | 41.094 | 427.1942 | 310.6571 | 479.0518 | 479.6541 | 20170217 space adjusted | 410.5732
0 | LS3 | BTS | PM | 4797 | LS3_BTS:PM_D4797 | PM3 | profile monitor |  | 310.6571 | 310.6571 | 0 | 310.6571 | 427.1942 | 41.094 | 41.094 | 427.1942 | 310.6571 | 479.6541 | 479.6541 |  | 411.1755
0 | LS3 | BTS | IP | 4797 | LS3_BTS:IP_D4797 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 310.6571 | 427.1942 | 41.094 | 41.094 | 427.1942 | 310.6571 | 479.6541 | 479.6541 |  | 411.1755
0 | [10 empty] | 0.79543 | 310.2594 | 427.1942 | 41.094 | 41.094 | 427.1942 | 309.8616 | 480.0518 | 480.4495 | 20170217 space adjusted | 411.5732
0 | LS3 | BTS | QV | 4806 | LS3_BTS:QV_D4806 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 309.7366 | 427.1942 | 41.094 | 41.094 | 427.1942 | 309.6116 | 480.5745 | 480.6995 |  | 412.0959
0 | [8 empty] | 3.176e-06 | -0.35958 | 2.0116 | 308.6058 | 427.1942 | 41.094 | 41.094 | 427.1942 | 307.6001 | 481.7053 | 482.7111 |  | 413.2267
0 | LS3 | BTS | PM | 4827 | LS3_BTS:PM_D4827 | PM3 | profile monitor |  | 307.6001 | 307.6001 | 0 | 307.6001 | 427.1942 | 41.094 | 41.094 | 427.1942 | 307.6001 | 482.7111 | 482.7111 |  | 414.2325
0 | LS3 | BTS | IP | 4827 | LS3_BTS:IP_D4827 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 307.6001 | 427.1942 | 41.094 | 41.094 | 427.1942 | 307.6001 | 482.7111 | 482.7111 |  | 414.2325
0 | [8 empty] | 3.176e-06 | 0.001019 | 1.2914 | 306.9544 | 427.1942 | 41.094 | 41.094 | 427.1942 | 306.3087 | 483.3568 | 484.0025 | 20170217 space adjusted | 414.8782
0 | LS3 | BTS | DCHV | 4841 | LS3_BTS:DCHV_D4841 | C2 | correctors-X,Y | 50 | 306.2087 | 306.2087 | 0.2 | 306.2087 | 427.1942 | 41.094 | 41.094 | 427.1942 | 306.1087 | 484.1025 | 484.2025 |  | 415.6239
0 | [8 empty] | 3.176e-06 | -0.05468 | 0.1332 | 306.0421 | 427.1942 | 41.094 | 41.094 | 427.1942 | 305.9755 | 484.2691 | 484.3357 | 20170217 space adjusted | 415.7905
0 | LS3 | BTS | BPM | 4843 | LS3_BTS:BPM_D4843 | BPM3 | position monitor | 47.5 | 305.9655 | 305.9655 | 0.02 | 305.9655 | 427.1942 | 41.094 | 41.094 | 427.1942 | 305.9555 | 484.3457 | 484.3557 | 20170217 changed from .137 to .02 m (button diameter) | 415.8671
0 | [8 empty] | 0.0585 |  | 0.04782 | 305.9316 | 427.1942 | 41.094 | 41.094 | 427.1942 | 305.9076 | 484.3796 | 484.4035 | 20170217 space adjusted | 415.901
0 | LS3 | BTS | QH | 4845 | LS3_BTS:QH_D4845 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 305.7826 | 427.1942 | 41.094 | 41.094 | 427.1942 | 305.6576 | 484.5285 | 484.6535 |  | 416.0499
0 | [10 empty] | 1.5 | 304.9076 | 427.1942 | 41.094 | 41.094 | 427.1942 | 304.1576 | 485.4035 | 486.1535 |  | 416.9249
0 | LS3 | BTS | PM | 4862 | LS3_BTS:PM_D4862 | PM3 | profile monitor |  |  |  | 0 | 304.1576 | 427.1942 | 41.094 | 41.094 | 427.1942 | 304.1576 | 486.1535 | 486.1535 |  | 417.6749
0 | LS3 | BTS | IP | 4862 | LS3_BTS:IP_D4862 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 304.1576 | 427.1942 | 41.094 | 41.094 | 427.1942 | 304.1576 | 486.1535 | 486.1535 |  | 417.6749
0 | [10 empty] | 0.5 | 303.9076 | 427.1942 | 41.094 | 41.094 | 427.1942 | 303.6576 | 486.4035 | 486.6535 |  | 417.9249
0 | LS3 | BTS | QV | 4868 | LS3_BTS:QV_D4868 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 303.5326 | 427.1942 | 41.094 | 41.094 | 427.1942 | 303.4076 | 486.7785 | 486.9035 |  | 418.2999
0 | [8 empty] | 3.176e-06 | 1.3481 | 0.12643 | 303.3444 | 427.1942 | 41.094 | 41.094 | 427.1942 | 303.2812 | 486.9667 | 487.0299 | 20170217 space adjusted | 418.4881
0 | LS3 | BTS | GV | 4871 | LS3_BTS:GV_D4871 | GV_M | gate valve | 63 | 303.2437 | 303.2437 | 0.075 | 303.2437 | 427.1942 | 41.094 | 41.094 | 427.1942 | 303.2062 | 487.0674 | 487.1049 |  | 418.5888
0 | [10 empty] | 1.3481 | 302.5322 | 427.1942 | 41.094 | 41.094 | 427.1942 | 301.8581 | 487.779 | 488.453 | 20170217 space adjusted | 419.3003
0 | [6 empty] | bellow |  | 3.176e-06 | -0.075651 | 0.17915 | 301.7686 | 427.1942 | 41.094 | 41.094 | 427.1942 | 301.679 | 488.5426 | 488.6322 | 20170217 space adjusted | 420.064
0 | LS3 | BTS | BPM | 4886 | LS3_BTS:BPM_D4886 | BPM3 | position monitor |  | 301.679 | 301.679 | 0 | 301.679 | 427.1942 | 41.094 | 41.094 | 427.1942 | 301.679 | 488.6322 | 488.6322 |  | 420.1535
0 | [6 empty] | diagnostic box |  | 3.176e-06 |  | 0.14528 | 301.6063 | 427.1942 | 41.094 | 41.094 | 427.1942 | 301.5337 | 488.7048 | 488.7774 |  | 420.2262
0 | LS3 | BTS | PORT | 4888 | LS3_BTS:PORT_D4888 |  | port for future diagnostics |  | 301.5337 | 301.5337 | 0 | 301.5337 | 427.1942 | 41.094 | 41.094 | 427.1942 | 301.5337 | 488.7774 | 488.7774 |  | 420.2988
0 | [6 empty] | diagnostic box |  |  |  | 0 | 301.5337 | 427.1942 | 41.094 | 41.094 | 427.1942 | 301.5337 | 488.7774 | 488.7774 |  | 420.2988
0 | LS3 | BTS | IP | 4888 | LS3_BTS:IP_D4888 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 301.5337 | 427.1942 | 41.094 | 41.094 | 427.1942 | 301.5337 | 488.7774 | 488.7774 |  | 420.2988
0 | [6 empty] | bellow |  |  |  | 0.13122 | 301.4681 | 427.1942 | 41.094 | 41.094 | 427.1942 | 301.4025 | 488.8431 | 488.9087 |  | 420.3644
0 | [10 empty] | 1.2988 | 300.7531 | 427.1942 | 41.094 | 41.094 | 427.1942 | 300.1036 | 489.5581 | 490.2075 | 20170217 space adjusted | 421.0795
0 | LS3 | BTS | DCHV | 4903 | LS3_BTS:DCHV_D4903 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 300.0036 | 427.1942 | 41.094 | 41.094 | 427.1942 | 299.9036 | 490.3075 | 490.4075 |  | 421.8289
0 | [10 empty] | 0.2 | 299.8036 | 427.1942 | 41.094 | 41.094 | 427.1942 | 299.7036 | 490.5075 | 490.6075 |  | 422.0289
0 | LS3 | BTS | QH | 4907 | LS3_BTS:QH_D4907 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 299.5786 | 427.1942 | 41.094 | 41.094 | 427.1942 | 299.4536 | 490.7325 | 490.8575 |  | 422.2539
0 | [10 empty] | 2 | 298.4536 | 427.1942 | 41.094 | 41.094 | 427.1942 | 297.4536 | 491.8575 | 492.8575 |  | 423.3789
0 | LS3 | BTS | QV | 4930 | LS3_BTS:QV_D4930 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 297.3286 | 427.1942 | 41.094 | 41.094 | 427.1942 | 297.2036 | 492.9825 | 493.1075 |  | 424.5039
0 | [10 empty] | 1.5495 | 296.4289 | 427.1942 | 41.094 | 41.094 | 427.1942 | 295.6541 | 493.8823 | 494.657 |  | 425.4036
0 | [10 empty] | 0.096 | 295.6061 | 427.1942 | 41.094 | 41.094 | 427.1942 | 295.5581 | 494.705 | 494.753 |  | 426.2264
0 | [10 empty] | 0 | 295.5581 | 427.1942 | 41.094 | 41.094 | 427.1942 | 295.5581 | 494.753 | 494.753 |  | 426.2744
0 | [6 empty] | bellow |  |  |  | 0.094 | 295.5111 | 427.1942 | 41.094 | 41.094 | 427.1942 | 295.4641 | 494.8 | 494.847 |  | 426.3214
0 | LS3 | BTS | PORT | 4948 | LS3_BTS:PORT_D4948 |  | port for future diagnostics |  |  |  | 0 | 295.4641 | 427.1942 | 41.094 | 41.094 | 427.1942 | 295.4641 | 494.847 | 494.847 |  | 426.3684
0 | [6 empty] | vacuum box |  | 2.53e-06 | 0.090994 | 6e-06 | 295.4641 | 427.1942 | 41.094 | 41.094 | 427.1942 | 295.4641 | 494.847 | 494.847 | 20170217 space adjusted | 426.3684
0 | LS3 | BTS | IP | 4948 | LS3_BTS:IP_D4948 | IP_S | ion pump, 75 l/s |  | 295.4641 | 295.4641 | 0 | 295.4641 | 427.1942 | 41.094 | 41.094 | 427.1942 | 295.4641 | 494.847 | 494.847 |  | 426.3684
0 | [6 empty] | bellow |  |  |  | 0.18999 | 295.3691 | 427.1942 | 41.094 | 41.094 | 427.1942 | 295.2741 | 494.942 | 495.037 | 20170217 space adjusted | 426.4634
0 | [8 empty] | 3.176e-06 | 0.001019 | 1.3735 | 294.5874 | 427.1942 | 41.094 | 41.094 | 427.1942 | 293.9007 | 495.7237 | 496.4105 | 20170217 space adjusted | 427.2451
0 | LS3 | BTS | DCHV | 4965 | LS3_BTS:DCHV_D4965 | C2 | correctors-X,Y | 50 | 293.8007 | 293.8007 | 0.2 | 293.8007 | 427.1942 | 41.094 | 41.094 | 427.1942 | 293.7007 | 496.5105 | 496.6105 |  | 428.0319
0 | [8 empty] | 3.176e-06 | -0.05468 | 0.1332 | 293.6341 | 427.1942 | 41.094 | 41.094 | 427.1942 | 293.5675 | 496.6771 | 496.7437 | 20170217 space adjusted | 428.1985
0 | LS3 | BTS | BPM | 4968 | LS3_BTS:BPM_D4968 | BPM3 | position monitor | 47.5 | 293.5575 | 293.5575 | 0.02 | 293.5575 | 427.1942 | 41.094 | 41.094 | 427.1942 | 293.5475 | 496.7537 | 496.7637 | 20170217 changed from .137 to .02 m (button diameter) | 428.2751
0 | [8 empty] | 0.0585 |  | 0.04782 | 293.5236 | 427.1942 | 41.094 | 41.094 | 427.1942 | 293.4996 | 496.7876 | 496.8115 | 20170217 space adjusted | 428.309
0 | LS3 | BTS | QH | 4969 | LS3_BTS:QH_D4969 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 293.3746 | 427.1942 | 41.094 | 41.094 | 427.1942 | 293.2496 | 496.9365 | 497.0615 |  | 428.4579
0 | [10 empty] | 2 | 292.2496 | 427.1942 | 41.094 | 41.094 | 427.1942 | 291.2496 | 498.0615 | 499.0615 |  | 429.5829
0 | LS3 | BTS | QV | 4992 | LS3_BTS:QV_D4992 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 291.1246 | 427.1942 | 41.094 | 41.094 | 427.1942 | 290.9996 | 499.1865 | 499.3115 |  | 430.7079
0 | [10 empty] | 1.6251 | 290.1871 | 427.1942 | 41.094 | 41.094 | 427.1942 | 289.3745 | 500.1241 | 500.9367 | 20170217 space adjusted | 431.6455
0 | [6 empty] | bellow |  | 3.176e-06 | -0.075649 | 0.1035 | 289.3227 | 427.1942 | 41.094 | 41.094 | 427.1942 | 289.271 | 500.9884 | 501.0402 |  | 432.5098
0 | LS3 | BTS | BPM | 5010 | LS3_BTS:BPM_D5010 | BPM3 | position monitor |  | 289.271 | 289.271 | 0 | 289.271 | 427.1942 | 41.094 | 41.094 | 427.1942 | 289.271 | 501.0402 | 501.0402 |  | 432.5615
0 | [6 empty] | diagnostic box |  | 3.176e-06 |  | 0.14528 | 289.1983 | 427.1942 | 41.094 | 41.094 | 427.1942 | 289.1257 | 501.1128 | 501.1854 |  | 432.6342
0 | LS3 | BTS | PORT | 5012 | LS3_BTS:PORT_D5012 |  | port for future diagnostics |  | 289.1257 | 289.1257 | 0 | 289.1257 | 427.1942 | 41.094 | 41.094 | 427.1942 | 289.1257 | 501.1854 | 501.1854 |  | 432.7068
0 | [6 empty] | diagnostic box |  |  |  | 0 | 289.1257 | 427.1942 | 41.094 | 41.094 | 427.1942 | 289.1257 | 501.1854 | 501.1854 |  | 432.7068
0 | LS3 | BTS | IP | 5012 | LS3_BTS:IP_D5012 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 289.1257 | 427.1942 | 41.094 | 41.094 | 427.1942 | 289.1257 | 501.1854 | 501.1854 |  | 432.7068
0 | [6 empty] | bellow |  |  |  | 0.13122 | 289.0601 | 427.1942 | 41.094 | 41.094 | 427.1942 | 288.9945 | 501.251 | 501.3167 |  | 432.7724
0 | [10 empty] | 1.2989 | 288.3451 | 427.1942 | 41.094 | 41.094 | 427.1942 | 287.6956 | 501.9661 | 502.6155 | 20170217 space adjusted | 433.4875
0 | LS3 | BTS | DCHV | 5027 | LS3_BTS:DCHV_D5027 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 287.5956 | 427.1942 | 41.094 | 41.094 | 427.1942 | 287.4956 | 502.7155 | 502.8155 |  | 434.2369
0 | [10 empty] | 0.2 | 287.3956 | 427.1942 | 41.094 | 41.094 | 427.1942 | 287.2956 | 502.9155 | 503.0155 |  | 434.4369
0 | LS3 | BTS | QH | 5031 | LS3_BTS:QH_D5031 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 287.1706 | 427.1942 | 41.094 | 41.094 | 427.1942 | 287.0456 | 503.1405 | 503.2655 |  | 434.6619
0 | [10 empty] | 2 | 286.0456 | 427.1942 | 41.094 | 41.094 | 427.1942 | 285.0456 | 504.2655 | 505.2655 |  | 435.7869
0 | LS3 | BTS | QV | 5054 | LS3_BTS:QV_D5054 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 284.9206 | 427.1942 | 41.094 | 41.094 | 427.1942 | 284.7956 | 505.3905 | 505.5155 |  | 436.9119
0 | [10 empty] | 1.5495 | 284.0209 | 427.1942 | 41.094 | 41.094 | 427.1942 | 283.2461 | 506.2903 | 507.065 |  | 437.8116
0 | [10 empty] | 0.096 | 283.1981 | 427.1942 | 41.094 | 41.094 | 427.1942 | 283.1501 | 507.113 | 507.161 |  | 438.6344
0 | [10 empty] | 0 | 283.1501 | 427.1942 | 41.094 | 41.094 | 427.1942 | 283.1501 | 507.161 | 507.161 |  | 438.6824
0 | [6 empty] | bellow |  |  |  | 0.094 | 283.1031 | 427.1942 | 41.094 | 41.094 | 427.1942 | 283.0561 | 507.208 | 507.255 |  | 438.7294
0 | LS3 | BTS | PORT | 5073 | LS3_BTS:PORT_D5073 |  | port for future diagnostics |  |  |  | 0 | 283.0561 | 427.1942 | 41.094 | 41.094 | 427.1942 | 283.0561 | 507.255 | 507.255 |  | 438.7764
0 | [6 empty] | vacuum box |  | 2.53e-06 | 0.090994 | 6e-06 | 283.0561 | 427.1942 | 41.094 | 41.094 | 427.1942 | 283.0561 | 507.255 | 507.255 | 20170217 space adjusted | 438.7764
0 | LS3 | BTS | IP | 5073 | LS3_BTS:IP_D5073 | IP_S | ion pump, 75 l/s |  | 283.0561 | 283.0561 | 0 | 283.0561 | 427.1942 | 41.094 | 41.094 | 427.1942 | 283.0561 | 507.255 | 507.255 |  | 438.7764
0 | [6 empty] | bellow |  |  |  | 0.18999 | 282.9611 | 427.1942 | 41.094 | 41.094 | 427.1942 | 282.8661 | 507.35 | 507.445 | 20170217 space adjusted | 438.8714
0 | [10 empty] | 1.3735 | 282.1794 | 427.1942 | 41.094 | 41.094 | 427.1942 | 281.4927 | 508.1317 | 508.8185 |  | 439.6531
0 | [8 empty] | 3.176e-06 | 0.001019 | 0 | 281.4927 | 427.1942 | 41.094 | 41.094 | 427.1942 | 281.4927 | 508.8185 | 508.8185 |  | 440.3399
0 | LS3 | BTS | DCHV | 5089 | LS3_BTS:DCHV_D5089 | C2 | correctors-X,Y | 50 | 281.3927 | 281.3927 | 0.2 | 281.3927 | 427.1942 | 41.094 | 41.094 | 427.1942 | 281.2927 | 508.9185 | 509.0185 |  | 440.4399
0 | [8 empty] | 3.176e-06 | -0.05468 | 0.1332 | 281.2261 | 427.1942 | 41.094 | 41.094 | 427.1942 | 281.1595 | 509.0851 | 509.1517 | 20170217 space adjusted | 440.6065
0 | LS3 | BTS | BPM | 5092 | LS3_BTS:BPM_D5092 | BPM3 | position monitor | 47.5 | 281.1495 | 281.1495 | 0.02 | 281.1495 | 427.1942 | 41.094 | 41.094 | 427.1942 | 281.1395 | 509.1617 | 509.1717 | 20170217 changed from .137 to .02 m (button diameter) | 440.6831
0 | [8 empty] | 0.0585 |  | 0.04782 | 281.1156 | 427.1942 | 41.094 | 41.094 | 427.1942 | 281.0916 | 509.1956 | 509.2195 | 20170217 space adjusted | 440.717
0 | LS3 | BTS | QH | 5093 | LS3_BTS:QH_D5093 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 280.9666 | 427.1942 | 41.094 | 41.094 | 427.1942 | 280.8416 | 509.3445 | 509.4695 |  | 440.8659
0 | [10 empty] | 2 | 279.8416 | 427.1942 | 41.094 | 41.094 | 427.1942 | 278.8416 | 510.4695 | 511.4695 |  | 441.9909
0 | LS3 | BTS | QV | 5116 | LS3_BTS:QV_D5116 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 278.7166 | 427.1942 | 41.094 | 41.094 | 427.1942 | 278.5916 | 511.5945 | 511.7195 |  | 443.1159
0 | [10 empty] | 1.6251 | 277.7791 | 427.1942 | 41.094 | 41.094 | 427.1942 | 276.9665 | 512.5321 | 513.3447 |  | 444.0535
0 | [6 empty] | bellow |  | -0.075645 |  | 0.1035 | 276.9147 | 427.1942 | 41.094 | 41.094 | 427.1942 | 276.863 | 513.3964 | 513.4482 |  | 444.9178
0 | LS3 | BTS | BPM | 5134 | LS3_BTS:BPM_D5134 | BPM3 | position monitor |  | 2.03e-06 | 276.863 | 0 | 276.863 | 427.1942 | 41.094 | 41.094 | 427.1942 | 276.863 | 513.4482 | 513.4482 |  | 444.9695
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 276.7904 | 427.1942 | 41.094 | 41.094 | 427.1942 | 276.7177 | 513.5208 | 513.5934 |  | 445.0422
0 | LS3 | BTS | PORT | 5136 | LS3_BTS:PORT_D5136 |  | port for future diagnostics |  |  |  | 0 | 276.7177 | 427.1942 | 41.094 | 41.094 | 427.1942 | 276.7177 | 513.5934 | 513.5934 |  | 445.1148
0 | [6 empty] | diagnostic box |  |  |  | 0 | 276.7177 | 427.1942 | 41.094 | 41.094 | 427.1942 | 276.7177 | 513.5934 | 513.5934 |  | 445.1148
0 | LS3 | BTS | IP | 5136 | LS3_BTS:IP_D5136 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 276.7177 | 427.1942 | 41.094 | 41.094 | 427.1942 | 276.7177 | 513.5934 | 513.5934 |  | 445.1148
0 | [6 empty] | bellow |  |  |  | 0.13122 | 276.6521 | 427.1942 | 41.094 | 41.094 | 427.1942 | 276.5865 | 513.659 | 513.7247 |  | 445.1804
0 | [10 empty] | 1.2989 | 275.9371 | 427.1942 | 41.094 | 41.094 | 427.1942 | 275.2876 | 514.3741 | 515.0235 |  | 445.8955
0 | LS3 | BTS | DCHV | 5151 | LS3_BTS:DCHV_D5151 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 275.1876 | 427.1942 | 41.094 | 41.094 | 427.1942 | 275.0876 | 515.1235 | 515.2235 |  | 446.6449
0 | [10 empty] | 0.2 | 274.9876 | 427.1942 | 41.094 | 41.094 | 427.1942 | 274.8876 | 515.3235 | 515.4235 |  | 446.8449
0 | LS3 | BTS | QH | 5155 | LS3_BTS:QH_D5155 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 274.7626 | 427.1942 | 41.094 | 41.094 | 427.1942 | 274.6376 | 515.5485 | 515.6735 |  | 447.0699
0 | [10 empty] | 2 | 273.6376 | 427.1942 | 41.094 | 41.094 | 427.1942 | 272.6376 | 516.6735 | 517.6735 |  | 448.1949
0 | LS3 | BTS | QV | 5178 | LS3_BTS:QV_D5178 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 272.5126 | 427.1942 | 41.094 | 41.094 | 427.1942 | 272.3876 | 517.7985 | 517.9235 |  | 449.3199
0 | [10 empty] | 1.5495 | 271.6129 | 427.1942 | 41.094 | 41.094 | 427.1942 | 270.8381 | 518.6983 | 519.473 |  | 450.2196
0 | [10 empty] | 0.096 | 270.7901 | 427.1942 | 41.094 | 41.094 | 427.1942 | 270.7421 | 519.521 | 519.569 |  | 451.0424
0 | [10 empty] | 0 | 270.7421 | 427.1942 | 41.094 | 41.094 | 427.1942 | 270.7421 | 519.569 | 519.569 |  | 451.0904
0 | [6 empty] | bellow |  |  |  | 0.094 | 270.6951 | 427.1942 | 41.094 | 41.094 | 427.1942 | 270.6481 | 519.616 | 519.663 |  | 451.1374
0 | LS3 | BTS | PORT | 5197 | LS3_BTS:PORT_D5197 |  | port for future diagnostics |  |  |  | 0 | 270.6481 | 427.1942 | 41.094 | 41.094 | 427.1942 | 270.6481 | 519.663 | 519.663 |  | 451.1844
0 | [6 empty] | vacuum box |  |  |  | 0.091 | 270.6026 | 427.1942 | 41.094 | 41.094 | 427.1942 | 270.5571 | 519.7085 | 519.754 |  | 451.2299
0 | LS3 | BTS | IP | 5198 | LS3_BTS:IP_D5198 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 270.5571 | 427.1942 | 41.094 | 41.094 | 427.1942 | 270.5571 | 519.754 | 519.754 |  | 451.2754
0 | [6 empty] | bellow |  |  |  | 0.099 | 270.5076 | 427.1942 | 41.094 | 41.094 | 427.1942 | 270.4581 | 519.8035 | 519.853 |  | 451.3249
0 | [10 empty] | 1.3745 | 269.7709 | 427.1942 | 41.094 | 41.094 | 427.1942 | 269.0836 | 520.5403 | 521.2275 |  | 452.0616
0 | [10 empty] | 0 | 269.0836 | 427.1942 | 41.094 | 41.094 | 427.1942 | 269.0836 | 521.2275 | 521.2275 |  | 452.7489
0 | LS3 | BTS | DCHV | 5213 | LS3_BTS:DCHV_D5213 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 268.9836 | 427.1942 | 41.094 | 41.094 | 427.1942 | 268.8836 | 521.3275 | 521.4275 |  | 452.8489
0 | [10 empty] | 0.14218 | 268.8126 | 427.1942 | 41.094 | 41.094 | 427.1942 | 268.7415 | 521.4986 | 521.5697 |  | 453.02
0 | LS3 | BTS | BPM | 5216 | LS3_BTS:BPM_D5216 | BPM3 | position monitor | 47.5 | -2.176e-06 | 268.7415 | 0 | 268.7415 | 427.1942 | 41.094 | 41.094 | 427.1942 | 268.7415 | 521.5697 | 521.5697 |  | 453.0911
0 | [10 empty] | 0.057821 | 268.7126 | 427.1942 | 41.094 | 41.094 | 427.1942 | 268.6836 | 521.5986 | 521.6275 |  | 453.12
0 | LS3 | BTS | QH | 5218 | LS3_BTS:QH_D5218 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 268.5586 | 427.1942 | 41.094 | 41.094 | 427.1942 | 268.4336 | 521.7525 | 521.8775 |  | 453.2739
0 | [10 empty] | 2 | 267.4336 | 427.1942 | 41.094 | 41.094 | 427.1942 | 266.4336 | 522.8775 | 523.8775 |  | 454.3989
0 | LS3 | BTS | QV | 5240 | LS3_BTS:QV_D5240 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 266.3086 | 427.1942 | 41.094 | 41.094 | 427.1942 | 266.1836 | 524.0025 | 524.1275 |  | 455.5239
0 | [10 empty] | 1.6252 | 265.3711 | 427.1942 | 41.094 | 41.094 | 427.1942 | 264.5585 | 524.9401 | 525.7527 |  | 456.4615
0 | [6 empty] | bellow |  | -0.075663 |  | 0.1035 | 264.5067 | 427.1942 | 41.094 | 41.094 | 427.1942 | 264.455 | 525.8044 | 525.8562 |  | 457.3258
0 | LS3 | BTS | BPM | 5259 | LS3_BTS:BPM_D5259 | BPM3 | position monitor |  | 2.03e-06 | 264.455 | 0 | 264.455 | 427.1942 | 41.094 | 41.094 | 427.1942 | 264.455 | 525.8562 | 525.8562 |  | 457.3775
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 264.3823 | 427.1942 | 41.094 | 41.094 | 427.1942 | 264.3097 | 525.9288 | 526.0015 |  | 457.4502
0 | LS3 | BTS | PORT | 5260 | LS3_BTS:PORT_D5260 |  | port for future diagnostics |  |  |  | 0 | 264.3097 | 427.1942 | 41.094 | 41.094 | 427.1942 | 264.3097 | 526.0015 | 526.0015 |  | 457.5228
0 | [6 empty] | diagnostic box |  | 2.03e-06 | 264.3097 | 0 | 264.3097 | 427.1942 | 41.094 | 41.094 | 427.1942 | 264.3097 | 526.0015 | 526.0015 |  | 457.5228
0 | LS3 | BTS | IP | 5260 | LS3_BTS:IP_D5260 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 264.3097 | 427.1942 | 41.094 | 41.094 | 427.1942 | 264.3097 | 526.0015 | 526.0015 |  | 457.5228
0 | [6 empty] | bellow |  |  |  | 0.13122 | 264.2441 | 427.1942 | 41.094 | 41.094 | 427.1942 | 264.1785 | 526.0671 | 526.1327 |  | 457.5884
0 | [10 empty] | 1.2988 | 263.5291 | 427.1942 | 41.094 | 41.094 | 427.1942 | 262.8796 | 526.7821 | 527.4315 |  | 458.3035
0 | LS3 | BTS | DCHV | 5275 | LS3_BTS:DCHV_D5275 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 262.7796 | 427.1942 | 41.094 | 41.094 | 427.1942 | 262.6796 | 527.5315 | 527.6315 |  | 459.0529
0 | [10 empty] | 0.2 | 262.5796 | 427.1942 | 41.094 | 41.094 | 427.1942 | 262.4796 | 527.7315 | 527.8315 |  | 459.2529
0 | LS3 | BTS | QH | 5280 | LS3_BTS:QH_D5280 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 262.3546 | 427.1942 | 41.094 | 41.094 | 427.1942 | 262.2296 | 527.9565 | 528.0815 |  | 459.4779
0 | [10 empty] | 2 | 261.2296 | 427.1942 | 41.094 | 41.094 | 427.1942 | 260.2296 | 529.0815 | 530.0815 |  | 460.6029
0 | LS3 | BTS | QV | 5302 | LS3_BTS:QV_D5302 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 260.1046 | 427.1942 | 41.094 | 41.094 | 427.1942 | 259.9796 | 530.2065 | 530.3315 |  | 461.7279
0 | [10 empty] | 1.5495 | 259.2049 | 427.1942 | 41.094 | 41.094 | 427.1942 | 258.4301 | 531.1063 | 531.881 |  | 462.6276
0 | [10 empty] | 0.096 | 258.3821 | 427.1942 | 41.094 | 41.094 | 427.1942 | 258.3341 | 531.929 | 531.977 |  | 463.4504
0 | [10 empty] | 0 | 258.3341 | 427.1942 | 41.094 | 41.094 | 427.1942 | 258.3341 | 531.977 | 531.977 |  | 463.4984
0 | [6 empty] | bellow |  |  |  | 0.094 | 258.2871 | 427.1942 | 41.094 | 41.094 | 427.1942 | 258.2401 | 532.024 | 532.071 |  | 463.5454
0 | LS3 | BTS | PORT | 5321 | LS3_BTS:PORT_D5321 |  | port for future diagnostics |  |  |  | 0 | 258.2401 | 427.1942 | 41.094 | 41.094 | 427.1942 | 258.2401 | 532.071 | 532.071 |  | 463.5924
0 | [6 empty] | vacuum box |  |  |  | 0 | 258.2401 | 427.1942 | 41.094 | 41.094 | 427.1942 | 258.2401 | 532.071 | 532.071 |  | 463.5924
0 | LS3 | BTS | IP | 5321 | LS3_BTS:IP_D5321 | IP_S | ion pump, 75 l/s |  | 2.176e-06 | 258.2401 | 0 | 258.2401 | 427.1942 | 41.094 | 41.094 | 427.1942 | 258.2401 | 532.071 | 532.071 |  | 463.5924
0 | [6 empty] | bellow |  |  |  | 0.19 | 258.1451 | 427.1942 | 41.094 | 41.094 | 427.1942 | 258.0501 | 532.166 | 532.261 |  | 463.6874
0 | [10 empty] | 1.3745 | 257.3629 | 427.1942 | 41.094 | 41.094 | 427.1942 | 256.6756 | 532.9483 | 533.6355 |  | 464.4696
0 | [10 empty] | 0 | 256.6756 | 427.1942 | 41.094 | 41.094 | 427.1942 | 256.6756 | 533.6355 | 533.6355 |  | 465.1569
0 | LS3 | BTS | DCHV | 5337 | LS3_BTS:DCHV_D5337 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 256.5756 | 427.1942 | 41.094 | 41.094 | 427.1942 | 256.4756 | 533.7355 | 533.8355 |  | 465.2569
0 | [10 empty] | 0.14219 | 256.4045 | 427.1942 | 41.094 | 41.094 | 427.1942 | 256.3334 | 533.9066 | 533.9777 |  | 465.428
0 | LS3 | BTS | BPM | 5340 | LS3_BTS:BPM_D5340 | BPM3 | position monitor | 47.5 | 2.176e-06 | 256.3335 | 0 | 256.3334 | 427.1942 | 41.094 | 41.094 | 427.1942 | 256.3334 | 533.9777 | 533.9777 |  | 465.4991
0 | [10 empty] | 0.057809 | 256.3045 | 427.1942 | 41.094 | 41.094 | 427.1942 | 256.2756 | 534.0066 | 534.0355 |  | 465.528
0 | LS3 | BTS | QH | 5342 | LS3_BTS:QH_D5342 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 256.1506 | 427.1942 | 41.094 | 41.094 | 427.1942 | 256.0256 | 534.1605 | 534.2855 |  | 465.6819
0 | [10 empty] | 2 | 255.0256 | 427.1942 | 41.094 | 41.094 | 427.1942 | 254.0256 | 535.2855 | 536.2855 |  | 466.8069
0 | LS3 | BTS | QV | 5364 | LS3_BTS:QV_D5364 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 253.9006 | 427.1942 | 41.094 | 41.094 | 427.1942 | 253.7756 | 536.4105 | 536.5355 |  | 467.9319
0 | [10 empty] | 1.4852 | 253.033 | 427.1942 | 41.094 | 41.094 | 427.1942 | 252.2904 | 537.2781 | 538.0207 |  | 468.7995
0 | [6 empty] | bellow |  |  |  | 0.1035 | 252.2387 | 427.1942 | 41.094 | 41.094 | 427.1942 | 252.1869 | 538.0725 | 538.1242 |  | 469.5938
0 | LS3 | BTS | BPM | 5381 | LS3_BTS:BPM_D5381 | BPM3 | position monitor |  | 2.176e-06 | 252.1869 | 0 | 252.1869 | 427.1942 | 41.094 | 41.094 | 427.1942 | 252.1869 | 538.1242 | 538.1242 |  | 469.6456
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 252.1143 | 427.1942 | 41.094 | 41.094 | 427.1942 | 252.0416 | 538.1969 | 538.2695 |  | 469.7182
0 | LS3 | BTS | PORT | 5383 | LS3_BTS:PORT_D5383 |  | port for future diagnostics |  |  |  | 0 | 252.0416 | 427.1942 | 41.094 | 41.094 | 427.1942 | 252.0416 | 538.2695 | 538.2695 |  | 469.7909
0 | [6 empty] | diagnostic box |  |  |  | 0 | 252.0416 | 427.1942 | 41.094 | 41.094 | 427.1942 | 252.0416 | 538.2695 | 538.2695 |  | 469.7909
0 | LS3 | BTS | IP | 5383 | LS3_BTS:IP_D5383 | IP_S | ion pump, 75 l/s |  | 2.176e-06 | 252.0416 | 0 | 252.0416 | 427.1942 | 41.094 | 41.094 | 427.1942 | 252.0416 | 538.2695 | 538.2695 |  | 469.7909
0 | [6 empty] | bellow |  |  |  | 0.13122 | 251.976 | 427.1942 | 41.094 | 41.094 | 427.1942 | 251.9104 | 538.3351 | 538.4007 |  | 469.8565
0 | [10 empty] | 0.35978 | 251.7305 | 427.1942 | 41.094 | 41.094 | 427.1942 | 251.5506 | 538.5806 | 538.7605 |  | 470.102
0 | LS3 | BTS | DCHV | 5389 | LS3_BTS:DCHV_D5389 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 251.4506 | 427.1942 | 41.094 | 41.094 | 427.1942 | 251.3506 | 538.8605 | 538.9605 |  | 470.3819
0 | [10 empty] | 0.2 | 251.2506 | 427.1942 | 41.094 | 41.094 | 427.1942 | 251.1506 | 539.0605 | 539.1605 |  | 470.5819
0 | LS3 | BTS | QH | 5393 | LS3_BTS:QH_D5393 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 251.0256 | 427.1942 | 41.094 | 41.094 | 427.1942 | 250.9006 | 539.2855 | 539.4105 |  | 470.8069
0 | [10 empty] | 1.25 | 250.2756 | 427.1942 | 41.094 | 41.094 | 427.1942 | 249.6506 | 540.0355 | 540.6605 |  | 471.5569
0 | LS3 | BTS | QV | 5408 | LS3_BTS:QV_D5408 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 249.5256 | 427.1942 | 41.094 | 41.094 | 427.1942 | 249.4006 | 540.7855 | 540.9105 |  | 472.3069
0 | [8 empty] | 3.176e-06 | 0.001019 | 1.779 | 248.5112 | 427.1942 | 41.094 | 41.094 | 427.1942 | 247.6217 | 541.8 | 542.6895 | 20170217 space adjusted | 473.3214
0 | LS3 | BTS | DCHV | 5428 | LS3_BTS:DCHV_D5428 | C2 | correctors-X,Y | 50 | 247.5217 | 247.5217 | 0.2 | 247.5217 | 427.1942 | 41.094 | 41.094 | 427.1942 | 247.4217 | 542.7895 | 542.8895 |  | 474.3109
0 | [8 empty] | 3.176e-06 | -0.05468 | 0.1332 | 247.3551 | 427.1942 | 41.094 | 41.094 | 427.1942 | 247.2885 | 542.9561 | 543.0227 | 20170217 space adjusted | 474.4775
0 | LS3 | BTS | BPM | 5430 | LS3_BTS:BPM_D5430 | BPM3 | position monitor | 47.5 | 247.2785 | 247.2785 | 0.02 | 247.2785 | 427.1942 | 41.094 | 41.094 | 427.1942 | 247.2685 | 543.0327 | 543.0427 | 20170217 changed from .137 to .02 m (button diameter) | 474.5541
0 | [8 empty] | 0.0585 |  | 0.04782 | 247.2446 | 427.1942 | 41.094 | 41.094 | 427.1942 | 247.2206 | 543.0666 | 543.0905 | 20170217 space adjusted | 474.588
0 | LS3 | BTS | QH | 5432 | LS3_BTS:QH_D5432 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 247.0956 | 427.1942 | 41.094 | 41.094 | 427.1942 | 246.9706 | 543.2155 | 543.3405 |  | 474.7369
0 | [10 empty] | 0.5 | 246.7206 | 427.1942 | 41.094 | 41.094 | 427.1942 | 246.4706 | 543.5905 | 543.8405 |  | 475.1119
0 | LS3 | BTS | QV | 5440 | LS3_BTS:QV_D5440 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 246.3456 | 427.1942 | 41.094 | 41.094 | 427.1942 | 246.2206 | 543.9655 | 544.0905 |  | 475.4869
0 | [10 empty] | 0.27322 | 246.084 | 427.1942 | 41.094 | 41.094 | 427.1942 | 245.9474 | 544.2271 | 544.3637 | 20170217 space adjusted | 475.7485
0 | [6 empty] | bellow |  | 2.99e-06 | -0.053068 | 0.1035 | 245.8957 | 427.1942 | 41.094 | 41.094 | 427.1942 | 245.8439 | 544.4155 | 544.4672 |  | 475.9369
0 | LS3 | BTS | BPM | 5445 | LS3_BTS:BPM_D5445 | BPM3 | position monitor |  | 245.8439 | 245.8439 | 0 | 245.8439 | 427.1942 | 41.094 | 41.094 | 427.1942 | 245.8439 | 544.4672 | 544.4672 |  | 475.9886
0 | [6 empty] | diagnostic box |  | 2.99e-06 |  | 0.14528 | 245.7713 | 427.1942 | 41.094 | 41.094 | 427.1942 | 245.6986 | 544.5399 | 544.6125 |  | 476.0612
0 | LS3 | BTS | PORT | 5446 | LS3_BTS:PORT_D5446 |  | port for future diagnostics |  | 245.6986 | 245.6986 | 0 | 245.6986 | 427.1942 | 41.094 | 41.094 | 427.1942 | 245.6986 | 544.6125 | 544.6125 |  | 476.1339
0 | [6 empty] | diagnostic box |  |  |  | 0 | 245.6986 | 427.1942 | 41.094 | 41.094 | 427.1942 | 245.6986 | 544.6125 | 544.6125 |  | 476.1339
0 | LS3 | BTS | IP | 5446 | LS3_BTS:IP_D5446 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 245.6986 | 427.1942 | 41.094 | 41.094 | 427.1942 | 245.6986 | 544.6125 | 544.6125 |  | 476.1339
0 | [6 empty] | bellow |  |  |  | 0.13122 | 245.633 | 427.1942 | 41.094 | 41.094 | 427.1942 | 245.5674 | 544.6781 | 544.7437 |  | 476.1995
0 | [10 empty] | 0.24678 | 245.444 | 427.1942 | 41.094 | 41.094 | 427.1942 | 245.3206 | 544.8671 | 544.9905 | 20170217 space adjusted | 476.3885
0 | [10 empty] | 1.6 | 244.5206 | 427.1942 | 41.094 | 41.094 | 427.1942 | 243.7206 | 545.7905 | 546.5905 |  | 477.3119
0 | BDS | BTS | DCHV | 5467 | BDS_BTS:DCHV_D5467 | C3 | correctors-X,Y | 50 |  |  | 0.2 | 243.6206 | 427.1942 | 41.094 | 41.094 | 427.1942 | 243.5206 | 546.6905 | 546.7905 | 20170217 C3 used to replace C2 | 478.2119
0 | [10 empty] | 0.2 | 243.4206 | 427.1942 | 41.094 | 41.094 | 427.1942 | 243.3206 | 546.8905 | 546.9905 |  | 478.4119
0 | BDS | BTS | QH | 5471 | BDS_BTS:QH_D5471 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 243.1956 | 427.1942 | 41.094 | 41.094 | 427.1942 | 243.0706 | 547.1155 | 547.2405 |  | 478.6369
0 | [10 empty] | 0.5 | 242.8206 | 427.1942 | 41.094 | 41.094 | 427.1942 | 242.5706 | 547.4905 | 547.7405 |  | 479.0119
0 | BDS | BTS | QV | 5479 | BDS_BTS:QV_D5479 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 242.4456 | 427.1942 | 41.094 | 41.094 | 427.1942 | 242.3206 | 547.8655 | 547.9905 |  | 479.3869
0 | [10 empty] | 1.55 | 241.5456 | 427.1942 | 41.094 | 41.094 | 427.1942 | 240.7706 | 548.7655 | 549.5405 |  | 480.2869
0 | BDS | BTS | DCHV | 5496 | BDS_BTS:DCHV_D5496 | C3 | correctors-X,Y | 50 |  |  | 0.2 | 240.6706 | 427.1942 | 41.094 | 41.094 | 427.1942 | 240.5706 | 549.6405 | 549.7405 | 20170217 C3 used to replace C2 | 481.1619
0 | [8 empty] | 2.976e-06 | -0.054676 | 0.18218 | 240.4796 | 427.1942 | 41.094 | 41.094 | 427.1942 | 240.3885 | 549.8316 | 549.9227 | 20170217 space adjusted | 481.353
0 | BDS | BTS | BPM | 5499 | BDS_BTS:BPM_D5499 | BPM3 | position monitor | 47.5 | 240.3785 | 240.3785 | 0.02 | 240.3785 | 427.1942 | 41.094 | 41.094 | 427.1942 | 240.3685 | 549.9327 | 549.9427 | 20170217 changed from .137 to .02 m (button diameter) | 481.4541
0 | [8 empty] | 0.0585 |  | 0.047824 | 240.3446 | 427.1942 | 41.094 | 41.094 | 427.1942 | 240.3206 | 549.9666 | 549.9905 | 20170217 space adjusted | 481.488
0 | BDS | BTS | QV | 5501 | BDS_BTS:QV_D5501 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 240.1956 | 427.1942 | 41.094 | 41.094 | 427.1942 | 240.0706 | 550.1155 | 550.2405 |  | 481.6369
0 | [10 empty] | 0.5 | 239.8206 | 427.1942 | 41.094 | 41.094 | 427.1942 | 239.5706 | 550.4905 | 550.7405 |  | 482.0119
0 | BDS | BTS | QH | 5509 | BDS_BTS:QH_D5509 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 239.4456 | 427.1942 | 41.094 | 41.094 | 427.1942 | 239.3206 | 550.8655 | 550.9905 |  | 482.3869
0 | [10 empty] | 0.17008 | 239.2356 | 427.1942 | 41.094 | 41.094 | 427.1942 | 239.1506 | 551.0755 | 551.1606 | 20170217 space adjusted | 482.5969
0 | [6 empty] | bellow |  | 2.99e-06 | -0.085079 | 0.1035 | 239.0988 | 427.1942 | 41.094 | 41.094 | 427.1942 | 239.0471 | 551.2123 | 551.2641 |  | 482.7337
0 | BDS | BTS | BPM | 5513 | BDS_BTS:BPM_D5513 | BPM3 | position monitor |  | 239.0471 | 239.0471 | 0 | 239.0471 | 427.1942 | 41.094 | 41.094 | 427.1942 | 239.0471 | 551.2641 | 551.2641 |  | 482.7855
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 238.9744 | 427.1942 | 41.094 | 41.094 | 427.1942 | 238.9018 | 551.3367 | 551.4094 |  | 482.8581
0 | BDS | BTS | PM | 5514 | BDS_BTS:PM_D5514 | PM3 | profile monitor |  |  |  | 0 | 238.9018 | 427.1942 | 41.094 | 41.094 | 427.1942 | 238.9018 | 551.4094 | 551.4094 |  | 482.9307
0 | [6 empty] | diagnostic box |  |  |  | 0 | 238.9018 | 427.1942 | 41.094 | 41.094 | 427.1942 | 238.9018 | 551.4094 | 551.4094 |  | 482.9307
0 | BDS | BTS | IP | 5514 | BDS_BTS:IP_D5514 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 238.9018 | 427.1942 | 41.094 | 41.094 | 427.1942 | 238.9018 | 551.4094 | 551.4094 |  | 482.9307
0 | [6 empty] | bellow |  |  |  | 0.13122 | 238.8362 | 427.1942 | 41.094 | 41.094 | 427.1942 | 238.7706 | 551.475 | 551.5406 |  | 482.9963
0 | [10 empty] | 0 | 238.7706 | 427.1942 | 41.094 | 41.094 | 427.1942 | 238.7706 | 551.5406 | 551.5406 |  | 483.062
0 | [8 empty] | 3.176e-06 | -0.19144 | 0.52464 | 238.5082 | 427.1942 | 41.094 | 41.094 | 427.1942 | 238.2459 | 551.8029 | 552.0652 | 20170217 space adjusted | 483.3243
0 | BDS | BTS | BCM | 5521 | BDS_BTS:BCM_D5521 | BCM3 | current monitor |  | 238.1959 | 238.1959 | 0.1 | 238.1959 | 427.1942 | 41.094 | 41.094 | 427.1942 | 238.1459 | 552.1152 | 552.1652 |  | 483.6366
0 | [8 empty] | 3.176e-06 | 0.32498 | 2.4836 | 236.9041 | 427.1942 | 41.094 | 41.094 | 427.1942 | 235.6623 | 553.407 | 554.6488 | 20170217 space adjusted | 484.9284
0 | BDS | BTS | GV | 5547 | BDS_BTS:GV_D5547 | GV_M | gate valve | 63 | 235.6248 | 235.6248 | 0.075 | 235.6248 | 427.1942 | 41.094 | 41.094 | 427.1942 | 235.5873 | 554.6863 | 554.7238 |  | 486.2077
0 | [10 empty] | 0.32498 | 235.4249 | 427.1942 | 41.094 | 41.094 | 427.1942 | 235.2624 | 554.8863 | 555.0488 | 20170217 space adjusted | 486.4077
0 | BDS | BTS | QV | 5552 | BDS_BTS:QV_D5552 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 235.1374 | 427.1942 | 41.094 | 41.094 | 427.1942 | 235.0124 | 555.1738 | 555.2988 |  | 486.6952
0 | [10 empty] | 0.5 | 234.7624 | 427.1942 | 41.094 | 41.094 | 427.1942 | 234.5124 | 555.5488 | 555.7988 |  | 487.0702
0 | BDS | BTS | QH | 5559 | BDS_BTS:QH_D5559 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 234.3874 | 427.1942 | 41.094 | 41.094 | 427.1942 | 234.2624 | 555.9238 | 556.0488 |  | 487.4452
0 | [10 empty] | 0.125 | 234.1999 | 427.1942 | 41.094 | 41.094 | 427.1942 | 234.1374 | 556.1113 | 556.1738 |  | 487.6327
0 | BDS | BTS | DCHV | 5563 | BDS_BTS:DCHV_D5563 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 234.0374 | 427.1942 | 41.094 | 41.094 | 427.1942 | 233.9374 | 556.2738 | 556.3738 |  | 487.7952
0 | [10 empty] | 0.03057 | 233.9221 | 427.1942 | 41.094 | 41.094 | 427.1942 | 233.9068 | 556.3891 | 556.4044 | 20170217 space adjusted | 487.9104
0 | [6 empty] | bellow |  |  | 0.058777 | 0.1035 | 233.855 | 427.1942 | 41.094 | 41.094 | 427.1942 | 233.8033 | 556.4561 | 556.5079 |  | 487.9775
0 | BDS | BTS | BPM | 5565 | BDS_BTS:BPM_D5565 | BPM3 | position monitor |  | 2.176e-06 | 233.8033 | 0 | 233.8033 | 427.1942 | 41.094 | 41.094 | 427.1942 | 233.8033 | 556.5079 | 556.5079 |  | 488.0292
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 233.7307 | 427.1942 | 41.094 | 41.094 | 427.1942 | 233.658 | 556.5805 | 556.6531 |  | 488.1019
0 | BDS | BTS | PM | 5567 | BDS_BTS:PM_D5567 | PM3 | profile monitor |  |  |  | 0 | 233.658 | 427.1942 | 41.094 | 41.094 | 427.1942 | 233.658 | 556.6531 | 556.6531 |  | 488.1745
0 | [6 empty] | diagnostic box |  |  |  | 0 | 233.658 | 427.1942 | 41.094 | 41.094 | 427.1942 | 233.658 | 556.6531 | 556.6531 |  | 488.1745
0 | BDS | BTS | IP | 5567 | BDS_BTS:IP_D5567 | IP_S | ion pump, 75 l/s |  | 2.176e-06 | 233.658 | 0 | 233.658 | 427.1942 | 41.094 | 41.094 | 427.1942 | 233.658 | 556.6531 | 556.6531 |  | 488.1745
0 | [10 empty] | 0 | 233.658 | 427.1942 | 41.094 | 41.094 | 427.1942 | 233.658 | 556.6531 | 556.6531 |  | 488.1745
0 | [6 empty] | bellow |  |  |  | 0.13122 | 233.5924 | 427.1942 | 41.094 | 41.094 | 427.1942 | 233.5268 | 556.7187 | 556.7844 |  | 488.2401
0 | Fourth beam dump line starts in "V" column | [9 empty] | 0.16443 | 233.4446 | 427.1942 | 41.094 | 41.094 | 427.1942 | 233.3624 | 556.8666 | 556.9488 | 20170217 space adjusted | 488.3879
0 | BDS | BBS | DH | 5578 | BDS_BBS:DH_D5578 | D4 | dipole-17.5deg, D4-1 | 35 x 50 |  |  | 1.6503 | 232.55 | 427.3193 | 41.094 | 41.094 | 427.4443 | 231.7377 | 557.7739 | 558.599 |  | 489.2953
0 | [10 empty] | 0.78617 | 231.3628 | 427.5625 | 41.094 | 41.094 | 427.6807 | 230.9879 | 558.9921 | 559.3852 |  | 490.5135
0 | [10 empty] | 0.3114 | 230.8394 | 427.7275 | 41.094 | 41.094 | 427.7743 | 230.6909 | 559.5409 | 559.6966 |  | 491.0623
0 | [10 empty] | 0.8 | 230.3094 | 427.8946 | 41.094 | 41.094 | 428.0149 | 229.9279 | 560.0966 | 560.4966 |  | 491.618
0 | BDS | BBS | S | 5606 | BDS_BBS:S_D5606 | H2 | hexapole-X+ | 100 |  |  | 0.2 | 229.8325 | 428.045 | 41.094 | 41.094 | 428.075 | 229.7372 | 560.5966 | 560.6966 |  | 492.118
0 | [10 empty] | 0.2 | 229.6418 | 428.1051 | 41.094 | 41.094 | 428.1352 | 229.5464 | 560.7966 | 560.8966 |  | 492.318
0 | BDS | BBS | QH | 5611 | BDS_BBS:QH_D5611 | Q5 | quadrupole-X,40cm,B3 | 100 |  |  | 0.4 | 229.3557 | 428.1953 | 41.094 | 41.094 | 428.2555 | 229.1649 | 561.0966 | 561.2966 |  | 492.618
0 | [10 empty] | 0.2 | 229.0696 | 428.2855 | 41.094 | 41.094 | 428.3156 | 228.9742 | 561.3966 | 561.4966 |  | 492.918
0 | BDS | BBS | DCHV | 5616 | BDS_BBS:DCHV_D5616 | C3 | correctors-X,Y | 100 |  |  | 0.2 | 228.8788 | 428.3457 | 41.094 | 41.094 | 428.3757 | 228.7834 | 561.5966 | 561.6966 |  | 493.118
0 | [10 empty] | 0.2 | 228.6881 | 428.4058 | 41.094 | 41.094 | 428.4359 | 228.5927 | 561.7966 | 561.8966 |  | 493.318
0 | [10 empty] | 0 | 228.5927 | 428.4359 | 41.094 | 41.094 | 428.4359 | 228.5927 | 561.8966 | 561.8966 |  | 493.418
0 | BDS | BBS | QV | 5621 | BDS_BBS:QV_D5621 | Q5 | quadrupole-Y,40cm,B3 | 100 | 0 | 0.2 | 0.4 | 228.402 | 428.496 | 41.094 | 41.094 | 428.5562 | 228.2112 | 562.0966 | 562.2966 |  | 493.618
0 | [8 empty] | 0.43015 | 0.19308 | 0.19308 | 228.1191 | 428.5852 | 41.094 | 41.094 | 428.6142 | 228.0271 | 562.3931 | 562.4897 | 20170216 space reduced by 0.79852 | 493.9145
0 | BDS | BBS | BPM | 5625 | BDS_BBS:BPM_D5625 | BPM4 | position monitor | 100 | 0 | 0.037069 | 0.074138 | 227.9917 | 428.6254 | 41.094 | 41.094 | 428.6365 | 227.9564 | 562.5268 | 562.5638 |  | 494.0481
0 | [8 empty] | 0.27004 | 0.23297 | 0.23297 | 227.8453 | 428.6715 | 41.094 | 41.094 | 428.7066 | 227.7342 | 562.6803 | 562.7968 | 20170216 0.049913 m added | 494.2017
0 | BDS | BBS | PM | 5628 | BDS_BBS:PM_D5628 | PM5 | profile monitor | 100 |  | 0 | 0 | 227.7342 | 428.7066 | 41.094 | 41.094 | 428.7066 | 227.7342 | 562.7968 | 562.7968 |  | 494.3182
0 | BDS | BBS | IP | 5628 | BDS_BBS:IP_D5628 | IP_L | ion pump, 300 l/s |  |  |  | 0 | 227.7342 | 428.7066 | 41.094 | 41.094 | 428.7066 | 227.7342 | 562.7968 | 562.7968 |  | 494.3182
0 | [10 empty] | 0.44982 | 227.5197 | 428.7742 | 41.094 | 41.094 | 428.8418 | 227.3052 | 563.0217 | 563.2466 | 20170216 0.029939 m added | 494.5431
0 | BDS | BBS | DH | 5641 | BDS_BBS:DH_D5641 | D4 | dipole-17.5deg, D4-2 | 35 x 50 |  |  | 1.6503 | 226.568 | 429.2054 | 41.094 | 41.094 | 429.5689 | 225.8309 | 564.0717 | 564.8969 |  | 495.5931
0 | [10 empty] | 0.33331 | 225.6943 | 429.6645 | 41.094 | 41.094 | 429.7601 | 225.5578 | 565.0635 | 565.2302 | 20170217 space adjusted | 496.5849
0 | BDS | BBS | BPM | 5653 | BDS_BBS:BPM_D5653 | BPM4 | position monitor | 100 |  | 0.037069 | 0.074138 | 225.5275 | 429.7813 | 41.094 | 41.094 | 429.8026 | 225.4971 | 565.2672 | 565.3043 |  | 496.7886
0 | [10 empty] | 0.00706 | 225.4942 | 429.8046 | 41.094 | 41.094 | 429.8066 | 225.4913 | 565.3078 | 565.3114 | 20170217 space adjusted | 496.8292
0 | BDS | BBS | PM | 5653 | BDS_BBS:PM_D5653 | PM5 | profile monitor | 100 | 0 |  | 0 | 225.4913 | 429.8066 | 41.094 | 41.094 | 429.8066 | 225.4913 | 565.3114 | 565.3114 |  | 496.8327
0 | BDS | BBS | IP | 5653 | BDS_BBS:IP_D5653 | IP_L | ion pump, 300 l/s |  | 0.34736 | 0.34736 | 0 | 225.4913 | 429.8066 | 41.094 | 41.094 | 429.8066 | 225.4913 | 565.3114 | 565.3114 |  | 496.8327
0 | [8 empty] | 2.9523e-06 |  | 0.24736 | 225.39 | 429.8776 | 41.094 | 41.094 | 429.9485 | 225.2887 | 565.4351 | 565.5587 | 20170217 space adjusted | 496.9564
0 | BDS | BBS | DCHV | 5657 | BDS_BBS:DCHV_D5657 | C3 | correctors-X,Y | 100 | 225.2068 | 225.2068 | 0.2 | 225.2068 | 430.0059 | 41.094 | 41.094 | 430.0632 | 225.1249 | 565.6587 | 565.7587 |  | 497.1801
0 | [7 empty] | C3 Y not match !!! | 430.0037 | 430.0059 | 0.23813 | 225.0273 | 430.1315 | 41.094 | 41.094 | 430.1998 | 224.9298 | 565.8778 | 565.9969 | 20170217 space adjusted | 497.3992
0 | BDS | BBS | DH | 5668 | BDS_BBS:DH_D5668 | D4 | dipole-17.5deg, D4-3 | 35 x 50 |  |  | 1.6503 | 224.3361 | 430.7682 | 41.094 | 41.094 | 431.3366 | 223.7423 | 566.822 | 567.6471 |  | 498.3434
0 | [8 empty] | 0 |  | 0.3069 | 223.6489 | 431.4583 | 41.094 | 41.094 | 431.58 | 223.5555 | 567.8006 | 567.954 | 20170217 space adjusted | 499.3219
0 | BDS | BBS | BPM | 5680 | BDS_BBS:BPM_D5680 | BPM4 | position monitor |  | 0.27004 | 0.27004 | 0.074138 | 223.5329 | 431.6095 | 41.094 | 41.094 | 431.6389 | 223.5104 | 567.9911 | 568.0281 |  | 499.5124
0 | [10 empty] | 0.23297 | 223.4395 | 431.7313 | 41.094 | 41.094 | 431.8237 | 223.3686 | 568.1446 | 568.2611 |  | 499.666
0 | BDS | BBS | PM | 5683 | BDS_BBS:PM_D5683 | PMx | profile monitor (not costed) | future PM5 |  |  | 0 | 223.3686 | 431.8237 | 41.094 | 41.094 | 431.8237 | 223.3686 | 568.2611 | 568.2611 |  | 499.7825
0 | BDS | BBS | IP | 5683 | BDS_BBS:IP_D5683 | IP_L | ion pump, 300 l/s |  | 0 | 0.093103 | 0 | 223.3686 | 431.8237 | 41.094 | 41.094 | 431.8237 | 223.3686 | 568.2611 | 568.2611 |  | 499.7825
0 | [8 empty] | 0.536 | 0.536 | 0.336 | 223.2663 | 431.957 | 41.094 | 41.094 | 432.0903 | 223.164 | 568.4291 | 568.5971 | 20170217 space adjusted | 499.9505
0 | BDS | BBS | QV | 5688 | BDS_BBS:QV_D5688 | Q5 | quadrupole-Y,40cm,B3 | 100 |  |  | 0.4 | 223.0423 | 432.2489 | 41.094 | 41.094 | 432.4076 | 222.9205 | 568.7971 | 568.9971 |  | 500.3185
0 | [10 empty] | 0.2 | 222.8596 | 432.4869 | 41.094 | 41.094 | 432.5663 | 222.7988 | 569.0971 | 569.1971 |  | 500.6185
0 | BDS | BBS | DCHV | 5693 | BDS_BBS:DCHV_D5693 | C3 | correctors-X,Y | 100 |  |  | 0.2 | 222.7379 | 432.6456 | 41.094 | 41.094 | 432.7249 | 222.677 | 569.2971 | 569.3971 |  | 500.8185
0 | [10 empty] | 0.2 | 222.6161 | 432.8043 | 41.094 | 41.094 | 432.8836 | 222.5553 | 569.4971 | 569.5971 |  | 501.0185
0 | BDS | BBS | QH | 5698 | BDS_BBS:QH_D5698 | Q5 | quadrupole-X,40cm,B3 | 100 |  |  | 0.4 | 222.4335 | 433.0423 | 41.094 | 41.094 | 433.2009 | 222.3118 | 569.7971 | 569.9971 |  | 501.3185
0 | [10 empty] | 0.2 | 222.2509 | 433.2803 | 41.094 | 41.094 | 433.3596 | 222.19 | 570.0971 | 570.1971 |  | 501.6185
0 | BDS | BBS | S | 5703 | BDS_BBS:S_D5703 | H2 | hexapole-X+ | 100 |  |  | 0.2 | 222.1291 | 433.439 | 41.094 | 41.094 | 433.5183 | 222.0682 | 570.2971 | 570.3971 |  | 501.8185
0 | [10 empty] | 1.8976 | 221.4907 | 434.271 | 41.094 | 41.094 | 435.0237 | 220.9131 | 571.3459 | 572.2947 |  | 502.8673
0 | BDS | BBS | DH | 5731 | BDS_BBS:DH_D5731 | D4 | dipole-17.5deg, D4-4 | 35 x 50 |  |  | 1.6503 | 220.5177 | 435.7443 | 41.094 | 41.094 | 436.4649 | 220.1224 | 573.1198 | 573.9449 |  | 504.6412
0 | [10 empty] | 0.15235 | 220.0964 | 436.5365 | 41.094 | 41.094 | 436.6081 | 220.0703 | 574.0211 | 574.0973 | 20170217 space adjusted | 505.5425
0 | [6 empty] | bellow |  | -3.5722e-06 |  | 0.1035 | 220.0526 | 436.6567 | 41.094 | 41.094 | 436.7054 | 220.0349 | 574.149 | 574.2008 |  | 505.6704
0 | BDS | FFS | BPM | 5742 | BDS_FFS:BPM_D5742 | BPM3 | position monitor | 50 | 220.0349 | 220.0349 | 0 | 220.0349 | 436.7054 | 41.094 | 41.094 | 436.7054 | 220.0349 | 574.2008 | 574.2008 |  | 505.7222
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 220.0101 | 436.7736 | 41.094 | 41.094 | 436.8419 | 219.9852 | 574.2734 | 574.3461 |  | 505.7948
0 | BDS | FFS | PM | 5743 | BDS_FFS:PM_D5743 | PM3 | profile monitor | 50 | 0 | -0.059133 | 0 | 219.9852 | 436.8419 | 41.094 | 41.094 | 436.8419 | 219.9852 | 574.3461 | 574.3461 |  | 505.8674
0 | [6 empty] | diagnostic box |  | 0.29927 | 0.29927 | 0 | 219.9852 | 436.8419 | 41.094 | 41.094 | 436.8419 | 219.9852 | 574.3461 | 574.3461 |  | 505.8674
0 | BDS | FFS | IP | 5743 | BDS_FFS:IP_D5743 | IP_S | ion pump, 75 l/s |  | -3.5435e-06 |  | 0 | 219.9852 | 436.8419 | 41.094 | 41.094 | 436.8419 | 219.9852 | 574.3461 | 574.3461 |  | 505.8674
0 | [6 empty] | bellow |  | 219.8828 | 219.8829 | 0.13122 | 219.9628 | 436.9035 | 41.094 | 41.094 | 436.9652 | 219.9403 | 574.4117 | 574.4773 |  | 505.9331
0 | [7 empty] | C2 Y matched! | 437.1231 | 437.1231 | 0.068049 | 219.9287 | 436.9972 | 41.094 | 41.094 | 437.0291 | 219.9171 | 574.5113 | 574.5453 | 20170217 space adjusted | 506.0327
0 | BDS | FFS | DCHV | 5746 | BDS_FFS:DCHV_D5746 | C2 | correctors-X,Y | 50 | 0 |  | 0.2 | 219.8829 | 437.1231 | 41.094 | 41.094 | 437.2171 | 219.8487 | 574.6453 | 574.7453 | moved 1" upstream | 506.1667
0 | [8 empty] | 0.3246 | 0.3246 | 0.0996 | 219.8316 | 437.2639 | 41.094 | 41.094 | 437.3107 | 219.8146 | 574.7951 | 574.8449 | 20170217 space adjusted | 506.3165
0 | BDS | FFS | QH | 5750 | BDS_FFS:QH_D5750 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 219.7718 | 437.4281 | 41.094 | 41.094 | 437.5456 | 219.7291 | 574.9699 | 575.0949 |  | 506.4913
0 | [10 empty] | 0.5 | 219.6436 | 437.7805 | 41.094 | 41.094 | 438.0154 | 219.5581 | 575.3449 | 575.5949 |  | 506.8663
0 | BDS | FFS | QV | 5757 | BDS_FFS:QV_D5757 | Q1 | quadrupole-Y | 50 |  |  | 0.25 | 219.5153 | 438.1329 | 41.094 | 41.094 | 438.2504 | 219.4726 | 575.7199 | 575.8449 |  | 507.2413
0 | [10 empty] | 1.2675 | 219.2558 | 438.8459 | 41.094 | 41.094 | 439.4414 | 219.0391 | 576.4787 | 577.1124 |  | 508
0 | [6 empty] | bellow |  |  |  | 0.1035 | 219.0214 | 439.49 | 41.094 | 41.094 | 439.5386 | 219.0037 | 577.1642 | 577.2159 |  | 508.6855
0 | BDS | FFS | BPM | 5772 | BDS_FFS:BPM_D5772 | BPM3 | position monitor | 50 |  | 219.0037 | 0 | 219.0037 | 439.5386 | 41.094 | 41.094 | 439.5386 | 219.0037 | 577.2159 | 577.2159 |  | 508.7373
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 218.9788 | 439.6069 | 41.094 | 41.094 | 439.6752 | 218.954 | 577.2886 | 577.3612 |  | 508.8099
0 | BDS | FFS | PM | 5774 | BDS_FFS:PM_D5774 | PM3 | profile monitor | 50 |  |  | 0 | 218.954 | 439.6752 | 41.094 | 41.094 | 439.6752 | 218.954 | 577.3612 | 577.3612 |  | 508.8826
0 | [6 empty] | diagnostic box |  |  |  | 0 | 218.954 | 439.6752 | 41.094 | 41.094 | 439.6752 | 218.954 | 577.3612 | 577.3612 |  | 508.8826
0 | BDS | FFS | IP | 5774 | BDS_FFS:IP_D5774 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 218.954 | 439.6752 | 41.094 | 41.094 | 439.6752 | 218.954 | 577.3612 | 577.3612 |  | 508.8826
0 | [6 empty] | bellow |  | 0 |  | 0.13122 | 218.9315 | 439.7368 | 41.094 | 41.094 | 439.7985 | 218.9091 | 577.4268 | 577.4924 |  | 508.9482
0 | [8 empty] | 0.33668 | 0.33668 | 0.080468 | 218.8953 | 439.8363 | 41.094 | 41.094 | 439.8741 | 218.8816 | 577.5326 | 577.5729 |  | 509.054
0 | BDS | FFS | QH | 5777 | BDS_FFS:QH_D5777 | Q1 | quadrupole-X | 50 |  |  | 0.25 | 218.8388 | 439.9915 | 41.094 | 41.094 | 440.109 | 218.7961 | 577.6979 | 577.8229 |  | 509.2192
0 | [10 empty] | 0.175 | 218.7661 | 440.1912 | 41.094 | 41.094 | 440.2735 | 218.7362 | 577.9104 | 577.9979 |  | 509.4317
0 | BDS | FFS | DCHV | 5781 | BDS_FFS:DCHV_D5781 | C2 | correctors-X,Y | 50 |  |  | 0.2 | 218.702 | 440.3674 | 41.094 | 41.094 | 440.4614 | 218.6678 | 578.0979 | 578.1979 |  | 509.6192
0 | [10 empty] | 0.125 | 218.6464 | 440.5201 | 41.094 | 41.094 | 440.5789 | 218.6251 | 578.2604 | 578.3229 |  | 509.7817
0 | BDS | FFS | QV | 5784 | BDS_FFS:QV_D5784 | Q1 | quadrupole-Y | 50 | 1e-07 |  | 0.25 | 218.5823 | 440.6963 | 41.094 | 41.094 | 440.8138 | 218.5396 | 578.4479 | 578.5729 |  | 509.9692
0 | [8 empty] | 0.31124 | 0.31124 | 0.14874 | 218.5141 | 440.8837 | 41.094 | 41.094 | 440.9535 | 218.4887 | 578.6472 | 578.7216 |  | 510.1686
0 | BDS | FFS | GV | 5788 | BDS_FFS:GV_D5788 | GV_M | gate valve | 63 | 1.249e-16 |  | 0.075 | 218.4759 | 440.9888 | 41.094 | 41.094 | 441.024 | 218.463 | 578.7591 | 578.7966 |  | 510.2805
0 | [8 empty] | 0.12363 | 0.12363 | 0.086132 | 218.4483 | 441.0645 | 41.094 | 41.094 | 441.105 | 218.4336 | 578.8397 | 578.8827 |  | 510.3611
0 | BDS | FFS | BCM | 5789 | BDS_FFS:BCM_D5789 | BCM3 | current monitor |  | 0 |  | 0 | 218.4336 | 441.105 | 41.094 | 41.094 | 441.105 | 218.4336 | 578.8827 | 578.8827 |  | 510.4041
0 | [8 empty] | 0.1235 | 0.1235 | 0.019997 | 218.4302 | 441.1144 | 41.094 | 41.094 | 441.1237 | 218.4267 | 578.8927 | 578.9027 |  | 510.4141
0 | [6 empty] | bellow |  |  |  | 0.1035 | 218.409 | 441.1724 | 41.094 | 41.094 | 441.221 | 218.3913 | 578.9545 | 579.0062 |  | 510.4759
0 | BDS | FFS | BPM | 5790 | BDS_FFS:BPM_D5790 | BPM3 | position monitor | 50 |  |  | 0 | 218.3913 | 441.221 | 41.094 | 41.094 | 441.221 | 218.3913 | 579.0062 | 579.0062 |  | 510.5276
0 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 218.3665 | 441.2893 | 41.094 | 41.094 | 441.3575 | 218.3416 | 579.0789 | 579.1515 |  | 510.6003
0 | BDS | FFS | PM | 5792 | BDS_FFS:PM_D5792 | PM3 | profile monitor | 50 |  |  | 0 | 218.3416 | 441.3575 | 41.094 | 41.094 | 441.3575 | 218.3416 | 579.1515 | 579.1515 |  | 510.6729
0 | [6 empty] | diagnostic box |  |  |  | 0 | 218.3416 | 441.3575 | 41.094 | 41.094 | 441.3575 | 218.3416 | 579.1515 | 579.1515 | 20170217 box (pump,PM), BPM, BCM, GV moved upstream | 510.6729
0 | BDS | FFS | TMP | 5792 | BDS_FFS:TMP_D5792 | TMP_M | turbo pump, 550 l/s |  |  |  | 0 | 218.3416 | 441.3575 | 41.094 | 41.094 | 441.3575 | 218.3416 | 579.1515 | 579.1515 | 20170217 pump 300 -> 550 l/s | 510.6729
0 | [6 empty] | bellow |  |  |  | 0.13122 | 218.3192 | 441.4192 | 41.094 | 41.094 | 441.4808 | 218.2968 | 579.2171 | 579.2827 |  | 510.7385
0 | [10 empty] | 1.0284 | 218.1209 | 441.964 | 41.094 | 41.094 | 442.4472 | 217.945 | 579.797 | 580.3112 |  | 511.3183
0 | BDS | FFS | BPM | 5803 | BDS_FFS:BPM_D5803 | BPM? | position monitor | 75 | 0 |  | 0 | 217.945 | 442.4472 | 41.094 | 41.094 | 442.4472 | 217.945 | 580.3112 | 580.3112 | 20170217 BPM added | 511.8325
0 | [8 empty] | 0.2367 | 0.2367 | 0.1367 | 217.9216 | 442.5115 | 41.094 | 41.094 | 442.5757 | 217.8983 | 580.3795 | 580.4479 |  | 511.9009
0 | BDS | FFS | DCHV | 5805 | BDS_FFS:DCHV_D5805 | C3R | correctors-X,Y | 100 | 0 |  | 0.2 | 217.8641 | 442.6697 | 41.094 | 41.094 | 442.7636 | 217.8299 | 580.5479 | 580.6479 |  | 512.0692
0 | [8 empty] | 0.40506 | 0.40506 | 0.17506 | 217.7999 | 442.8459 | 41.094 | 41.094 | 442.9281 | 217.77 | 580.7354 | 580.8229 |  | 512.2568
0 | BDS | FFS | QH | 5810 | BDS_FFS:QH_D5810 | Q2R | quadrupole-X | 75 |  |  | 0.26 | 217.7255 | 443.0503 | 41.094 | 41.094 | 443.1725 | 217.6811 | 580.9529 | 581.0829 |  | 512.4743
0 | [10 empty] | 0.24 | 217.64 | 443.2852 | 41.094 | 41.094 | 443.398 | 217.599 | 581.2029 | 581.3229 |  | 512.7243
0 | BDS | FFS | QV | 5815 | BDS_FFS:QV_D5815 | Q3R | quadrupole-Y | 75 |  |  | 0.4 | 217.5306 | 443.5859 | 41.094 | 41.094 | 443.7739 | 217.4622 | 581.5229 | 581.7229 |  | 513.0443
0 | [10 empty] | 0.12288 | 217.4412 | 443.8316 | 41.094 | 41.094 | 443.8893 | 217.4201 | 581.7844 | 581.8458 |  | 513.3058
0 | BDS | FFS | BPM | 5818 | BDS_FFS:BPM_D5818 | BPM? | position monitor | 75 | 0 |  | 0 | 217.4201 | 443.8893 | 41.094 | 41.094 | 443.8893 | 217.4201 | 581.8458 | 581.8458 | 20170217 BPM added | 513.3672
0 | [8 empty] | 0.24712 | 0.24712 | 0.11712 | 217.4001 | 443.9444 | 41.094 | 41.094 | 443.9994 | 217.3801 | 581.9044 | 581.9629 |  | 513.4258
0 | BDS | FFS | QH | 5821 | BDS_FFS:QH_D5821 | Q2R | quadrupole-X | 75 |  |  | 0.26 | 217.3356 | 444.1216 | 41.094 | 41.094 | 444.2437 | 217.2912 | 582.0929 | 582.2229 |  | 513.6143
0 | [10 empty] | 0.7264 | 217.1669 | 444.585 | 41.094 | 41.094 | 444.9263 | 217.0427 | 582.5861 | 582.9493 |  | 514.1075
0 | [10 empty] | 1.0236 | 216.8677 | 445.4072 | 41.094 | 41.094 | 445.8882 | 216.6926 | 583.4611 | 583.9729 |  | 514.9825
| Target Center | [9 empty] | 0 | 216.6926 | 445.8882 | 41.094 | 41.094 | 445.8882 | 216.6926 | 583.9729 | 583.9729 |  | 515.4943
1 | FE | ISRC2 | ELC1 | 664 | FE_ISRC2:ELC1_D0664 |  | VENUS electrode |  |  |  | 0 | 368.5533 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 368.5533 | 66.4319 | 66.4319
1 | [10 empty] | 0.025397 | 368.5406 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 368.5279 | 66.4446 | 66.4573
1 | FE | ISRC2 | ELC2 | 665 | FE_ISRC2:ELC2_D0665 |  | VENUS extraction/puller | 150 |  |  | 0 | 368.5279 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 368.5279 | 66.4573 | 66.4573
1 | [10 empty] | 0.33074 | 368.3626 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 368.1972 | 66.6227 | 66.788
1 | [10 empty] | 0 | 368.1972 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 368.1972 | 66.788 | 66.788
1 | [10 empty] | 0 | 368.1972 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 368.1972 | 66.788 | 66.788
1 | FE | ISRC2 | SOLR | 670 | FE_ISRC2:SOLR_D0670 | S4s | solenoid | 150 |  |  | 0.33815 | 368.0281 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 367.8591 | 66.9571 | 67.1262
1 | [10 empty] | 0.17229 | 367.7729 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 367.6868 | 67.2123 | 67.2985
1 | FE | ISRC2 | TMP | 673 | FE_ISRC2:TMP_D0673 | TMP_L | turbo pump, 2050 l/s |  |  |  | 0 | 367.6868 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 367.6868 | 67.2985 | 67.2985
1 | [10 empty] | 0.20348 | 367.585 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 367.4833 | 67.4002 | 67.502
1 | FE | ISRC2 | GV | 675 | FE_ISRC2:GV_D0675 | GV_L | gate valve, 10" x 8" | 150 |  |  | 0.080264 | 367.4432 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 367.403 | 67.5421 | 67.5822
1 | [10 empty] | 0 | 367.403 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 367.403 | 67.5822 | 67.5822
1 | FE | ISRC2 | ACC | 677 | FE_ISRC2:ACC_D0677 | ACC | acceleration column | 150 |  |  | 0.26772 | 367.2692 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 367.1353 | 67.7161 | 67.8499
1 | [6 empty] | suppression electrode | 150 |  |  | 0.024638 | 367.123 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 367.1107 | 67.8623 | 67.8746
1 | [6 empty] | bellows | 150 |  |  | 0.016422 | 367.1025 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 367.0942 | 67.8828 | 67.891
1 | [10 empty] | 0 | 367.0942 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 367.0942 | 67.891 | 67.891
1 | FE | SCS2 | DCHV | 680 | FE_SCS2:DCHV_D0680 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 366.999 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 366.9037 | 67.9862 | 68.0815
1 | [7 empty] | 150 |  |  | 0.0254 | 366.891 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 366.8783 | 68.0942 | 68.1069
1 | FE | SCS2 | SOLR | 683 | FE_SCS2:SOLR_D0683 | S4b | solenoid | 150 |  |  | 0.4 | 366.6783 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 366.4783 | 68.3069 | 68.5069
1 | [7 empty] | 150 |  |  | 0.125 | 366.4158 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 366.3533 | 68.5694 | 68.6319
1 | [10 empty] | 0 | 366.3533 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 366.3533 | 68.6319 | 68.6319
1 | FE | SCS2 | TMP | 686 | FE_SCS2:TMP_D0686 | TMP_M | turbo pump, 700 l/s | 150 |  |  | 0 | 366.3533 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 366.3533 | 68.6319 | 68.6319
1 | [10 empty] | 0 | 366.3533 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 366.3533 | 68.6319 | 68.6319
1 | [7 empty] | 150 |  |  | 0.11975 | 366.2935 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 366.2336 | 68.6918 | 68.7516
1 | [7 empty] | 150 |  |  | 0 | 366.2336 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 366.2336 | 68.7516 | 68.7516
1 | FE | SCS2 | DCHV | 688 | FE_SCS2:DCHV_D0688 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 366.1383 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 366.0431 | 68.8469 | 68.9421
1 | [10 empty] | 0 | 366.0431 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 366.0431 | 68.9421 | 68.9421
1 | [10 empty] | 0 | 366.0431 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 366.0431 | 68.9421 | 68.9421
1 | [7 empty] | 150 |  |  | 0.13975 | 365.9732 | 418.6264 | 51.6647 | 51.6647 | 418.6264 | 365.9033 | 69.012 | 69.0819
1 | FE | SCS2 | DH | 696 | FE_SCS2:DH_D0696 | D5 | dipole-90deg, D5-1 | 120 X 150 |  |  | 1 | 365.585 | 418.9447 | 51.6647 | 51.6647 | 419.263 | 365.2667 | 69.5819 | 70.0819
1 | FE | SRC2 | TMP | 701 | FE_SRC2:TMP_D0701 | TMP_M | turbo pump, 700 l/s |  |  |  | 0 | 365.2667 | 419.263 | 51.6647 | 51.6647 | 419.263 | 365.2667 | 70.0819 | 70.0819
1 | [10 empty] | 0.06705 | 365.2667 | 419.2965 | 51.6647 | 51.6647 | 419.33 | 365.2667 | 70.1154 | 70.1489
1 | FE | SCS2 | DCHV | 702 | FE_SCS2:DCHV_D0702 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 365.2667 | 419.4253 | 51.6647 | 51.6647 | 419.5205 | 365.2667 | 70.2442 | 70.3394
1 | [7 empty] | 150 |  |  | 0.09245 | 365.2667 | 419.5668 | 51.6647 | 51.6647 | 419.613 | 365.2667 | 70.3857 | 70.4319
1 | FE | SCS2 | QHE | 705 | FE_SCS2:QHE_D0705 | Q7 | ES_quad_X | 150 |  |  | 0.2 | 365.2667 | 419.713 | 51.6647 | 51.6647 | 419.813 | 365.2667 | 70.5319 | 70.6319 | TRIPLET HOUSING
1 | [7 empty] | 150 |  |  | 0.135 | 365.2667 | 419.8805 | 51.6647 | 51.6647 | 419.948 | 365.2667 | 70.6994 | 70.7669 | TRIPLET HOUSING
1 | FE | SCS2 | QVE | 709 | FE_SCS2:QVE_D0709 | Q7 | ES_quad_Y | 150 |  |  | 0.2 | 365.2667 | 420.048 | 51.6647 | 51.6647 | 420.148 | 365.2667 | 70.8669 | 70.9669 | TRIPLET HOUSING
1 | [7 empty] | 150 |  |  | 0.135 | 365.2667 | 420.2155 | 51.6647 | 51.6647 | 420.283 | 365.2667 | 71.0344 | 71.1019 | TRIPLET HOUSING
1 | FE | SCS2 | QHE | 712 | FE_SCS2:QHE_D0712 | Q7 | ES_quad_X | 150 |  |  | 0.2 | 365.2667 | 420.383 | 51.6647 | 51.6647 | 420.483 | 365.2667 | 71.2019 | 71.3019 | TRIPLET HOUSING
1 | [10 empty] | 0 | 365.2667 | 420.483 | 51.6647 | 51.6647 | 420.483 | 365.2667 | 71.3019 | 71.3019
1 | [6 empty] | bellow+box | 150 |  | 0.28 | 0.22821 | 365.2667 | 420.5971 | 51.6647 | 51.6647 | 420.7112 | 365.2667 | 71.416 | 71.5301
1 | FE | SCS2 | SLT | 715 | FE_SCS2:SLT_D0715 |  | four jaw collimator | 150 |  |  | 0 | 365.2667 | 420.7112 | 51.6647 | 51.6647 | 420.7112 | 365.2667 | 71.5301 | 71.5301
1 | [7 empty] | 150 |  | 0.2159 | 0.021 | 365.2667 | 420.7217 | 51.6647 | 51.6647 | 420.7322 | 365.2667 | 71.5406 | 71.5511
1 | FE | SCS2 | TMP | 716 | FE_SCS2:TMP_D0716 | TMP_M | turbo pump, 700 l/s |  |  |  | 0 | 365.2667 | 420.7322 | 51.6647 | 51.6647 | 420.7322 | 365.2667 | 71.5511 | 71.5511
1 | [10 empty] | 0.11811 | 365.2667 | 420.7913 | 51.6647 | 51.6647 | 420.8503 | 365.2667 | 71.6102 | 71.6692
1 | FE | SCS2 | FC | 717 | FE_SCS2:FC_D0717 | FC2 | Faraday cup | 150 |  |  | 0 | 365.2667 | 420.8503 | 51.6647 | 51.6647 | 420.8503 | 365.2667 | 71.6692 | 71.6692
1 | [7 empty] | 150 |  |  | 0.18034 | 365.2667 | 420.9405 | 51.6647 | 51.6647 | 421.0306 | 365.2667 | 71.7594 | 71.8496
1 | FE | SCS2 | VD | 718 | FE_SCS2:VD_D0718 | V1 | viewer | 150 |  |  | 0 | 365.2667 | 421.0306 | 51.6647 | 51.6647 | 421.0306 | 365.2667 | 71.8496 | 71.8496
1 | FE | SCS2 | EMS | 718 | FE_SCS2:EMS_D0718 | EM1 | emittance scanner (x & y) | 150 |  |  | 0 | 365.2667 | 421.0306 | 51.6647 | 51.6647 | 421.0306 | 365.2667 | 71.8496 | 71.8496
1 | [7 empty] | 150 |  | 0.2641 | 0.21234 | 365.2667 | 421.1368 | 51.6647 | 51.6647 | 421.243 | 365.2667 | 71.9557 | 72.0619
1 | FE | SCS2 | QHE | 722 | FE_SCS2:QHE_D0722 | Q7 | ES_quad_X | 150 |  |  | 0.2 | 365.2667 | 421.343 | 51.6647 | 51.6647 | 421.443 | 365.2667 | 72.1619 | 72.2619 | TRIPLET HOUSING
1 | [7 empty] | 150 |  |  | 0.135 | 365.2667 | 421.5105 | 51.6647 | 51.6647 | 421.578 | 365.2667 | 72.3294 | 72.3969 | TRIPLET HOUSING
1 | FE | SCS2 | QVE | 725 | FE_SCS2:QVE_D0725 | Q7 | ES_quad_Y | 150 |  |  | 0.2 | 365.2667 | 421.678 | 51.6647 | 51.6647 | 421.778 | 365.2667 | 72.4969 | 72.5969 | TRIPLET HOUSING
1 | [7 empty] | 150 |  |  | 0.135 | 365.2667 | 421.8455 | 51.6647 | 51.6647 | 421.913 | 365.2667 | 72.6644 | 72.7319 | TRIPLET HOUSING
1 | FE | SCS2 | QHE | 728 | FE_SCS2:QHE_D0728 | Q7 | ES_quad_X | 150 |  |  | 0.2 | 365.2667 | 422.013 | 51.6647 | 51.6647 | 422.113 | 365.2667 | 72.8319 | 72.9319 | TRIPLET HOUSING
1 | [7 empty] | 150 |  | 0.06975 | 0.09245 | 365.2667 | 422.1592 | 51.6647 | 51.6647 | 422.2054 | 365.2667 | 72.9781 | 73.0243
1 | [10 empty] | 0 | 365.2667 | 422.2054 | 51.6647 | 51.6647 | 422.2054 | 365.2667 | 73.0243 | 73.0243
1 | [10 empty] | 0 | 365.2667 | 422.2054 | 51.6647 | 51.6647 | 422.2054 | 365.2667 | 73.0243 | 73.0243
1 | [10 empty] | 0 | 365.2667 | 422.2054 | 51.6647 | 51.6647 | 422.2054 | 365.2667 | 73.0243 | 73.0243
1 | FE | SCS2 | DCHV | 731 | FE_SCS2:DCHV_D0731 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 365.2667 | 422.3007 | 51.6647 | 51.6647 | 422.3959 | 365.2667 | 73.1196 | 73.2148
1 | [9 empty] | 0.08975 | 0.06705 | 365.2667 | 422.4295 | 51.6647 | 51.6647 | 422.463 | 365.2667 | 73.2484 | 73.2819
1 | FE | SCS2 | DH | 738 | FE_SCS2:DH_D0738 | D5 | dipole-90deg, D5-2 | 120 X 150 |  |  | 1 | 365.585 | 422.7813 | 51.6647 | 51.6647 | 423.0996 | 365.9033 | 73.7819 | 74.2819
1 | FE | SCS2 | TMP | 743 | FE_SCS2:TMP_D0743 | TMP_M | turbo pump, 700 l/s | 150 |  |  | 0 | 365.9033 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 365.9033 | 74.2819 | 74.2819
1 | [9 empty] | 0.225 | 0.225 | 366.0158 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 366.1283 | 74.3944 | 74.5069
1 | FE | SCS2 | QVE | 746 | FE_SCS2:QVE_D0746 | Q7 | ES_quad_Y | 150 |  |  | 0.2 | 366.2283 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 366.3283 | 74.6069 | 74.7069 | DOUBLET HOUSING
1 | [7 empty] | 150 |  |  | 0.135 | 366.3958 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 366.4633 | 74.7744 | 74.8419 | DOUBLET HOUSING
1 | FE | SCS2 | QHE | 749 | FE_SCS2:QHE_D0749 | Q7 | ES_quad_X | 150 |  |  | 0.2 | 366.5633 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 366.6633 | 74.9419 | 75.0419 | DOUBLET HOUSING
1 | [9 empty] | 0.20331 | 0.19159 | 366.7591 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 366.8549 | 75.1377 | 75.2335
1 | FE | SCS2 | PM | 752 | FE_SCS2:PM_D0752 | PM1 | profile monitor (large fork) | 150 |  |  | 0 | 366.8549 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 366.8549 | 75.2335 | 75.2335 | and 700 l/s turbo pump
1 | FE | SCS2 | TMP | 752 | FE_SCS2:TMP_D0752 | TMP_M | turbo pump, 700 l/s | 150 |  |  | 0 | 366.8549 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 366.8549 | 75.2335 | 75.2335
1 | [7 empty] | 150 |  | 0.10475 | 0.12327 | 366.9166 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 366.9782 | 75.2951 | 75.3567
1 | [7 empty] | 150 |  |  | 0 | 366.9782 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 366.9782 | 75.3567 | 75.3567
1 | FE | SCS2 | DCHV | 755 | FE_SCS2:DCHV_D0755 | C4 | corrector magnet & bellow | 150 |  |  | 0.1905 | 367.0734 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 367.1687 | 75.452 | 75.5472
1 | [9 empty] | 0.0067937 | 0 | 367.1687 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 367.1687 | 75.5472 | 75.5472
1 | FE | SCS2 | GV | 756 | FE_SCS2:GV_D0756 | GV_L | gate valve | 150 |  |  | 0.080264 | 367.2088 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 367.249 | 75.5874 | 75.6275
1 | [6 empty] | reducer / tie-in to LEBT | 150 |  |  | 0.75438 | 367.6262 | 423.0996 | 51.6647 | 51.6647 | 423.0996 | 368.0033 | 76.0047 | 76.3819
1
1 | FE | BTS | DH | 1089 | FE_BTS:DH_D1089 |  | dipole-45deg, D6 (bent leg) | 30 |  |  | 0.50921 | 359.5565 | 422.7395 | 41.094 | 41.094 | 422.7395 | 359.5565 | 108.8514 | 109.106
1 | [10 empty] | 0 | 359.5565 | 422.7395 | 41.094 | 41.094 | 422.7395 | 359.5565 | 109.106 | 109.106
1 | FE | BTS | GV | 1091 | FE_BTS:GV_D1091 | GV_S | gate valve | 30 |  |  | 0.060706 | 359.535 | 422.7181 | 41.094 | 41.094 | 422.6966 | 359.5135 | 109.1364 | 109.1667
1 | [10 empty] | 0.94143 | 359.1807 | 422.3638 | 41.094 | 41.094 | 422.0309 | 358.8479 | 109.6374 | 110.1081
1 | FE | BTS | SLT | 1101 | FE_BTS:SLT_D1101 |  | 2-jaw collimator | 30 |  |  | 0 | 358.8479 | 422.0309 | 41.094 | 41.094 | 422.0309 | 358.8479 | 110.1081 | 110.1081
1 | FE | BTS | VD | 1101 | FE_BTS:VD_D1101 | V2 | viewer | 30 |  |  | 0 | 358.8479 | 422.0309 | 41.094 | 41.094 | 422.0309 | 358.8479 | 110.1081 | 110.1081
1 | FE | BTS | IP | 1101 | FE_BTS:IP_D1101 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 358.8479 | 422.0309 | 41.094 | 41.094 | 422.0309 | 358.8479 | 110.1081 | 110.1081
1 | [7 empty] | 40 |  |  | 0.18857 | 358.7812 | 421.9643 | 41.094 | 41.094 | 421.8976 | 358.7145 | 110.2024 | 110.2967
1 | FE | BTS | FC | 1104 | FE_BTS:FC_D1104 | FC1 | beam stopper & faraday cup | 40 |  |  | 0.21346 | 358.6391 | 421.8221 | 41.094 | 41.094 | 421.7466 | 358.5636 | 110.4034 | 110.5102
1
exit valve of the last beta=0.041 module, diagnstic box, and D-station
2 | LS1 | CA03 | GV | 1221 | LS1_CA03:GV_D1221 |  | gate valve | 40 |  | M403_GV2 | 0.057404 | 346.4308 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.4021 | 122.0826 | 122.1113
2 | [6 empty] | bellow | 40 |  |  | 0.1035 | 346.3503 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.2986 | 122.163 | 122.2148
2 | LS1 | WA03 | BPM | 1222 | LS1_WA03:BPM_D1222 | BPM2 | position monitor |  |  | M403_BPM3 | 0 | 346.2986 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.2986 | 122.2148 | 122.2148
2 | [6 empty] | diagnostic box |  |  |  | 0.14528 | 346.2259 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.1533 | 122.2874 | 122.3601
2 | LS1 | WA03 | PM | 1224 | LS1_WA03:PM_D1224 | PMx | profile monitor (not costed) | future PM |  |  | 0 | 346.1533 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.1533 | 122.3601 | 122.3601
2 | [10 empty] | 0 | 346.1533 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.1533 | 122.3601 | 122.3601
2 | [6 empty] | diagnostic box |  |  |  | 0 | 346.1533 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.1533 | 122.3601 | 122.3601
2 | LS1 | WA03 | IP | 1224 | LS1_WA03:IP_D1224 | IP_S | ion pump, 75 l/s |  |  |  | 0 | 346.1533 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.1533 | 122.3601 | 122.3601
2 | [10 empty] | 0 | 346.1533 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.1533 | 122.3601 | 122.3601
2 | [10 empty] | 0 | 346.1533 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.1533 | 122.3601 | 122.3601
2 | [10 empty] | 0.036116 | 346.1352 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.1172 | 122.3781 | 122.3962
2 | LS1 | WA03 | BLM | 1224 | LS1_WA03:BLM_D1224 | HR_L | Halo ring, large aperture | 33 |  | M403_HR | 0 | 346.1172 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.1172 | 122.3962 | 122.3962
2 | [6 empty] | bellow | 40 |  |  | 0.0951 | 346.0696 | 423.0996 | 41.094 | 41.094 | 423.0996 | 346.0221 | 122.4437 | 122.4913
2 | D-station start | [9 empty] | 0.028661 | 346.0077 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.9934 | 122.5056 | 122.5199
2 | LS1 | D | GV | 1225 | LS1_D:GV_D1225 | GV_M | gate valve |  |  | D_GV | 0 | 345.9934 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.9934 | 122.5199 | 122.5199 | borrowed from BDS
2 | [10 empty] | 0.12629 | 345.9303 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.8671 | 122.5831 | 122.6462
2 | LS1 | D | FAV | 1226 | LS1_D:FAV_D1226 | FV | fast valve |  |  | D_FV | 0 | 345.8671 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.8671 | 122.6462 | 122.6462 | borrowed from BDS
2 | [10 empty] | 0.15123 | 345.7915 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.7159 | 122.7218 | 122.7975
2 | LS1 | D | LPM | 1228 | LS1_D:LPM_D1228 | LPM | longitudial profile monitor |  |  | D_LPM | 0 | 345.7159 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.7159 | 122.7975 | 122.7975 | Not installed
2 | [10 empty] | 0.13376 | 345.649 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.5821 | 122.8643 | 122.9312
2 | LS1 | D | BCM | 1229 | LS1_D:BCM_D1229 | BCM | beam current monitor |  |  | D_BCM1 | 0 | 345.5821 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.5821 | 122.9312 | 122.9312 | borrowed from BDS
2 | [10 empty] | 0.12349 | 345.5204 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.4586 | 122.993 | 123.0547
2 | [10 empty] | 0 | 345.4586 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.4586 | 123.0547 | 123.0547
2 | LS1 | D | BPM | 1231 | LS1_D:BPM_D1231 | BPM | beam position monitor |  |  | D_BPM1 | 0 | 345.4586 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.4586 | 123.0547 | 123.0547 | borrowed from BDS
2 | [10 empty] | 0.14529 | 345.386 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.3133 | 123.1274 | 123.2
2 | LS1 | D | PM | 1232 | LS1_D:PM_D1232 | PM | wire scanner |  |  | D_PM1 | 0 | 345.3133 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.3133 | 123.2 | 123.2 | borrowed from BDS
2 | LS1 | D | PM | 1232 | LS1_D:PM_D1232 | IP_S | ion pump, 75 l/s |  |  | D_IP | 0 | 345.3133 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.3133 | 123.2 | 123.2 | borrowed from BDS
2 | LS1 | D | CCG | 1232 | LS1_D:CCG_D1232 | CCG | cold cathde gauge |  |  | D_CCG1 | 0 | 345.3133 | 423.0996 | 41.094 | 41.094 | 423.0996 | 345.3133 | 123.2 | 123.2 | borrowed from BDS
2 | [10 empty] | 0.37186 | 345.1274 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.9415 | 123.3859 | 123.5719
2 | LS1 | D | SiD | 1236 | LS1_D:SiD_D1236 | SiD | silicon detector |  |  | D_SiD | 0 | 344.9415 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.9415 | 123.5719 | 123.5719 | borrowed from NSCL
2 | LS1 | D | TMP | 1236 | LS1_D:TMP_D1236 | TMP | turbo pump 500 l/s |  |  | D_TMP | 0 | 344.9415 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.9415 | 123.5719 | 123.5719 | borrowed from BDS
2 | [10 empty] | 0.23619 | 344.8234 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.7053 | 123.69 | 123.8081
2 | LS1 | D | BPM | 1238 | LS1_D:BPM_D1238 | BPM | beam position monitor |  |  | D_BPM2 | 0 | 344.7053 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.7053 | 123.8081 | 123.8081 | borrowed from BDS
2 | [10 empty] | 0.14529 | 344.6326 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.56 | 123.8807 | 123.9533
2 | LS1 | D | PM | 1240 | LS1_D:PM_D1240 | PM | wire scanner |  |  | D_PM2 | 0 | 344.56 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.56 | 123.9533 | 123.9533 | Not installed
2 | [10 empty] | 0.036144 | 344.5419 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.5239 | 123.9714 | 123.9895
2 | LS1 | D | BLM | 1240 | LS1_D:BLM_D1240 | HR_S | Halo ring, small aperture | 30 |  | D_HR | 0 | 344.5239 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.5239 | 123.9895 | 123.9895
2 | [10 empty] | 0.11504 | 344.4663 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.4088 | 124.047 | 124.1045
2 | LS1 | D | BCM | 1241 | LS1_D:BCM_D1241 | BCM | beam current monitor |  |  | D_BCM2 | 0 | 344.4088 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.4088 | 124.1045 | 124.1045 | borrowed from BDS
2 | [10 empty] | 0.26302 | 344.2773 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.1458 | 124.236 | 124.3675
2 | LS1 | D | FC | 1244 | LS1_D:FC_D1244 | FC | Faraday cup |  |  | D_FC | 0 | 344.1458 | 423.0996 | 41.094 | 41.094 | 423.0996 | 344.1458 | 124.3675 | 124.3675 | borrowed from FS2
2 | D-station end
| dump1 line start (FS1 straight line dump) | [13 empty] | 41.094 | 423.0996 | 229.662 |  |  | 2013/06/20 length updated from 5.892 to 5.922 Nathan Bultman
3 | FS1 | BTS | DH | 2394 | FS1_BTS:DH_D2394 | D2 | dipole-45deg, D2-1 | 40 x 100 | 1.83e-06 |  | 1.0974 | 228.5646 | 423.0996 | 41.094 | 41.094 | 423.0996 | 228.5646 | 239.4069 | 239.9556
3 | FS1 | BTS | GV | 2400 | FS1_BTS:GV_D2400 | GV_M | gate valve | 63 | 228.5271 | 228.5271 | 0.075 | 228.5271 | 423.0996 | 41.094 | 41.094 | 423.0996 | 228.4896 | 239.9931 | 240.0306
3 | [10 empty] | 0.7826 | 228.0983 | 423.0996 | 41.094 | 41.094 | 423.0996 | 227.707 | 240.4219 | 240.8132
3 | [8 empty] | 1.83e-06 |  | 0 | 227.707 | 423.0996 | 41.094 | 41.094 | 423.0996 | 227.707 | 240.8132 | 240.8132
3 | FS1 | BTS | Q | 2409 | FS1_BTS:Q_D2409 | Q1 | quadrupole-25cm |  | 227.582 | 227.582 | 0.25 | 227.582 | 423.0996 | 41.094 | 41.094 | 423.0996 | 227.457 | 240.9382 | 241.0632
3 | [8 empty] | 1.83e-06 |  | 0.09 | 227.412 | 423.0996 | 41.094 | 41.094 | 423.0996 | 227.367 | 241.1082 | 241.1532
3 | FS1 | BTS | Q | 2413 | FS1_BTS:Q_D2413 | Q1 | quadrupole-25cm |  | 227.242 | 227.242 | 0.25 | 227.242 | 423.0996 | 41.094 | 41.094 | 423.0996 | 227.117 | 241.2782 | 241.4032
3 | [10 empty] | 0.58999 | 226.822 | 423.0996 | 41.094 | 41.094 | 423.0996 | 226.527 | 241.6982 | 241.9932
3 | FS1 | BTS | Q | 2421 | FS1_BTS:Q_D2421 | Q2 | quadrupole-26cm |  |  |  | 0.26 | 226.397 | 423.0996 | 41.094 | 41.094 | 423.0996 | 226.267 | 242.1232 | 242.2532
3 | [10 empty] | 1.2438 | 225.6451 | 423.0996 | 41.094 | 41.094 | 423.0996 | 225.0232 | 242.8751 | 243.4969
3 | [6 empty] | bellow |  | 1.83e-06 |  | 0.1035 | 224.9715 | 423.0996 | 41.094 | 41.094 | 423.0996 | 224.9197 | 243.5487 | 243.6004
3 | FS1 | BTS | BPM | 2436 | FS1_BTS:BPM_D2436 | BPM4 | position monitor | 100 | 224.9197 | 224.9197 | 0 | 224.9197 | 423.0996 | 41.094 | 41.094 | 423.0996 | 224.9197 | 243.6004 | 243.6004 | 20170427 adjusted based on ME
3 | [6 empty] | diagnostic box |  |  |  | 0.2966 | 224.7714 | 423.0996 | 41.094 | 41.094 | 423.0996 | 224.6231 | 243.7487 | 243.8971
3 | FS1 | BTS | PM | 2439 | FS1_BTS:PM_D2439 | PM5 | profile monitor | 100 | defer installation |  | 0 | 224.6231 | 423.0996 | 41.094 | 41.094 | 423.0996 | 224.6231 | 243.8971 | 243.8971
3 | [6 empty] | diagnostic box |  | 1.83e-06 | 0.2966 | 0 | 224.6231 | 423.0996 | 41.094 | 41.094 | 423.0996 | 224.6231 | 243.8971 | 243.8971
3 | FS1 | BTS | IP | 2439 | FS1_BTS:IP_D2439 | IP_L | ion pump 300 l/s |  | 224.6231 | 224.6231 | 0 | 224.6231 | 423.0996 | 41.094 | 41.094 | 423.0996 | 224.6231 | 243.8971 | 243.8971
3 | [6 empty] | bellow |  |  |  | 0.13122 | 224.5575 | 423.0996 | 41.094 | 41.094 | 423.0996 | 224.4919 | 243.9627 | 244.0283 | 20170427 adjusted based on ME
3 | [10 empty] | 0.75191 | 224.116 | 423.0996 | 41.094 | 41.094 | 423.0996 | 223.74 | 244.4042 | 244.7802
3 | FS1 | BTS | BD | 2448 | FS1_BTS:BD_D2448 |  | dump1 |  |  |  | 0 | 223.74 | 423.0996 | 41.094 | 41.094 | 423.0996 | 223.74 | 244.7802 | 244.7802
| dump1 line end (FS1 straight line dump) | [18 empty] | 45-deg line
| dump2 line start (FS1 45-deg line dump) | [13 empty] | 41.094 | 420.5455 | 226.5591 |  |  | 2013/06/20 length updated from 5.222 to 4.922 Nathan Bultman
4 | FS1 | BTS | DH | 2435 | FS1_BTS:DH_D2435 | D2 | dipole-45deg, D2-2 | 40 x 100 |  |  | 1.0974 | 225.7831 | 419.7695 | 41.094 | 41.094 | 419.7695 | 225.7831 | 243.5108 | 244.0595
4 | FS1 | BTS | GV | 2441 | FS1_BTS:GV_D2441 | GV_M | gate valve | 63 |  |  | 0.075 | 225.7566 | 419.743 | 41.094 | 41.094 | 419.7165 | 225.7301 | 244.097 | 244.1345
4 | [10 empty] | 0.26759 | 225.6354 | 419.6219 | 41.094 | 41.094 | 419.5273 | 225.5408 | 244.2683 | 244.4021
4 | FS1 | BTS | Q | 2445 | FS1_BTS:Q_D2445 | Q1 | quadrupole-25cm | 50 |  |  | 0.25 | 225.4525 | 419.4389 | 41.094 | 41.094 | 419.3505 | 225.3641 | 244.5271 | 244.6521
4 | [10 empty] | 0.25 | 225.2757 | 419.2621 | 41.094 | 41.094 | 419.1737 | 225.1873 | 244.7771 | 244.9021
4 |  |  | BCM | 2449 | FS1_BTS:BCM_D2449 | BCM? | temporary BCM until ARR04 | 75 |  |  | 0 | 225.1873 | 419.1737 | 41.094 | 41.094 | 419.1737 | 225.1873 | 244.9021 | 244.9021
4 | [10 empty] | 0.26 | 225.0954 | 419.0818 | 41.094 | 41.094 | 418.9899 | 225.0034 | 245.0321 | 245.1621
4 | FS1 | BTS | Q | 2453 | FS1_BTS:Q_D2453 | Q2 | quadrupole-26cm | 75 |  |  | 0.26 | 224.9115 | 418.8979 | 41.094 | 41.094 | 418.806 | 224.8196 | 245.2921 | 245.4221
4 | [10 empty] | 1.1978 | 224.3961 | 418.3825 | 41.094 | 41.094 | 417.959 | 223.9726 | 246.021 | 246.6199
4 | [6 empty] | bellow |  |  |  | 0.1035 | 223.936 | 417.9224 | 41.094 | 41.094 | 417.8858 | 223.8994 | 246.6716 | 246.7234
4 | FS1 | BTS | BPM | 2467 | FS1_BTS:BPM_D2467 | BPM4 | position monitor | 100 |  |  | 0 | 223.8994 | 417.8858 | 41.094 | 41.094 | 417.8858 | 223.8994 | 246.7234 | 246.7234 | 20170427 adjusted based on ME
4 | [6 empty] | diagnostic box |  |  |  | 0.29661 | 223.7946 | 417.781 | 41.094 | 41.094 | 417.6761 | 223.6897 | 246.8717 | 247.02
4 | [10 empty] | 0 | 223.6897 | 417.6761 | 41.094 | 41.094 | 417.6761 | 223.6897 | 247.02 | 247.02
4 | FS1 | BTS | PM | 2470 | FS1_BTS:PM_D2470 | PM5 | profile monitor | 100 | defer installation |  | 0 | 223.6897 | 417.6761 | 41.094 | 41.094 | 417.6761 | 223.6897 | 247.02 | 247.02
4 | [6 empty] | diagnostic box |  |  |  | 0 | 223.6897 | 417.6761 | 41.094 | 41.094 | 417.6761 | 223.6897 | 247.02 | 247.02
4 | FS1 | BTS | IP | 2470 | FS1_BTS:IP_D2470 | IP_L | ion pump 300 l/s |  |  |  | 0 | 223.6897 | 417.6761 | 41.094 | 41.094 | 417.6761 | 223.6897 | 247.02 | 247.02
4 | [6 empty] | bellow |  |  |  | 0.13122 | 223.6433 | 417.6297 | 41.094 | 41.094 | 417.5833 | 223.5969 | 247.0856 | 247.1512 | 20170427 adjusted based on ME
4 | [10 empty] | 0.73286 | 223.3378 | 417.3242 | 41.094 | 41.094 | 417.0651 | 223.0787 | 247.5176 | 247.8841
4 | FS1 | BTS | BD | 2479 | FS1_BTS:BD_D2479 |  | dump2 |  |  |  | 0 | 223.0787 | 417.0651 | 41.094 | 41.094 | 417.0651 | 223.0787 | 247.8841 | 247.8841
dump2 line end (FS1 45-deg line dump)
| dump3 line start (FS2 straight line dump) | [13 empty] | 41.094 | 415.3476 | 376.475 |  |  | 2013/06/20 length updated from 5.283 to 5.0992 Nathan Bultman
5 | FS2 | BTS | DH | 3979 | FS2_BTS:DH_D3979 | D3 | dipole-45deg, D3-1 | 40 x 100 | 2.03e-06 |  | 2.2419 | 378.717 | 415.3476 | 41.094 | 41.094 | 415.3476 | 378.717 | 397.8512 | 398.9722
5 | FS2 | BTS | GV | 3990 | FS2_BTS:GV_D3990 | GV_M | gate valve | 63 | 378.7545 | 378.7545 | 0.075 | 378.7545 | 415.3476 | 41.094 | 41.094 | 415.3476 | 378.792 | 399.0097 | 399.0472
5 | [8 empty] | 2.03e-06 |  | 0.16876 | 378.8764 | 415.3476 | 41.094 | 41.094 | 415.3476 | 378.9607 | 399.1316 | 399.2159
5 | FS2 | BTS | Q | 3994 | FS2_BTS:Q_D3994 | Q3 | quadrupole-40cm |  | 379.1607 | 379.1607 | 0.4 | 379.1607 | 415.3476 | 41.094 | 41.094 | 415.3476 | 379.3607 | 399.4159 | 399.6159
5 | [10 empty] | 0 | 379.3607 | 415.3476 | 41.094 | 41.094 | 415.3476 | 379.3607 | 399.6159 | 399.6159
5 | [10 empty] | 0 | 379.3607 | 415.3476 | 41.094 | 41.094 | 415.3476 | 379.3607 | 399.6159 | 399.6159
5 | [10 empty] | 0.22731 | 379.4744 | 415.3476 | 41.094 | 41.094 | 415.3476 | 379.588 | 399.7296 | 399.8433
5 | FS2 | BTS | Q | 4000 | FS2_BTS:Q_D4000 | Q2 | quadrupole-26cm |  |  |  | 0.26 | 379.718 | 415.3476 | 41.094 | 41.094 | 415.3476 | 379.848 | 399.9733 | 400.1033
5 | [10 empty] | 0 | 379.848 | 415.3476 | 41.094 | 41.094 | 415.3476 | 379.848 | 400.1033 | 400.1033
5 | [10 empty] | 0.38824 | 380.0422 | 415.3476 | 41.094 | 41.094 | 415.3476 | 380.2363 | 400.2974 | 400.4915
5 | [10 empty] | 0 | 380.2363 | 415.3476 | 41.094 | 41.094 | 415.3476 | 380.2363 | 400.4915 | 400.4915
5 | [6 empty] | bellow |  | 2.03e-06 |  | 0.1035 | 380.288 | 415.3476 | 41.094 | 41.094 | 415.3476 | 380.3398 | 400.5432 | 400.595
5 | FS2 | BTS | BPM | 4006 | FS2_BTS:BPM_D4006 | BPM4 | position monitor | 100 | 380.3398 | 380.3398 | 0 | 380.3398 | 415.3476 | 41.094 | 41.094 | 415.3476 | 380.3398 | 400.595 | 400.595 | 20170427 adjusted based on ME, but .1657 is different from .1513 other areas
5 | [6 empty] | diagnostic box |  |  |  | 0.31096 | 380.4953 | 415.3476 | 41.094 | 41.094 | 415.3476 | 380.6507 | 400.7505 | 400.906
5 | FS2 | BTS | PM | 4009 | FS2_BTS:PM_D4009 | PM5 | profile monitor | 100 |  |  | 0 | 380.6507 | 415.3476 | 41.094 | 41.094 | 415.3476 | 380.6507 | 400.906 | 400.906
5 | [6 empty] | diagnostic box |  |  |  | 0 | 380.6507 | 415.3476 | 41.094 | 41.094 | 415.3476 | 380.6507 | 400.906 | 400.906
5 | [8 empty] | 2.03e-06 | 0.16567 | 0 | 380.6507 | 415.3476 | 41.094 | 41.094 | 415.3476 | 380.6507 | 400.906 | 400.906
5 | FS2 | BTS | IP | 4009 | FS2_BTS:IP_D4009 | IP_L | ion pump 300 l/s |  | 380.6507 | 380.6507 | 0 | 380.6507 | 415.3476 | 41.094 | 41.094 | 415.3476 | 380.6507 | 400.906 | 400.906
5 | [6 empty] | bellow |  |  |  | 0.13122 | 380.7163 | 415.3476 | 41.094 | 41.094 | 415.3476 | 380.782 | 400.9716 | 401.0372 | 20170427 adjusted based on ME
5 | [10 empty] | 0.79228 | 381.1781 | 415.3476 | 41.094 | 41.094 | 415.3476 | 381.5742 | 401.4333 | 401.8295
5 | FS2 | BTS | BD | 4018 | FS2_BTS:BD_D4018 |  | dump3 |  |  |  | 0 | 381.5742 | 415.3476 | 41.094 | 41.094 | 415.3476 | 381.5742 | 401.8295 | 401.8295
dump3 lind end (FS2 straight line dump)
[15 empty] | 41.094 | 427.1942 | 231.5562 |  |  | straight line
| dump4 line start (BDS straight line dump) | [13 empty] | 41.094 | 427.1942 | 233.3624 |  |  | 2013/06/20 length updated from 10.47 to 9.1406 Nathan Bultman
6 | BDS | BTS | DH | 5579 | BDS_BTS:DH_D5579 | D4 | dipole-17.5deg, D4-1 | 35 x 50 | 2.03e-06 | 1.8061 | 1.8061 | 231.5562 | 427.1942 | 41.094 | 41.094 | 427.1942 | 231.5562 | 557.8519 | 558.7549
6 | BDS | BTS | GV | 5588 | BDS_BTS:GV_D5588 | GV_M | gate valve | 63 | 231.5187 | 231.5187 | 0.075 | 231.5187 | 427.1942 | 41.094 | 41.094 | 427.1942 | 231.4812 | 558.7924 | 558.8299
6 | [9 empty] | 0.98885 | 0 | 231.4812 | 427.1942 | 41.094 | 41.094 | 427.1942 | 231.4812 | 558.8299 | 558.8299
6 | [10 empty] | 0 | 231.4812 | 427.1942 | 41.094 | 41.094 | 427.1942 | 231.4812 | 558.8299 | 558.8299
6 | [10 empty] | 0 | 231.4812 | 427.1942 | 41.094 | 41.094 | 427.1942 | 231.4812 | 558.8299 | 558.8299
6 | [10 empty] | 0 | 231.4812 | 427.1942 | 41.094 | 41.094 | 427.1942 | 231.4812 | 558.8299 | 558.8299
6 | [10 empty] | 0.98885 | 230.9868 | 427.1942 | 41.094 | 41.094 | 427.1942 | 230.4924 | 559.3244 | 559.8188
6 | BDS | BTS | Q | 5599 | BDS_BTS:Q_D5599 | Q1 | quadrupole-25cm |  |  |  | 0.25 | 230.3674 | 427.1942 | 41.094 | 41.094 | 427.1942 | 230.2424 | 559.9438 | 560.0688
6 | [10 empty] | 0.75 | 229.8674 | 427.1942 | 41.094 | 41.094 | 427.1942 | 229.4924 | 560.4438 | 560.8188
6 | BDS | BTS | Q | 5609 | BDS_BTS:Q_D5609 | Q2 | quadrupole-26cm |  |  |  | 0.26 | 229.3624 | 427.1942 | 41.094 | 41.094 | 427.1942 | 229.2324 | 560.9488 | 561.0788
6 | [10 empty] | 3.7274 | 227.3687 | 427.1942 | 41.094 | 41.094 | 427.1942 | 225.505 | 562.9425 | 564.8062
6 | [6 empty] | bellow |  | 2.03e-06 |  | 0.1035 | 225.4532 | 427.1942 | 41.094 | 41.094 | 427.1942 | 225.4015 | 564.8579 | 564.9097
6 | BDS | BTS | BPM | 5649 | BDS_BTS:BPM_D5649 | BPM4 | position monitor | 100 | 225.4015 | 225.4015 | 0 | 225.4015 | 427.1942 | 41.094 | 41.094 | 427.1942 | 225.4015 | 564.9097 | 564.9097 | 20170427 adjusted based on ME
6 | [6 empty] | diagnostic box |  |  |  | 0.2966 | 225.2532 | 427.1942 | 41.094 | 41.094 | 427.1942 | 225.1049 | 565.058 | 565.2063 | PM5 NOT COSTED
6 | BDS | BTS | PM | 5652 | BDS_BTS:PM_D5652 | PMx | profile monitor (not costed) | 100 |  |  | 0 | 225.1049 | 427.1942 | 41.094 | 41.094 | 427.1942 | 225.1049 | 565.2063 | 565.2063
6 | [6 empty] | diagnostic box |  | -0.15132 |  | 0 | 225.1049 | 427.1942 | 41.094 | 41.094 | 427.1942 | 225.1049 | 565.2063 | 565.2063
6 | BDS | BTS | IP | 5652 | BDS_BTS:IP_D5652 | IP_L | ion pump 300 l/s |  | 2.03e-06 | 225.1049 | 0 | 225.1049 | 427.1942 | 41.094 | 41.094 | 427.1942 | 225.1049 | 565.2063 | 565.2063
6 | [10 empty] | 0 | 225.1049 | 427.1942 | 41.094 | 41.094 | 427.1942 | 225.1049 | 565.2063 | 565.2063
6 | [6 empty] | bellow |  |  |  | 0.13122 | 225.0393 | 427.1942 | 41.094 | 41.094 | 427.1942 | 224.9737 | 565.2719 | 565.3375 | 20170427 adjusted based on ME
6 | [10 empty] | 0.75191 | 224.5977 | 427.1942 | 41.094 | 41.094 | 427.1942 | 224.2218 | 565.7134 | 566.0894
6 | BDS | BTS | BD | 5661 | BDS_BTS:BD_D5661 |  | dump4 |  |  |  | 0 | 224.2218 | 427.1942 | 41.094 | 41.094 | 427.1942 | 224.2218 | 566.0894 | 566.0894
dump4 line end (BDS straight line dump)
Note on G28: Author: S0 is an existing solenoid from Artemis B source
Note on D277: Author: to be determined as "QH" or "QV" at a later time
Note on D279: Author: to be determined as "QH" or "QV" at a later time
Note on D281: Author: to be determined as "QH" or "QV" at a later time
Note on D293: Author: to be determined as "QH" or "QV" at a later time
Note on D295: Author: to be determined as "QH" or "QV" at a later time
Note on D297: Author: to be determined as "QH" or "QV" at a later time
Note on D309: Author: to be determined as "QH" or "QV" at a later time
Note on D311: Author: to be determined as "QH" or "QV" at a later time
Note on D313: Author: to be determined as "QH" or "QV" at a later time
Note on D328: Author: to be determined as "QH" or "QV" at a later time
Note on D330: Author: to be determined as "QH" or "QV" at a later time
Note on D332: Author: to be determined as "QH" or "QV" at a later time

## FE_N0105
FE_N0105
Long Run Cable | [5 empty] | Inter-Rack Connections | [9 empty] | Terminating Device
Diagnostics Device | Lattice Beamline Location | CableDB Location | Cable Number | Cable Type | Function | Termination Device | Port Number In | Pin Number | CPC pinout | CPC pinout | CPC pinout | Signal Pin | Shield Pin | Port Number Out | Patch Cable | MTCA Chassis | MTCA Slot | PICO Port | FGPDB Port | Bias Supply Chassis | Bias Supply Slot | Bias Supply Channel | Notes
Fast Faraday Cup | FE_LEBT:FFC_D0998 | D0998 | 24E000058 | 2C_5/8in_Hardline_095 | Signal | N-type Patch panel | [16 empty] | Terminates into FFC scope
Fast Faraday Cup | FE_MEBT:FFC_D1102 | D1120 | 24E000062 | 2C_5/8in_Hardline_095 | Signal | N-type Patch panel | [16 empty] | Terminates into FFC scope
Allison Scanner | FE_SCS1:EMS_D0739 | D0739 | 24H000014 | 2C_5kVDC_Coax_035 | HV deflection plate : y top | Matsusada #1
Allison Scanner | FE_SCS1:EMS_D0739 | D0739 | 24H000010 | 2C_5kVDC_Coax_035 | HV deflection plate: x bot | Matsusada #1
Allison Scanner | FE_SCS1:EMS_D0739 | D0739 | 24H000001 | 2C_5kVDC_Coax_035 | HV deflection plate: x top | Matsusada #2
Allison Scanner | FE_SCS1:EMS_D0739 | D0739 | 24H000000 | 2C_5kVDC_Coax_035 | HV deflection plate: y bot | Matsusada #2
Allison Scanner | FE_SCS2:EMS_D0718 | D0718 | 24H000029 | 2C_5kVDC_Coax_035 | HV deflection plate | Matsusada #3
Allison Scanner | FE_SCS2:EMS_D0718 | D0718 | 24H000030 | 2C_5kVDC_Coax_035 | HV deflection plate | Matsusada #3
Allison Scanner | FE_SCS2:EMS_D0718 | D0718 | 24H000002 | 2C_5kVDC_Coax_035 | HV deflection plate | Matsusada #4
Allison Scanner | FE_SCS2:EMS_D0718 | D0718 | 24H000003 | 2C_5kVDC_Coax_035 | HV deflection plate | Matsusada #5
Pepper Pot Camera | FE_LEBT:EMS_D0812 | D0811 | 24F000026 | 8C_23AWG_Cat6_046 | video signal | POE Switch | [16 empty] | FRIB Diag CPU N0105
Viewer Camera | FE_LEBT:VD_D0812 | D0812 | 24F000186 | 8C_23AWG_Cat6_046 | video signal | POE Switch | [16 empty] | FRIB Diag CPU N0105
Viewer Camera | FE_SCS1:VD_D0739 | D0739 | 24F000000 | 8C_23AWG_Cat6_046 | video signal | POE Switch | [16 empty] | FRIB Diag CPU N0105
Viewer Camera | FE_SCS2:VD_D0718 | D0718 | 24F000009 | 8C_23AWG_Cat6_046 | video signal | POE Switch | [16 empty] | FRIB Diag CPU N0105

## FE_N0106
FE_N0106
Long Run Cable | [6 empty] | Inter-Rack Connections | [9 empty] | Terminating Device
Diagnostics Device | Lattice Beamline Location | Device Type | CableDB Location | Cable Number | Cable Type | Function | Termination Device | Port Number In | Pin Number | CPC pinout | CPC pinout | CPC pinout | Signal Pin | Shield Pin | Port Number Out | Patch Cable | MTCA Chassis | MTCA Slot | PICO Port | FGPDB Port | Bias Supply Chassis | Bias Supply Slot | Bias Supply Channel |  |  |  | search for row | device name | D number
Profile Monitors | [23 empty] | PICO | CH[0-7] | Device
PM Flapper | FE_LEBT:PM_D0856 | PM | D0856 | 24F000084 | 4C_22AWG_Multi_006 | X Signal | CPC to LEMO650 Box | CPC-1 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-1A | LEMO650/LEMO650 | 1 | 4 | CH6 - fmc1-3 |  | 1 | 1: EBS8005 | Ch0 | 1 | 4 | FE_LEBT:PM_D0856 | 141 | FE_LEBT:PM_D0856 | 856
PM Flapper | FE_LEBT:PM_D0856 | PM | D0856 | 24F000080 | 4C_22AWG_Multi_006 | Y Signal | CPC to LEMO650 Box | CPC-1 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | LEMO-1B | LEMO650/LEMO650 | 1 | 4 | CH5 - fmc1-2 |  | 1 | 1: EBS8005 | Ch0 | 1 | 4 | FE_LEBT:PM_D0856 | 141 | FE_LEBT:PM_D0856 | 856
PM Flapper | FE_LEBT:PM_D0885 | PM | D0884 | 24F000128 | 4C_22AWG_Multi_006 | X Signal | CPC to LEMO650 Box | CPC-2 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-2A | LEMO650/LEMO650 | 1 | 4 | CH4 - fmc1-1 |  | 1 | 1: EBS8005 | Ch0 | 1 | 4 | FE_LEBT:PM_D0885 | 156 | FE_LEBT:PM_D0885 | 885
PM Flapper | FE_LEBT:PM_D0885 | PM | D0884 | 24F000124 | 4C_22AWG_Multi_006 | Y Signal | CPC to LEMO650 Box | CPC-2 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | LEMO-2B | LEMO650/LEMO650 | 1 | 4 | CH2 - fmc2-3 |  | 1 | 1: EBS8005 | Ch0 | 1 | 4 | FE_LEBT:PM_D0885 | 156 | FE_LEBT:PM_D0885 | 885
PM Flapper | FE_LEBT:PM_D0912 | PM | D0913 | 24F000120 | 4C_22AWG_Multi_006 | X Signal | CPC to LEMO650 Box | CPC-3 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-3A | LEMO650/LEMO650 | 1 | 4 | CH1 - fmc2-2 |  | 1 | 1: EBS8005 | Ch0 | 1 | 4 | FE_LEBT:PM_D0912 | 174 | FE_LEBT:PM_D0912 | 912
PM Flapper | FE_LEBT:PM_D0912 | PM | D0913 | 24F000116 | 4C_22AWG_Multi_006 | Y Signal | CPC to LEMO650 Box | CPC-3 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | LEMO-3B | LEMO650/LEMO650 | 1 | 4 | CH0 - fmc2-1 |  | 1 | 1: EBS8005 | Ch0 | 1 | 4 | FE_LEBT:PM_D0912 | 174 | FE_LEBT:PM_D0912 | 912
PM 6" | FE_SCS2:PM_D0752 | PM | D0752 | 24F000086 | 4C_22AWG_Multi_006 | C Signal | CPC to LEMO650 Box | CPC-4 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-4A | LEMO650/LEMO650 | 1 | 5 | CH6 - fmc1-3 |  | 1 | 1: EBS8005 | Ch2 | 1 | 5 | FE_SCS2:PM_D0752 | 3459 | FE_SCS2:PM_D0752 | 752
PM 12" | FE_SCS2:PM_D0752 | PM | D0752 | 24F000140 | 6C_22AWG_Multi_010 | A Signal | CPC to LEMO650 Box | CPC-4 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | LEMO-4B | LEMO650/LEMO650 | 1 | 5 | CH4 - fmc1-1 |  | 1 | 1: EBS8005 | Ch2 | 1 | 5 | FE_SCS2:PM_D0752 | 3459 | FE_SCS2:PM_D0752 | 752
PM 12" | FE_SCS2:PM_D0752 | PM | D0752 | 24F000140 | 6C_22AWG_Multi_010 | B Signal | CPC to LEMO650 Box | CPC-4 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-4C | LEMO650/LEMO650 | 1 | 5 | CH5 - fmc1-2 |  | 1 | 1: EBS8005 | Ch2 | 1 | 5 | FE_SCS2:PM_D0752 | 3459 | FE_SCS2:PM_D0752 | 752
PM 6" | FE_LEBT:PM_D0783 | PM | D0783 | 24F000060 | 4C_22AWG_Multi_006 | C Signal | CPC to LEMO650 Box | CPC-5 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-5A | LEMO650/LEMO650 | 1 | 5 | CH2 - fmc2-3 |  | 1 | 1: EBS8005 | Ch3 | 1 | 5 | FE_LEBT:PM_D0783 | 89 | FE_LEBT:PM_D0783 | 783
PM 12" | FE_LEBT:PM_D0783 | PM | D0783 | 24F000067 | 6C_22AWG_Multi_010 | A Signal | CPC to LEMO650 Box | CPC-5 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | LEMO-5B | LEMO650/LEMO650 | 1 | 5 | CH0 - fmc2-1 |  | 1 | 1: EBS8005 | Ch3 | 1 | 5 | FE_LEBT:PM_D0783 | 89 | FE_LEBT:PM_D0783 | 783
PM 12" | FE_LEBT:PM_D0783 | PM | D0783 | 24F000067 | 6C_22AWG_Multi_010 | B Signal | CPC to LEMO650 Box | CPC-5 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-5C | LEMO650/LEMO650 | 1 | 5 | CH1 - fmc2-2 |  | 1 | 1: EBS8005 | Ch3 | 1 | 5 | FE_LEBT:PM_D0783 | 89 | FE_LEBT:PM_D0783 | 783
PM 6" | FE_LEBT:PM_D0808 | PM | D0811 | 24F000131 | 4C_22AWG_Multi_006 | C Signal | CPC to LEMO650 Box | CPC-6 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-6A | LEMO650/LEMO650 | 1 | 6 | CH6 - fmc1-3 |  | 1 | 1: EBS8005 | Ch4 | 1 | 6 | FE_LEBT:PM_D0808 | 111 | FE_LEBT:PM_D0808 | 808
PM 12" | FE_LEBT:PM_D0808 | PM | D0811 | 24F000130 | 6C_22AWG_Multi_010 | A Signal | CPC to LEMO650 Box | CPC-6 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | LEMO-6B | LEMO650/LEMO650 | 1 | 6 | CH4 - fmc1-1 |  | 1 | 1: EBS8005 | Ch4 | 1 | 6 | FE_LEBT:PM_D0808 | 111 | FE_LEBT:PM_D0808 | 808
PM 12" | FE_LEBT:PM_D0808 | PM | D0811 | 24F000130 | 6C_22AWG_Multi_010 | B Signal | CPC to LEMO650 Box | CPC-6 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-6C | LEMO650/LEMO650 | 1 | 6 | CH5 - fmc1-2 |  | 1 | 1: EBS8005 | Ch4 | 1 | 6 | FE_LEBT:PM_D0808 | 111 | FE_LEBT:PM_D0808 | 808
PM 6" | FE_LEBT:PM_D0824 | PM | D0824 | 24F000072 | 4C_22AWG_Multi_006 | C Signal | CPC to LEMO650 Box | CPC-7 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-7A | LEMO650/LEMO650 | 1 | 6 | CH2 - fmc2-3 |  | 1 | 1: EBS8005 | Ch5 | 1 | 6 | FE_LEBT:PM_D0824 | 125 | FE_LEBT:PM_D0824 | 824
PM 12" | FE_LEBT:PM_D0824 | PM | D0824 | 24F000066 | 6C_22AWG_Multi_010 | A Signal | CPC to LEMO650 Box | CPC-7 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | LEMO-7B | LEMO650/LEMO650 | 1 | 6 | CH0 - fmc2-1 |  | 1 | 1: EBS8005 | Ch5 | 1 | 6 | FE_LEBT:PM_D0824 | 125 | FE_LEBT:PM_D0824 | 824
PM 12" | FE_LEBT:PM_D0824 | PM | D0824 | 24F000066 | 6C_22AWG_Multi_010 | B Signal | CPC to LEMO650 Box | CPC-7 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-7C | LEMO650/LEMO650 | 1 | 6 | CH1 - fmc2-2 |  | 1 | 1: EBS8005 | Ch5 | 1 | 6 | FE_LEBT:PM_D0824 | 125 | FE_LEBT:PM_D0824 | 824
PM 8" | FE_LEBT:PM_D0998 | PM | D0998 | 24F000142 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-8 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-8A | LEMO650/LEMO650 | 1 | 7 | CH6 - fmc1-3 |  | 1 | 1: EBS8005 | Ch6 | 1 | 7 | FE_LEBT:PM_D0998 | 244 | FE_LEBT:PM_D0998 | 998
PM 8" | FE_LEBT:PM_D0998 | PM | D0998 | 24F000142 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-8 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield-6 | 4 | 6 | LEMO-8B | LEMO650/LEMO650 | 1 | 7 | CH5 - fmc1-2 |  | 1 | 1: EBS8005 | Ch6 | 1 | 7 | FE_LEBT:PM_D0998 | 244 | FE_LEBT:PM_D0998 | 998
PM 8" | FE_LEBT:PM_D0998 | PM | D0998 | 24F000142 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-8 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-8C | LEMO650/LEMO650 | 1 | 7 | CH4 - fmc1-1 |  | 1 | 1: EBS8005 | Ch6 | 1 | 7 | FE_LEBT:PM_D0998 | 244 | FE_LEBT:PM_D0998 | 998
PM 8" | FE_MEBT:PM_D1092 | PM | D1092 | 24F000134 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-9 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-9A | LEMO650/LEMO650 | 1 | 7 | CH2 - fmc2-3 |  | 1 | 1: EBS8005 | Ch7 | 1 | 7 | FE_MEBT:PM_D1092 | 303 | FE_MEBT:PM_D1092 | 1092
PM 8" | FE_MEBT:PM_D1092 | PM | D1092 | 24F000134 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-9 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | LEMO-9B | LEMO650/LEMO650 | 1 | 7 | CH1 - fmc2-2 |  | 1 | 1: EBS8005 | Ch7 | 1 | 7 | FE_MEBT:PM_D1092 | 303 | FE_MEBT:PM_D1092 | 1092
PM 8" | FE_MEBT:PM_D1092 | PM | D1092 | 24F000134 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-9 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-9C | LEMO650/LEMO650 | 1 | 7 | CH0 - fmc2-1 |  | 1 | 1: EBS8005 | Ch7 | 1 | 7 | FE_MEBT:PM_D1092 | 303 | FE_MEBT:PM_D1092 | 1092
Allison Scanner Signal
Allison Scanner | FE_SCS1:EMS_D0739 | EMS | D0739 | 24E000050 | 2C_75LC_Coax_089 | X Signal | N-Type to SMA Panel | J1 | NA | [5 empty] | J1 | SMA/LEMO650 | 1 | 3 | CH5 - fmc1-2 |  |  |  |  | 1 | 3 | FE_SCS1:EMS_D0739 | 66 | FE_SCS1:EMS_D0739 | 739
Allison Scanner | FE_SCS1:EMS_D0739 | EMS | D0739 | 24E000051 | 2C_75LC_Coax_089 | Y Signal | N-Type to SMA Panel | J2 | NA | [5 empty] | J2 | SMA/LEMO650 | 1 | 3 | CH4 - fmc1-1 |  |  |  |  | 1 | 3 | FE_SCS1:EMS_D0739 | 66 | FE_SCS1:EMS_D0739 | 739
Allison Scanner | FE_SCS2:EMS_D0718 | EMS | D0718 | 24E000052 | 2C_75LC_Coax_089 | X Signal | N-Type to SMA Panel | J3 | NA | [5 empty] | J3 | SMA/LEMO650 | 1 | 3 | CH1 - fmc2-2 |  |  |  |  | 1 | 3 | FE_SCS2:EMS_D0718 | 3439 | FE_SCS2:EMS_D0718 | 718
Allison Scanner | FE_SCS2:EMS_D0718 | EMS | D0718 | 24E000053 | 2C_75LC_Coax_089 | Y Signal | N-Type to SMA Panel | J4 | NA | [5 empty] | J4 | SMA/LEMO650 | 1 | 3 | CH0 - fmc2-1 |  |  |  |  | 1 | 3 | FE_SCS2:EMS_D0718 | 3439 | FE_SCS2:EMS_D0718 | 718
Faraday Cup Signal
Faraday Cup | FE_SCS2:FC_D0717 | FC | D0718 | 24E000054 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | J5 | NA | [5 empty] | J5 | SMA/LEMO650 | 1 | 9 | CH0 - fmc2-1 |  |  |  |  | 1 | 9 | FE_SCS2:FC_D0717 | 3436 | FE_SCS2:FC_D0717 | 717
Faraday Cup | FE_SCS1:FC_D0738 | FC | D0739 | 24E000049 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | J6 | NA | [5 empty] | J6 | SMA/LEMO650 | 1 | 9 | CH2 - fmc2-3 |  |  |  |  | 1 | 9 | FE_SCS1:FC_D0738 | 63 | FE_SCS1:FC_D0738 | 738
Faraday Cup | FE_LEBT:FC_D0796 | FC | D0794 | 24E000055 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | J7 | NA | [5 empty] | J7 | SMA/LEMO650 | 1 | 8 | CH0 - fmc2-1 |  |  |  |  | 1 | 8 | FE_LEBT:FC_D0796 | 103 | FE_LEBT:FC_D0796 | 796
Faraday Cup | FE_LEBT:FC_D0814 | FC | D0811 | 24E000056 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | J8 | NA | [5 empty] | J8 | SMA/LEMO650 | 1 | 8 | CH1 - fmc2-2 |  |  |  |  | 1 | 8 | FE_LEBT:FC_D0814 | 117 | FE_LEBT:FC_D0814 | 814
Faraday Cup | FE_LEBT:FC_D0977 | FC | D0976 | 24E000057 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | J9 | NA | [5 empty] | J9 | SMA/LEMO650 | 1 | 8 | CH2 - fmc2-3 |  |  |  |  | 1 | 8 | FE_LEBT:FC_D0977 | 219 | FE_LEBT:FC_D0977 | 977
Faraday Cup | FE_LEBT:FC_D0998 | FC | D0998 | 24E000059 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | J10 | NA | [5 empty] | J10 | SMA/LEMO650 | 1 | 8 | CH4 - fmc1-1 |  |  |  |  | 1 | 8 | FE_LEBT:FC_D0998 | 241 | FE_LEBT:FC_D0998 | 998
Faraday Cup | FE_BTS:FC_D1104 | FC | D1100 | 24E000063 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | J11 | NA | [5 empty] | J11 | SMA/LEMO650 | 1 | 8 | CH5 - fmc1-2 |  |  |  |  | 1 | 8 | FE_BTS:FC_D1104 | 3476 | FE_BTS:FC_D1104 | 1104
Faraday Cup | FE_MEBT:FC_D1102 | FC | D1123 | 24E000064 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | J12 | NA | [5 empty] | J12 | SMA/LEMO650 | 1 | 8 | CH6 - fmc1-3 |  |  |  |  | 1 | 8 | FE_MEBT:FC_D1102 | 316 | FE_MEBT:FC_D1102 | 1102
Allison Scanner Suppressor Ring Bias
Allison Scanner | FE_SCS1:EMS_D0739 |  | D0739 | 24H000011 | 2C_5kVDC_Coax_035 | HV Bias | SHV 6-> 12 Box | 1-A | [6 empty] | 1 IN | SHV/SHV |  |  |  |  | 1 | 2: EHS8020n | Ch0 |  |  | FE_SCS1:EMS_D0739 | 66 | FE_SCS1:EMS_D0739 | 739
Allison Scanner | FE_SCS1:EMS_D0739 |  | D0739 | 24H000012 | 2C_5kVDC_Coax_035 | HV Bias | SHV 6-> 12 Box | 1-B | [6 empty] | 1 IN | SHV/SHV |  |  |  |  | 1 | 2: EHS8020n | Ch0 |  |  | FE_SCS1:EMS_D0739 | 66 | FE_SCS1:EMS_D0739 | 739
Allison Scanner | FE_SCS2:EMS_D0718 |  | D0718 | 24H000002 | 2C_5kVDC_Coax_035 | HV Bias | SHV 6-> 12 Box | [8 empty] | SHV/SHV |  |  |  |  | 1 | 2: EHS8020n | Ch1 |  |  | FE_SCS2:EMS_D0718 | 3439 | FE_SCS2:EMS_D0718 | 718
Allison Scanner | FE_SCS2:EMS_D0718 |  | D0718 | 24H000003 | 2C_5kVDC_Coax_035 | HV Bias | SHV 6-> 12 Box | [8 empty] | SHV/SHV |  |  |  |  | 1 | 2: EHS8020n | Ch1 |  |  | FE_SCS2:EMS_D0718 | 3439 | FE_SCS2:EMS_D0718 | 718
Faraday Cup Suppressor Ring Bias
Faraday Cup | FE_SCS2:FC_D0717 |  | D0718 | 24H000007 | 2C_5kVDC_Coax_035 | HV Bias | ISEG - EHS8020n | Ch2 | [12 empty] | 1 | 2: EHS8020n | Ch2 |  |  | FE_SCS2:FC_D0717 | 3436 | FE_SCS2:FC_D0717 | 717
Faraday Cup | FE_SCS2:FC_D0717 |  | D0718 | 24H000008 | 2C_5kVDC_Coax_035 | HV Bias |  | no connection | [17 empty] | FE_SCS2:FC_D0717 | 3436 | FE_SCS2:FC_D0717 | 717
Faraday Cup | FE_SCS1:FC_D0738 |  | D0739 | 24H000009 | 2C_5kVDC_Coax_035 | HV Bias | SHV 6-> 12 Box | 4-A | [6 empty] | 4 IN | SHV/SHV |  |  |  |  | 1 | 2: EHS8020n | Ch3 |  |  | FE_SCS1:FC_D0738 | 63 | FE_SCS1:FC_D0738 | 738
Faraday Cup | FE_LEBT:FC_D0814 |  | D0811 | 24H000013 | 2C_5kVDC_Coax_035 | HV Bias | SHV 6-> 12 Box | 4-B | [6 empty] | 4 IN | SHV/SHV |  |  |  |  | 1 | 2: EHS8020n | Ch3 |  |  | FE_LEBT:FC_D0814 | 117 | FE_LEBT:FC_D0814 | 814
Faraday Cup | FE_LEBT:FC_D0977 |  | D0976 | 24H000015 | 2C_5kVDC_Coax_035 | HV Bias | SHV 6-> 12 Box | 5-A | [6 empty] | 5 IN | SHV/SHV |  |  |  |  | 1 | 2: EHS8020n | Ch4 |  |  | FE_LEBT:FC_D0977 | 219 | FE_LEBT:FC_D0977 | 977
Faraday Cup | FE_LEBT:FC_D0977 |  | D0976 | 24H000016 | 2C_5kVDC_Coax_035 | HV Bias |  | no connection | [17 empty] | FE_LEBT:FC_D0977 | 219 | FE_LEBT:FC_D0977 | 977
Faraday Cup | FE_LEBT:FC_D0998 |  | D0998 | 24H000019 | 2C_5kVDC_Coax_035 | HV Bias | SHV 6-> 12 Box | 3-A | [6 empty] | 3 IN | SHV/SHV |  |  |  |  | 1 | 2: EHS8020n | Ch5 |  |  | FE_LEBT:FC_D0998 | 241 | FE_LEBT:FC_D0998 | 998
Faraday Cup | FE_LEBT:FC_D0998 |  | D0998 | 24H000020 | 2C_5kVDC_Coax_035 | HV Bias |  | no connection | [17 empty] | FE_LEBT:FC_D0998 | 241 | FE_LEBT:FC_D0998 | 998
Faraday Cup | FE_BTS:FC_D1104 |  | D1100 | 24H000021 | 2C_5kVDC_Coax_035 | HV Bias | SHV 6-> 12 Box | 5-B | [6 empty] | 5 IN | SHV/SHV |  |  |  |  | 1 | 2: EHS8020n | Ch4 |  |  | FE_BTS:FC_D1104 | 3476 | FE_BTS:FC_D1104 | 1104
Faraday Cup | FE_BTS:FC_D1104 |  | D1100 | 24H000022 | 2C_5kVDC_Coax_035 | HV Bias |  | no connection | [17 empty] | FE_BTS:FC_D1104 | 3476 | FE_BTS:FC_D1104 | 1104
Faraday Cup | FE_MEBT:FC_D1102 |  | D1123 | 24H000023 | 2C_5kVDC_Coax_035 | HV Bias | SHV 6-> 12 Box | 3-B | [6 empty] | 3 IN | SHV/SHV |  |  |  |  | 1 | 2: EHS8020n | Ch5 |  |  | FE_MEBT:FC_D1102 | 316 | FE_MEBT:FC_D1102 | 1102
Faraday Cup | FE_MEBT:FC_D1102 |  | D1123 | 24H000026 | 2C_5kVDC_Coax_035 | HV Bias |  | no connection | [17 empty] | FE_MEBT:FC_D1102 | 316 | FE_MEBT:FC_D1102 | 1102
Fast Faraday Cup Suppressor Ring Bias
Fast Faraday Cup | FE_LEBT:FFC_D0998 |  | D0998 | 24H000017 | 2C_5kVDC_Coax_035 | HV Bias | [19 empty] | FE_LEBT:FFC_D0998 | 242 | FE_LEBT:FFC_D0998 | 998
Fast Faraday Cup | FE_LEBT:FFC_D0998 |  | D0998 | 24H000018 | 2C_5kVDC_Coax_035 | HV Bias | [19 empty] | FE_LEBT:FFC_D0998 | 242 | FE_LEBT:FFC_D0998 | 998
Fast Faraday Cup | FE_MEBT:FFC_D1102 |  | D1120 | 24H000024 | 2C_5kVDC_Coax_035 | HV Bias | [19 empty] | FE_MEBT:FFC_D1102 | 317 | FE_MEBT:FFC_D1102 | 1102
Fast Faraday Cup | FE_MEBT:FFC_D1102 |  | D1120 | 24H000025 | 2C_5kVDC_Coax_035 | HV Bias | SHV 6-> 12 Box | 2-A | [6 empty] | 2IN | Lemo250/Lemo250 |  |  |  |  | 1 | 2: EBS8003 | Ch1 |  |  | FE_MEBT:FFC_D1102 | 317 | FE_MEBT:FFC_D1102 | 1102
Lemo to Lemo Panel
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | SHV 6-> 12 Box (modded) | 6IN | [6 empty] | 6-A | Lemo250/Lemo250 |  |  |  |  | 1 | 2: EBS8003 | Ch0 |  |  | PM bias for PICO
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | SHV 6-> 12 Box (modded) | 6IN | [6 empty] | 6-B | Lemo250/Lemo250 |  |  |  |  | 1 | 2: EBS8003 | Ch0
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 6-A | [6 empty] | 6-A | SHV/Lemo650 | 1 | 4 | CH7 - fmc1-4 |  | 1 | 2: EBS8003 | Ch0 | 1 | 4 | PM bias for PICO
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 6-B | [6 empty] | 6-B | SHV/Lemo650 | 1 | 4 | CH3 - fmc2-4 |  | 1 | 2: EBS8003 | Ch0 | 1 | 4 | PM bias for PICO
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 3_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 1 | 2: EBS8003 | Ch2 |  |  | PM bias for PICO
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 3_OUT Port | [7 empty] | Lemo650/Lemo650 | 1 | 5 | CH7 - fmc1-4 |  |  |  |  | 1 | 5 | PM bias for PICO
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 4_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 1 | 2: EBS8003 | Ch3 |  |  | PM bias for PICO
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 4_OUT Port | [7 empty] | Lemo650/Lemo650 | 1 | 5 | CH3 - fmc2-4 |  |  |  |  | 1 | 5 | PM bias for PICO
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 5_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 1 | 2: EBS8003 | Ch4 |  |  | PM bias for PICO
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 5_OUT Port | [7 empty] | Lemo650/Lemo650 | 1 | 6 | CH7 - fmc1-4 |  |  |  |  | 1 | 6 | PM bias for PICO
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 6_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 1 | 2: EBS8003 | Ch5 |  |  | PM bias for PICO
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 6_OUT Port | [7 empty] | Lemo650/Lemo650 | 1 | 6 | CH3 - fmc2-4 |  |  |  |  | 1 | 6 | PM bias for PICO
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 7_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 1 | 2: EBS8003 | Ch6 |  |  | PM bias for PICO
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 7_OUT Port | [7 empty] | Lemo650/Lemo650 | 1 | 7 | CH7 - fmc1-4 |  |  |  |  | 1 | 7 | PM bias for PICO
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 8_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 1 | 2: EBS8003 | Ch7 |  |  | PM bias for PICO
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 8_OUT Port | [7 empty] | Lemo650/Lemo650 | 1 | 7 | CH3 - fmc2-4 |  |  |  |  | 1 | 7 | PM bias for PICO
Ground Plugs
Grounding Plug |  | ground |  | - | Lemo650 Plug | Ground | Pico-8 | [8 empty] | Ground Plug | 1 | 3 | CH3 - fmc2-4
Grounding Plug |  | ground |  | - | Lemo650 Plug | Ground | Pico-8 | [8 empty] | Ground Plug | 1 | 3 | CH7 - fmc1-4
Grounding Plug |  | ground |  | - | Lemo650 Plug | Ground | Pico-8 | [8 empty] | Ground Plug | 1 | 8 | CH3 - fmc2-4
Grounding Plug |  | ground |  | - | Lemo650 Plug | Ground | Pico-8 | [8 empty] | Ground Plug | 1 | 8 | CH7 - fmc1-4
Grounding Plug |  | ground |  | - | Lemo650 Plug | Ground | Pico-8 | [8 empty] | Ground Plug | 1 | 9 | CH3 - fmc2-4
Grounding Plug |  | ground |  | - | Lemo650 Plug | Ground | Pico-8 | [8 empty] | Ground Plug | 1 | 9 | CH7 - fmc1-4

## FE_N0403
FE_N0403
Long Run Cable | [5 empty] | Inter-Rack Connections | [9 empty] | Terminating Device
Diagnostics Device | Lattice Beamline Location | CableDB Location | Cable Number | Cable Type | Function | Termination Device | Port Number In | Pin Number | CPC pinout | CPC pinout | CPC pinout | Signal Pin | Shield Pin | Port Number Out | Patch Cable | MTCA Chassis | MTCA Slot | PICO Port | FGPDB Port | Bias Supply Chassis | Bias Supply Slot | Bias Supply Channel | Notes
Viewer Camera | FE_EML:VD_D1098 | D1098 | 24F000194 | 8C_23AWG_Cat6_046 | video signal | POE Switch | [16 empty] | FRIB Diag CPU N0105
Viewer Camera | FE_LEBT:VD_D0976 | D0976 | 24F000035 | 8C_23AWG_Cat6_046 | video signal | POE Switch | [16 empty] | FRIB Diag CPU N0105
Viewer Camera | FE_MEBT:VD_D1105? | D1105 | 24F000189 | 8C_23AWG_Cat6_046 | video signal | POE Switch | [16 empty] | FRIB Diag CPU N0105

## FE_N0404
FE_N0106 / FE_N0404 (temp)
Long Run Cable | [5 empty] | Inter-Rack Connections | [9 empty] | Terminating Device
Diagnostics Device | Lattice Beamline Location | CableDB Location | Cable Number | Cable Type | Function | Termination Device | Port Number In | Pin Number | CPC pinout | CPC pinout | CPC pinout | Signal Pin | Shield Pin | Port Number Out | Patch Cable | MTCA Chassis | MTCA Slot | Struck RTM Port | FGPDB Port | Bias Supply Chassis | Bias Supply Slot | Bias Supply Channel | Notes
BCM | FE_LEBT:BCM_D0989 | D0998 | 24E000044 | SILTEM 2C_22AWG_Multi_063 | signal | [10 empty] | 6 | 4 | 1
BCM | FE_MEBT:BCM_D1055 | D1055 | 24E000042 | SILTEM 2C_22AWG_Multi_063 | signal | [10 empty] | 6 | 4 | 2
BCM | FE_MEBT:BCM_D1120 | D1121 | 24E000043 | SILTEM 2C_22AWG_Multi_063 | signal | [10 empty] | 6 | 4 | 3
BCM | FS2_BMS:BCM_D4176 | D4175 | 44E000226 | SILTEM 2C_22AWG_Multi_063 | signal | [10 empty] | 6 | 4 | 4

## LS1_N0206
LS1_N0206
Long Run Cable | [6 empty] | Inter-Rack Connections | [9 empty] | Terminating Device
Diagnostics Device | Lattice Beamline Location | Device Type | CableDB Location | Cable Number | Cable Type | Function | Termination Device | Port Number In | Pin Number | CPC pinout | CPC pinout | CPC pinout | Signal Pin | Shield Pin | Port Number Out | Patch Cable | MTCA Chassis | MTCA Slot | PICO Port | FGPDB Port | Bias Supply Chassis | Bias Supply Slot | Bias Supply Channel | Notes | search for row | device name | D number
Panel 1 | [24 empty] | 275 | FE_MEBT:BPM_D1056 | 1056
BPM | FE_MEBT:BPM_D1056 | BPM | D1056 | 24E000015 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 2 | 3 |  | 1 |  |  |  |  | 275 | FE_MEBT:BPM_D1056 | 1056
BPM | FE_MEBT:BPM_D1056 | BPM | D1056 | 24E000016 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 2 | 3 |  | 2 |  |  |  |  | 275 | FE_MEBT:BPM_D1056 | 1056
BPM | FE_MEBT:BPM_D1056 | BPM | D1056 | 24E000017 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 2 | 3 |  | 3 |  |  |  |  | 275 | FE_MEBT:BPM_D1056 | 1056
BPM | FE_MEBT:BPM_D1056 | BPM | D1056 | 24E000018 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 2 | 3 |  | 4
BPM | FE_MEBT:RFS_D1056 | BPM | D1056 | 24E000019 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 2 | 3 |  | Ref |  |  |  |  | 291 | FE_MEBT:BPM_D1072 | 1072
BPM | FE_MEBT:BPM_D1072 | BPM | D1073 | 24E000011 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 2 | 3 |  | 5 |  |  |  |  | 291 | FE_MEBT:BPM_D1072 | 1072
BPM | FE_MEBT:BPM_D1072 | BPM | D1073 | 24E000012 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 2 | 3 |  | 6 |  |  |  |  | 291 | FE_MEBT:BPM_D1072 | 1072
BPM | FE_MEBT:BPM_D1072 | BPM | D1073 | 24E000013 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 2 | 3 |  | 7 |  |  |  |  | 291 | FE_MEBT:BPM_D1072 | 1072
BPM | FE_MEBT:BPM_D1072 | BPM | D1073 | 24E000014 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 2 | 3 |  | 8
BPM | FE_MEBT:RFS_D1072 | BPM | D1073 | 24E000010 | 2C_1/4in_Hardline_036 | Ref | No termination | No termination | [16 empty] | 306 | FE_MEBT:BPM_D1094 | 1094
BPM | FE_MEBT:BPM_D1094 | BPM | D1095 | 24E000005 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 2 | 4 |  | 1 |  |  |  |  | 306 | FE_MEBT:BPM_D1094 | 1094
BPM | FE_MEBT:BPM_D1094 | BPM | D1095 | 24E000006 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 2 | 4 |  | 2 |  |  |  |  | 306 | FE_MEBT:BPM_D1094 | 1094
BPM | FE_MEBT:BPM_D1094 | BPM | D1095 | 24E000007 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 2 | 4 |  | 3 |  |  |  |  | 306 | FE_MEBT:BPM_D1094 | 1094
BPM | FE_MEBT:BPM_D1094 | BPM | D1095 | 24E000008 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 2 | 4 |  | 4
BPM | FE_MEBT:RFS_D1094 | BPM | D1095 | 24E000009 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 2 | 4 |  | Ref |  |  |  |  | 325 | FE_MEBT:BPM_D1111 | 1111
BPM | FE_MEBT:BPM_D1111 | BPM | D1113 | 24E000001 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 2 | 4 |  | 5 |  |  |  |  | 325 | FE_MEBT:BPM_D1111 | 1111
BPM | FE_MEBT:BPM_D1111 | BPM | D1113 | 24E000002 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 2 | 4 |  | 6 |  |  |  |  | 325 | FE_MEBT:BPM_D1111 | 1111
BPM | FE_MEBT:BPM_D1111 | BPM | D1113 | 24E000003 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 2 | 4 |  | 7 |  |  |  |  | 325 | FE_MEBT:BPM_D1111 | 1111
BPM | FE_MEBT:BPM_D1111 | BPM | D1113 | 24E000004 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 2 | 4 |  | 8
BPM | FE_MEBT:RFS_D1111 | BPM | D1113 | 24E000000 | 2C_1/4in_Hardline_036 | Ref | No termination | No termination
Panel 2 | [24 empty] | 353 | LS1_CA01:BPM_D1129 | 1129
BPM | LS1_CA01:BPM_D1129 | BPM | D1129 | 34E000104 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 2 | 5 |  | 1 |  |  |  |  | 353 | LS1_CA01:BPM_D1129 | 1129
BPM | LS1_CA01:BPM_D1129 | BPM | D1129 | 34E000105 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 2 | 5 |  | 2 |  |  |  |  | 353 | LS1_CA01:BPM_D1129 | 1129
BPM | LS1_CA01:BPM_D1129 | BPM | D1129 | 34E000106 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 2 | 5 |  | 3 |  |  |  |  | 353 | LS1_CA01:BPM_D1129 | 1129
BPM | LS1_CA01:BPM_D1129 | BPM | D1129 | 34E000107 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 2 | 5 |  | 4
BPM | LS1_CA01:RFS_D1129 | BPM | D1129 | 34E000109 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 2 | 5 |  | Ref |  |  |  |  | 368 | LS1_CA01:BPM_D1144 | 1144
BPM | LS1_CA01:BPM_D1144 | BPM | D1144 | 34E000100 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 2 | 5 |  | 5 |  |  |  |  | 368 | LS1_CA01:BPM_D1144 | 1144
BPM | LS1_CA01:BPM_D1144 | BPM | D1144 | 34E000101 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 2 | 5 |  | 6 |  |  |  |  | 368 | LS1_CA01:BPM_D1144 | 1144
BPM | LS1_CA01:BPM_D1144 | BPM | D1144 | 34E000102 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 2 | 5 |  | 7 |  |  |  |  | 368 | LS1_CA01:BPM_D1144 | 1144
BPM | LS1_CA01:BPM_D1144 | BPM | D1144 | 34E000103 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 2 | 5 |  | 8
BPM | LS1_CA01:RFS_D1144 | BPM | D1144 | 34E000108 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No terminiation | [16 empty] | 380 | LS1_WA01:BPM_D1155 | 1155
BPM | LS1_WA01:BPM_D1155 | BPM | D1155 | 34E000391 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 2 | 6 |  | 1 |  |  |  |  | 380 | LS1_WA01:BPM_D1155 | 1155
BPM | LS1_WA01:BPM_D1155 | BPM | D1155 | 34E000392 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 2 | 6 |  | 2 |  |  |  |  | 380 | LS1_WA01:BPM_D1155 | 1155
BPM | LS1_WA01:BPM_D1155 | BPM | D1155 | 34E000393 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 2 | 6 |  | 3 |  |  |  |  | 380 | LS1_WA01:BPM_D1155 | 1155
BPM | LS1_WA01:BPM_D1155 | BPM | D1155 | 34E000394 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 2 | 6 |  | 4
BPM | LS1_WA01:RFS_D1155 | BPM | D1155 | 34E000306 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 2 | 6 |  | Ref |  |  |  |  | 392 | LS1_CA02:BPM_D1163 | 1163
BPM | LS1_CA02:BPM_D1163 | BPM | D1163 | 34E000343 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 2 | 6 |  | 5 |  |  |  |  | 392 | LS1_CA02:BPM_D1163 | 1163
BPM | LS1_CA02:BPM_D1163 | BPM | D1163 | 34E000344 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 2 | 6 |  | 6 |  |  |  |  | 392 | LS1_CA02:BPM_D1163 | 1163
BPM | LS1_CA02:BPM_D1163 | BPM | D1163 | 34E000345 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 2 | 6 |  | 7 |  |  |  |  | 392 | LS1_CA02:BPM_D1163 | 1163
BPM | LS1_CA02:BPM_D1163 | BPM | D1163 | 34E000346 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 2 | 6 |  | 8
BPM | LS1_CA02:RFS_D1163 | BPM | D1163 | 34E000305 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No terminiation
Panel 3 | [24 empty] | 407 | LS1_CA02:BPM_D1177 | 1177
BPM | LS1_CA02:BPM_D1177 | BPM | D1178 | 34E000339 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 2 | 7 |  | 1 |  |  |  |  | 407 | LS1_CA02:BPM_D1177 | 1177
BPM | LS1_CA02:BPM_D1177 | BPM | D1178 | 34E000340 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 2 | 7 |  | 2 |  |  |  |  | 407 | LS1_CA02:BPM_D1177 | 1177
BPM | LS1_CA02:BPM_D1177 | BPM | D1178 | 34E000341 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 2 | 7 |  | 3 |  |  |  |  | 407 | LS1_CA02:BPM_D1177 | 1177
BPM | LS1_CA02:BPM_D1177 | BPM | D1178 | 34E000342 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 2 | 7 |  | 4
BPM | LS1_CA02:RFS_D1177 | BPM | D1178 | 34E000304 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No terminiation | [16 empty] | 419 | LS1_WA02:BPM_D1188 | 1188
BPM | LS1_WA02:BPM_D1188 | BPM | D1189 | 34E000387 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 2 | 7 |  | 5 |  |  |  |  | 419 | LS1_WA02:BPM_D1188 | 1188
BPM | LS1_WA02:BPM_D1188 | BPM | D1189 | 34E000388 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 2 | 7 |  | 6 |  |  |  |  | 419 | LS1_WA02:BPM_D1188 | 1188
BPM | LS1_WA02:BPM_D1188 | BPM | D1189 | 34E000389 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 2 | 7 |  | 7 |  |  |  |  | 419 | LS1_WA02:BPM_D1188 | 1188
BPM | LS1_WA02:BPM_D1188 | BPM | D1189 | 34E000390 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 2 | 7 |  | 8
BPM | LS1_WA02:RFS_D1188 | BPM | D1189 | 34E000303 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 2 | 7 |  | Ref |  |  |  |  | 431 | LS1_CA03:BPM_D1196 | 1196
BPM | LS1_CA03:BPM_D1196 | BPM | D1196 | 34E000335 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 2 | 8 |  | 1 |  |  |  |  | 431 | LS1_CA03:BPM_D1196 | 1196
BPM | LS1_CA03:BPM_D1196 | BPM | D1196 | 34E000336 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 2 | 8 |  | 2 |  |  |  |  | 431 | LS1_CA03:BPM_D1196 | 1196
BPM | LS1_CA03:BPM_D1196 | BPM | D1196 | 34E000337 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 2 | 8 |  | 3 |  |  |  |  | 431 | LS1_CA03:BPM_D1196 | 1196
BPM | LS1_CA03:BPM_D1196 | BPM | D1196 | 34E000338 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 2 | 8 |  | 4
BPM | LS1_CA03:RFS_D1196 | BPM | D1196 | 34E000302 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 2 | 8 |  | Ref |  |  |  |  | 446 | LS1_CA03:BPM_D1211 | 1211
BPM | LS1_CA03:BPM_D1211 | BPM | D1212 | 34E000331 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 2 | 8 |  | 5 |  |  |  |  | 446 | LS1_CA03:BPM_D1211 | 1211
BPM | LS1_CA03:BPM_D1211 | BPM | D1212 | 34E000332 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 2 | 8 |  | 6 |  |  |  |  | 446 | LS1_CA03:BPM_D1211 | 1211
BPM | LS1_CA03:BPM_D1211 | BPM | D1212 | 34E000333 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 2 | 8 |  | 7 |  |  |  |  | 446 | LS1_CA03:BPM_D1211 | 1211
BPM | LS1_CA03:BPM_D1211 | BPM | D1212 | 34E000334 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 2 | 8 |  | 8
BPM | LS1_CA03:RFS_D1211 | BPM | D1212 | 34E000301 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No terminiation
Panel 4 | [24 empty] | 2741 | FS2_BMS:BPM_D4216 | 4216
BPM | FS2_BMS:BPM_D4216 | BPM | D4223 | 44E000065 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 2 | 9 |  | 1 |  |  |  |  | 2741 | FS2_BMS:BPM_D4216 | 4216
BPM | FS2_BMS:BPM_D4216 | BPM | D4223 | 44E000066 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 2 | 9 |  | 2 |  |  |  |  | 2741 | FS2_BMS:BPM_D4216 | 4216
BPM | FS2_BMS:BPM_D4216 | BPM | D4223 | 44E000067 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 2 | 9 |  | 3 |  |  |  |  | 2741 | FS2_BMS:BPM_D4216 | 4216
BPM | FS2_BMS:BPM_D4216 | BPM | D4223 | 44E000068 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 2 | 9 |  | 4
BPM | FS2_BMS:RFS_D4216 | BPM | D4223 | 44E000073 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No terminiation | [16 empty] | 2760 | FS2_BMS:BPM_D4283 | 4283
BPM | FS2_BMS:BPM_D4283 | BPM | D4283 | 44E000061 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 2 | 9 |  | 5 |  |  |  |  | 2760 | FS2_BMS:BPM_D4283 | 4283
BPM | FS2_BMS:BPM_D4283 | BPM | D4283 | 44E000062 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 2 | 9 |  | 6 |  |  |  |  | 2760 | FS2_BMS:BPM_D4283 | 4283
BPM | FS2_BMS:BPM_D4283 | BPM | D4283 | 44E000063 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 2 | 9 |  | 7 |  |  |  |  | 2760 | FS2_BMS:BPM_D4283 | 4283
BPM | FS2_BMS:BPM_D4283 | BPM | D4283 | 44E000064 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 2 | 9 |  | 8
BPM | FS2_BMS:RFS_D4283 | BPM | D4283 | 44E000072 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 2 | 9 |  | Ref |  |  |  |  | 2769 | FS2_BMS:BPM_D4326 | 4326
BPM | FS2_BMS:BPM_D4326 | BPM | D4327 | 44E000059 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 2 | 10 |  | 1 |  |  |  |  | 2769 | FS2_BMS:BPM_D4326 | 4326
BPM | FS2_BMS:BPM_D4326 | BPM | D4327 | 44E000069 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 2 | 10 |  | 2 |  |  |  |  | 2769 | FS2_BMS:BPM_D4326 | 4326
BPM | FS2_BMS:BPM_D4326 | BPM | D4327 | 44E000070 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 2 | 10 |  | 3 |  |  |  |  | 2769 | FS2_BMS:BPM_D4326 | 4326
BPM | FS2_BMS:BPM_D4326 | BPM | D4327 | 44E000071 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 2 | 10 |  | 4
BPM | FS2_BMS:RFS_D4326 | BPM | D4327 | 44E000060 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No terminiation | [16 empty] | 2814 | LS3_WD01:BPM_D4389 | 4389
BPM | LS3_WD01:BPM_D4389 | BPM | D4389 | 34E000092 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 2 | 10 |  | 5 |  |  |  |  | 2814 | LS3_WD01:BPM_D4389 | 4389
BPM | LS3_WD01:BPM_D4389 | BPM | D4389 | 34E000093 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 2 | 10 |  | 6 |  |  |  |  | 2814 | LS3_WD01:BPM_D4389 | 4389
BPM | LS3_WD01:BPM_D4389 | BPM | D4389 | 34E000094 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 2 | 10 |  | 7 |  |  |  |  | 2814 | LS3_WD01:BPM_D4389 | 4389
BPM | LS3_WD01:BPM_D4389 | BPM | D4389 | 34E000095 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 2 | 10 |  | 8
BPM | LS3_WD01:RFS_D4389 | BPM | D4389 | 34E000099 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 2 | 10 |  | Ref
Temp Pico | LS1_CA01:HMR_D0001 | HMR |  |  |  | Signal | [9 empty] | SMA to lemo | 2 | 11 | CH0 - fmc2-1
Temp Pico | LS1_CA01:HMR_D0002 | HMR |  |  |  | Signal | [9 empty] | SMA to lemo | 2 | 11 | CH1 - fmc2-2
Temp Pico | LS1_CA01:HMR_D0003 | HMR |  |  |  | Signal | [9 empty] | SMA to lemo | 2 | 11 | CH2 - fmc2-3
Temp Pico | LS1_CA01:HMR_D0004 | HMR |  |  |  | Signal | [9 empty] | SMA to lemo | 2 | 11 | CH3 - fmc2-4
Temp Pico | LS1_CA01:HMR_D0005 | HMR |  |  |  | Signal | [9 empty] | SMA to lemo | 2 | 11 | CH4 - fmc1-1
Temp Pico | LS1_CA01:HMR_D0006 | HMR |  |  |  | Signal | [9 empty] | SMA to lemo | 2 | 11 | CH5 - fmc1-2
Temp Pico | LS1_CA01:HMR_D0007 | HMR |  |  |  | Signal | [9 empty] | SMA to lemo | 2 | 11 | CH6 - fmc1-3
Temp Pico | LS1_CA01:HMR_D0008 | HMR |  |  |  | Signal | [9 empty] | SMA to lemo | 2 | 11 | CH7 - fmc1-4
Ground Plugs
Grounding Plug |  | ground |  | - | Lemo650 Plug | Ground | Pico-8 | [8 empty] | Ground Plug | 1 | 11 | CH3 - fmc2-4
Grounding Plug |  | ground |  | - | Lemo650 Plug | Ground | Pico-8 | [8 empty] | Ground Plug | 1 | 11 | CH7 - fmc1-4

## LS1_N0306
LS1_N0306
Long Run Cable | [6 empty] | Inter-Rack Connections | [9 empty] | Terminating Device
Diagnostics Device | Lattice Beamline Location | Device Type | CableDB Location | Cable Number | Cable Type | Function | Termination Device | Port Number In | Pin Number | CPC pinout | CPC pinout | CPC pinout | Signal Pin | Shield Pin | Port Number Out | Patch Cable | MTCA Chassis | MTCA Slot | PICO Port | FGPDB Port | Bias Supply Chassis | Bias Supply Slot | Bias Supply Channel |  |  |  |  | search for row | device name | D number
Ion Chamber
Ion Chamber | FS2_BMS:IC_D4141 | IC | D4141 | 44E000237 | 2C_1kV75LC_Coax_090 | Signal | N-Type to Lemo Box | J1 | [6 empty] | J1 | Lemo to Lemo | 3 | 3 | CH0 - fmc2-1 | [5 empty] | 3 | 3 | FS2_BMS:IC_D4141 | #N/A | #N/A | #N/A
Ion Chamber | FS2_BMS:IC_D4180 | IC | D4180 | 44E000238. | 2C_1kV75LC_Coax_090 | Signal | N-Type to Lemo Box | J2 | [6 empty] | J2 | Lemo to Lemo | 3 | 3 | CH1 - fmc2-2 | [5 empty] | 3 | 3 | FS2_BMS:IC_D4180 | #N/A | #N/A | #N/A
Ion Chamber | FS2_BMS:IC_D4221 | IC | D4221 | 44E000239 | 2C_1kV75LC_Coax_090 | Signal | N-Type to Lemo Box | J3 | [6 empty] | J3 | Lemo to Lemo | 3 | 3 | CH2 - fmc2-3 | [5 empty] | 3 | 3 | FS2_BMS:IC_D4221 | #N/A | #N/A | #N/A
Ion Chamber | FS2_BMS:IC_D4277 | IC | D4277 | 44E000240 | 2C_1kV75LC_Coax_090 | Signal | N-Type to Lemo Box | J4 | [6 empty] | J4 | Lemo to Lemo | 3 | 3 | CH3 - fmc2-4 | [5 empty] | 3 | 3 | FS2_BMS:IC_D4277 | #N/A | #N/A | #N/A
Ion Chamber | FS2_BMS:IC_D4305 | IC | D4305 | 44E000241 | 2C_1kV75LC_Coax_090 | Signal | N-Type to Lemo Box | J5 | [6 empty] | J5 | Lemo to Lemo | 3 | 3 | CH4 - fmc1-1 | [5 empty] | 3 | 3 | FS2_BMS:IC_D4305 | #N/A | #N/A | #N/A
Ion Chamber | LS3_CD01:IC_D4385 | IC | D4385 | 34E000558 | 2C_1kV75LC_Coax_090 | Signal | N-Type to Lemo Box | J6 | [6 empty] | J6 | Lemo to Lemo | 3 | 3 | CH5 - fmc1-2 | [5 empty] | 3 | 3 | LS3_CD01:IC_D4385 | #N/A | #N/A | #N/A
Ion Chamber | LS3_CD03:IC_D4503 | IC | D4503 | 34E000559 | 2C_1kV75LC_Coax_090 | Signal | N-Type to Lemo Box | J7 | [6 empty] | J7 | Lemo to Lemo | 3 | 3 | CH6 - fmc1-3 | [5 empty] | 3 | 3 | LS3_CD03:IC_D4503 | #N/A | #N/A | #N/A
Ion Chamber Bias
Ion Chamber | FS2_BMS:IC_D4141 |  | D4141 | 44H000021 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 1 | [6 empty] | 1 | SHV to SHV |  |  |  |  | 2 | 1 : EHS8020n | Ch0 |  |  |  | FS2_BMS:IC_D4141
Ion Chamber | FS2_BMS:IC_D4181 |  | D4181 | 44H000022 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 2 | [6 empty] | 2 | SHV to SHV |  |  |  |  | 2 | 1 : EHS8020n | Ch1 |  |  |  | FS2_BMS:IC_D4181
Ion Chamber | FS2_BMS:IC_D4221 |  | D4221 | 44H000023 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 3 | [6 empty] | 3 | SHV to SHV |  |  |  |  | 2 | 1 : EHS8020n | Ch2 |  |  |  | FS2_BMS:IC_D4221
Ion Chamber | FS2_BMS:HMR_D4277 |  | D4277 | 44H000024 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 4 | [6 empty] | 4 | SHV to SHV |  |  |  |  | 2 | 1 : EHS8020n | Ch3 |  |  |  | FS2_BMS:HMR_D4277
Ion Chamber | FS2_BMS:HMR_D4305 |  | D4305 | 44H000025 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 5 | [6 empty] | 5 | SHV to SHV |  |  |  |  | 2 | 1 : EHS8020n | Ch4 |  |  |  | FS2_BMS:HMR_D4305
Ion Chamber | LS3_CD01:HMR_D4385 |  | D4385 | 34H000076 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 6 | [6 empty] | 6 | SHV to SHV |  |  |  |  | 2 | 1 : EHS8020n | Ch5 |  |  |  | LS3_CD01:HMR_D4385
Ion Chamber | LS3_CD03:HMR_D4503 |  | D4503 | 34H000077 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 7 | [6 empty] | 7 | SHV to SHV |  |  |  |  | 2 | 1 : EHS8020n | Ch6
faraday Cup Bias
Faraday Cup | FS2_BMS:FC_D4253 |  | D4253 | 34H000077 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 8 | [6 empty] | 8 | SHV to SHV |  |  |  |  | 2 | 1 : EHS8020n | Ch7 |  |  |  | FS2_BMS:FC_D4253
Faraday Cup
Faraday Cup | LS1_WA03:FC_D1224 | FC | D1224 | 34H000007 | 2C_1kV75LC_Coax_090 | Signal | N-Type to SMA Panel | J1 | [6 empty] | J1 | Lemo250/Lemo650 | 3 | 4 | CH0 - fmc2-1 |  | 2 | 1: EBS8003 | Ch0 | 1:8 current divider installed | 3 | 4 | LS1_WA03:FC_D1224
Faraday Cup | FS2_BMS:FC_D4253 | FC | D4253 | 44H000001 | 2C_1kV75LC_Coax_090 | Signal | N-Type to SMA Panel | J2 | [6 empty] | J2 | Lemo250/Lemo650 | 3 | 5 | CH0 - fmc2-1 |  | 2 | 1 : EHS8020n | Ch7 |  | 3 | 5 | FS2_BMS:FC_D4253
Halo Monitor Ring
Halo Monitor Ring,Spare | FS2_BMS:HMR_D4144 | HMR - Unused | D4144 | 44H000000 | 2C_1kV75LC_Coax_090 | Signal | [20 empty] | FS2_BMS:HMR_D4144
| FS2_BMS:HMR_D4144 |  | D4144 | 44H000055 | 2C_5kVDC_Coax_035 | [21 empty] | FS2_BMS:HMR_D4144
Halo Monitor Ring,Spare | FS2_BMS:HMR_D4178 | HMR - Unused | D4179 | 44H000001 | 2C_1kV75LC_Coax_090 | Signal | [20 empty] | FS2_BMS:HMR_D4178
| FS2_BMS:HMR_D4178 |  | D4179 | 44H000054 | 2C_5kVDC_Coax_035 | [21 empty] | FS2_BMS:HMR_D4178
Halo Monitor Ring | FS2_BMS:HMR_D4328 | HMR | D4328 | 44H000002 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J2 | [6 empty] | J2 | SHV to Lemo | 3 | 4 | CH4 - fmc1-1 |  | 2 | 1: EBS8003 | Ch1 |  | 3 | 4 | FS2_BMS:HMR_D4328
| FS2_BMS:HMR_D4328 |  | D4328 | 44H000057 | 2C_5kVDC_Coax_035 | [21 empty] | FS2_BMS:HMR_D4328
Halo Monitor Ring,Spare | LS3_WD01:HMR_D4390 | HMR - Unused | D4390 | 34H000009 | 2C_1kV75LC_Coax_090 | Signal | [20 empty] | LS3_WD01:HMR_D4390
| LS3_WD01:HMR_D4390 |  | D4390 | 34H000100 | 2C_5kVDC_Coax_035 | [21 empty] | LS3_WD01:HMR_D4390
Halo Monitor Ring,Spare | LS3_WD02:HMR_D4452 | HMR - Unused | D4452 | 34H000008 | 2C_1kV75LC_Coax_090 | Signal | [20 empty] | LS3_WD02:HMR_D4452
| LS3_WD02:HMR_D4452 |  | D4452 | 34H000101 | 2C_5kVDC_Coax_035 | [21 empty] | LS3_WD02:HMR_D4452
| LS3_WD03:HMR_D4514 |  | D4514 | 34H000086 | 2C_5kVDC_Coax_035 | [21 empty] | LS3_WD03:HMR_D4514
Halo Monitor Ring | LS3_WD04:HMR_D4576 | HMR | D4576 | 34H000010 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J3 | [6 empty] | J3 | SHV to Lemo | 3 | 4 | CH5 - fmc1-2 |  | 2 | 1: EBS8003 | Ch1 |  | 3 | 4 | LS3_WD04:HMR_D4576
| LS3_WD04:HMR_D4576 |  | D4576 | 34H000099 | 2C_5kVDC_Coax_035 | [21 empty] | LS3_WD04:HMR_D4576
Halo Monitor Ring,Spare | LS3_WD05:HMR_D4638 | HMR - Unused | D4638 | 34H000011 | 2C_1kV75LC_Coax_090 | Signal | [20 empty] | LS3_WD05:HMR_D4638
| LS3_WD05:HMR_D4638 |  | D4638 | 34H000095 | 2C_5kVDC_Coax_035 | [21 empty] | LS3_WD05:HMR_D4638
Lemo to Lemo Panel
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 1_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 2 | 1: EBS8003 | Ch0 |  |  |  | PM bias for PICO
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 1_OUT Port | [7 empty] | Lemo650/Lemo650 | 3 | 4 | CH3 - fmc2-4 | [5 empty] | 3 | 4 | PM bias for PICO
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 2_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 2 | 1: EBS8003 | Ch1 |  |  |  | PM bias for PICO
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 2_OUT Port | [7 empty] | Lemo650/Lemo650 | 3 | 4 | CH7 - fmc1-4 | [5 empty] | 3 | 4 | PM bias for PICO
Ground Plugs
Grounding Plug |  | ground |  | - | Lemo650 Plug | Ground | Pico-8 | [8 empty] | Ground Plug | 3 | 5 | CH3 - fmc2-4 | [5 empty] | 3 | 5

## LS1_N0604
LS1_N0604
Long Run Cable | [6 empty] | Inter-Rack Connections | [9 empty] | Terminating Device
Diagnostics Device | Lattice Beamline Location | Device Type | CableDB Location | Cable Number | Cable Type | Function | Termination Device | Port Number In | Pin Number | CPC pinout | CPC pinout | CPC pinout | Signal Pin | Shield Pin | Port Number Out | Patch Cable | MTCA Chassis | MTCA Slot | PICO Port | FGPDB Port | Bias Supply Chassis | Bias Supply Slot | Bias Supply Channel |  |  |  |  | search for row | device name | D number
Neuton Detector | [24 empty] | PICO | CH[0-7] | Device
Neutron Detector | FE_RFQ:ND_D1025 | ND | D1025 | 24E000066 | 2C_75LC_Coax_089 | Signal | N-type to Lemo Box | J1 | [6 empty] | J1 | Lemo to Lemo | 4 | 3 | CH0 - fmc2-1 | [8 empty] | #N/A | #N/A | #N/A
Neutron Detector | FE_RFQ:ND_D1025 |  | D1025 | 24E000067 | 2C_75LC_Coax_089 | Test Signal | N-type to SMA Panel | J1 | [6 empty] | J1 | SMA to Ferrule | [10 empty] | FE_RFQ:ND_D1025 | #N/A | #N/A | #N/A
Neutron Detector | LS1_CA01:ND_D1128 | ND | D1128 | 34E000522 | 2C_75LC_Coax_089 | Signal | N-type to Lemo Box | J2 | [6 empty] | J2 | Lemo to Lemo | 4 | 3 | CH1 - fmc2-2 | [5 empty] | 4 | 3 | LS1_CA01:ND_D1128 | #N/A | #N/A | #N/A
Neutron Detector | LS1_CA01:ND_D1128 |  | D1128 | 34E000521 | 2C_75LC_Coax_089 | Test signal | N-type to SMA Panel | J2 | [6 empty] | J2 | SMA to Ferrule | [11 empty] | #N/A | #N/A | #N/A
Neutron Detector | LS1_CA03:ND_D1195 | ND | D1195 | 34E000524 | 2C_75LC_Coax_089 | Signal | N-type to Lemo Box | J3 | [6 empty] | J3 | Lemo to Lemo | 4 | 3 | CH2 - fmc2-3 | [5 empty] | 4 | 3 | LS1_CA03:ND_D1195 | #N/A | #N/A | #N/A
Neutron Detector | LS1_CA03:ND_D1195 |  | D1195 | 34E000523 | 2C_75LC_Coax_089 | Test signal | N-type to SMA Panel | J3 | [6 empty] | J3 | SMA to Ferrule | [10 empty] | LS1_CA03:ND_D1195 | #N/A | #N/A | #N/A
Neutron Detector | LS1_CB01:ND_D1250 | ND | D1250 | 34E000520 | 2C_75LC_Coax_089 | Signal | N-type to Lemo Box | J4 | [6 empty] | J4 | Lemo to Lemo | 4 | 3 | CH4 - fmc1-1 | [8 empty] | #N/A | #N/A | #N/A
Neutron Detector | LS1_CB01:ND_D1250 |  | D1250 | 34E000519 | 2C_75LC_Coax_089 | Test signal | N-type to SMA Panel | J4 | [6 empty] | J4 | SMA to Ferrule | [10 empty] | LS1_CB01:ND_D1250 | #N/A | #N/A | #N/A
Neutron Detector | LS1_CB03:ND_D1378 | ND | D1378 | 34E000518 | 2C_75LC_Coax_089 | Signal | N-type to Lemo Box | J5 | [6 empty] | J5 | Lemo to Lemo | 4 | 3 | CH5 - fmc1-2 | [5 empty] | 4 | 3 | LS1_CB03:ND_D1378 | #N/A | #N/A | #N/A
Neutron Detector | LS1_CB03:ND_D1378 |  | D1378 | 34E000517 | 2C_75LC_Coax_089 | Test signal | N-type to SMA Panel | J5 | [6 empty] | J5 | SMA to Ferrule | [10 empty] | LS1_CB03:ND_D1378 | #N/A | #N/A | #N/A
Neutron Detector | LS3_CD06:ND_D4647 | ND | D4647 | 34E000516 | 2C_75LC_Coax_089 | Signal | N-type to Lemo Box | J6 | [6 empty] | J6 | Lemo to Lemo | 4 | 3 | CH6 - fmc1-3 | [5 empty] | 4 | 3 | LS3_CD06:ND_D4647 | #N/A | #N/A | #N/A
Neutron Detector | LS3_CD06:ND_D4647 |  | D4647 | 34E000515 | 2C_75LC_Coax_089 | Test signal | N-type to SMA Panel | J6 | [6 empty] | J6 | SMA to Ferrule | [10 empty] | LS3_CD06:ND_D4647 | #N/A | #N/A | #N/A
Neuton Detector Bias
Neutron Detector | FE_RFQ:ND_D1025 |  | D1025 | 24H000006 | 2C_5kVDC_Coax_035 | HV Bias | Direct to ISEG | [13 empty] | 3 | 2: EHS8020n | Ch0
Neutron Detector | LS1_CA01:ND_D1128 |  | D1128 | 34H000057 | 2C_5kVDC_Coax_035 | HV Bias | Direct to ISEG | [13 empty] | 3 | 2: EHS8020n | Ch1
Neutron Detector | LS1_CA03:ND_D1195 |  | D1195 | 34H000058 | 2C_5kVDC_Coax_035 | HV Bias | Direct to ISEG | [13 empty] | 3 | 2: EHS8020n | Ch2
Neutron Detector | LS1_CB01:ND_D1250 |  | D1250 | 34H000056 | 2C_5kVDC_Coax_035 | HV Bias | Direct to ISEG | [13 empty] | 3 | 2: EHS8020n | Ch3
Neutron Detector | LS1_CB03:ND_D1378 |  | D1378 | 34H000055 | 2C_5kVDC_Coax_035 | HV Bias | Direct to ISEG | [13 empty] | 3 | 2: EHS8020n | Ch4
Neutron Detector | LS3_CD06:ND_D4647 |  | D4647 | 34H000054 | 2C_5kVDC_Coax_035 | HV Bias | Direct to ISEG | [13 empty] | 3 | 2: EHS8020n | Ch5
Halo Monitor Ring
Halo Monitor Ring | LS1_WA01:HMR_D1156 | HMR | D1156 | 34H000029 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J1 | [6 empty] | J1 | SHV to Lemo | 4 | 4 | CH0 - fmc2-1 |  | 3 | 1: EBS8003 | Ch0 |  | 4 | 4 | LS1_WA01:HMR_D1156 | 382 | LS1_WA01:HMR_D1156 | 1156
| LS1_WA01:HMR_D1156 |  | D1156 | 34H000110 | 2C_5kVDC_Coax_035 | [21 empty] | LS1_WA01:HMR_D1156 | 382 | LS1_WA01:HMR_D1156 | 1156
Halo Monitor Ring | LS1_WA02:HMR_D1190 | HMR | D1190 | 34H000024 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J2 | [6 empty] | J2 | SHV to Lemo | 4 | 4 | CH1 - fmc2-2 |  | 3 | 1: EBS8003 | Ch0 |  | 4 | 4 | LS1_WA02:HMR_D1190 | 421 | LS1_WA02:HMR_D1190 | 1190
| LS1_WA02:HMR_D1190 |  | D1190 | 34H000114 | 2C_5kVDC_Coax_035 | [21 empty] | LS1_WA02:HMR_D1190 | 421 | LS1_WA02:HMR_D1190 | 1190
Halo Monitor Ring | LS1_WA03:HMR_D1224 | HMR | D1223 | 34H000026 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J3 | [6 empty] | J3 | SHV to Lemo | 4 | 4 | CH2 - fmc2-3 |  | 3 | 1: EBS8003 | Ch0 | Spliter Installed | 4 | 4 | LS1_WA03:HMR_D1224 | 460 | LS1_WA03:HMR_D1224 | 1224
| LS1_WA03:HMR_D1224 |  | D1223 | 34H000109 | 2C_5kVDC_Coax_035 | [21 empty] | LS1_WA03:HMR_D1224 | 460 | LS1_WA03:HMR_D1224 | 1224
Halo Monitor Ring | LS1_WB01:HMR_D1287 | HMR | D1287 | 34H000027 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J4 | [6 empty] | J4 | SHV to Lemo | 4 | 4 | CH4 - fmc1-1 |  | 3 | 1: EBS8003 | Ch1 |  | 4 | 4 | LS1_WB01:HMR_D1287 | 523 | LS1_WB01:HMR_D1287 | 1287
| LS1_WB01:HMR_D1287 |  | D1287 | 34H000111 | 2C_5kVDC_Coax_035 | [21 empty] | LS1_WB01:HMR_D1287 | 523 | LS1_WB01:HMR_D1287 | 1287
Halo Monitor Ring | LS1_WB02:HMR_D1351 | HMR | D1351 | 34H000028 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J5 | [6 empty] | J5 | SHV to Lemo | 4 | 4 | CH5 - fmc1-2 |  | 3 | 1: EBS8003 | Ch1 |  |  |  |  | 583 | LS1_WB02:HMR_D1351 | 1351
| LS1_WB02:HMR_D1351 |  | D1351 | 34H000112 | 2C_5kVDC_Coax_035 | [22 empty] | 583 | LS1_WB02:HMR_D1351 | 1351
Halo Monitor Ring | LS1_WB03:HMR_D1415 | HMR | D1415 | 34H000025 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J6 | [6 empty] | J6 | SHV to Lemo | 4 | 4 | CH6 - fmc1-3 |  | 3 | 1: EBS8003 | Ch1 | Spliter Installed |  |  |  | 643 | LS1_WB03:HMR_D1415 | 1415
| LS1_WB03:HMR_D1415 |  | D1415 | 34H000113 | 2C_5kVDC_Coax_035 | [22 empty] | 643 | LS1_WB03:HMR_D1415 | 1415
Beam Position Monitors
BPM | LS1_WA03:BPM_D1222 | BPM | D1222 | 34E000383 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 4 | 5 |  | 1 |  |  |  |  | 4 | 5 | LS1_WA03:BPM_D1222 | 458 | LS1_WA03:BPM_D1222 | 1222
BPM | LS1_WA03:BPM_D1222 | BPM | D1222 | 34E000384 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 4 | 5 |  | 2 |  |  |  |  | 4 | 5 | LS1_WA03:BPM_D1222 | 458 | LS1_WA03:BPM_D1222 | 1222
BPM | LS1_WA03:BPM_D1222 | BPM | D1222 | 34E000385 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 4 | 5 |  | 3 |  |  |  |  | 4 | 5 | LS1_WA03:BPM_D1222 | 458 | LS1_WA03:BPM_D1222 | 1222
BPM | LS1_WA03:BPM_D1222 | BPM | D1222 | 34E000386 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 4 | 5 |  | 4 |  |  |  |  | 4 | 5 | LS1_WA03:BPM_D1222 | 458 | LS1_WA03:BPM_D1222 | 1222
BPM | LS1_WA03:BPM_D1222 | BPM | D1222 | 34E000300 | 2C_1/4in_Hardline_036 | RF Pick-up |  | Ref | [6 empty] | Ref | SMA to SMA | 4 | 5 |  | Ref |  |  |  |  | 4 | 5 | LS1_WA03:BPM_D1222 | 458 | LS1_WA03:BPM_D1222 | 1222
BPM | LS1_CB01:BPM_D1231 | BPM | D1231 | 34E000323 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 4 | 5 |  | 5 |  |  |  |  | 4 | 5 | LS1_CB01:BPM_D1231 | 473 | LS1_CB01:BPM_D1231 | 1231
BPM | LS1_CB01:BPM_D1231 | BPM | D1231 | 34E000324 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 4 | 5 |  | 6 |  |  |  |  | 4 | 5 | LS1_CB01:BPM_D1231 | 473 | LS1_CB01:BPM_D1231 | 1231
BPM | LS1_CB01:BPM_D1231 | BPM | D1231 | 34E000325 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 4 | 5 |  | 7 |  |  |  |  | 4 | 5 | LS1_CB01:BPM_D1231 | 473 | LS1_CB01:BPM_D1231 | 1231
BPM | LS1_CB01:BPM_D1231 | BPM | D1231 | 34E000326 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 4 | 5 |  | 8 |  |  |  |  | 4 | 5 | LS1_CB01:BPM_D1231 | 473 | LS1_CB01:BPM_D1231 | 1231
BPM | LS1_CB01:BPM_D1231 | BPM | D1231 | 34E000299 | 2C_1/4in_Hardline_036 | RF Pick-up |  | Ref | [18 empty] | LS1_CB01:BPM_D1231 | 473 | LS1_CB01:BPM_D1231 | 1231
BPM | LS1_CB01:BPM_D1251 | BPM | D1251 | 34E000312 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 4 | 6 |  | 1 |  |  |  |  | 4 | 6 | LS1_CB01:BPM_D1251 | 491 | LS1_CB01:BPM_D1251 | 1251
BPM | LS1_CB01:BPM_D1251 | BPM | D1251 | 34E000313 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 4 | 6 |  | 2 |  |  |  |  | 4 | 6 | LS1_CB01:BPM_D1251 | 491 | LS1_CB01:BPM_D1251 | 1251
BPM | LS1_CB01:BPM_D1251 | BPM | D1251 | 34E000314 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 4 | 6 |  | 3 |  |  |  |  | 4 | 6 | LS1_CB01:BPM_D1251 | 491 | LS1_CB01:BPM_D1251 | 1251
BPM | LS1_CB01:BPM_D1251 | BPM | D1251 | 34E000315 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 4 | 6 |  | 4 |  |  |  |  | 4 | 6 | LS1_CB01:BPM_D1251 | 491 | LS1_CB01:BPM_D1251 | 1251
BPM | LS1_CB01:BPM_D1251 | BPM | D1251 | 34E000298 | 2C_1/4in_Hardline_036 | RF Pick-up |  | Ref | [6 empty] | Ref | SMA to SMA | 4 | 6 |  | Ref |  |  |  |  | 4 | 6 | LS1_CB01:BPM_D1251 | 491 | LS1_CB01:BPM_D1251 | 1251
BPM | LS1_CB01:BPM_D1271 | BPM | D1271 | 34E000308 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 4 | 6 |  | 5 |  |  |  |  | 4 | 6 | LS1_CB01:BPM_D1271 | 509 | LS1_CB01:BPM_D1271 | 1271
BPM | LS1_CB01:BPM_D1271 | BPM | D1271 | 34E000309 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 4 | 6 |  | 6 |  |  |  |  | 4 | 6 | LS1_CB01:BPM_D1271 | 509 | LS1_CB01:BPM_D1271 | 1271
BPM | LS1_CB01:BPM_D1271 | BPM | D1271 | 34E000310 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 4 | 6 |  | 7 |  |  |  |  | 4 | 6 | LS1_CB01:BPM_D1271 | 509 | LS1_CB01:BPM_D1271 | 1271
BPM | LS1_CB01:BPM_D1271 | BPM | D1271 | 34E000311 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 4 | 6 |  | 8 |  |  |  |  | 4 | 6 | LS1_CB01:BPM_D1271 | 509 | LS1_CB01:BPM_D1271 | 1271
BPM | LS1_CB01:BPM_D1271 | BPM | D1271 | 34E000297 | 2C_1/4in_Hardline_036 | RF Pick-up |  | Ref | [18 empty] | LS1_CB01:BPM_D1271 | 509 | LS1_CB01:BPM_D1271 | 1271
BPM | LS1_WB01:BPM_D1286 | BPM | D1286 | 34E000379 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 4 | 7 |  | 1 |  |  |  |  | 4 | 7 | LS1_WB01:BPM_D1286 | 521 | LS1_WB01:BPM_D1286 | 1286
BPM | LS1_WB01:BPM_D1286 | BPM | D1286 | 34E000380 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 4 | 7 |  | 2 |  |  |  |  | 4 | 7 | LS1_WB01:BPM_D1286 | 521 | LS1_WB01:BPM_D1286 | 1286
BPM | LS1_WB01:BPM_D1286 | BPM | D1286 | 34E000381 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 4 | 7 |  | 3 |  |  |  |  | 4 | 7 | LS1_WB01:BPM_D1286 | 521 | LS1_WB01:BPM_D1286 | 1286
BPM | LS1_WB01:BPM_D1286 | BPM | D1286 | 34E000382 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 4 | 7 |  | 4 |  |  |  |  | 4 | 7 | LS1_WB01:BPM_D1286 | 521 | LS1_WB01:BPM_D1286 | 1286
BPM | LS1_WB01:BPM_D1286 | BPM | D1286 | 34E000296 | 2C_1/4in_Hardline_036 | RF Pick-up |  | Ref | [6 empty] | Ref | SMA to SMA | 4 | 7 |  | Ref |  |  |  |  | 4 | 7 | LS1_WB01:BPM_D1286 | 521 | LS1_WB01:BPM_D1286 | 1286
BPM | LS1_CB02:BPM_D1295 | BPM | D1295 | 34E000316 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 4 | 7 |  | 5 |  |  |  |  | 4 | 7 | LS1_CB02:BPM_D1295 | 533 | LS1_CB02:BPM_D1295 | 1295
BPM | LS1_CB02:BPM_D1295 | BPM | D1295 | 34E000317 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 4 | 7 |  | 6 |  |  |  |  | 4 | 7 | LS1_CB02:BPM_D1295 | 533 | LS1_CB02:BPM_D1295 | 1295
BPM | LS1_CB02:BPM_D1295 | BPM | D1295 | 34E000318 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 4 | 7 |  | 7 |  |  |  |  | 4 | 7 | LS1_CB02:BPM_D1295 | 533 | LS1_CB02:BPM_D1295 | 1295
BPM | LS1_CB02:BPM_D1295 | BPM | D1295 | 34E000319 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 4 | 7 |  | 8 |  |  |  |  | 4 | 7 | LS1_CB02:BPM_D1295 | 533 | LS1_CB02:BPM_D1295 | 1295
BPM | LS1_CB02:BPM_D1295 | BPM | D1295 | 34E000295 | 2C_1/4in_Hardline_036 | RF Pick-up |  | Ref | [18 empty] | LS1_CB02:BPM_D1295 | 533 | LS1_CB02:BPM_D1295 | 1295
BPM | LS1_CB02:BPM_D1315 | BPM | D1315 | 34E000327 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 4 | 8 |  | 1 |  |  |  |  | 4 | 8 | LS1_CB02:BPM_D1315 | 551 | LS1_CB02:BPM_D1315 | 1315
BPM | LS1_CB02:BPM_D1315 | BPM | D1315 | 34E000328 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 4 | 8 |  | 2 |  |  |  |  | 4 | 8 | LS1_CB02:BPM_D1315 | 551 | LS1_CB02:BPM_D1315 | 1315
BPM | LS1_CB02:BPM_D1315 | BPM | D1315 | 34E000329 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 4 | 8 |  | 3 |  |  |  |  | 4 | 8 | LS1_CB02:BPM_D1315 | 551 | LS1_CB02:BPM_D1315 | 1315
BPM | LS1_CB02:BPM_D1315 | BPM | D1315 | 34E000330 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 4 | 8 |  | 4 |  |  |  |  | 4 | 8 | LS1_CB02:BPM_D1315 | 551 | LS1_CB02:BPM_D1315 | 1315
BPM | LS1_CB02:BPM_D1315 | BPM | D1315 | 34E000294 | 2C_1/4in_Hardline_036 | RF Pick-up |  | Ref | [6 empty] | Ref | SMA to SMA | 4 | 8 |  | Ref |  |  |  |  | 4 | 8 | LS1_CB02:BPM_D1315 | 551 | LS1_CB02:BPM_D1315 | 1315
BPM | LS1_CB02:BPM_D1335 | BPM | D1335 | 34E000307 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 4 | 8 |  | 5 |  |  |  |  | 4 | 8 | LS1_CB02:BPM_D1335 | 569 | LS1_CB02:BPM_D1335 | 1335
BPM | LS1_CB02:BPM_D1335 | BPM | D1335 | 34E000320 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 4 | 8 |  | 6 |  |  |  |  | 4 | 8 | LS1_CB02:BPM_D1335 | 569 | LS1_CB02:BPM_D1335 | 1335
BPM | LS1_CB02:BPM_D1335 | BPM | D1335 | 34E000321 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 4 | 8 |  | 7 |  |  |  |  | 4 | 8 | LS1_CB02:BPM_D1335 | 569 | LS1_CB02:BPM_D1335 | 1335
BPM | LS1_CB02:BPM_D1335 | BPM | D1335 | 34E000322 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 4 | 8 |  | 8 |  |  |  |  | 4 | 8 | LS1_CB02:BPM_D1335 | 569 | LS1_CB02:BPM_D1335 | 1335
BPM | LS1_CB02:BPM_D1335 | BPM | D1335 | 34E000293 | 2C_1/4in_Hardline_036 | RF Pick-up |  | Ref | [19 empty] | 569 | LS1_CB02:BPM_D1335 | 1335
BPM | LS1_WB02:BPM_D1350 | BPM | D1350 | 34E000375 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 4 | 9 |  | 1 |  |  |  |  | 4 | 9 | LS1_WB02:BPM_D1350 | 581 | LS1_WB02:BPM_D1350 | 1350
BPM | LS1_WB02:BPM_D1350 | BPM | D1350 | 34E000376 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 4 | 9 |  | 2 |  |  |  |  | 4 | 9 | LS1_WB02:BPM_D1350 | 581 | LS1_WB02:BPM_D1350 | 1350
BPM | LS1_WB02:BPM_D1350 | BPM | D1350 | 34E000377 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 4 | 9 |  | 3 |  |  |  |  | 4 | 9 | LS1_WB02:BPM_D1350 | 581 | LS1_WB02:BPM_D1350 | 1350
BPM | LS1_WB02:BPM_D1350 | BPM | D1350 | 34E000378 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 4 | 9 |  | 4 |  |  |  |  | 4 | 9 | LS1_WB02:BPM_D1350 | 581 | LS1_WB02:BPM_D1350 | 1350
BPM | LS1_WB02:BPM_D1350 | BPM | D1350 | 34E000292 | 2C_1/4in_Hardline_036 | RF Pick-up |  | Ref | [6 empty] | Ref | SMA to SMA | 4 | 9 |  | Ref |  |  |  |  | 4 | 9 | LS1_WB02:BPM_D1350 | 581 | LS1_WB02:BPM_D1350 | 1350
BPM | LS1_CB03:BPM_D1359 | BPM | D1359 | 34E000371 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 4 | 9 |  | 5 | [7 empty] | 593 | LS1_CB03:BPM_D1359 | 1359
BPM | LS1_CB03:BPM_D1359 | BPM | D1359 | 34E000372 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 4 | 9 |  | 6 |  |  |  |  | 4 | 9 | LS1_CB03:BPM_D1359 | 593 | LS1_CB03:BPM_D1359 | 1359
BPM | LS1_CB03:BPM_D1359 | BPM | D1359 | 34E000373 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 4 | 9 |  | 7 |  |  |  |  | 4 | 9 | LS1_CB03:BPM_D1359 | 593 | LS1_CB03:BPM_D1359 | 1359
BPM | LS1_CB03:BPM_D1359 | BPM | D1359 | 34E000374 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 4 | 9 |  | 8 |  |  |  |  | 4 | 9 | LS1_CB03:BPM_D1359 | 593 | LS1_CB03:BPM_D1359 | 1359
BPM | LS1_CB03:BPM_D1359 | BPM | D1359 | 34E000291 | 2C_1/4in_Hardline_036 | RF Pick-up |  | Ref | [18 empty] | LS1_CB03:BPM_D1359 | 593 | LS1_CB03:BPM_D1359 | 1359
BPM | LS1_CB03:BPM_D1379 | BPM | D1379 | 34E000352 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 4 | 10 |  | 1 | [7 empty] | 611 | LS1_CB03:BPM_D1379 | 1379
BPM | LS1_CB03:BPM_D1379 | BPM | D1379 | 34E000353 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 4 | 10 |  | 2 | [7 empty] | 611 | LS1_CB03:BPM_D1379 | 1379
BPM | LS1_CB03:BPM_D1379 | BPM | D1379 | 34E000354 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 4 | 10 |  | 3 | [7 empty] | 611 | LS1_CB03:BPM_D1379 | 1379
BPM | LS1_CB03:BPM_D1379 | BPM | D1379 | 34E000355 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 4 | 10 |  | 4 | [7 empty] | 611 | LS1_CB03:BPM_D1379 | 1379
BPM | LS1_CB03:BPM_D1379 | BPM | D1379 | 34E000287 | 2C_1/4in_Hardline_036 | RF Pick-up |  | Ref | [6 empty] | Ref | SMA to SMA | 4 | 10 |  | Ref | [7 empty] | 611 | LS1_CB03:BPM_D1379 | 1379
BPM | LS1_CB03:BPM_D1399 | BPM | D1399 | 34E000348 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 4 | 10 |  | 5 | [7 empty] | 629 | LS1_CB03:BPM_D1399 | 1399
BPM | LS1_CB03:BPM_D1399 | BPM | D1399 | 34E000349 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 4 | 10 |  | 6 | [7 empty] | 629 | LS1_CB03:BPM_D1399 | 1399
BPM | LS1_CB03:BPM_D1399 | BPM | D1399 | 34E000350 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 4 | 10 |  | 7 | [7 empty] | 629 | LS1_CB03:BPM_D1399 | 1399
BPM | LS1_CB03:BPM_D1399 | BPM | D1399 | 34E000351 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 4 | 10 |  | 8 | [7 empty] | 629 | LS1_CB03:BPM_D1399 | 1399
BPM | LS1_CB03:BPM_D1399 | BPM | D1399 | 34E000286 | 2C_1/4in_Hardline_036 | RF Pick-up |  | Ref | [19 empty] | 629 | LS1_CB03:BPM_D1399 | 1399
Lemo to Lemo Panel
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 1_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 3 | 1: EBS8003 | Ch0
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 1_OUT Port | [7 empty] | Lemo650/Lemo650 | 4 | 4 | CH3 - fmc2-4
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 2_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 3 | 1: EBS8003 | Ch1
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 2_OUT Port | [7 empty] | Lemo650/Lemo650 | 4 | 4 | CH7 - fmc1-4

## LS1_N1102
LS1_N1102
Long Run Cable | [6 empty] | Inter-Rack Connections | [9 empty] | Terminating Device
Diagnostics Device | Lattice Beamline Location | Device Type | CableDB Location | Cable Number | Cable Type | Function | Termination Device | Port Number In | Pin Number | CPC pinout | CPC pinout | CPC pinout | Signal Pin | Shield Pin | Port Number Out | Patch Cable | MTCA Chassis | MTCA Slot | PICO Port | FGPDB Port | Bias Supply Chassis | Bias Supply Slot | Bias Supply Channel |  |  |  | search for row | device name | D number
Beam Position Monitors
BPM | LS1_WB03:BPM_D1413 | BPM | D1413 | 34E000367 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 5 | 3 |  | 1 |  |  |  | 5 | 3 | LS1_WB03:BPM_D1413 | 641 | LS1_WB03:BPM_D1413 | 1413
BPM | LS1_WB03:BPM_D1413 | BPM | D1413 | 34E000368 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 5 | 3 |  | 2 |  |  |  | 5 | 3 | LS1_WB03:BPM_D1413 | 641 | LS1_WB03:BPM_D1413 | 1413
BPM | LS1_WB03:BPM_D1413 | BPM | D1413 | 34E000369 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 5 | 3 |  | 3 |  |  |  | 5 | 3 | LS1_WB03:BPM_D1413 | 641 | LS1_WB03:BPM_D1413 | 1413
BPM | LS1_WB03:BPM_D1413 | BPM | D1413 | 34E000370 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 5 | 3 |  | 4 |  |  |  | 5 | 3 | LS1_WB03:BPM_D1413 | 641 | LS1_WB03:BPM_D1413 | 1413
BPM | LS1_WB03:BPM_D1413 | BPM | D1413 | 34E000289 | 2C_1/4in_Hardline_036 | Ref | [8 empty] | Ref | SMA to SMA | 5 | 3 |  | Ref |  |  |  | 5 | 3 | LS1_WB03:BPM_D1413 | 641 | LS1_WB03:BPM_D1413 | 1413
BPM | LS1_CB04:BPM_D1423 | BPM | D1423 | 34E000356 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 5 | 3 |  | 5 |  |  |  | 5 | 3 | LS1_CB04:BPM_D1423 | 653 | LS1_CB04:BPM_D1423 | 1423
BPM | LS1_CB04:BPM_D1423 | BPM | D1423 | 34E000357 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 5 | 3 |  | 6 |  |  |  | 5 | 3 | LS1_CB04:BPM_D1423 | 653 | LS1_CB04:BPM_D1423 | 1423
BPM | LS1_CB04:BPM_D1423 | BPM | D1423 | 34E000358 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 5 | 3 |  | 7 |  |  |  | 5 | 3 | LS1_CB04:BPM_D1423 | 653 | LS1_CB04:BPM_D1423 | 1423
BPM | LS1_CB04:BPM_D1423 | BPM | D1423 | 34E000359 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 5 | 3 |  | 8 |  |  |  | 5 | 3 | LS1_CB04:BPM_D1423 | 653 | LS1_CB04:BPM_D1423 | 1423
BPM | LS1_CB04:BPM_D1423 | BPM | D1423 | 34E000288 | 2C_1/4in_Hardline_036 | Ref | [19 empty] | LS1_CB04:BPM_D1423 | 653 | LS1_CB04:BPM_D1423 | 1423
BPM | LS1_CB04:BPM_D1442 | BPM | D1442 | 34E000360 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 5 | 4 |  | 1 |  |  |  | 5 | 4 | LS1_CB04:BPM_D1442 | 671 | LS1_CB04:BPM_D1442 | 1442
BPM | LS1_CB04:BPM_D1442 | BPM | D1442 | 34E000361 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 5 | 4 |  | 2 |  |  |  | 5 | 4 | LS1_CB04:BPM_D1442 | 671 | LS1_CB04:BPM_D1442 | 1442
BPM | LS1_CB04:BPM_D1442 | BPM | D1442 | 34E000362 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 5 | 4 |  | 3 |  |  |  | 5 | 4 | LS1_CB04:BPM_D1442 | 671 | LS1_CB04:BPM_D1442 | 1442
BPM | LS1_CB04:BPM_D1442 | BPM | D1442 | 34E000363 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 5 | 4 |  | 4 |  |  |  | 5 | 4 | LS1_CB04:BPM_D1442 | 671 | LS1_CB04:BPM_D1442 | 1442
BPM | LS1_CB04:BPM_D1442 | BPM | D1442 | 34E000285 | 2C_1/4in_Hardline_036 | Ref | [8 empty] | Ref | SMA to SMA | 5 | 4 |  | Ref |  |  |  | 5 | 4 | LS1_CB04:BPM_D1442 | 671 | LS1_CB04:BPM_D1442 | 1442
BPM | LS1_CB04:BPM_D1462 | BPM | D1462 | 34E000347 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 5 | 4 |  | 5 |  |  |  | 5 | 4 | LS1_CB04:BPM_D1462 | 689 | LS1_CB04:BPM_D1462 | 1462
BPM | LS1_CB04:BPM_D1462 | BPM | D1462 | 34E000364 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 5 | 4 |  | 6 |  |  |  | 5 | 4 | LS1_CB04:BPM_D1462 | 689 | LS1_CB04:BPM_D1462 | 1462
BPM | LS1_CB04:BPM_D1462 | BPM | D1462 | 34E000365 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 5 | 4 |  | 7 |  |  |  | 5 | 4 | LS1_CB04:BPM_D1462 | 689 | LS1_CB04:BPM_D1462 | 1462
BPM | LS1_CB04:BPM_D1462 | BPM | D1462 | 34E000366 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 5 | 4 |  | 8 |  |  |  | 5 | 4 | LS1_CB04:BPM_D1462 | 689 | LS1_CB04:BPM_D1462 | 1462
BPM | LS1_CB04:BPM_D1462 | BPM | D1462 | 34E000290 | 2C_1/4in_Hardline_036 | Ref | [19 empty] | LS1_CB04:BPM_D1462 | 689 | LS1_CB04:BPM_D1462 | 1462
BPM | LS1_WB04:BPM_D1477 | BPM | D1477 | 34E000216 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 5 | 5 |  | 1 |  |  |  | 5 | 5 | LS1_WB04:BPM_D1477 | 701 | LS1_WB04:BPM_D1477 | 1477
BPM | LS1_WB04:BPM_D1477 | BPM | D1477 | 34E000217 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 5 | 5 |  | 2 |  |  |  | 5 | 5 | LS1_WB04:BPM_D1477 | 701 | LS1_WB04:BPM_D1477 | 1477
BPM | LS1_WB04:BPM_D1477 | BPM | D1477 | 34E000218 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 5 | 5 |  | 3 |  |  |  | 5 | 5 | LS1_WB04:BPM_D1477 | 701 | LS1_WB04:BPM_D1477 | 1477
BPM | LS1_WB04:BPM_D1477 | BPM | D1477 | 34E000219 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel Panel | 4B | [6 empty] | 4B | SMA to SMA | 5 | 5 |  | 4 |  |  |  | 5 | 5 | LS1_WB04:BPM_D1477 | 701 | LS1_WB04:BPM_D1477 | 1477
BPM | LS1_WB04:BPM_D1477 | BPM | D1477 | 34E000119 | 2C_1/4in_Hardline_036 | Ref | [8 empty] | Ref | SMA to SMA | 5 | 5 |  | Ref |  |  |  | 5 | 5 | LS1_WB04:BPM_D1477 | 701 | LS1_WB04:BPM_D1477 | 1477
BPM | LS1_CB05:BPM_D1486 | BPM | D1486 | 34E000212 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel Panel | 1T | [6 empty] | 1T | SMA to SMA | 5 | 5 |  | 5 |  |  |  | 5 | 5 | LS1_CB05:BPM_D1486 | 713 | LS1_CB05:BPM_D1486 | 1486
BPM | LS1_CB05:BPM_D1486 | BPM | D1486 | 34E000213 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel Panel | 2L | [6 empty] | 2L | SMA to SMA | 5 | 5 |  | 6 |  |  |  | 5 | 5 | LS1_CB05:BPM_D1486 | 713 | LS1_CB05:BPM_D1486 | 1486
BPM | LS1_CB05:BPM_D1486 | BPM | D1486 | 34E000214 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 5 | 5 |  | 7 |  |  |  | 5 | 5 | LS1_CB05:BPM_D1486 | 713 | LS1_CB05:BPM_D1486 | 1486
BPM | LS1_CB05:BPM_D1486 | BPM | D1486 | 34E000215 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 5 | 5 |  | 8 |  |  |  | 5 | 5 | LS1_CB05:BPM_D1486 | 713 | LS1_CB05:BPM_D1486 | 1486
BPM | LS1_CB05:BPM_D1486 | BPM | D1486 | 34E000118 | 2C_1/4in_Hardline_036 | Ref | [19 empty] | LS1_CB05:BPM_D1486 | 713 | LS1_CB05:BPM_D1486 | 1486
BPM | LS1_CB05:BPM_D1506 | BPM | D1506 | 34E000208 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 5 | 6 |  | 1 |  |  |  | 5 | 6 | LS1_CB05:BPM_D1506 | 731 | LS1_CB05:BPM_D1506 | 1506
BPM | LS1_CB05:BPM_D1506 | BPM | D1506 | 34E000209 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 5 | 6 |  | 2 |  |  |  | 5 | 6 | LS1_CB05:BPM_D1506 | 731 | LS1_CB05:BPM_D1506 | 1506
BPM | LS1_CB05:BPM_D1506 | BPM | D1506 | 34E000210 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 5 | 6 |  | 3 |  |  |  | 5 | 6 | LS1_CB05:BPM_D1506 | 731 | LS1_CB05:BPM_D1506 | 1506
BPM | LS1_CB05:BPM_D1506 | BPM | D1506 | 34E000211 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 5 | 6 |  | 4 |  |  |  | 5 | 6 | LS1_CB05:BPM_D1506 | 731 | LS1_CB05:BPM_D1506 | 1506
BPM | LS1_CB05:BPM_D1506 | BPM | D1506 | 34E000117 | 2C_1/4in_Hardline_036 | Ref | [8 empty] | Ref | SMA to SMA | 5 | 6 |  | Ref |  |  |  | 5 | 6 | LS1_CB05:BPM_D1506 | 731 | LS1_CB05:BPM_D1506 | 1506
BPM | LS1_CB05:BPM_D1526 | BPM | D1526 | 34E000204 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 5 | 6 |  | 5 |  |  |  | 5 | 6 | LS1_CB05:BPM_D1526 | 749 | LS1_CB05:BPM_D1526 | 1526
BPM | LS1_CB05:BPM_D1526 | BPM | D1526 | 34E000205 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 5 | 6 |  | 6 |  |  |  | 5 | 6 | LS1_CB05:BPM_D1526 | 749 | LS1_CB05:BPM_D1526 | 1526
BPM | LS1_CB05:BPM_D1526 | BPM | D1526 | 34E000206 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 5 | 6 |  | 7 |  |  |  | 5 | 6 | LS1_CB05:BPM_D1526 | 749 | LS1_CB05:BPM_D1526 | 1526
BPM | LS1_CB05:BPM_D1526 | BPM | D1526 | 34E000207 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 5 | 6 |  | 8 |  |  |  | 5 | 6 | LS1_CB05:BPM_D1526 | 749 | LS1_CB05:BPM_D1526 | 1526
BPM | LS1_CB05:BPM_D1526 | BPM | D1526 | 34E000116 | 2C_1/4in_Hardline_036 | Ref | [19 empty] | LS1_CB05:BPM_D1526 | 749 | LS1_CB05:BPM_D1526 | 1526
BPM | LS1_WB05:BPM_D1541 | BPM | D1541 | 34E000200 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 5 | 7 |  | 1 |  |  |  | 5 | 7 | LS1_WB05:BPM_D1541 | 761 | LS1_WB05:BPM_D1541 | 1541
BPM | LS1_WB05:BPM_D1541 | BPM | D1541 | 34E000201 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 5 | 7 |  | 2 |  |  |  | 5 | 7 | LS1_WB05:BPM_D1541 | 761 | LS1_WB05:BPM_D1541 | 1541
BPM | LS1_WB05:BPM_D1541 | BPM | D1541 | 34E000202 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 5 | 7 |  | 3 |  |  |  | 5 | 7 | LS1_WB05:BPM_D1541 | 761 | LS1_WB05:BPM_D1541 | 1541
BPM | LS1_WB05:BPM_D1541 | BPM | D1541 | 34E000203 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 5 | 7 |  | 4 |  |  |  | 5 | 7 | LS1_WB05:BPM_D1541 | 761 | LS1_WB05:BPM_D1541 | 1541
BPM | LS1_WB05:BPM_D1541 | BPM | D1541 | 34E000112 | 2C_1/4in_Hardline_036 | Ref | [8 empty] | Ref | SMA to SMA | 5 | 7 |  | Ref |  |  |  | 5 | 7 | LS1_WB05:BPM_D1541 | 761 | LS1_WB05:BPM_D1541 | 1541

## LS1_N1501
LS1_N1501
Long Run Cable | [6 empty] | Inter-Rack Connections | [9 empty] | Terminating Device
Diagnostics Device | Lattice Beamline Location | Device Type | CableDB Location | Cable Number | Cable Type | Function | Termination Device | Port Number In | Pin Number | CPC pinout | CPC pinout | CPC pinout | Signal Pin | Shield Pin | Port Number Out | Patch Cable | MTCA Chassis | MTCA Slot | PICO Port | RTM Port | Bias Supply Chassis | Bias Supply Slot | Bias Supply Channel |  | search for row | device name | D number
Halo Monitor Ring | LS1_WB04:HMR_D1478 | HMR | D1478 | 34H000018 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J1 | [6 empty] | J1 | SHV to Lemo | 6 | 3 | CH0 - fmc2-1 |  | 4 | 1: EBS8003 | Ch0 |  | 703 | LS1_WB04:HMR_D1478 | 1478
Halo Monitor Ring | LS1_WB05:HMR_D1542 | HMR | D1542 | 34H000020 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J2 | [6 empty] | J2 | SHV to Lemo | 6 | 3 | CH1 - fmc2-2 |  | 4 | 1: EBS8003 | Ch0 |  | 763 | LS1_WB05:HMR_D1542 | 1542
Halo Monitor Ring | LS1_WB06:HMR_D1606 | HMR | D1606 | 34H000021 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J3 | [6 empty] | J3 | SHV to Lemo | 6 | 3 | CH2 - fmc2-3 |  | 4 | 1: EBS8003 | Ch0 |  | 823 | LS1_WB06:HMR_D1606 | 1606
Halo Monitor Ring | LS1_WB07:HMR_D1670 | HMR | D1670 | 34H000023 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J4 | [6 empty] | J4 | SHV to Lemo | 6 | 3 | CH4 - fmc1-1 |  | 4 | 1: EBS8003 | Ch1 |  | 883 | LS1_WB07:HMR_D1670 | 1670
Halo Monitor Ring | LS1_WB08:HMR_D1733 | HMR | D1733 | 34H000022 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J5 | [6 empty] | J5 | SHV to Lemo | 6 | 3 | CH5 - fmc1-2 |  | 4 | 1: EBS8003 | Ch1 |  | 943 | LS1_WB08:HMR_D1733 | 1733
Halo Monitor Ring | LS1_WB09:HMR_D1797 | HMR | D1797 | 34H000019 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J6 | [6 empty] | J6 | SHV to Lemo | 6 | 3 | CH6 - fmc1-3 |  | 4 | 1: EBS8003 | Ch1 |  | 1003 | LS1_WB09:HMR_D1797 | 1797
| LS1_WB04:HMR_D1478 |  | D1478 | 34H000105 | 2C_5kVDC_Coax_035 |  | SHV to SHV Panel | [17 empty] | 703 | LS1_WB04:HMR_D1478 | 1478
| LS1_WB05:HMR_D1542 |  | D1542 | 34H000103 | 2C_5kVDC_Coax_035 |  | SHV to SHV Panel | [17 empty] | 763 | LS1_WB05:HMR_D1542 | 1542
| LS1_WB06:HMR_D1606 |  | D1606 | 34H000102 | 2C_5kVDC_Coax_035 |  | SHV to SHV Panel | [17 empty] | 823 | LS1_WB06:HMR_D1606 | 1606
| LS1_WB07:HMR_D1670 |  | D1670 | 34H000115 | 2C_5kVDC_Coax_035 |  | SHV to SHV Panel | [17 empty] | 883 | LS1_WB07:HMR_D1670 | 1670
| LS1_WB08:HMR_D1733 |  | D1733 | 34H000116 | 2C_5kVDC_Coax_035 |  | SHV to SHV Panel | [17 empty] | 943 | LS1_WB08:HMR_D1733 | 1733
| LS1_WB09:HMR_D1797 |  | D1797 | 34H000104 | 2C_5kVDC_Coax_035 |  | SHV to SHV Panel | [17 empty] | 1003 | LS1_WB09:HMR_D1797 | 1797
Panel 1
BPM | LS1_CB06:BPM_D1550 | BPM | D1550 | 34E000196 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 6 | 7 |  | 1 |  |  |  |  | 773 | LS1_CB06:BPM_D1550 | 1550
BPM | LS1_CB06:BPM_D1550 | BPM | D1550 | 34E000197 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 6 | 7 |  | 2 |  |  |  |  | 773 | LS1_CB06:BPM_D1550 | 1550
BPM | LS1_CB06:BPM_D1550 | BPM | D1550 | 34E000198 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 6 | 7 |  | 3 |  |  |  |  | 773 | LS1_CB06:BPM_D1550 | 1550
BPM | LS1_CB06:BPM_D1550 | BPM | D1550 | 34E000199 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 6 | 7 |  | 4 |  |  |  |  | 773 | LS1_CB06:BPM_D1550 | 1550
BPM | LS1_CB06:BPM_D1550 | BPM | D1550 | 34E000111 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No terminiation | [16 empty] | 773 | LS1_CB06:BPM_D1550 | 1550
BPM | LS1_CB06:BPM_D1570 | BPM | D1570 | 34E000192 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 6 | 7 |  | 5 |  |  |  |  | 791 | LS1_CB06:BPM_D1570 | 1570
BPM | LS1_CB06:BPM_D1570 | BPM | D1570 | 34E000193 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 6 | 7 |  | 6 |  |  |  |  | 791 | LS1_CB06:BPM_D1570 | 1570
BPM | LS1_CB06:BPM_D1570 | BPM | D1570 | 34E000194 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 6 | 7 |  | 7 |  |  |  |  | 791 | LS1_CB06:BPM_D1570 | 1570
BPM | LS1_CB06:BPM_D1570 | BPM | D1570 | 34E000195 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 6 | 7 |  | 8 |  |  |  |  | 791 | LS1_CB06:BPM_D1570 | 1570
BPM | LS1_CB06:BPM_D1570 | BPM | D1570 | 34E000113 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 6 | 7 |  | Ref |  |  |  |  | 791 | LS1_CB06:BPM_D1570 | 1570
BPM | LS1_CB06:BPM_D1590 | BPM | D1590 | 34E000188 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 6 | 8 |  | 1 |  |  |  |  | 809 | LS1_CB06:BPM_D1590 | 1590
BPM | LS1_CB06:BPM_D1590 | BPM | D1590 | 34E000189 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 6 | 8 |  | 2 |  |  |  |  | 809 | LS1_CB06:BPM_D1590 | 1590
BPM | LS1_CB06:BPM_D1590 | BPM | D1590 | 34E000190 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 6 | 8 |  | 3 |  |  |  |  | 809 | LS1_CB06:BPM_D1590 | 1590
BPM | LS1_CB06:BPM_D1590 | BPM | D1590 | 34E000191 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 6 | 8 |  | 4 |  |  |  |  | 809 | LS1_CB06:BPM_D1590 | 1590
BPM | LS1_CB06:BPM_D1590 | BPM | D1590 | 34E000115 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No terminiation | [16 empty] | 809 | LS1_CB06:BPM_D1590 | 1590
BPM | LS1_WB06:BPM_D1604 | BPM | D1605 | 34E000184 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 6 | 8 |  | 5 |  |  |  |  | 821 | LS1_WB06:BPM_D1604 | 1604
BPM | LS1_WB06:BPM_D1604 | BPM | D1605 | 34E000185 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 6 | 8 |  | 6 |  |  |  |  | 821 | LS1_WB06:BPM_D1604 | 1604
BPM | LS1_WB06:BPM_D1604 | BPM | D1605 | 34E000186 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 6 | 8 |  | 7 |  |  |  |  | 821 | LS1_WB06:BPM_D1604 | 1604
BPM | LS1_WB06:BPM_D1604 | BPM | D1605 | 34E000187 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 6 | 8 |  | 8 |  |  |  |  | 821 | LS1_WB06:BPM_D1604 | 1604
BPM | LS1_WB06:BPM_D1604 | BPM | D1605 | 34E000114 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 6 | 8 |  | Ref |  |  |  |  | 821 | LS1_WB06:BPM_D1604 | 1604
BPM | LS1_CB07:BPM_D1614 | BPM | D1614 | 34E000180 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 6 | 9 |  | 1 |  |  |  |  | 833 | LS1_CB07:BPM_D1614 | 1614
BPM | LS1_CB07:BPM_D1614 | BPM | D1614 | 34E000181 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 6 | 9 |  | 2 |  |  |  |  | 833 | LS1_CB07:BPM_D1614 | 1614
BPM | LS1_CB07:BPM_D1614 | BPM | D1614 | 34E000182 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 6 | 9 |  | 3 |  |  |  |  | 833 | LS1_CB07:BPM_D1614 | 1614
BPM | LS1_CB07:BPM_D1614 | BPM | D1614 | 34E000183 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 6 | 9 |  | 4 |  |  |  |  | 833 | LS1_CB07:BPM_D1614 | 1614
BPM | LS1_CB07:BPM_D1614 | BPM | D1614 | 34E000110 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No terminiation | [16 empty] | 833 | LS1_CB07:BPM_D1614 | 1614
BPM | LS1_CB07:BPM_D1634 | BPM | D1634 | 34E000176 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 6 | 9 |  | 5 |  |  |  |  | 851 | LS1_CB07:BPM_D1634 | 1634
BPM | LS1_CB07:BPM_D1634 | BPM | D1634 | 34E000177 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 6 | 9 |  | 6 |  |  |  |  | 851 | LS1_CB07:BPM_D1634 | 1634
BPM | LS1_CB07:BPM_D1634 | BPM | D1634 | 34E000178 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 6 | 9 |  | 7 |  |  |  |  | 851 | LS1_CB07:BPM_D1634 | 1634
BPM | LS1_CB07:BPM_D1634 | BPM | D1634 | 34E000179 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 6 | 9 |  | 8 |  |  |  |  | 851 | LS1_CB07:BPM_D1634 | 1634
BPM | LS1_CB07:BPM_D1634 | BPM | D1634 | 34E000129 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 6 | 9 |  | Ref |  |  |  |  | 851 | LS1_CB07:BPM_D1634 | 1634
BPM | LS1_CB07:BPM_D1654 | BPM | D1654 | 34E000172 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 2 | 10 |  | 1 |  |  |  |  | 869 | LS1_CB07:BPM_D1654 | 1654
BPM | LS1_CB07:BPM_D1654 | BPM | D1654 | 34E000173 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 6 | 10 |  | 2 |  |  |  |  | 869 | LS1_CB07:BPM_D1654 | 1654
BPM | LS1_CB07:BPM_D1654 | BPM | D1654 | 34E000174 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 6 | 10 |  | 3 |  |  |  |  | 869 | LS1_CB07:BPM_D1654 | 1654
BPM | LS1_CB07:BPM_D1654 | BPM | D1654 | 34E000175 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 6 | 10 |  | 4 |  |  |  |  | 869 | LS1_CB07:BPM_D1654 | 1654
BPM | LS1_CB07:BPM_D1654 | BPM | D1654 | 34E000128 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 6 | 10 |  | Ref |  |  |  |  | 869 | LS1_CB07:BPM_D1654 | 1654
BCM Siltem Cables
BCM | FE_LEBT:BCM_D0989 | BCM | D0998 | 24E000044 | 2C_22AWG_Multi_063 | Signal | Amplifier | FE_LEBT:AIM_D0998 | [15 empty] | Amplifier in FE_N0404 | 235 | FE_LEBT:BCM_D0989 | 989
BCM | FE_MEBT:BCM_D1055 | BCM | D1055 | 24E000042 | 2C_22AWG_Multi_063 | Signal | Amplifier | FE_MEBT:AIM_D1055 | [15 empty] | Amplifier in FE_N0404 | 273 | FE_MEBT:BCM_D1055 | 1055
BCM | FE_MEBT:BCM_D1120 | BCM | D1121 | 24E000043 | 2C_22AWG_Multi_063 | Signal | Amplifier | FE_MEBT:AIM_D1121 | [15 empty] | Amplifier in FE_N0404 | 339 | FE_MEBT:BCM_D1120 | 1120
BCM | FS1_CSS:BCM_D2183 | BCM | D2183 | 44E000222 | 2C_22AWG_Multi_063 | Signal | Amplifier | FS1_CSS:AIM_D2183 | [15 empty] | Amplifier in BDS_N0305 | 1274 | FS1_CSS:BCM_D2183 | 2183
BCM | FS1_CSS:BCM_D2264 | BCM | D2264 | 44E000221 | 2C_22AWG_Multi_063 | Signal | Amplifier | FS1_CSS:AIM_D2264 | [15 empty] | Amplifier in BDS_N0305 | 1348 | FS1_CSS:BCM_D2264 | 2264
BCM | FS1_CSS:BCM_D2353 | BCM | D2353 | 44E000223 | 2C_22AWG_Multi_063 | Signal | Amplifier | FS1_CSS:AIM_D2359 | [15 empty] | Amplifier in BDS_N0305 | 1426 | FS1_CSS:BCM_D2353 | 2353
BCM | FS1_BTS:BCM_D2449 | BCM | D2464 | 44E000219 | 2C_22AWG_Multi_063 | Signal | Amplifier | BDS_FFS:AIM_D5798 | [15 empty] | Amplifier in BDS_N0305 | 3549 | FS1_BTS:BCM_D2449 | 2449
BCM | FS1_BMS:BCM_D2519 | BCM | D2519 | 44E000224 | 2C_22AWG_Multi_063 | Signal | Amplifier | FS1_BMS:AIM_D2519 | [15 empty] | Amplifier in FS1_N0506 | 1528 | FS1_BMS:BCM_D2519 | 2519
BCM | FS1_BMS:BCM_D2675 | BCM | D2675 | 44E000225 | 2C_22AWG_Multi_063 | Signal | Amplifier | FS1_BMS:AIM_D2675 | [15 empty] | Amplifier in FS1_N0506 | 1608 | FS1_BMS:BCM_D2675 | 2675
BCM | FS2_BTS:BCM_D3936 | BCM | D3952 | 44E000227 | 2C_22AWG_Multi_063 | Signal | Amplifier | FS2_BTS:AIM_D3952 | [15 empty] | Amplifier in LS2_N4201 | 2600 | FS2_BTS:BCM_D3936 | 3936
BCM | FS2_BBS:BCM_D4169 | BCM | D4175 | 44E000226 | 2C_22AWG_Multi_063 | Signal | Amplifier | FS2_BMS:AIM_D4175 | [15 empty] | Amplifier in FE_N0404 | 2702 | FS2_BBS:BCM_D4169 | 4169
BCM | BDS_BTS:BCM_D5521 | BCM | D5550 | 44E000220 | 2C_22AWG_Multi_063 | Signal | Amplifier | BDS_BTS:AIM_D5550 | [15 empty] | Amplifier in BDS_N0305 | 3271 | BDS_BTS:BCM_D5521 | 5521
BCM | BDS_FFS:BCM_D5789 | BCM | D5798 | [20 empty] | CABLE MOVED TO: FS1_BTS:BCM_D2449 | 3364 | BDS_FFS:BCM_D5789 | 5789
BCM Amplifier Cable
BCM | FE_LEBT:BCM_D0989 | BCM | D0998 | 24E000041 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | J1 | [7 empty] | SMA to SMB | 6 | 4 |  | 1 |  |  |  | FE_LEBT:AIM_D0998 in FE_N0404
BCM | FE_MEBT:BCM_D1055 | BCM | D1055 | 24E000039 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | J2 | [7 empty] | SMA to SMB | 6 | 4 |  | 2 |  |  |  | FE_MEBT:AIM_D1055 in FE_N0404
BCM | FE_MEBT:BCM_D1120 | BCM | D1121 | 24E000040 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | J3 | [7 empty] | SMA to SMB | 6 | 4 |  | 3 |  |  |  | FE_MEBT:AIM_D1121 in FE_N0404
BCM | FS1_CSS:BCM_D2183 | BCM | D2228 | 44E000214 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | J5 | [7 empty] | SMA to SMB | 6 | 4 |  | 4 |  |  |  | FS1_CSS:AIM_D2228 in BDS_N0305
BCM | FS1_CSS:BCM_D2264 | BCM | D2265 | 44E000213 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | J4 | [7 empty] | SMA to SMB | 6 | 5 |  | 1 |  |  |  | FS1_CSS:AIM_D2265 in BDS_N0305
BCM | FS1_BTS:BCM_D2449 | BCM | D2464 | 44E000210 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | J12 | [7 empty] | SMA to SMB | 6 | 5 |  | 2 |  |  |  | BDS_FFS:AIM_D5798 in BDS_N0305
BCM | FS1_CSS:BCM_D2353 | BCM | D2359 | 44E000212 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | J6 | [7 empty] | SMA to SMB | 6 | 5 |  | 3 |  |  |  | FS1_CSS:AIM_D2359 in BDS_N0305
BCM | FS1_BMS:BCM_D2519 | BCM | D2525 | 44E000215 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | J7 | [7 empty] | SMA to SMB | 6 | 5 |  | 4 |  |  |  | FS1_BMS:AIM_D2525 in FS1_N0506
BCM | FS1_BMS:BCM_D2675 | BCM | D2686 | 44E000216 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | J8 | [7 empty] | SMA to SMB | 6 | 6 |  | 1 |  |  |  | FS1_BMS:AIM_D2686 in FS1_N0506
BCM | FS2_BTS:BCM_D3936 | BCM | D3952 | 44E000209 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | J9 | [7 empty] | SMA to SMB | 6 | 6 |  | 2 |  |  |  | FS2_BTS:AIM_D3952 in LS2_N4201
BCM | FS2_BBS:BCM_D4169 | BCM | D4175 | 44E000217 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | J10 | [7 empty] | SMA to SMB | 6 | 6 |  | 3 |  |  |  | FS2_BMS:AIM_D4175 in FE_N0404
BCM | BDS_BTS:BCM_D5521 | BCM | D5550 | 44E000211 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | J11 | [7 empty] | SMA to SMB | 6 | 6 |  | 4 |  |  |  | BDS_BTS:AIM_D5550 in BDS_N0305
BCM | BDS_FFS:BCM_D5789 | BCM | D5798 | [20 empty] | CABLE MOVED TO: FS1_BTS:BCM_D2449
Lemo to Lemo Panel
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 1_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 4 | 1: EBS8003 | Ch0
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 1_OUT Port | [7 empty] | Lemo650/Lemo650 | 6 | 3 | CH3 - fmc2-4
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 2_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 4 | 1: EBS8003 | Ch1
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 2_OUT Port | [7 empty] | Lemo650/Lemo650 | 6 | 3 | CH7 - fmc1-4
BCM

## LS1_N1902
LS1_N1902
Long Run Cable | [6 empty] | Inter-Rack Connections | [9 empty] | Terminating Device
Diagnostics Device | Lattice Beamline Location | Device Type | CableDB Location | Cable Number | Cable Type | Function | Termination Device | Port Number In | Pin Number | CPC pinout | CPC pinout | CPC pinout | Signal Pin | Shield Pin | Port Number Out | Patch Cable | MTCA Chassis | MTCA Slot | PICO Port | FGPDB Port | Bias Supply Chassis | Bias Supply Slot | Bias Supply Channel |  |  |  | search for row | device name | D number
BPMs | [23 empty] | PICO | CH[0-7] | Device
BPM | LS1_WB07:BPM_D1668 | BPM | D1668 | 34E000168 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 7 | 3 |  | 1 |  |  |  | 7 | 3 | LS1_WB07:BPM_D1668 | 881 | LS1_WB07:BPM_D1668 | 1668
BPM | LS1_WB07:BPM_D1668 | BPM | D1668 | 34E000169 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 7 | 3 |  | 2 |  |  |  | 7 | 3 | LS1_WB07:BPM_D1668 | 881 | LS1_WB07:BPM_D1668 | 1668
BPM | LS1_WB07:BPM_D1668 | BPM | D1668 | 34E000170 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 7 | 3 |  | 3 |  |  |  | 7 | 3 | LS1_WB07:BPM_D1668 | 881 | LS1_WB07:BPM_D1668 | 1668
BPM | LS1_WB07:BPM_D1668 | BPM | D1668 | 34E000171 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 7 | 3 |  | 4 |  |  |  | 7 | 3 | LS1_WB07:BPM_D1668 | 881 | LS1_WB07:BPM_D1668 | 1668
BPM | LS1_WB07:RFS_D1668 | BPM | D1668 | 34E000127 | 2C_1/4in_Hardline_036 | Ref | No terminiation | Ref | [6 empty] | Ref | SMA to SMA | 7 | 3 |  | Ref |  |  |  | 7 | 3 | LS1_WB07:RFS_D1668
BPM | LS1_CB08:BPM_D1677 | BPM | D1677 | 34E000164 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 7 | 3 |  | 5 |  |  |  | 7 | 3 | LS1_CB08:BPM_D1677 | 893 | LS1_CB08:BPM_D1677 | 1677
BPM | LS1_CB08:BPM_D1677 | BPM | D1677 | 34E000165 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 7 | 3 |  | 6 |  |  |  | 7 | 3 | LS1_CB08:BPM_D1677 | 893 | LS1_CB08:BPM_D1677 | 1677
BPM | LS1_CB08:BPM_D1677 | BPM | D1677 | 34E000166 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 7 | 3 |  | 7 |  |  |  | 7 | 3 | LS1_CB08:BPM_D1677 | 893 | LS1_CB08:BPM_D1677 | 1677
BPM | LS1_CB08:BPM_D1677 | BPM | D1677 | 34E000167 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 7 | 3 |  | 8 |  |  |  | 7 | 3 | LS1_CB08:BPM_D1677 | 893 | LS1_CB08:BPM_D1677 | 1677
BPM | LS1_CB08:RFS_D1677 | BPM | D1677 | 34E000126 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | No term | [17 empty] | LS1_CB08:RFS_D1677
BPM | LS1_CB08:BPM_D1697 | BPM | D1697 | 34E000160 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 7 | 4 |  | 1 |  |  |  | 7 | 4 | LS1_CB08:BPM_D1697 | 911 | LS1_CB08:BPM_D1697 | 1697
BPM | LS1_CB08:BPM_D1697 | BPM | D1697 | 34E000161 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 7 | 4 |  | 2 |  |  |  | 7 | 4 | LS1_CB08:BPM_D1697 | 911 | LS1_CB08:BPM_D1697 | 1697
BPM | LS1_CB08:BPM_D1697 | BPM | D1697 | 34E000162 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 7 | 4 |  | 3 |  |  |  | 7 | 4 | LS1_CB08:BPM_D1697 | 911 | LS1_CB08:BPM_D1697 | 1697
BPM | LS1_CB08:BPM_D1697 | BPM | D1697 | 34E000163 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 7 | 4 |  | 4 |  |  |  | 7 | 4 | LS1_CB08:BPM_D1697 | 911 | LS1_CB08:BPM_D1697 | 1697
BPM | LS1_CB08:RFS_D1697 | BPM | D1697 | 34E000124 | 2C_1/4in_Hardline_036 | Ref | No terminiation | Ref | [6 empty] | Ref | SMA to SMA | 7 | 4 |  | Ref |  |  |  | 7 | 4 | LS1_CB08:RFS_D1697
BPM | LS1_CB08:BPM_D1717 | BPM | D1717 | 34E000156 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 7 | 4 |  | 5 |  |  |  | 7 | 4 | LS1_CB08:BPM_D1717 | 929 | LS1_CB08:BPM_D1717 | 1717
BPM | LS1_CB08:BPM_D1717 | BPM | D1717 | 34E000157 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 7 | 4 |  | 6 |  |  |  | 7 | 4 | LS1_CB08:BPM_D1717 | 929 | LS1_CB08:BPM_D1717 | 1717
BPM | LS1_CB08:BPM_D1717 | BPM | D1717 | 34E000158 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 7 | 4 |  | 7 |  |  |  | 7 | 4 | LS1_CB08:BPM_D1717 | 929 | LS1_CB08:BPM_D1717 | 1717
BPM | LS1_CB08:BPM_D1717 | BPM | D1717 | 34E000159 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 7 | 4 |  | 8 |  |  |  | 7 | 4 | LS1_CB08:BPM_D1717 | 929 | LS1_CB08:BPM_D1717 | 1717
BPM | LS1_CB08:RFS_D1717 | BPM | D1717 | 34E000121 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | No term | [17 empty] | LS1_CB08:RFS_D1717
BPM | LS1_WB08:BPM_D1732 | BPM | D1732 | 34E000140 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 7 | 5 |  | 1 |  |  |  | 7 | 5 | LS1_WB08:BPM_D1732 | 941 | LS1_WB08:BPM_D1732 | 1732
BPM | LS1_WB08:BPM_D1732 | BPM | D1732 | 34E000141 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 7 | 5 |  | 2 |  |  |  | 7 | 5 | LS1_WB08:BPM_D1732 | 941 | LS1_WB08:BPM_D1732 | 1732
BPM | LS1_WB08:BPM_D1732 | BPM | D1732 | 34E000142 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 7 | 5 |  | 3 |  |  |  | 7 | 5 | LS1_WB08:BPM_D1732 | 941 | LS1_WB08:BPM_D1732 | 1732
BPM | LS1_WB08:BPM_D1732 | BPM | D1732 | 34E000143 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 7 | 5 |  | 4 |  |  |  | 7 | 5 | LS1_WB08:BPM_D1732 | 941 | LS1_WB08:BPM_D1732 | 1732
BPM | LS1_WB08:RFS_D1732 | BPM | D1732 | 34E000122 | 2C_1/4in_Hardline_036 | Ref | No terminiation | Ref | [6 empty] | Ref | SMA to SMA | 7 | 5 |  | Ref |  |  |  | 7 | 5 | LS1_WB08:RFS_D1732
BPM | LS1_CB09:BPM_D1741 | BPM | D1741 | 34E000136 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 7 | 5 |  | 5 |  |  |  | 7 | 5 | LS1_CB09:BPM_D1741 | 953 | LS1_CB09:BPM_D1741 | 1741
BPM | LS1_CB09:BPM_D1741 | BPM | D1741 | 34E000137 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 7 | 5 |  | 6 |  |  |  | 7 | 5 | LS1_CB09:BPM_D1741 | 953 | LS1_CB09:BPM_D1741 | 1741
BPM | LS1_CB09:BPM_D1741 | BPM | D1741 | 34E000138 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 7 | 5 |  | 7 |  |  |  | 7 | 5 | LS1_CB09:BPM_D1741 | 953 | LS1_CB09:BPM_D1741 | 1741
BPM | LS1_CB09:BPM_D1741 | BPM | D1741 | 34E000139 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 7 | 5 |  | 8 |  |  |  | 7 | 5 | LS1_CB09:BPM_D1741 | 953 | LS1_CB09:BPM_D1741 | 1741
BPM | LS1_CB09:RFS_D1741 | BPM | D1741 | 34E000123 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | No term | [17 empty] | LS1_CB09:RFS_D1741
BPM | LS1_CB09:BPM_D1761 | BPM | D1761 | 34E000152 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 7 | 6 |  | 1 |  |  |  | 7 | 6 | LS1_CB09:BPM_D1761 | 971 | LS1_CB09:BPM_D1761 | 1761
BPM | LS1_CB09:BPM_D1761 | BPM | D1761 | 34E000153 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 7 | 6 |  | 2 |  |  |  | 7 | 6 | LS1_CB09:BPM_D1761 | 971 | LS1_CB09:BPM_D1761 | 1761
BPM | LS1_CB09:BPM_D1761 | BPM | D1761 | 34E000154 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 7 | 6 |  | 3 |  |  |  | 7 | 6 | LS1_CB09:BPM_D1761 | 971 | LS1_CB09:BPM_D1761 | 1761
BPM | LS1_CB09:BPM_D1761 | BPM | D1761 | 34E000155 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 7 | 6 |  | 4 |  |  |  | 7 | 6 | LS1_CB09:BPM_D1761 | 971 | LS1_CB09:BPM_D1761 | 1761
BPM | LS1_CB09:RFS_D1761 | BPM | D1761 | 34E000125 | 2C_1/4in_Hardline_036 | Ref | No terminiation | Ref | [6 empty] | Ref | SMA to SMA | 7 | 6 |  | Ref |  |  |  | 7 | 6 | LS1_CB09:RFS_D1761
BPM | LS1_CB09:BPM_D1781 | BPM | D1781 | 34E000148 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 7 | 6 |  | 5 |  |  |  | 7 | 6 | LS1_CB09:BPM_D1781 | 989 | LS1_CB09:BPM_D1781 | 1781
BPM | LS1_CB09:BPM_D1781 | BPM | D1781 | 34E000149 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 7 | 6 |  | 6 |  |  |  | 7 | 6 | LS1_CB09:BPM_D1781 | 989 | LS1_CB09:BPM_D1781 | 1781
BPM | LS1_CB09:BPM_D1781 | BPM | D1781 | 34E000150 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 7 | 6 |  | 7 |  |  |  | 7 | 6 | LS1_CB09:BPM_D1781 | 989 | LS1_CB09:BPM_D1781 | 1781
BPM | LS1_CB09:BPM_D1781 | BPM | D1781 | 34E000151 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 7 | 6 |  | 8 |  |  |  | 7 | 6 | LS1_CB09:BPM_D1781 | 989 | LS1_CB09:BPM_D1781 | 1781
BPM | LS1_CB09:RFS_D1781 | BPM | D1781 | 34E000120 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | No term | [17 empty] | LS1_CB09:RFS_D1781
BPM | LS1_WB09:BPM_D1796 | BPM | D1796 | 34E000132 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 7 | 7 |  | 1 |  |  |  | 7 | 7 | LS1_WB09:BPM_D1796 | 1001 | LS1_WB09:BPM_D1796 | 1796
BPM | LS1_WB09:BPM_D1796 | BPM | D1796 | 34E000133 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 7 | 7 |  | 2 |  |  |  | 7 | 7 | LS1_WB09:BPM_D1796 | 1001 | LS1_WB09:BPM_D1796 | 1796
BPM | LS1_WB09:BPM_D1796 | BPM | D1796 | 34E000134 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 7 | 7 |  | 3 |  |  |  | 7 | 7 | LS1_WB09:BPM_D1796 | 1001 | LS1_WB09:BPM_D1796 | 1796
BPM | LS1_WB09:BPM_D1796 | BPM | D1796 | 34E000135 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 7 | 7 |  | 4 |  |  |  | 7 | 7 | LS1_WB09:BPM_D1796 | 1001 | LS1_WB09:BPM_D1796 | 1796
BPM | LS1_WB09:RFS_D1796 | BPM | D1796 | 34E000131 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 7 | 7 |  | Ref |  |  |  | 7 | 7 | LS1_WB09:RFS_D1796
Neutron Detector Signal
Neutron Detector | LS1_CB05:ND_D1505 | ND | D1505 | 34E000526 | 2C_75LC_Coax_089 | Signal | N-Type to Lemo Box | 1 | [6 empty] | 1 | SMA to Lemo650 | 7 | 8 | CH0 - fmc2-1 |  |  |  |  | 7 | 8 | LS1_CB05:ND_D1505 | #N/A | #N/A | #N/A
Neutron Detector | LS1_CB07:ND_D1633 | ND | D1633 | 34E000528 | 2C_75LC_Coax_089 | Signal | N-Type to Lemo Box | 2 | [6 empty] | 2 | SMA to Lemo650 | 7 | 8 | CH1 - fmc2-2 |  |  |  |  | 7 | 8 | LS1_CB07:ND_D1633 | #N/A | #N/A | #N/A
Neutron Detector | LS1_CB09:ND_D1760 | ND | D1760 | 34E000530 | 2C_75LC_Coax_089 | Signal | N-Type to Lemo Box | 3 | [6 empty] | 3 | SMA to Lemo650 | 7 | 8 | CH2 - fmc2-3 |  |  |  |  | 7 | 8 | LS1_CB09:ND_D1760 | #N/A | #N/A | #N/A
Neutron Detector | LS1_CB11:ND_D1908 | ND | D1908 | 34E000532 | 2C_75LC_Coax_089 | Signal | N-Type to Lemo Box | 4 | [6 empty] | 4 | SMA to Lemo650 | 7 | 8 | CH3 - fmc2-4 |  |  |  |  | 7 | 8 | LS1_CB11:ND_D1908 | #N/A | #N/A | #N/A
Neutron Detector | LS1_BTS:ND_D2076 | ND | D2076 | 34E000534 | 2C_75LC_Coax_089 | Signal | N-Type to Lemo Box | 5 | [6 empty] | 5 | SMA to Lemo650 | 7 | 8 | CH4 - fmc1-1 |  |  |  |  | 7 | 8 | LS1_BTS:ND_D2076 | #N/A | #N/A | #N/A
Neutron Detector Test signal
Neutron Detector | LS1_CB05:ND_D1505 |  | D1505 | 34E000525 | 2C_75LC_Coax_089 | Test Signal | N-Type to SMA Panel | [8 empty] | SMA to Ferrule | [9 empty] | LS1_CB05:ND_D1505 | #N/A | #N/A | #N/A
Neutron Detector | LS1_CB07:ND_D1633 |  | D1633 | 34E000527 | 2C_75LC_Coax_089 | Test Signal | N-Type to SMA Panel | [8 empty] | SMA to Ferrule | [9 empty] | LS1_CB07:ND_D1633 | #N/A | #N/A | #N/A
Neutron Detector | LS1_CB09:ND_D1760 |  | D1760 | 34E000529 | 2C_75LC_Coax_089 | Test Signal | N-Type to SMA Panel | [8 empty] | SMA to Ferrule | [9 empty] | LS1_CB09:ND_D1760 | #N/A | #N/A | #N/A
Neutron Detector | LS1_CB11:ND_D1908 |  | D1908 | 34E000531 | 2C_75LC_Coax_089 | Test Signal | N-Type to SMA Panel | [8 empty] | SMA to Ferrule | [9 empty] | LS1_CB11:ND_D1908 | #N/A | #N/A | #N/A
Neutron Detector | LS1_BTS:ND_D2076 |  | D2076 | 34E000533 | 2C_75LC_Coax_089 | Test Signal | N-Type to SMA Panel | [8 empty] | SMA to Ferrule | [9 empty] | LS1_BTS:ND_D2076 | #N/A | #N/A | #N/A
Neutron Detector Bias
Neutron Detector | LS1_CB05:ND_D1505 |  | D1505 | 34H000059 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 1 | [6 empty] | 1 | SHV to SHV |  |  |  |  | 5 | 1:EHS8020n | Ch0 |  |  | LS1_CB05:ND_D1505
Neutron Detector | LS1_CB07:ND_D1633 |  | D1633 | 34H000060 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 2 | [6 empty] | 2 | SHV to SHV |  |  |  |  | 5 | 1:EHS8020n | Ch1 |  |  | LS1_CB07:ND_D1633
Neutron Detector | LS1_CB09:ND_D1760 |  | D1760 | 34H000061 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 3 | [6 empty] | 3 | SHV to SHV |  |  |  |  | 5 | 1:EHS8020n | Ch2 |  |  | LS1_CB09:ND_D1760
Neutron Detector | LS1_CB11:ND_D1908 |  | D1908 | 34H000062 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 4 | [6 empty] | 4 | SHV to SHV |  |  |  |  | 5 | 1:EHS8020n | Ch3 |  |  | LS1_CB11:ND_D1908
Neutron Detector | LS1_BTS:ND_D2076 |  | D2076 | 34H000063 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 5 | [6 empty] | 5 | SHV to SHV |  |  |  |  | 5 | 1:EHS8020n | Ch4 |  |  | LS1_BTS:ND_D2076

## LS1_N2302
LS1_N2302
Long Run Cable | [6 empty] | Inter-Rack Connections | [9 empty] | Terminating Device
Diagnostics Device | Lattice Beamline Location | Device Type | CableDB Location | Cable Number | Cable Type | Function | Termination Device | Port Number In | Pin Number | CPC pinout | CPC pinout | CPC pinout | Signal Pin | Shield Pin | Port Number Out | Patch Cable | MTCA Chassis | MTCA Slot | PICO Port | FGPDB Port | Bias Supply Chassis | Bias Supply Slot | Bias Supply Channel |  |  |  | search for row | device name | D number
BPMs | [23 empty] | PICO | CH[0-7] | Device
BPM | LS1_CB10:BPM_D1805 | BPM | D1805 | 34E000144 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 8 | 3 |  | 1 |  |  |  | 8 | 3 | LS1_CB10:BPM_D1805 | 1013 | LS1_CB10:BPM_D1805 | 1805
BPM | LS1_CB10:BPM_D1805 | BPM | D1805 | 34E000145 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 8 | 3 |  | 2 |  |  |  | 8 | 3 | LS1_CB10:BPM_D1805 | 1013 | LS1_CB10:BPM_D1805 | 1805
BPM | LS1_CB10:BPM_D1805 | BPM | D1805 | 34E000146 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 8 | 3 |  | 3 |  |  |  | 8 | 3 | LS1_CB10:BPM_D1805 | 1013 | LS1_CB10:BPM_D1805 | 1805
BPM | LS1_CB10:BPM_D1805 | BPM | D1805 | 34E000147 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 8 | 3 |  | 4 |  |  |  | 8 | 3 | LS1_CB10:BPM_D1805 | 1013 | LS1_CB10:BPM_D1805 | 1805
BPM | LS1_CB10:RFS_D1805 | BPM | D1805 | 34E000130 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | [17 empty] | LS1_CB10:RFS_D1805
BPM | LS1_CB10:BPM_D1825 | BPM | D1825 | 34E000233 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 8 | 3 |  | 5 |  |  |  | 8 | 3 | LS1_CB10:BPM_D1825 | 1031 | LS1_CB10:BPM_D1825 | 1825
BPM | LS1_CB10:BPM_D1825 | BPM | D1825 | 34E000250 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 8 | 3 |  | 6 |  |  |  | 8 | 3 | LS1_CB10:BPM_D1825 | 1031 | LS1_CB10:BPM_D1825 | 1825
BPM | LS1_CB10:BPM_D1825 | BPM | D1825 | 34E000251 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 8 | 3 |  | 7 |  |  |  | 8 | 3 | LS1_CB10:BPM_D1825 | 1031 | LS1_CB10:BPM_D1825 | 1825
BPM | LS1_CB10:BPM_D1825 | BPM | D1825 | 34E000252 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 8 | 3 |  | 8 |  |  |  | 8 | 3 | LS1_CB10:BPM_D1825 | 1031 | LS1_CB10:BPM_D1825 | 1825
BPM | LS1_CB10:RFS_D1825 | BPM | D1825 | 34E000224 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 8 | 3 |  | Ref |  |  |  | 8 | 3 | LS1_CB10:RFS_D1825
BPM | LS1_CB10:BPM_D1845 | BPM | D1845 | 34E000246 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 8 | 4 |  | 1 |  |  |  | 8 | 4 | LS1_CB10:BPM_D1845 | 1049 | LS1_CB10:BPM_D1845 | 1845
BPM | LS1_CB10:BPM_D1845 | BPM | D1845 | 34E000247 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 8 | 4 |  | 2 |  |  |  | 8 | 4 | LS1_CB10:BPM_D1845 | 1049 | LS1_CB10:BPM_D1845 | 1845
BPM | LS1_CB10:BPM_D1845 | BPM | D1845 | 34E000248 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 8 | 4 |  | 3 |  |  |  | 8 | 4 | LS1_CB10:BPM_D1845 | 1049 | LS1_CB10:BPM_D1845 | 1845
BPM | LS1_CB10:BPM_D1845 | BPM | D1845 | 34E000249 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 8 | 4 |  | 4 |  |  |  | 8 | 4 | LS1_CB10:BPM_D1845 | 1049 | LS1_CB10:BPM_D1845 | 1845
BPM | LS1_CB10:RFS_D1845 | BPM | D1845 | 34E000225 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | [17 empty] | LS1_CB10:RFS_D1845
BPM | LS1_WB10:BPM_D1859 | BPM | D1860 | 34E000253 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 8 | 4 |  | 5 |  |  |  | 8 | 4 | LS1_WB10:BPM_D1859 | 1061 | LS1_WB10:BPM_D1859 | 1859
BPM | LS1_WB10:BPM_D1859 | BPM | D1860 | 34E000254 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 8 | 4 |  | 6 |  |  |  | 8 | 4 | LS1_WB10:BPM_D1859 | 1061 | LS1_WB10:BPM_D1859 | 1859
BPM | LS1_WB10:BPM_D1859 | BPM | D1860 | 34E000255 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 8 | 4 |  | 7 |  |  |  | 8 | 4 | LS1_WB10:BPM_D1859 | 1061 | LS1_WB10:BPM_D1859 | 1859
BPM | LS1_WB10:BPM_D1859 | BPM | D1860 | 34E000256 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 8 | 4 |  | 8 |  |  |  | 8 | 4 | LS1_WB10:BPM_D1859 | 1061 | LS1_WB10:BPM_D1859 | 1859
BPM | LS1_WB10:RFS_D1859 | BPM | D1860 | 34E000220 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 8 | 4 |  | Ref |  |  |  | 8 | 4 | LS1_WB10:RFS_D1859
BPM | LS1_CB11:BPM_D1869 | BPM | D1869 | 34E000234 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 8 | 5 |  | 1 |  |  |  | 8 | 5 | LS1_CB11:BPM_D1869 | 1073 | LS1_CB11:BPM_D1869 | 1869
BPM | LS1_CB11:BPM_D1869 | BPM | D1869 | 34E000235 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 8 | 5 |  | 2 |  |  |  | 8 | 5 | LS1_CB11:BPM_D1869 | 1073 | LS1_CB11:BPM_D1869 | 1869
BPM | LS1_CB11:BPM_D1869 | BPM | D1869 | 34E000236 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 8 | 5 |  | 3 |  |  |  | 8 | 5 | LS1_CB11:BPM_D1869 | 1073 | LS1_CB11:BPM_D1869 | 1869
BPM | LS1_CB11:BPM_D1869 | BPM | D1869 | 34E000237 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 8 | 5 |  | 4 |  |  |  | 8 | 5 | LS1_CB11:BPM_D1869 | 1073 | LS1_CB11:BPM_D1869 | 1869
BPM | LS1_CB11:RFS_D1869 | BPM | D1869 | 34E000222 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | [17 empty] | LS1_CB11:RFS_D1869
BPM | LS1_CB11:BPM_D1889 | BPM | D1889 | 34E000242 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 8 | 5 |  | 5 |  |  |  | 8 | 5 | LS1_CB11:BPM_D1889 | 1091 | LS1_CB11:BPM_D1889 | 1889
BPM | LS1_CB11:BPM_D1889 | BPM | D1889 | 34E000243 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 8 | 5 |  | 6 |  |  |  | 8 | 5 | LS1_CB11:BPM_D1889 | 1091 | LS1_CB11:BPM_D1889 | 1889
BPM | LS1_CB11:BPM_D1889 | BPM | D1889 | 34E000244 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 8 | 5 |  | 7 |  |  |  | 8 | 5 | LS1_CB11:BPM_D1889 | 1091 | LS1_CB11:BPM_D1889 | 1889
BPM | LS1_CB11:BPM_D1889 | BPM | D1889 | 34E000245 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 8 | 5 |  | 8 |  |  |  | 8 | 5 | LS1_CB11:BPM_D1889 | 1091 | LS1_CB11:BPM_D1889 | 1889
BPM | LS1_CB11:RFS_D1889 | BPM | D1889 | 34E000223 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 8 | 5 |  | Ref |  |  |  | 8 | 5 | LS1_CB11:RFS_D1889
BPM | LS1_CB11:BPM_D1909 | BPM | D1909 | 34E000238 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 8 | 6 |  | 1 |  |  |  | 8 | 6 | LS1_CB11:BPM_D1909 | 1109 | LS1_CB11:BPM_D1909 | 1909
BPM | LS1_CB11:BPM_D1909 | BPM | D1909 | 34E000239 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 8 | 6 |  | 2 |  |  |  | 8 | 6 | LS1_CB11:BPM_D1909 | 1109 | LS1_CB11:BPM_D1909 | 1909
BPM | LS1_CB11:BPM_D1909 | BPM | D1909 | 34E000240 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 8 | 6 |  | 3 |  |  |  | 8 | 6 | LS1_CB11:BPM_D1909 | 1109 | LS1_CB11:BPM_D1909 | 1909
BPM | LS1_CB11:BPM_D1909 | BPM | D1909 | 34E000241 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 8 | 6 |  | 4 |  |  |  | 8 | 6 | LS1_CB11:BPM_D1909 | 1109 | LS1_CB11:BPM_D1909 | 1909
BPM | LS1_CB11:RFS_D1909 | BPM | D1909 | 34E000221 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | [17 empty] | LS1_CB11:RFS_D1909
BPM | LS1_WB11:BPM_D1923 | BPM | D1923 | 34E000257 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 8 | 6 |  | 5 |  |  |  | 8 | 6 | LS1_WB11:BPM_D1923 | 1121 | LS1_WB11:BPM_D1923 | 1923
BPM | LS1_WB11:BPM_D1923 | BPM | D1923 | 34E000258 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 8 | 6 |  | 6 |  |  |  | 8 | 6 | LS1_WB11:BPM_D1923 | 1121 | LS1_WB11:BPM_D1923 | 1923
BPM | LS1_WB11:BPM_D1923 | BPM | D1923 | 34E000259 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 8 | 6 |  | 7 |  |  |  | 8 | 6 | LS1_WB11:BPM_D1923 | 1121 | LS1_WB11:BPM_D1923 | 1923
BPM | LS1_WB11:BPM_D1923 | BPM | D1923 | 34E000260 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 8 | 6 |  | 8 |  |  |  | 8 | 6 | LS1_WB11:BPM_D1923 | 1121 | LS1_WB11:BPM_D1923 | 1923
BPM | LS1_WB11:RFS_D1923 | BPM | D1923 | 34E000226 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 8 | 6 |  | Ref |  |  |  | 8 | 6 | LS1_WB11:RFS_D1923
BPM | LS1_BTS:BPM_D1967 | BPM | D1966 | 34E000261 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 8 | 7 |  | 1 |  |  |  | 8 | 7 | LS1_BTS:BPM_D1967 | 1140 | LS1_BTS:BPM_D1967 | 1967
BPM | LS1_BTS:BPM_D1967 | BPM | D1966 | 34E000262 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 8 | 7 |  | 2 |  |  |  | 8 | 7 | LS1_BTS:BPM_D1967 | 1140 | LS1_BTS:BPM_D1967 | 1967
BPM | LS1_BTS:BPM_D1967 | BPM | D1966 | 34E000263 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 8 | 7 |  | 3 |  |  |  | 8 | 7 | LS1_BTS:BPM_D1967 | 1140 | LS1_BTS:BPM_D1967 | 1967
BPM | LS1_BTS:BPM_D1967 | BPM | D1966 | 34E000264 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 8 | 7 |  | 4 |  |  |  | 8 | 7 | LS1_BTS:BPM_D1967 | 1140 | LS1_BTS:BPM_D1967 | 1967
BPM | LS1_BTS:RFS_D1967 | BPM | D1966 | 34E000227 | 2C_1/4in_Hardline_036 | Ref | No terminiation | Ref | [6 empty] | Ref | SMA to SMA | 8 | 7 |  | Ref |  |  |  | 8 | 7 | LS1_BTS:RFS_D1967
BPM | LS1_BTS:BPM_D1980 | BPM | D1980 | 34E000265 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 8 | 7 |  | 5 |  |  |  | 8 | 7 | LS1_BTS:BPM_D1980 | 1147 | LS1_BTS:BPM_D1980 | 1980
BPM | LS1_BTS:BPM_D1980 | BPM | D1980 | 34E000266 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 8 | 7 |  | 6 |  |  |  | 8 | 7 | LS1_BTS:BPM_D1980 | 1147 | LS1_BTS:BPM_D1980 | 1980
BPM | LS1_BTS:BPM_D1980 | BPM | D1980 | 34E000267 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 8 | 7 |  | 7 |  |  |  | 8 | 7 | LS1_BTS:BPM_D1980 | 1147 | LS1_BTS:BPM_D1980 | 1980
BPM | LS1_BTS:BPM_D1980 | BPM | D1980 | 34E000268 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 8 | 7 |  | 8 |  |  |  | 8 | 7 | LS1_BTS:BPM_D1980 | 1147 | LS1_BTS:BPM_D1980 | 1980
BPM | LS1_BTS:RFS_D1980 | BPM | D1980 | 34E000228 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | No term | [17 empty] | LS1_BTS:RFS_D1980
BPM | LS1_BTS:BPM_D2027 | BPM | D2033 | 34E000269 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 8 | 8 |  | 1 |  |  |  | 8 | 8 | LS1_BTS:BPM_D2027 | 1181 | LS1_BTS:BPM_D2027 | 2027
BPM | LS1_BTS:BPM_D2027 | BPM | D2033 | 34E000270 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 8 | 8 |  | 2 |  |  |  | 8 | 8 | LS1_BTS:BPM_D2027 | 1181 | LS1_BTS:BPM_D2027 | 2027
BPM | LS1_BTS:BPM_D2027 | BPM | D2033 | 34E000271 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 8 | 8 |  | 3 |  |  |  | 8 | 8 | LS1_BTS:BPM_D2027 | 1181 | LS1_BTS:BPM_D2027 | 2027
BPM | LS1_BTS:BPM_D2027 | BPM | D2033 | 34E000272 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 8 | 8 |  | 4 |  |  |  | 8 | 8 | LS1_BTS:BPM_D2027 | 1181 | LS1_BTS:BPM_D2027 | 2027
BPM | LS1_BTS:RFS_D2027 | BPM | D2033 | 34E000229 | 2C_1/4in_Hardline_036 | Ref | No terminiation | Ref | [6 empty] | Ref | SMA to SMA | 8 | 8 |  | Ref |  |  |  | 8 | 8 | LS1_BTS:RFS_D2027
BPM | LS1_BTS:BPM_D2054 | BPM | D2056 | 34E000273 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 8 | 8 |  | 5 |  |  |  | 8 | 8 | LS1_BTS:BPM_D2054 | 1189 | LS1_BTS:BPM_D2054 | 2054
BPM | LS1_BTS:BPM_D2054 | BPM | D2056 | 34E000282 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 8 | 8 |  | 6 |  |  |  | 8 | 8 | LS1_BTS:BPM_D2054 | 1189 | LS1_BTS:BPM_D2054 | 2054
BPM | LS1_BTS:BPM_D2054 | BPM | D2056 | 34E000283 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 8 | 8 |  | 7 |  |  |  | 8 | 8 | LS1_BTS:BPM_D2054 | 1189 | LS1_BTS:BPM_D2054 | 2054
BPM | LS1_BTS:BPM_D2054 | BPM | D2056 | 34E000284 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 8 | 8 |  | 8 |  |  |  | 8 | 8 | LS1_BTS:BPM_D2054 | 1189 | LS1_BTS:BPM_D2054 | 2054
BPM | LS1_BTS:RFS_D2054 | BPM | D2056 | 34E000230 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | No term | [17 empty] | LS1_BTS:RFS_D2054

## FS1_N0302
FS1_N0302
Long Run Cable | [6 empty] | Inter-Rack Connections | [9 empty] | Terminating Device
Diagnostics Device | Lattice Beamline Location | Device Type | CableDB Location | Cable Number | Cable Type | Function | Termination Device | Port Number In | Pin Number | CPC pinout | CPC pinout | CPC pinout | Signal Pin | Shield Pin | Port Number Out | Patch Cable | MTCA Chassis | MTCA Slot | PICO Port | FGPDB Port | Bias Supply Chassis | Bias Supply Slot | Bias Supply Channel |  |  |  | search for row | device name | D number
Profile Monitors | [23 empty] | PICO | CH[0-7] | Device
PM 8" | LS1_BTS:PM_D2056 | PM | D2056 | 34F000019 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-1 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-1A | LEMO650/LEMO650 | 9 | 3 | CH0 - fmc2-1 |  | 6 | 1: EBS8005 | Ch0 | 9 | 3 | LS1_BTS:PM_D2056 | 1191 | LS1_BTS:PM_D2056 | 2056
PM 8" | LS1_BTS:PM_D2056 | PM | D2056 | 34F000019 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-1 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield-6 | 4 | 6 | LEMO-1B | LEMO650/LEMO650 | 9 | 3 | CH1 - fmc2-2 |  | 6 | 1: EBS8005 | Ch0 | 9 | 3 | LS1_BTS:PM_D2056 | 1191 | LS1_BTS:PM_D2056 | 2056
PM 8" | LS1_BTS:PM_D2056 | PM | D2056 | 34F000019 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-1 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-1C | LEMO650/LEMO650 | 9 | 3 | CH2 - fmc2-3 |  | 6 | 1: EBS8005 | Ch0 | 9 | 3 | LS1_BTS:PM_D2056 | 1191 | LS1_BTS:PM_D2056 | 2056
PM 8" | LS1_BTS:PM_D2131 | PM | D2131 | 34F000012 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-2 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-2A | LEMO650/LEMO650 | 9 | 3 | CH4 - fmc1-1 |  | 6 | 1: EBS8005 | Ch1 | 9 | 3 | LS1_BTS:PM_D2131 | 1236 | LS1_BTS:PM_D2131 | 2131
PM 8" | LS1_BTS:PM_D2131 | PM | D2131 | 34F000012 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-2 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | LEMO-2B | LEMO650/LEMO650 | 9 | 3 | CH5 - fmc1-2 |  | 6 | 1: EBS8005 | Ch1 | 9 | 3 | LS1_BTS:PM_D2131 | 1236 | LS1_BTS:PM_D2131 | 2131
PM 8" | LS1_BTS:PM_D2131 | PM | D2131 | 34F000012 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-2 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-2C | LEMO650/LEMO650 | 9 | 3 | CH6 - fmc1-3 |  | 6 | 1: EBS8005 | Ch1 | 9 | 3 | LS1_BTS:PM_D2131 | 1236 | LS1_BTS:PM_D2131 | 2131
PM 8" | FS1_CSS:PM_D2198 | PM | D2198 | 44F000056 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-3 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-3A | LEMO650/LEMO650 | 9 | 4 | CH0 - fmc2-1 |  | 6 | 1: EBS8005 | Ch2 | 9 | 4 | FS1_CSS:PM_D2198 | 1287 | FS1_CSS:PM_D2198 | 2198
PM 8" | FS1_CSS:PM_D2198 | PM | D2198 | 44F000056 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-3 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield-6 | 4 | 6 | LEMO-3B | LEMO650/LEMO650 | 9 | 4 | CH1 - fmc2-2 |  | 6 | 1: EBS8005 | Ch2 | 9 | 4 | FS1_CSS:PM_D2198 | 1287 | FS1_CSS:PM_D2198 | 2198
PM 8" | FS1_CSS:PM_D2198 | PM | D2198 | 44F000056 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-3 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-3C | LEMO650/LEMO650 | 9 | 4 | CH2 - fmc2-3 |  | 6 | 1: EBS8005 | Ch2 | 9 | 4 | FS1_CSS:PM_D2198 | 1287 | FS1_CSS:PM_D2198 | 2198
PM 8" | FS1_CSS:PM_D2225 | PM | D2225 | 44F000052 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-4 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-4A | LEMO650/LEMO650 | 9 | 4 | CH4 - fmc1-1 |  | 6 | 1: EBS8005 | Ch3 | 9 | 4 | FS1_CSS:PM_D2225 | 1311 | FS1_CSS:PM_D2225 | 2225
PM 8" | FS1_CSS:PM_D2225 | PM | D2225 | 44F000052 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-4 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | LEMO-4B | LEMO650/LEMO650 | 9 | 4 | CH5 - fmc1-2 |  | 6 | 1: EBS8005 | Ch3 | 9 | 4 | FS1_CSS:PM_D2225 | 1311 | FS1_CSS:PM_D2225 | 2225
PM 8" | FS1_CSS:PM_D2225 | PM | D2225 | 44F000052 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-4 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-4C | LEMO650/LEMO650 | 9 | 4 | CH6 - fmc1-3 |  | 6 | 1: EBS8005 | Ch3 | 9 | 4 | FS1_CSS:PM_D2225 | 1311 | FS1_CSS:PM_D2225 | 2225
PM 8" | FS1_CSS:PM_D2249 | PM | D2249 | 44F000048 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-5 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-5A | LEMO650/LEMO650 | 9 | 5 | CH0 - fmc2-1 |  | 6 | 1: EBS8005 | Ch4 | 9 | 5 | FS1_CSS:PM_D2249 | 1330 | FS1_CSS:PM_D2249 | 2249
PM 8" | FS1_CSS:PM_D2249 | PM | D2249 | 44F000048 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-5 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield-6 | 4 | 6 | LEMO-5B | LEMO650/LEMO650 | 9 | 5 | CH1 - fmc2-2 |  | 6 | 1: EBS8005 | Ch4 | 9 | 5 | FS1_CSS:PM_D2249 | 1330 | FS1_CSS:PM_D2249 | 2249
PM 8" | FS1_CSS:PM_D2249 | PM | D2249 | 44F000048 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-5 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-5C | LEMO650/LEMO650 | 9 | 5 | CH2 - fmc2-3 |  | 6 | 1: EBS8005 | Ch4 | 9 | 5 | FS1_CSS:PM_D2249 | 1330 | FS1_CSS:PM_D2249 | 2249
PM 8" | FS1_CSS:PM_D2385 | PM | D2385 | 44F000044 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-6 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-6A | LEMO650/LEMO650 | 9 | 5 | CH4 - fmc1-1 |  | 6 | 1: EBS8005 | Ch5 | 9 | 5 | FS1_CSS:PM_D2385 | 1451 | FS1_CSS:PM_D2385 | 2385
PM 8" | FS1_CSS:PM_D2385 | PM | D2385 | 44F000044 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-6 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | LEMO-6B | LEMO650/LEMO650 | 9 | 5 | CH5 - fmc1-2 |  | 6 | 1: EBS8005 | Ch5 | 9 | 5 | FS1_CSS:PM_D2385 | 1451 | FS1_CSS:PM_D2385 | 2385
PM 8" | FS1_CSS:PM_D2385 | PM | D2385 | 44F000044 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-6 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-6C | LEMO650/LEMO650 | 9 | 5 | CH6 - fmc1-3 |  | 6 | 1: EBS8005 | Ch5 | 9 | 5 | FS1_CSS:PM_D2385 | 1451 | FS1_CSS:PM_D2385 | 2385
PM 12" | FS1_BTS:PM_D2439 | PM | D2439 | 44F000036 | 6C_22AWG_Multi_010 | A Signal | CPC to LEMO650 Box | CPC-7 |  | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-7A | LEMO650/LEMO650 | 9 | 6 | CH0 - fmc2-1 |  | 6 | 1: EBS8005 | Ch6 | 9 | 6 | FS1_BTS:PM_D2439 | 3536 | FS1_BTS:PM_D2439 | 2439
PM 12" | FS1_BTS:PM_D2439 | PM | D2439 | 44F000036 | 6C_22AWG_Multi_010 | B Signal | CPC to LEMO650 Box | CPC-7 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | LEMO-7B | LEMO650/LEMO650 | 9 | 6 | CH1 - fmc2-2 |  | 6 | 1: EBS8005 | Ch6 | 9 | 6 | FS1_BTS:PM_D2439 | 3536 | FS1_BTS:PM_D2439 | 2439
PM 6" | FS1_BTS:PM_D2439 | PM | D2439 | 44F000040 | 4C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-7 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-7C | LEMO650/LEMO650 | 9 | 6 | CH2 - fmc2-3 |  | 6 | 1: EBS8005 | Ch6 | 9 | 6 | FS1_BTS:PM_D2439 | 3536 | FS1_BTS:PM_D2439 | 2439
Ion Chamber
Ion Chamber | FS1_BTS:IC_D2448 | IC | D2448 | 44E000256 | 2C_1kV75LC_Coax_090 | Signal | N-type to Lemo Box | J1 | [6 empty] | J1 | Lemo to Lemo | 9 | 7 | CH0 - fmc2-1 | [7 empty] | #N/A | #N/A | #N/A
Ion Chamber | FS1_BTS:IC_D2449 | IC | D2449 | 44E000257 | 2C_1kV75LC_Coax_090 | Signal | N-type to Lemo Box | J2 | [6 empty] | J2 | Lemo to Lemo | 9 | 7 | CH1 - fmc2-2 | [7 empty] | #N/A | #N/A | #N/A
Ion Chamber | FS1_BTS:IC_D2478 | IC | D2478 | 44E000254 | 2C_1kV75LC_Coax_090 | Signal | N-type to Lemo Box | J3 | [6 empty] | J3 | Lemo to Lemo | 9 | 7 | CH2 - fmc2-3 | [7 empty] | #N/A | #N/A | #N/A
Ion Chamber | FS1_BTS:IC_D2479 | IC | D2479 | 44E000255 | 2C_1kV75LC_Coax_090 | Signal | N-type to Lemo Box | J4 | [6 empty] | J4 | Lemo to Lemo | 9 | 7 | CH3 - fmc2-4 | [7 empty] | #N/A | #N/A | #N/A
Ion Chamber | FS1_STRL:IC_D2233 | IC | D2233 | 44E000258 | 2C_1kV75LC_Coax_090 | Signal | N-type to Lemo Box | J5 | [6 empty] | J5 | Lemo to Lemo | 9 | 7 | CH4 - fmc1-1 | [7 empty] | #N/A | #N/A | #N/A
Ion Chamber | FS1_STRL:IC_D2234 | IC | D2234 | 44E000259 | 2C_1kV75LC_Coax_090 | Signal | N-type to Lemo Box | J6 | [6 empty] | J6 | Lemo to Lemo | 9 | 7 | CH5 - fmc1-2 | [7 empty] | #N/A | #N/A | #N/A
Ion Chamber | FS1_STRL:IC_D2253 | IC | D2253 | 44E000261 | 2C_1kV75LC_Coax_090 | Signal | N-type to Lemo Box | J7 | [6 empty] | J7 | Lemo to Lemo | 9 | 7 | CH6 - fmc1-3 | [7 empty] | #N/A | #N/A | #N/A
Ion Chamber | FS1_STRL:IC_D2254 | IC | D2254 | 44E000260 | 2C_1kV75LC_Coax_090 | Signal | N-type to Lemo Box | J8 | [6 empty] | J8 | Lemo to Lemo | 9 | 7 | CH7 - fmc1-4 | [7 empty] | #N/A | #N/A | #N/A
Ion Chamber Bias
Ion Chamber | FS1_BTS:IC_D2448 |  | D2448 | 44H000041 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 1 | [6 empty] | 1 | SHV to SHV |  |  |  |  | 6 | 2: EHS8020n | Ch0 |  |  |  | #N/A | #N/A | #N/A
Ion Chamber | FS1_BTS:IC_D2449 |  | D2449 | 44H000040 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 2 | [6 empty] | 2 | SHV to SHV |  |  |  |  | 6 | 2: EHS8020n | Ch1 |  |  |  | #N/A | #N/A | #N/A
Ion Chamber | FS1_BTS:IC_D2478 |  | D2478 | 44H000039 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 3 | [6 empty] | 3 | SHV to SHV |  |  |  |  | 6 | 2: EHS8020n | Ch2 |  |  |  | #N/A | #N/A | #N/A
Ion Chamber | FS1_BTS:IC_D2479 |  | D2479 | 44H000038 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 4 | [6 empty] | 4 | SHV to SHV |  |  |  |  | 6 | 2: EHS8020n | Ch3 |  |  |  | #N/A | #N/A | #N/A
Ion Chamber | FS1_STRL:IC_D2233 |  | D2233 | 44H000044 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 5 | [6 empty] | 5 | SHV to SHV |  |  |  |  | 6 | 2: EHS8020n | Ch4 |  |  |  | #N/A | #N/A | #N/A
Ion Chamber | FS1_STRL:IC_D2234 |  | D2234 | 44H000045 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 6 | [6 empty] | 6 | SHV to SHV |  |  |  |  | 6 | 2: EHS8020n | Ch5 |  |  |  | #N/A | #N/A | #N/A
Ion Chamber | FS1_STRL:IC_D2253 |  | D2253 | 44H000042 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 7 | [6 empty] | 7 | SHV to SHV |  |  |  |  | 6 | 2: EHS8020n | Ch6 |  |  |  | #N/A | #N/A | #N/A
Ion Chamber | FS1_STRL:IC_D2254 |  | D2254 | 44H000043 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 8 | [6 empty] | 8 | SHV to SHV |  |  |  |  | 6 | 2: EHS8020n | Ch7 |  |  |  | #N/A | #N/A | #N/A
Halo Ring Monitor
Halo Monitor Ring | LS1_WB10:HMR_D1861 | HMR | D1861 | 34H000013 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J1 | [6 empty] | J1 | SHV to Lemo | 9 | 9 | CH0 - fmc2-1 |  | 6 | 1: EBS8005 | Ch7 |  |  |  | 1063 | LS1_WB10:HMR_D1861 | 1861
Halo Monitor Ring | LS1_WB11:HMR_D1925 | HMR | D1924 | 34H000017 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J2 | [6 empty] | J2 | SHV to Lemo | 9 | 9 | CH1 - fmc2-2 |  | 6 | 1: EBS8005 | Ch7 |  |  |  | 1123 | LS1_WB11:HMR_D1925 | 1925
| LS1_BTS:HMR_D1981 |  | D1981 | 34H000014 | 2C_1kV75LC_Coax_090 | Signal | [20 empty] | #N/A | #N/A | #N/A
| LS1_BTS:HMR_D2057 |  | D2057 | 34H000016 | 2C_1kV75LC_Coax_090 | Signal | [20 empty] | #N/A | #N/A | #N/A
| LS1_BTS:HMR_D2150 |  | D2150 | 34H000015 | 2C_1kV75LC_Coax_090 | Signal | [20 empty] | #N/A | #N/A | #N/A
| FS1_CSS:HMR_D2204 |  | D2204 | 44H000003 | 2C_1kV75LC_Coax_090 | Signal | [20 empty] | #N/A | #N/A | #N/A
| FS1_CSS:HMR_D2232 |  | D2232 | 44H000004 | 2C_1kV75LC_Coax_090 | Signal | [20 empty] | #N/A | #N/A | #N/A
| FS1_CSS:HMR_D2255 |  | D2255 | 44H000005 | 2C_1kV75LC_Coax_090 | Signal | [20 empty] | #N/A | #N/A | #N/A
| FS1_CSS:HMR_D2318 |  | D2318 | 44H000006 | 2C_1kV75LC_Coax_090 | Signal | [20 empty] | #N/A | #N/A | #N/A
| FS1_CSS:HMR_D2385 |  | D2385 | 44H000007 | 2C_1kV75LC_Coax_090 | Signal | [20 empty] | #N/A | #N/A | #N/A
| FS1_BMS:HMR_D2504 |  | D2504 | 44H000008 | 2C_1kV75LC_Coax_090 | Signal | [20 empty] | #N/A | #N/A | #N/A
[27 empty] | #N/A | #N/A | #N/A
| LS1_WB10:HMR_D1861 |  | D1861 | 34H000098 | 2C_5kVDC_Coax_035 | HV Bias | [20 empty] | 1063 | LS1_WB10:HMR_D1861 | 1861
| LS1_WB11:HMR_D1924 |  | D1924 | 34H000106 | 2C_5kVDC_Coax_035 | HV Bias | [20 empty] | #N/A | #N/A | #N/A
| LS1_BTS:HMR_D1981 |  | D1981 | 34H000096 | 2C_5kVDC_Coax_035 | HV Bias | [20 empty] | #N/A | #N/A | #N/A
| LS1_BTS:HMR_D2057 |  | D2057 | 34H000107 | 2C_5kVDC_Coax_035 | HV Bias | [20 empty] | #N/A | #N/A | #N/A
| LS1_BTS:HMR_D2150 |  | D2150 | 34H000094 | 2C_5kVDC_Coax_035 | HV Bias | [20 empty] | #N/A | #N/A | #N/A
| FS1_CSS:HMR_D2204 |  | D2204 | 44H000056 | 2C_5kVDC_Coax_035 | HV Bias | [20 empty] | #N/A | #N/A | #N/A
| FS1_CSS:HMR_D2232 |  | D2232 | 44H000061 | 2C_5kVDC_Coax_035 | HV Bias | [20 empty] | #N/A | #N/A | #N/A
| FS1_CSS:HMR_D2255 |  | D2255 | 44H000060 | 2C_5kVDC_Coax_035 | HV Bias | [20 empty] | #N/A | #N/A | #N/A
| FS1_CSS:HMR_D2318 |  | D2318 | 44H000058 | 2C_5kVDC_Coax_035 | HV Bias | [20 empty] | #N/A | #N/A | #N/A
| FS1_CSS:HMR_D2385 |  | D2385 | 44H000059 | 2C_5kVDC_Coax_035 | HV Bias | [20 empty] | #N/A | #N/A | #N/A
| FS1_BMS:HMR_D2504 |  | D2504 | 44H000062 | 2C_5kVDC_Coax_035 | HV Bias | [20 empty] | #N/A | #N/A | #N/A
[27 empty] | #N/A | #N/A | #N/A
BPM | [26 empty] | #N/A | #N/A | #N/A
BPM | LS1_BTS:BPM_D2116 | BPM | D2116 | 34E000274 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA | 1T | [6 empty] | 1T | SMA to SMA | 9 | 10 |  | 1 | [6 empty] | 1227 | LS1_BTS:BPM_D2116 | 2116
BPM | LS1_BTS:BPM_D2116 | BPM | D2116 | 34E000275 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA | 2L | [6 empty] | 2L | SMA to SMA | 9 | 10 |  | 2 | [6 empty] | 1227 | LS1_BTS:BPM_D2116 | 2116
BPM | LS1_BTS:BPM_D2116 | BPM | D2116 | 34E000276 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA | 3R | [6 empty] | 3R | SMA to SMA | 9 | 10 |  | 3 | [6 empty] | 1227 | LS1_BTS:BPM_D2116 | 2116
BPM | LS1_BTS:BPM_D2116 | BPM | D2116 | 34E000277 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA | 4B | [6 empty] | 4B | SMA to SMA | 9 | 10 |  | 4 | [6 empty] | 1227 | LS1_BTS:BPM_D2116 | 2116
BPM | LS1_BTS:BPM_D2133 | BPM | D2133 | 34E000232 | 2C_1/4in_Hardline_036 | Ref | [13 empty] | 4
BPM | LS1_BTS:BPM_D2130 | BPM | D2130 | 34E000278 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA | 1T | [6 empty] | 1T | SMA to SMA | 9 | 10 |  | 5 | [6 empty] | 1234 | LS1_BTS:BPM_D2130 | 2130
BPM | LS1_BTS:BPM_D2130 | BPM | D2130 | 34E000279 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA | 2L | [6 empty] | 2L | SMA to SMA | 9 | 10 |  | 6 | [6 empty] | 1234 | LS1_BTS:BPM_D2130 | 2130
BPM | LS1_BTS:BPM_D2130 | BPM | D2130 | 34E000280 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA | 3R | [6 empty] | 3R | SMA to SMA | 9 | 10 |  | 7 | [6 empty] | 1234 | LS1_BTS:BPM_D2130 | 2130
BPM | LS1_BTS:BPM_D2130 | BPM | D2130 | 34E000281 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA | 4B | [6 empty] | 4B | SMA to SMA | 9 | 10 |  | 8 | [6 empty] | 1234 | LS1_BTS:BPM_D2130 | 2130
BPM | LS1_BTS:RFS_D2149 | BPM | D2149 | 34E000275 | 2C_1/4in_Hardline_036 | No Term | [13 empty] | 8
BSM
BSM | FS1_CSS:LPM_D2263 | BSM | D2263 | 44F000181 | [12 empty] | 9 | 9 | CH4 - fmc1-1
Lemo to Lemo Panel
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 1_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 6 | 2: EBS8003 | Ch0
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 1_OUT Port | [7 empty] | Lemo650/Lemo650 | 9 | 3 | CH3 - fmc2-4
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 2_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 6 | 2: EBS8003 | Ch1
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 2_OUT Port | [7 empty] | Lemo650/Lemo650 | 9 | 3 | CH7 - fmc1-4
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 3_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 6 | 2: EBS8003 | Ch2
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 3_OUT Port | [7 empty] | Lemo650/Lemo650 | 9 | 4 | CH3 - fmc2-4
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 4_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 6 | 2: EBS8003 | Ch3
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 4_OUT Port | [7 empty] | Lemo650/Lemo650 | 9 | 4 | CH7 - fmc1-4
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 5_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 6 | 2: EBS8003 | Ch4
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 5_OUT Port | [7 empty] | Lemo650/Lemo650 | 9 | 5 | CH3 - fmc2-4
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 6_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 6 | 2: EBS8003 | Ch5
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 6_OUT Port | [7 empty] | Lemo650/Lemo650 | 9 | 5 | CH7 - fmc1-4
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 7_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 6 | 2: EBS8003 | Ch6
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 7_OUT Port | [7 empty] | Lemo650/Lemo650 | 9 | 6 | CH3 - fmc2-4
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 8_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 6 | 2: EBS8003 | Ch7
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 8_OUT Port | [7 empty] | Lemo650/Lemo650 | 9 | 9 | CH3 - fmc2-4
Ground Plugs | [23 empty] | #VALUE! | #VALUE! | #VALUE!
Grounding Plug |  | ground |  | - | Lemo650 Plug | Ground | Pico-8 | [8 empty] | Ground Plug | 9 | 9 | CH7 - fmc1-4

## FS1_N0506
FS1_N0506
Long Run Cable | [6 empty] | Inter-Rack Connections | [9 empty] | Terminating Device
Diagnostics Device | Lattice Beamline Location | Device Type | CableDB Location | Cable Number | Cable Type | Function | Termination Device | Port Number In | Pin Number | CPC pinout | CPC pinout | CPC pinout | Signal Pin | Shield Pin | Port Number Out | Patch Cable | MTCA Chassis | MTCA Slot | PICO Port | FGPDB Port | Bias Supply Chassis | Bias Supply Slot | Bias Supply Channel |  |  |  | search for row | device name | D number
BPMs | [23 empty] | PICO | CH[0-7] | Device
BPM | FS1_BBS:BPM_D2466 | BPM | D2466 | 44E000134 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 10 | 3 |  | 1 |  |  |  | 10 | 3 | FS1_BBS:BPM_D2466 | 1496 | FS1_BBS:BPM_D2466 | 2466
BPM | FS1_BBS:BPM_D2466 | BPM | D2466 | 44E000140 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 10 | 3 |  | 2 |  |  |  | 10 | 3 | FS1_BBS:BPM_D2466 | 1496 | FS1_BBS:BPM_D2466 | 2466
BPM | FS1_BBS:BPM_D2466 | BPM | D2466 | 44E000141 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 10 | 3 |  | 3 |  |  |  | 10 | 3 | FS1_BBS:BPM_D2466 | 1496 | FS1_BBS:BPM_D2466 | 2466
BPM | FS1_BBS:BPM_D2466 | BPM | D2466 | 44E000142 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 10 | 3 |  | 4 |  |  |  | 10 | 3 | FS1_BBS:BPM_D2466 | 1496 | FS1_BBS:BPM_D2466 | 2466
BPM | FS1_BBS:BPM_D2466 | BPM | D2466 | 44E000139 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | [17 empty] | FS1_BBS:BPM_D2466 | 1496 | FS1_BBS:BPM_D2466 | 2466
BPM | FS1_BTS:BPM_D2467 | BPM | D2467 | 44E000174 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 10 | 3 |  | 5 |  |  |  | 10 | 3 | FS1_BTS:BPM_D2467 | 3554 | FS1_BTS:BPM_D2467 | 2467
BPM | FS1_BTS:BPM_D2467 | BPM | D2467 | 44E000175 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 10 | 3 |  | 6 |  |  |  | 10 | 3 | FS1_BTS:BPM_D2467 | 3554 | FS1_BTS:BPM_D2467 | 2467
BPM | FS1_BTS:BPM_D2467 | BPM | D2467 | 44E000176 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 10 | 3 |  | 7 |  |  |  | 10 | 3 | FS1_BTS:BPM_D2467 | 3554 | FS1_BTS:BPM_D2467 | 2467
BPM | FS1_BTS:BPM_D2467 | BPM | D2467 | 44E000177 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 10 | 3 |  | 8 |  |  |  | 10 | 3 | FS1_BTS:BPM_D2467 | 3554 | FS1_BTS:BPM_D2467 | 2467
BPM | FS1_BTS:BPM_D2467 | BPM | D2467 | 44E000178 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 10 | 3 |  | Ref |  |  |  | 10 | 3 | FS1_BTS:BPM_D2467 | 3554 | FS1_BTS:BPM_D2467 | 2467
BPM | FS1_BMS:BPM_D2502 | BPM | D2502 | 44E000147 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 10 | 4 |  | 1 |  |  |  | 10 | 4 | FS1_BMS:BPM_D2502 | 1512 | FS1_BMS:BPM_D2502 | 2502
BPM | FS1_BMS:BPM_D2502 | BPM | D2502 | 44E000148 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 10 | 4 |  | 2 |  |  |  | 10 | 4 | FS1_BMS:BPM_D2502 | 1512 | FS1_BMS:BPM_D2502 | 2502
BPM | FS1_BMS:BPM_D2502 | BPM | D2502 | 44E000149 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 10 | 4 |  | 3 |  |  |  | 10 | 4 | FS1_BMS:BPM_D2502 | 1512 | FS1_BMS:BPM_D2502 | 2502
BPM | FS1_BMS:BPM_D2502 | BPM | D2502 | 44E000150 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 10 | 4 |  | 4 |  |  |  | 10 | 4 | FS1_BMS:BPM_D2502 | 1512 | FS1_BMS:BPM_D2502 | 2502
BPM | FS1_BMS:BPM_D2502 | BPM | D2502 | 44E000155 | 2C_1/4in_Hardline_036 | Ref | No terminiation | Ref | [6 empty] | Ref | SMA to SMA | 10 | 4 |  | Ref |  |  |  | 10 | 4 | FS1_BMS:BPM_D2502 | 1512 | FS1_BMS:BPM_D2502 | 2502
BPM | FS1_BMS:BPM_D2537 | BPM | D2537 | 44E000151 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 10 | 4 |  | 5 |  |  |  | 10 | 4 | FS1_BMS:BPM_D2537 | 1533 | FS1_BMS:BPM_D2537 | 2537
BPM | FS1_BMS:BPM_D2537 | BPM | D2537 | 44E000152 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 10 | 4 |  | 6 |  |  |  | 10 | 4 | FS1_BMS:BPM_D2537 | 1533 | FS1_BMS:BPM_D2537 | 2537
BPM | FS1_BMS:BPM_D2537 | BPM | D2537 | 44E000153 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 10 | 4 |  | 7 |  |  |  | 10 | 4 | FS1_BMS:BPM_D2537 | 1533 | FS1_BMS:BPM_D2537 | 2537
BPM | FS1_BMS:BPM_D2537 | BPM | D2537 | 44E000154 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 10 | 4 |  | 8 |  |  |  | 10 | 4 | FS1_BMS:BPM_D2537 | 1533 | FS1_BMS:BPM_D2537 | 2537
BPM | FS1_BMS:BPM_D2537 | BPM | D2537 | 44E000156 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | No term | [17 empty] | FS1_BMS:BPM_D2537 | 1533 | FS1_BMS:BPM_D2537 | 2537
BPM | FS1_BMS:BPM_D2587 | BPM | D2587 | 44E000143 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 10 | 5 |  | 1 |  |  |  | 10 | 5 | FS1_BMS:BPM_D2587 | 1547 | FS1_BMS:BPM_D2587 | 2587
BPM | FS1_BMS:BPM_D2587 | BPM | D2587 | 44E000144 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 10 | 5 |  | 2 |  |  |  | 10 | 5 | FS1_BMS:BPM_D2587 | 1547 | FS1_BMS:BPM_D2587 | 2587
BPM | FS1_BMS:BPM_D2587 | BPM | D2587 | 44E000145 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 10 | 5 |  | 3 |  |  |  | 10 | 5 | FS1_BMS:BPM_D2587 | 1547 | FS1_BMS:BPM_D2587 | 2587
BPM | FS1_BMS:BPM_D2587 | BPM | D2587 | 44E000146 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 10 | 5 |  | 4 |  |  |  | 10 | 5 | FS1_BMS:BPM_D2587 | 1547 | FS1_BMS:BPM_D2587 | 2587
BPM | FS1_BMS:BPM_D2587 | BPM | D2587 | 44E000157 | 2C_1/4in_Hardline_036 | Ref | No terminiation | Ref | [6 empty] | Ref | SMA to SMA | 10 | 5 |  | Ref |  |  |  | 10 | 5 | FS1_BMS:BPM_D2587 | 1547 | FS1_BMS:BPM_D2587 | 2587
BPM | FS1_BMS:BPM_D2600 | BPM | D2600 | 44E000135 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 10 | 5 |  | 5 |  |  |  | 10 | 5 | FS1_BMS:BPM_D2600 | 1556 | FS1_BMS:BPM_D2600 | 2600
BPM | FS1_BMS:BPM_D2600 | BPM | D2600 | 44E000136 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 10 | 5 |  | 6 |  |  |  | 10 | 5 | FS1_BMS:BPM_D2600 | 1556 | FS1_BMS:BPM_D2600 | 2600
BPM | FS1_BMS:BPM_D2600 | BPM | D2600 | 44E000137 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 10 | 5 |  | 7 |  |  |  | 10 | 5 | FS1_BMS:BPM_D2600 | 1556 | FS1_BMS:BPM_D2600 | 2600
BPM | FS1_BMS:BPM_D2600 | BPM | D2600 | 44E000138 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 10 | 5 |  | 8 |  |  |  | 10 | 5 | FS1_BMS:BPM_D2600 | 1556 | FS1_BMS:BPM_D2600 | 2600
BPM | FS1_BMS:BPM_D2600 | BPM | D2600 | 44E000158 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | No term | [17 empty] | FS1_BMS:BPM_D2600 | 1556 | FS1_BMS:BPM_D2600 | 2600
BPM | FS1_BMS:BPM_D2665 | BPM | D2665 | 44E000159 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 10 | 6 |  | 1 |  |  |  | 10 | 6 | FS1_BMS:BPM_D2665 | 1601 | FS1_BMS:BPM_D2665 | 2665
BPM | FS1_BMS:BPM_D2665 | BPM | D2665 | 44E000170 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 10 | 6 |  | 2 |  |  |  | 10 | 6 | FS1_BMS:BPM_D2665 | 1601 | FS1_BMS:BPM_D2665 | 2665
BPM | FS1_BMS:BPM_D2665 | BPM | D2665 | 44E000171 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 10 | 6 |  | 3 |  |  |  | 10 | 6 | FS1_BMS:BPM_D2665 | 1601 | FS1_BMS:BPM_D2665 | 2665
BPM | FS1_BMS:BPM_D2665 | BPM | D2665 | 44E000172 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 10 | 6 |  | 4 |  |  |  | 10 | 6 | FS1_BMS:BPM_D2665 | 1601 | FS1_BMS:BPM_D2665 | 2665
BPM | FS1_BMS:BPM_D2665 | BPM | D2665 | 44E000173 | 2C_1/4in_Hardline_036 | Ref | No terminiation | Ref | [6 empty] | Ref | SMA to SMA | 10 | 6 |  | Ref |  |  |  | 10 | 6 | FS1_BMS:BPM_D2665 | 1601 | FS1_BMS:BPM_D2665 | 2665
BPM | FS1_BMS:BPM_D2690 | BPM | D2690 | 44E000166 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 10 | 6 |  | 5 |  |  |  | 10 | 6 | FS1_BMS:BPM_D2690 | 1616 | FS1_BMS:BPM_D2690 | 2690
BPM | FS1_BMS:BPM_D2690 | BPM | D2690 | 44E000167 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 10 | 6 |  | 6 |  |  |  | 10 | 6 | FS1_BMS:BPM_D2690 | 1616 | FS1_BMS:BPM_D2690 | 2690
BPM | FS1_BMS:BPM_D2690 | BPM | D2690 | 44E000168 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 10 | 6 |  | 7 |  |  |  | 10 | 6 | FS1_BMS:BPM_D2690 | 1616 | FS1_BMS:BPM_D2690 | 2690
BPM | FS1_BMS:BPM_D2690 | BPM | D2690 | 44E000169 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 10 | 6 |  | 8 |  |  |  | 10 | 6 | FS1_BMS:BPM_D2690 | 1616 | FS1_BMS:BPM_D2690 | 2690
BPM | FS1_BMS:BPM_D2690 | BPM | D2690 | 44E000160 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | No term | [17 empty] | FS1_BMS:BPM_D2690 | 1616 | FS1_BMS:BPM_D2690 | 2690
BPM | FS1_BMS:BPM_D2702 | BPM | D2702 | 44E000162 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 10 | 7 |  | 1 |  |  |  | 10 | 7 | FS1_BMS:BPM_D2702 | 1623 | FS1_BMS:BPM_D2702 | 2702
BPM | FS1_BMS:BPM_D2702 | BPM | D2702 | 44E000163 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 10 | 7 |  | 2 |  |  |  | 10 | 7 | FS1_BMS:BPM_D2702 | 1623 | FS1_BMS:BPM_D2702 | 2702
BPM | FS1_BMS:BPM_D2702 | BPM | D2702 | 44E000164 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 10 | 7 |  | 3 |  |  |  | 10 | 7 | FS1_BMS:BPM_D2702 | 1623 | FS1_BMS:BPM_D2702 | 2702
BPM | FS1_BMS:BPM_D2702 | BPM | D2702 | 44E000165 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 10 | 7 |  | 4 |  |  |  | 10 | 7 | FS1_BMS:BPM_D2702 | 1623 | FS1_BMS:BPM_D2702 | 2702
BPM | FS1_BMS:BPM_D2702 | BPM | D2702 | 44E000161 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | [17 empty] | FS1_BMS:BPM_D2702 | 1623 | FS1_BMS:BPM_D2702 | 2702
BPM | LS2_WC01:BPM_D2742 | BPM | D2742 | 34E000395 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 10 | 7 |  | 5 |  |  |  | 10 | 7 | LS2_WC01:BPM_D2742 | 1661 | LS2_WC01:BPM_D2742 | 2742
BPM | LS2_WC01:BPM_D2742 | BPM | D2742 | 34E000396 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 10 | 7 |  | 6 |  |  |  | 10 | 7 | LS2_WC01:BPM_D2742 | 1661 | LS2_WC01:BPM_D2742 | 2742
BPM | LS2_WC01:BPM_D2742 | BPM | D2742 | 34E000397 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 10 | 7 |  | 7 |  |  |  | 10 | 7 | LS2_WC01:BPM_D2742 | 1661 | LS2_WC01:BPM_D2742 | 2742
BPM | LS2_WC01:BPM_D2742 | BPM | D2742 | 34E000398 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 10 | 7 |  | 8 |  |  |  | 10 | 7 | LS2_WC01:BPM_D2742 | 1661 | LS2_WC01:BPM_D2742 | 2742
BPM | LS2_WC01:BPM_D2742 | BPM | D2742 | 34E000411 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 10 | 7 |  | Ref |  |  |  | 10 | 7 | LS2_WC01:BPM_D2742 | 1661 | LS2_WC01:BPM_D2742 | 2742
Profile Monitors
PM 12" | FS1_BBS:PM_D2444 | PM | D2444 | 44F000084 | 6C_22AWG_Multi_010 | A Signal | CPC to LEMO650 Box | CPC-1 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-1A | LEMO650/LEMO650 | 10 | 8 | CH0 - fmc2-1 |  | 7 | 1: EBS8005 | Ch0 | 10 | 8 | FS1_BBS:PM_D2444 | 1483 | FS1_BBS:PM_D2444 | 2444
PM 12" | FS1_BBS:PM_D2444 | PM | D2444 | 44F000084 | 6C_22AWG_Multi_010 | B Signal | CPC to LEMO650 Box | CPC-1 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | LEMO-1B | LEMO650/LEMO650 | 10 | 8 | CH1 - fmc2-2 |  | 7 | 1: EBS8005 | Ch0 | 10 | 8 | FS1_BBS:PM_D2444 | 1483 | FS1_BBS:PM_D2444 | 2444
PM 6" | FS1_BBS:PM_D2444 | PM | D2444 | 44F000080 | 4C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-1 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-1C | LEMO650/LEMO650 | 10 | 8 | CH2 - fmc2-3 |  | 7 | 1: EBS8005 | Ch0 | 10 | 8 | FS1_BBS:PM_D2444 | 1483 | FS1_BBS:PM_D2444 | 2444
PM 12" | FS1_BBS:PM_D2482 | PM | D2482 | 44F000088 | 6C_22AWG_Multi_010 | A Signal | CPC to LEMO650 Box | CPC-2 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-2A | LEMO650/LEMO650 | 10 | 8 | CH4 - fmc1-1 |  | 7 | 1: EBS8005 | Ch1 | 10 | 8 | FS1_BBS:PM_D2482 | 1504 | FS1_BBS:PM_D2482 | 2482
PM 12" | FS1_BBS:PM_D2482 | PM | D2482 | 44F000088 | 6C_22AWG_Multi_010 | B Signal | CPC to LEMO650 Box | CPC-2 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | LEMO-2B | LEMO650/LEMO650 | 10 | 8 | CH5 - fmc1-2 |  | 7 | 1: EBS8005 | Ch1 | 10 | 8 | FS1_BBS:PM_D2482 | 1504 | FS1_BBS:PM_D2482 | 2482
PM 6" | FS1_BBS:PM_D2482 | PM | D2482 | 44F000091 | 4C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-2 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-2C | LEMO650/LEMO650 | 10 | 8 | CH6 - fmc1-3 |  | 7 | 1: EBS8005 | Ch1 | 10 | 8 | FS1_BBS:PM_D2482 | 1504 | FS1_BBS:PM_D2482 | 2482
PM 12" | FS1_BTS:PM_D2470 | PM | D2470 | 44F000072 | 6C_22AWG_Multi_010 | A Signal | CPC to LEMO650 Box | CPC-3 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-3A | LEMO650/LEMO650 | 10 | 9 | CH0 - fmc2-1 |  | 7 | 1: EBS8005 | Ch2 | 10 | 9 | FS1_BTS:PM_D2470 | 3557 | FS1_BTS:PM_D2470 | 2470
PM 12" | FS1_BTS:PM_D2470 | PM | D2470 | 44F000072 | 6C_22AWG_Multi_010 | B Signal | CPC to LEMO650 Box | CPC-3 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | LEMO-3B | LEMO650/LEMO650 | 10 | 9 | CH1 - fmc2-2 |  | 7 | 1: EBS8005 | Ch2 | 10 | 9 | FS1_BTS:PM_D2470 | 3557 | FS1_BTS:PM_D2470 | 2470
PM 6" | FS1_BTS:PM_D2470 | PM | D2470 | 44F000076 | 4C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-3 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-3C | LEMO650/LEMO650 | 10 | 9 | CH2 - fmc2-3 |  | 7 | 1: EBS8005 | Ch2 | 10 | 9 | FS1_BTS:PM_D2470 | 3557 | FS1_BTS:PM_D2470 | 2470
PM 8" | FS1_BMS:PM_D2552 | PM | D2504 | 44F000063 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-4 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-4A | LEMO650/LEMO650 | 10 | 9 | CH4 - fmc1-1 |  | 7 | 1: EBS8005 | Ch3 | 10 | 9 | FS1_BMS:PM_D2552 | #N/A | #N/A | #N/A
PM 8" | FS1_BMS:PM_D2552 | PM | D2504 | 44F000063 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-4 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | LEMO-4B | LEMO650/LEMO650 | 10 | 9 | CH5 - fmc1-2 |  | 7 | 1: EBS8005 | Ch3 | 10 | 9 | FS1_BMS:PM_D2552 | #N/A | #N/A | #N/A
PM 8" | FS1_BMS:PM_D2552 | PM | D2504 | 44F000063 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-4 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-4C | LEMO650/LEMO650 | 10 | 9 | CH6 - fmc1-3 |  | 7 | 1: EBS8005 | Ch3 | 10 | 9 | FS1_BMS:PM_D2552 | #N/A | #N/A | #N/A
PM 8" | FS1_BMS:PM_D2602 | PM | D2602 | 44F000064 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-5 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-5A | LEMO650/LEMO650 | 10 | 10 | CH0 - fmc2-1 |  | 7 | 1: EBS8005 | Ch4 | 10 | 10 | FS1_BMS:PM_D2602 | 1558 | FS1_BMS:PM_D2602 | 2602
PM 8" | FS1_BMS:PM_D2602 | PM | D2602 | 44F000064 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-5 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield-6 | 4 | 6 | LEMO-5B | LEMO650/LEMO650 | 10 | 10 | CH1 - fmc2-2 |  | 7 | 1: EBS8005 | Ch4 | 10 | 10 | FS1_BMS:PM_D2602 | 1558 | FS1_BMS:PM_D2602 | 2602
PM 8" | FS1_BMS:PM_D2602 | PM | D2602 | 44F000064 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-5 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-6C | LEMO650/LEMO650 | 10 | 10 | CH2 - fmc2-3 |  | 7 | 1: EBS8005 | Ch4 | 10 | 10 | FS1_BMS:PM_D2602 | 1558 | FS1_BMS:PM_D2602 | 2602
PM 8" | FS1_BMS:PM_D2703 | PM | D2703 | 44F000068 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-6 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | LEMO-7A | LEMO650/LEMO650 | 10 | 10 | CH4 - fmc1-1 |  | 7 | 1: EBS8005 | Ch5 | 10 | 10 | FS1_BMS:PM_D2703 | 1625 | FS1_BMS:PM_D2703 | 2703
PM 8" | FS1_BMS:PM_D2703 | PM | D2703 | 44F000068 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-6 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | LEMO-7B | LEMO650/LEMO650 | 10 | 10 | CH5 - fmc1-2 |  | 7 | 1: EBS8005 | Ch5 | 10 | 10 | FS1_BMS:PM_D2703 | 1625 | FS1_BMS:PM_D2703 | 2703
PM 8" | FS1_BMS:PM_D2703 | PM | D2703 | 44F000068 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-6 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | LEMO-7C | LEMO650/LEMO650 | 10 | 10 | CH6 - fmc1-3 |  | 7 | 1: EBS8005 | Ch5 | 10 | 10 | FS1_BMS:PM_D2703 | 1625 | FS1_BMS:PM_D2703 | 2703
Lemo to Lemo Panel
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 1_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 7 | 2: EBS8003 | Ch0 |  |  | PM bias for PICO
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 1_OUT Port | [7 empty] | Lemo650/Lemo650 | 10 | 8 | CH3 - fmc2-4 |  |  |  |  | 10 | 8 | PM bias for PICO
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 2_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 7 | 2: EBS8003 | Ch1 |  |  | PM bias for PICO
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 2_OUT Port | [7 empty] | Lemo650/Lemo650 | 10 | 8 | CH7 - fmc1-4 |  |  |  |  | 10 | 8 | PM bias for PICO
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 3_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 7 | 2: EBS8003 | Ch2 |  |  | PM bias for PICO
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 3_OUT Port | [7 empty] | Lemo650/Lemo650 | 10 | 9 | CH3 - fmc2-4
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 4_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 7 | 2: EBS8003 | Ch3
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 4_OUT Port | [7 empty] | Lemo650/Lemo650 | 10 | 9 | CH7 - fmc1-4
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 5_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 7 | 2: EBS8003 | Ch4
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 5_OUT Port | [7 empty] | Lemo650/Lemo650 | 10 | 10 | CH3 - fmc2-4
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 6_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 1 | 2: EBS8003 | Ch5
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 6_OUT Port | [7 empty] | Lemo650/Lemo650 | 10 | 10 | CH7 - fmc1-4

## LS2_N0808
LS2_N0808
Long Run Cable | [6 empty] | Inter-Rack Connections | [9 empty] | Terminating Device
Diagnostics Device | Lattice Beamline Location | Device Type | CableDB Location | Cable Number | Cable Type | Function | Termination Device | Port Number In | Pin Number | CPC pinout | CPC pinout | CPC pinout | Signal Pin | Shield Pin | Port Number Out | Patch Cable | MTCA Chassis | MTCA Slot | PICO Port | FGPDB Port | Bias Supply Chassis | Bias Supply Slot | Bias Supply Channel |  |  |  | search for row | device name | D number
BPMs | [23 empty] | PICO | CH[0-7] | Device
BPM | LS2_WC02:BPM_D2782 | BPM | D2782 | 34E000399 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 11 | 3 |  | 1 |  |  |  | 11 | 3 | LS2_WC02:BPM_D2782 | 1697 | LS2_WC02:BPM_D2782 | 2782
BPM | LS2_WC02:BPM_D2782 | BPM | D2782 | 34E000400 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 11 | 3 |  | 2 |  |  |  | 11 | 3 | LS2_WC02:BPM_D2782 | 1697 | LS2_WC02:BPM_D2782 | 2782
BPM | LS2_WC02:BPM_D2782 | BPM | D2782 | 34E000401 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 11 | 3 |  | 3 |  |  |  | 11 | 3 | LS2_WC02:BPM_D2782 | 1697 | LS2_WC02:BPM_D2782 | 2782
BPM | LS2_WC02:BPM_D2782 | BPM | D2782 | 34E000402 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 11 | 3 |  | 4 |  |  |  | 11 | 3 | LS2_WC02:BPM_D2782 | 1697 | LS2_WC02:BPM_D2782 | 2782
BPM | LS2_WC02:BPM_D2782 | BPM | D2782 | 34E000412 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | [17 empty] | LS2_WC02:BPM_D2782 | 1697 | LS2_WC02:BPM_D2782 | 2782
BPM | LS2_WC03:BPM_D2821 | BPM | D2821 | 34E000407 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 11 | 3 |  | 5 |  |  |  | 11 | 3 | LS2_WC03:BPM_D2821 | 1733 | LS2_WC03:BPM_D2821 | 2821
BPM | LS2_WC03:BPM_D2821 | BPM | D2821 | 34E000408 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 11 | 3 |  | 6 |  |  |  | 11 | 3 | LS2_WC03:BPM_D2821 | 1733 | LS2_WC03:BPM_D2821 | 2821
BPM | LS2_WC03:BPM_D2821 | BPM | D2821 | 34E000409 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 11 | 3 |  | 7 |  |  |  | 11 | 3 | LS2_WC03:BPM_D2821 | 1733 | LS2_WC03:BPM_D2821 | 2821
BPM | LS2_WC03:BPM_D2821 | BPM | D2821 | 34E000410 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 11 | 3 |  | 8 |  |  |  | 11 | 3 | LS2_WC03:BPM_D2821 | 1733 | LS2_WC03:BPM_D2821 | 2821
BPM | LS2_WC03:BPM_D2821 | BPM | D2821 | 34E000413 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 11 | 3 |  | Ref |  |  |  | 11 | 3 | LS2_WC03:BPM_D2821 | 1733 | LS2_WC03:BPM_D2821 | 2821
BPM | LS2_WC04:BPM_D2861 | BPM | D2861 | 34E000403 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 11 | 4 |  | 1 |  |  |  | 11 | 4 | LS2_WC04:BPM_D2861 | 1769 | LS2_WC04:BPM_D2861 | 2861
BPM | LS2_WC04:BPM_D2861 | BPM | D2861 | 34E000404 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 11 | 4 |  | 2 |  |  |  | 11 | 4 | LS2_WC04:BPM_D2861 | 1769 | LS2_WC04:BPM_D2861 | 2861
BPM | LS2_WC04:BPM_D2861 | BPM | D2861 | 34E000405 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 11 | 4 |  | 3 |  |  |  | 11 | 4 | LS2_WC04:BPM_D2861 | 1769 | LS2_WC04:BPM_D2861 | 2861
BPM | LS2_WC04:BPM_D2861 | BPM | D2861 | 34E000406 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 11 | 4 |  | 4 |  |  |  | 11 | 4 | LS2_WC04:BPM_D2861 | 1769 | LS2_WC04:BPM_D2861 | 2861
BPM | LS2_WC04:BPM_D2861 | BPM | D2861 | 34E000414 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | [17 empty] | LS2_WC04:BPM_D2861 | 1769 | LS2_WC04:BPM_D2861 | 2861
BPM | LS2_WC05:BPM_D2901 | BPM | D2901 | 34E000415 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 11 | 4 |  | 5 |  |  |  | 11 | 4 | LS2_WC05:BPM_D2901 | 1805 | LS2_WC05:BPM_D2901 | 2901
BPM | LS2_WC05:BPM_D2901 | BPM | D2901 | 34E000416 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 11 | 4 |  | 6 |  |  |  | 11 | 4 | LS2_WC05:BPM_D2901 | 1805 | LS2_WC05:BPM_D2901 | 2901
BPM | LS2_WC05:BPM_D2901 | BPM | D2901 | 34E000417 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 11 | 4 |  | 7 |  |  |  | 11 | 4 | LS2_WC05:BPM_D2901 | 1805 | LS2_WC05:BPM_D2901 | 2901
BPM | LS2_WC05:BPM_D2901 | BPM | D2901 | 34E000418 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 11 | 4 |  | 8 |  |  |  | 11 | 4 | LS2_WC05:BPM_D2901 | 1805 | LS2_WC05:BPM_D2901 | 2901
BPM | LS2_WC05:BPM_D2901 | BPM | D2901 | 34E000420 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 11 | 4 |  | Ref |  |  |  | 11 | 4 | LS2_WC05:BPM_D2901 | 1805 | LS2_WC05:BPM_D2901 | 2901
BPM | LS2_WC06:BPM_D2941 | BPM | D2941 | 34E000421 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 11 | 5 |  | 1 |  |  |  | 11 | 5 | LS2_WC06:BPM_D2941 | 1841 | LS2_WC06:BPM_D2941 | 2941
BPM | LS2_WC06:BPM_D2941 | BPM | D2941 | 34E000422 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 11 | 5 |  | 2 |  |  |  | 11 | 5 | LS2_WC06:BPM_D2941 | 1841 | LS2_WC06:BPM_D2941 | 2941
BPM | LS2_WC06:BPM_D2941 | BPM | D2941 | 34E000423 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 11 | 5 |  | 3 |  |  |  | 11 | 5 | LS2_WC06:BPM_D2941 | 1841 | LS2_WC06:BPM_D2941 | 2941
BPM | LS2_WC06:BPM_D2941 | BPM | D2941 | 34E000424 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 11 | 5 |  | 4 |  |  |  | 11 | 5 | LS2_WC06:BPM_D2941 | 1841 | LS2_WC06:BPM_D2941 | 2941
BPM | LS2_WC06:BPM_D2941 | BPM | D2941 | 34E000419 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | [17 empty] | LS2_WC06:BPM_D2941 | 1841 | LS2_WC06:BPM_D2941 | 2941
BPM | LS2_WC07:BPM_D2981 | BPM | D2981 | 34E000425 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 11 | 5 |  | 5 |  |  |  | 11 | 5 | LS2_WC07:BPM_D2981 | 1877 | LS2_WC07:BPM_D2981 | 2981
BPM | LS2_WC07:BPM_D2981 | BPM | D2981 | 34E000426 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 11 | 5 |  | 6 |  |  |  | 11 | 5 | LS2_WC07:BPM_D2981 | 1877 | LS2_WC07:BPM_D2981 | 2981
BPM | LS2_WC07:BPM_D2981 | BPM | D2981 | 34E000427 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 11 | 5 |  | 7 |  |  |  | 11 | 5 | LS2_WC07:BPM_D2981 | 1877 | LS2_WC07:BPM_D2981 | 2981
BPM | LS2_WC07:BPM_D2981 | BPM | D2981 | 34E000428 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 11 | 5 |  | 8 |  |  |  | 11 | 5 | LS2_WC07:BPM_D2981 | 1877 | LS2_WC07:BPM_D2981 | 2981
BPM | LS2_WC07:BPM_D2981 | BPM | D2981 | 34E000469 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 11 | 5 |  | Ref |  |  |  | 11 | 5 | LS2_WC07:BPM_D2981 | 1877 | LS2_WC07:BPM_D2981 | 2981
BPM | LS2_WC08:BPM_D3020 | BPM | D3020 | 34E000433 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 11 | 6 |  | 1 |  |  |  | 11 | 6 | LS2_WC08:BPM_D3020 | 1913 | LS2_WC08:BPM_D3020 | 3020
BPM | LS2_WC08:BPM_D3020 | BPM | D3020 | 34E000434 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 11 | 6 |  | 2 |  |  |  | 11 | 6 | LS2_WC08:BPM_D3020 | 1913 | LS2_WC08:BPM_D3020 | 3020
BPM | LS2_WC08:BPM_D3020 | BPM | D3020 | 34E000435 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 11 | 6 |  | 3 |  |  |  | 11 | 6 | LS2_WC08:BPM_D3020 | 1913 | LS2_WC08:BPM_D3020 | 3020
BPM | LS2_WC08:BPM_D3020 | BPM | D3020 | 34E000436 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 11 | 6 |  | 4 |  |  |  | 11 | 6 | LS2_WC08:BPM_D3020 | 1913 | LS2_WC08:BPM_D3020 | 3020
BPM | LS2_WC08:BPM_D3020 | BPM | D3020 | 34E000470 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | [17 empty] | LS2_WC08:BPM_D3020 | 1913 | LS2_WC08:BPM_D3020 | 3020
BPM | LS2_WC09:BPM_D3060 | BPM | D3060 | 34E000437 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 11 | 6 |  | 5 |  |  |  | 11 | 6 | LS2_WC09:BPM_D3060 | 1949 | LS2_WC09:BPM_D3060 | 3060
BPM | LS2_WC09:BPM_D3060 | BPM | D3060 | 34E000438 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 11 | 6 |  | 6 |  |  |  | 11 | 6 | LS2_WC09:BPM_D3060 | 1949 | LS2_WC09:BPM_D3060 | 3060
BPM | LS2_WC09:BPM_D3060 | BPM | D3060 | 34E000439 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 11 | 6 |  | 7 |  |  |  | 11 | 6 | LS2_WC09:BPM_D3060 | 1949 | LS2_WC09:BPM_D3060 | 3060
BPM | LS2_WC09:BPM_D3060 | BPM | D3060 | 34E000440 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 11 | 6 |  | 8 |  |  |  | 11 | 6 | LS2_WC09:BPM_D3060 | 1949 | LS2_WC09:BPM_D3060 | 3060
BPM | LS2_WC09:BPM_D3060 | BPM | D3060 | 34E000471 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 11 | 6 |  | Ref |  |  |  | 11 | 6 | LS2_WC09:BPM_D3060 | 1949 | LS2_WC09:BPM_D3060 | 3060
Neutron Detector Signal
Neutron Detector | FS1_BMS:ND_D2588 | ND | D2588 | 44E000229 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | 1 | [6 empty] | 1 | SMA to Lemo650 | 11 | 9 | CH0 - fmc2-1 |  |  |  |  | 11 | 9 | FS1_BMS:ND_D2588 | #N/A | #N/A | #N/A
Neutron Detector | LS2_CC02:ND_D2748 | ND | D2748 | 34E000536 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | 2 | [6 empty] | 2 | SMA to Lemo650 | 11 | 9 | CH1 - fmc2-2 |  |  |  |  | 11 | 9 | LS2_CC02:ND_D2748 | #N/A | #N/A | #N/A
Neutron Detector | LS2_CC04:ND_D2828 | ND | D2828 | 34E000538 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | 3 | [6 empty] | 3 | SMA to Lemo650 | 11 | 9 | CH2 - fmc2-3 |  |  |  |  | 11 | 9 | LS2_CC04:ND_D2828 | #N/A | #N/A | #N/A
Neutron Detector | LS2_CC06:ND_D2916 | ND | D2916 | 34E000540 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | 4 | [6 empty] | 4 | SMA to Lemo650 | 11 | 9 | CH3 - fmc2-4 |  |  |  |  | 11 | 9 | LS2_CC06:ND_D2916 | #N/A | #N/A | #N/A
Neutron Detector Test signal
Neutron Detector | FS1_BMS:ND_D2588 |  | D2588 | 44E000228 | 2C_75LC_Coax_089 | Test Signal | N-Type to SMA Panel | [8 empty] | SMA to Ferrule | [9 empty] | FS1_BMS:ND_D2588 | #N/A | #N/A | #N/A
Neutron Detector | LS2_CC02:ND_D2748 |  | D2748 | 34E000535 | 2C_75LC_Coax_089 | Test Signal | N-Type to SMA Panel | [8 empty] | SMA to Ferrule | [9 empty] | LS2_CC02:ND_D2748 | #N/A | #N/A | #N/A
Neutron Detector | LS2_CC04:ND_D2828 |  | D2828 | 34E000537 | 2C_75LC_Coax_089 | Test Signal | N-Type to SMA Panel | [8 empty] | SMA to Ferrule | [9 empty] | LS2_CC04:ND_D2828 | #N/A | #N/A | #N/A
Neutron Detector | LS2_CC06:ND_D2916 |  | D2916 | 34E000539 | 2C_75LC_Coax_089 | Test Signal | N-Type to SMA Panel | [8 empty] | SMA to Ferrule | [9 empty] | LS2_CC06:ND_D2916 | #N/A | #N/A | #N/A
Neutron Detector Bias
Neutron Detector | FS1_BMS:ND_D2588 |  | D2588 | 44H000013 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 1 | [6 empty] | 1 | SHV to SHV |  |  |  |  | 8 | 2 : EHS8020n | Ch0 |  |  | FS1_BMS:ND_D2588
Neutron Detector | LS2_CC02:ND_D2748 |  | D2748 | 34H000064 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 2 | [6 empty] | 2 | SHV to SHV |  |  |  |  | 8 | 2 : EHS8020n | Ch1 |  |  | LS2_CC02:ND_D2748
Neutron Detector | LS2_CC04:ND_D2828 |  | D2828 | 34H000065 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 3 | [6 empty] | 3 | SHV to SHV |  |  |  |  | 8 | 2 : EHS8020n | Ch2 |  |  | LS2_CC04:ND_D2828
Neutron Detector | LS2_CC06:ND_D2916 |  | D2916 | 34H000066 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 4 | [6 empty] | 4 | SHV to SHV |  |  |  |  | 8 | 2 : EHS8020n | Ch3 |  |  | LS2_CC06:ND_D2916
Faraday Cup
Faraday Cup | FS1_BMS:FC_D2634 | FC | D2605 | 44H000011 | 2C_1kV75LC_Coax_090 | Signal | N-Type to SMA Panel | J1 | [6 empty] | J1 | SMA/LEMO650 | 11 | 7 | CH4 - fmc1-1 |  |  |  |  | 11 | 7 | FS1_BMS:FC_D2634
Faraday Cup | FS1_BMS:FC_D2634 |  | D2605 | 44H000063 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 5 | [6 empty] | 5 | SHV to SHV |  |  |  |  | 8 | 2 : EHS8020n | Ch4 |  |  | FS1_BMS:FC_D2634
Halo Monitor Ring
Halo Monitor Ring | FS1_BMS:HMR_D2703 | HMR | D2703 | 44H000009 | 2C_1kV75LC_Coax_090 | Signal | SHV to Lemo | J1 | [6 empty] | J1 | Lemo to Lemo | 11 | 7 | CH0 - fmc2-1 |  | 8 | 1: EBS8003 | Ch0 | 11 | 7 | FS1_BMS:HMR_D2703
| FS1_BMS:HMR_D2703 |  | D2703 | 44H000065 | 2C_5kVDC_Coax_035 | HV Bias | [19 empty] | FS1_BMS:HMR_D2703
Halo Monitor Ring,Spare | LS2_WC01:HMR_D2743 | HMR - Unused | D2743 | 34H000030 | 2C_1kV75LC_Coax_090 | Signal | [19 empty] | LS2_WC01:HMR_D2743
| LS2_WC01:HMR_D2743 |  | D2743 | 34H000108 | 2C_5kVDC_Coax_035 | HV Bias | [19 empty] | LS2_WC01:HMR_D2743
Halo Monitor Ring,Spare | LS2_WC02:HMR_D2783 | HMR - Unused | D2783 | 34H000031 | 2C_1kV75LC_Coax_090 | Signal | [19 empty] | LS2_WC02:HMR_D2783
| LS2_WC02:HMR_D2783 |  | D2783 | 34H000123 | 2C_5kVDC_Coax_035 | HV Bias | [19 empty] | LS2_WC02:HMR_D2783
Halo Monitor Ring | LS2_WC03:HMR_D2823 | HMR | D2823 | 34H000032 | 2C_1kV75LC_Coax_090 | Signal | SHV to Lemo | J2 | [6 empty] | J2 | Lemo to Lemo | 11 | 7 | CH1 - fmc2-2 |  | 8 | 1: EBS8003 | Ch0 | 11 | 7 | LS2_WC03:HMR_D2823
| LS2_WC03:HMR_D2823 |  | D2823 | 34H000118 | 2C_5kVDC_Coax_035 | HV Bias | [19 empty] | LS2_WC03:HMR_D2823
Halo Monitor Ring,Spare | LS2_WC04:HMR_D2862 | HMR - Unused | D2862 | 34H000033 | 2C_1kV75LC_Coax_090 | Signal | [19 empty] | LS2_WC04:HMR_D2862
| LS2_WC04:HMR_D2862 |  | D2862 | 34H000121 | 2C_5kVDC_Coax_035 | HV Bias
Halo Monitor Ring,Spare | LS2_WC05:HMR_D2902 | HMR - Unused | D2902 | 34H000034 | 2C_1kV75LC_Coax_090 | Signal
| LS2_WC05:HMR_D2902 |  | D2902 | 34H000122 | 2C_5kVDC_Coax_035 | HV Bias
Halo Monitor Ring | LS2_WC06:HMR_D2942 | HMR | D2942 | 34H000035 | 2C_1kV75LC_Coax_090 | Signal | SHV to Lemo | J3 | [6 empty] | J3 | Lemo to Lemo | 11 | 7 | CH2 - fmc2-3 |  | 8 | 1: EBS8003 | Ch0
| LS2_WC06:HMR_D2942 |  | D2942 | 34H000120 | 2C_5kVDC_Coax_035 | HV Bias
Halo Monitor Ring,Spare | LS2_WC07:HMR_D2982 | HMR - Unused | D2982 | 34H000036 | 2C_1kV75LC_Coax_090 | Signal
| LS2_WC07:HMR_D2982 |  | D2982 | 34H000119 | 2C_5kVDC_Coax_035 | HV Bias
Ground Plugs
Grounding Plug |  | ground |  | - | Lemo650 Plug | Ground | Pico-8 | [8 empty] | Ground Plug | 11 | 7 | CH7 - fmc1-4
Lemo to Lemo Panel
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 1_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 8 | 2: EBS8003 | Ch0
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 1_OUT Port | [7 empty] | Lemo650/Lemo650 | 11 | 7 | CH3 - fmc2-4

## LS2_N1708
LS2_N1708
Long Run Cable | [6 empty] | Inter-Rack Connections | [9 empty] | Terminating Device
Diagnostics Device | Lattice Beamline Location | Device Type | CableDB Location | Cable Number | Cable Type | Function | Termination Device | Port Number In | Pin Number | CPC pinout | CPC pinout | CPC pinout | Signal Pin | Shield Pin | Port Number Out | Patch Cable | MTCA Chassis | MTCA Slot | PICO Port | FGPDB Port | Bias Supply Chassis | Bias Supply Slot | Bias Supply Channel |  |  |  | search for row | device name | D number
BPMs | [23 empty] | PICO | CH[0-7] | Device
BPM | LS2_WC10:BPM_D3100 | BPM | D3100 | 34E000441 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 12 | 6 |  | 1 |  |  |  | 12 | 6 | LS2_WC10:BPM_D3100 | 1985 | LS2_WC10:BPM_D3100 | 3100
BPM | LS2_WC10:BPM_D3100 | BPM | D3100 | 34E000442 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 12 | 6 |  | 2 |  |  |  | 12 | 6 | LS2_WC10:BPM_D3100 | 1985 | LS2_WC10:BPM_D3100 | 3100
BPM | LS2_WC10:BPM_D3100 | BPM | D3100 | 34E000443 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 12 | 6 |  | 3 |  |  |  | 12 | 6 | LS2_WC10:BPM_D3100 | 1985 | LS2_WC10:BPM_D3100 | 3100
BPM | LS2_WC10:BPM_D3100 | BPM | D3100 | 34E000444 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 12 | 6 |  | 4 |  |  |  | 12 | 6 | LS2_WC10:BPM_D3100 | 1985 | LS2_WC10:BPM_D3100 | 3100
BPM | LS2_WC10:BPM_D3100 | BPM | D3100 | 34E000472 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | [17 empty] | LS2_WC10:BPM_D3100 | 1985 | LS2_WC10:BPM_D3100 | 3100
BPM | LS2_WC11:BPM_D3140 | BPM | D3140 | 34E000445 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 12 | 6 |  | 5 |  |  |  | 12 | 6 | LS2_WC11:BPM_D3140 | 2021 | LS2_WC11:BPM_D3140 | 3140
BPM | LS2_WC11:BPM_D3140 | BPM | D3140 | 34E000446 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 12 | 6 |  | 6 |  |  |  | 12 | 6 | LS2_WC11:BPM_D3140 | 2021 | LS2_WC11:BPM_D3140 | 3140
BPM | LS2_WC11:BPM_D3140 | BPM | D3140 | 34E000447 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 12 | 6 |  | 7 |  |  |  | 12 | 6 | LS2_WC11:BPM_D3140 | 2021 | LS2_WC11:BPM_D3140 | 3140
BPM | LS2_WC11:BPM_D3140 | BPM | D3140 | 34E000448 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 12 | 6 |  | 8 |  |  |  | 12 | 6 | LS2_WC11:BPM_D3140 | 2021 | LS2_WC11:BPM_D3140 | 3140
BPM | LS2_WC11:BPM_D3140 | BPM | D3140 | 34E000473 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 12 | 6 |  | Ref |  |  |  | 12 | 6 | LS2_WC11:BPM_D3140 | 2021 | LS2_WC11:BPM_D3140 | 3140
BPM | LS2_WC12:BPM_D3180 | BPM | D3180 | 34E000429 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 12 | 7 |  | 1 |  |  |  | 12 | 7 | LS2_WC12:BPM_D3180 | 2057 | LS2_WC12:BPM_D3180 | 3180
BPM | LS2_WC12:BPM_D3180 | BPM | D3180 | 34E000430 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 12 | 7 |  | 2 |  |  |  | 12 | 7 | LS2_WC12:BPM_D3180 | 2057 | LS2_WC12:BPM_D3180 | 3180
BPM | LS2_WC12:BPM_D3180 | BPM | D3180 | 34E000431 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 12 | 7 |  | 3 |  |  |  | 12 | 7 | LS2_WC12:BPM_D3180 | 2057 | LS2_WC12:BPM_D3180 | 3180
BPM | LS2_WC12:BPM_D3180 | BPM | D3180 | 34E000432 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 12 | 7 |  | 4 |  |  |  | 12 | 7 | LS2_WC12:BPM_D3180 | 2057 | LS2_WC12:BPM_D3180 | 3180
BPM | LS2_WC12:BPM_D3180 | BPM | D3180 | 34E000474 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | [17 empty] | LS2_WC12:BPM_D3180 | 2057 | LS2_WC12:BPM_D3180 | 3180
BPM | LS2_WD01:BPM_D3242 | BPM | D3242 | 34E000449 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 12 | 7 |  | 5 |  |  |  | 12 | 7 | LS2_WD01:BPM_D3242 | 2101 | LS2_WD01:BPM_D3242 | 3242
BPM | LS2_WD01:BPM_D3242 | BPM | D3242 | 34E000450 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 12 | 7 |  | 6 |  |  |  | 12 | 7 | LS2_WD01:BPM_D3242 | 2101 | LS2_WD01:BPM_D3242 | 3242
BPM | LS2_WD01:BPM_D3242 | BPM | D3242 | 34E000451 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 12 | 7 |  | 7 |  |  |  | 12 | 7 | LS2_WD01:BPM_D3242 | 2101 | LS2_WD01:BPM_D3242 | 3242
BPM | LS2_WD01:BPM_D3242 | BPM | D3242 | 34E000452 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 12 | 7 |  | 8 |  |  |  | 12 | 7 | LS2_WD01:BPM_D3242 | 2101 | LS2_WD01:BPM_D3242 | 3242
BPM | LS2_WD01:BPM_D3242 | BPM | D3242 | 34E000475 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 12 | 7 |  | Ref |  |  |  | 12 | 7 | LS2_WD01:BPM_D3242 | 2101 | LS2_WD01:BPM_D3242 | 3242
Neutron Detector Signal
Neutron Detector | LS2_CC08:ND_D2987 | ND | D2987 | 34E000542 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | 1 | [6 empty] | 1 | SMA to Lemo650 | 12 | 5 | CH0 - fmc2-1 |  |  |  |  | 12 | 5 | LS2_CC08:ND_D2987 | #N/A | #N/A | #N/A
Neutron Detector | LS2_CC10:ND_D3075 | ND | D3075 | 34E000543 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | 2 | [6 empty] | 2 | SMA to Lemo650 | 12 | 5 | CH1 - fmc2-2 |  |  |  |  | 12 | 5 | LS2_CC10:ND_D3075 | #N/A | #N/A | #N/A
Neutron Detector | LS2_CC12:ND_D3155 | ND | D3155 | 34E000619 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | 3 | [6 empty] | 3 | SMA to Lemo650 | 12 | 5 | CH2 - fmc2-3 |  |  |  |  | 12 | 5 | LS2_CC12:ND_D3155 | #N/A | #N/A | #N/A
Neutron Detector | LS2_CD02:ND_D3252 | ND | D3252 | 34E000547 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | 4 | [6 empty] | 4 | SMA to Lemo650 | 12 | 5 | CH3 - fmc2-4 |  |  |  |  | 12 | 5 | LS2_CD02:ND_D3252 | #N/A | #N/A | #N/A
Neutron Detector | LS2_CD04:ND_D3376 | ND | D3376 | 34E000549 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | 5 | [6 empty] | 5 | SMA to Lemo650 | 12 | 5 | CH4 - fmc1-1 | [7 empty] | #N/A | #N/A | #N/A
Neutron Detector | LS2_CD06:ND_D3500 | ND | D3500 | 34E000551 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | 6 | [6 empty] | 6 | SMA to Lemo650 | 12 | 5 | CH5 - fmc1-2 | [7 empty] | #N/A | #N/A | #N/A
Neutron Detector | LS2_CD08:ND_D3624 | ND | D3624 | 34E000553 | 2C_75LC_Coax_089 | Signal | N-Type to SMA Panel | 7 | [6 empty] | 7 | SMA to Lemo650 | 12 | 5 | CH6 - fmc1-3 | [7 empty] | #N/A | #N/A | #N/A
Neutron Detector Test signal
Neutron Detector | LS2_CC08:ND_D2987 |  | D2987 | 34E000541 | 2C_75LC_Coax_089 | Test Signal | N-Type to SMA Panel | 1 | [7 empty] | SMA to Ferrule | [9 empty] | LS2_CC08:ND_D2987 | #N/A | #N/A | #N/A
Neutron Detector | LS2_CC10:ND_D3075 |  | D3075 | 34E000544 | 2C_75LC_Coax_089 | Test Signal | N-Type to SMA Panel | 2 | [7 empty] | SMA to Ferrule | [9 empty] | LS2_CC10:ND_D3075 | #N/A | #N/A | #N/A
Neutron Detector | LS2_CC12:ND_D3155 |  | D3155 | 34E000545 | 2C_75LC_Coax_089 | Test Signal | N-Type to SMA Panel | 3 | [7 empty] | SMA to Ferrule | [9 empty] | LS2_CC12:ND_D3155 | #N/A | #N/A | #N/A
Neutron Detector | LS2_CD02:ND_D3252 |  | D3252 | 34E000546 | 2C_75LC_Coax_089 | Test Signal | N-Type to SMA Panel | 4 | [7 empty] | SMA to Ferrule | [9 empty] | LS2_CD02:ND_D3252 | #N/A | #N/A | #N/A
Neutron Detector | LS2_CD04:ND_D3376 |  | D3376 | 34E000548 | 2C_75LC_Coax_089 | Test Signal | N-Type to SMA Panel | 5 | [7 empty] | SMA to Ferrule | [10 empty] | #N/A | #N/A | #N/A
Neutron Detector | LS2_CD06:ND_D3500 |  | D3500 | 34E000550 | 2C_75LC_Coax_089 | Test Signal | N-Type to SMA Panel | 6 | [7 empty] | SMA to Ferrule | [10 empty] | #N/A | #N/A | #N/A
Neutron Detector | LS2_CD08:ND_D3624 |  | D3624 | 34E000552 | 2C_75LC_Coax_089 | Test Signal | N-Type to SMA Panel | 7 | [7 empty] | SMA to Ferrule | [10 empty] | #N/A | #N/A | #N/A
SMA to Ferrule
Neutron Detector Bias
Neutron Detector | LS2_CC08:ND_D2987 |  | D2987 | 34H000067 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 1 | [6 empty] | 1 | SHV to SHV |  |  |  |  | 9 | 2 : EHS8020n | Ch0 |  |  | LS2_CC08:ND_D2987
Neutron Detector | LS2_CC10:ND_D3075 |  | D3075 | 34H000068 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 2 | [6 empty] | 2 | SHV to SHV |  |  |  |  | 9 | 2 : EHS8020n | Ch1 |  |  | LS2_CC10:ND_D3075
Neutron Detector | LS2_CC12:ND_D3155 |  | D3155 | 34H000069 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 3 | [6 empty] | 3 | SHV to SHV |  |  |  |  | 9 | 2 : EHS8020n | Ch2 |  |  | LS2_CC12:ND_D3155
Neutron Detector | LS2_CD02:ND_D3252 |  | D3252 | 34H000070 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 4 | [6 empty] | 4 | SHV to SHV |  |  |  |  | 9 | 2 : EHS8020n | Ch3 |  |  | LS2_CD02:ND_D3252
Neutron Detector | LS2_CD04:ND_D3376 |  | D3376 | 34H000071 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 5 | [6 empty] | 5 | SHV to SHV |  |  |  |  | 9 | 2 : EHS8020n | Ch4 |  |  | LS2_CD04:ND_D3376
Neutron Detector | LS2_CD06:ND_D3500 |  | D3500 | 34H000072 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 6 | [6 empty] | 6 | SHV to SHV |  |  |  |  | 9 | 2 : EHS8020n | Ch5 |  |  | LS2_CD06:ND_D3500
Neutron Detector | LS2_CD08:ND_D3624 |  | D3624 | 34H000073 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV Panel | 7 | [6 empty] | 7 | SHV to SHV |  |  |  |  | 9 | 2 : EHS8020n | Ch6 |  |  | LS2_CD08:ND_D3624
Halo Monitor Rings
Halo Monitor Ring,Spare | LS2_WC08:HMR_D3022 | HMR - Unused | D3022 | 34H000037 | 2C_1kV75LC_Coax_090 | Signal | [19 empty] | LS2_WC08:HMR_D3022
| LS2_WC08:HMR_D3022 |  | D3022 | 34H000117 | 2C_5kVDC_Coax_035 | HV Bias | [19 empty] | LS2_WC08:HMR_D3022
Halo Monitor Ring | LS2_WC09:HMR_D3062 | HMR | D3061 | 34H000039 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J1 | [6 empty] | J1 | SHV to Lemo | 12 | 3 | CH0 - fmc2-1 |  | 9 | 1:EHS8003 | Ch0 | 12 | 3 | LS2_WC09:HMR_D3062
| LS2_WC09:HMR_D3062 |  | D3061 | 34H000132 | 2C_5kVDC_Coax_035 | HV Bias | [19 empty] | LS2_WC09:HMR_D3062
Halo Monitor Ring,Spare | LS2_WC10:HMR_D3101 | HMR - Unused | D3101 | 34H000040 | 2C_1kV75LC_Coax_090 | Signal | [19 empty] | LS2_WC10:HMR_D3101
| LS2_WC10:HMR_D3101 |  | D3101 | 34H000131 | 2C_5kVDC_Coax_035 | HV Bias | [19 empty] | LS2_WC10:HMR_D3101
Halo Monitor Ring,Spare | LS2_WC11:HMR_D3141 | HMR - Unused | D3141 | 34H000041 | 2C_1kV75LC_Coax_090 | Signal | [19 empty] | LS2_WC11:HMR_D3141
| LS2_WC11:HMR_D3141 |  | D3141 | 34H000130 | 2C_5kVDC_Coax_035 | HV Bias
Halo Monitor Ring | LS2_WC12:HMR_D3181 | HMR | D3181 | 34H000042 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J2 | [6 empty] | J2 | SHV to Lemo | 12 | 3 | CH1 - fmc2-2 |  | 9 | 1:EHS8003 | Ch0
| LS2_WC12:HMR_D3181 |  | D3181 | 34H000126 | 2C_5kVDC_Coax_035 | HV Bias
Halo Monitor Ring,Spare | LS2_WD01:HMR_D3243 | HMR - Unused | D3243 | 34H000038 | 2C_1kV75LC_Coax_090 | Signal
| LS2_WD01:HMR_D3243 |  | D3243 | 34H000133 | 2C_5kVDC_Coax_035 | HV Bias
Halo Monitor Ring,Spare | LS2_WD02:HMR_D3305 | HMR - Unused | D3305 | 34H000043 | 2C_1kV75LC_Coax_090 | Signal
| LS2_WD02:HMR_D3305 |  | D3305 | 34H000128 | 2C_5kVDC_Coax_035 | HV Bias
Halo Monitor Ring | LS2_WD03:HMR_D3367 | HMR | D3367 | 34H000044 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J3 | [6 empty] | J3 | SHV to Lemo | 12 | 3 | CH4 - fmc1-1 |  | 9 | 1:EHS8003 | Ch1
| LS2_WD03:HMR_D3367 |  | D3367 | 34H000127 | 2C_5kVDC_Coax_035 | HV Bias
Halo Monitor Ring,Spare | LS2_WD04:HMR_D3429 | HMR - Unused | D3429 | 34H000045 | 2C_1kV75LC_Coax_090 | Signal
| LS2_WD04:HMR_D3429 |  | D3429 | 34H000125 | 2C_5kVDC_Coax_035 | HV Bias
Halo Monitor Ring,Spare | LS2_WD05:HMR_D3491 | HMR - Unused | D3491 | 34H000046 | 2C_1kV75LC_Coax_090 | Signal
| LS2_WD05:HMR_D3491 |  | D3491 | 34H000129 | 2C_5kVDC_Coax_035 | HV Bias
Halo Monitor Ring | LS2_WD06:HMR_D3553 | HMR | D3553 | 34H000047 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J4 | [6 empty] | J4 | SHV to Lemo | 12 | 3 | CH5 - fmc1-2 |  | 9 | 1:EHS8003 | Ch1
| LS2_WD06:HMR_D3553 |  | D3553 | 34H000124 | 2C_5kVDC_Coax_035 | HV Bias
Halo Monitor Ring,Spare | LS2_WD07:HMR_D3615 | HMR - Unused | D3615 | 34H000049 | 2C_1kV75LC_Coax_090 | Signal
| LS2_WD07:HMR_D3615 |  | D3615 | 34H000134 | 2C_5kVDC_Coax_035 | HV Bias
Lemo to Lemo Panel
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 1_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 9 | 2: EBS8003 | Ch0
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 1_OUT Port | [7 empty] | Lemo650/Lemo650 | 12 | 3 | CH3 - fmc2-4
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 2_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 9 | 2: EBS8003 | Ch1
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 2_OUT Port | [7 empty] | Lemo650/Lemo650 | 12 | 3 | CH7 - fmc1-4

## LS2_N2308
LS2_N2308
Long Run Cable | [6 empty] | Inter-Rack Connections | [9 empty] | Terminating Device
Diagnostics Device | Lattice Beamline Location | Device Type | CableDB Location | Cable Number | Cable Type | Function | Termination Device | Port Number In | Pin Number | CPC pinout | CPC pinout | CPC pinout | Signal Pin | Shield Pin | Port Number Out | Patch Cable | MTCA Chassis | MTCA Slot | PICO Port | FGPDB Port | Bias Supply Chassis | Bias Supply Slot | Bias Supply Channel |  |  |  | search for row | device name | D number
BPMs | [23 empty] | PICO | CH[0-7] | Device
BPM | LS2_WD02:BPM_D3304 | BPM | D3304 | 34E000453 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 13 | 3 |  | 1 |  |  |  | 13 | 3 | LS2_WD02:BPM_D3304 | 2145 | LS2_WD02:BPM_D3304 | 3304
BPM | LS2_WD02:BPM_D3304 | BPM | D3304 | 34E000454 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 13 | 3 |  | 2 |  |  |  | 13 | 3 | LS2_WD02:BPM_D3304 | 2145 | LS2_WD02:BPM_D3304 | 3304
BPM | LS2_WD02:BPM_D3304 | BPM | D3304 | 34E000455 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 13 | 3 |  | 3 |  |  |  | 13 | 3 | LS2_WD02:BPM_D3304 | 2145 | LS2_WD02:BPM_D3304 | 3304
BPM | LS2_WD02:BPM_D3304 | BPM | D3304 | 34E000456 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 13 | 3 |  | 4 |  |  |  | 13 | 3 | LS2_WD02:BPM_D3304 | 2145 | LS2_WD02:BPM_D3304 | 3304
BPM | LS2_WD02:BPM_D3304 | BPM | D3304 | 34E000476 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | [17 empty] | LS2_WD02:BPM_D3304 | 2145 | LS2_WD02:BPM_D3304 | 3304
BPM | LS2_WD03:BPM_D3366 | BPM | D3366 | 34E000457 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 13 | 3 |  | 5 |  |  |  | 13 | 3 | LS2_WD03:BPM_D3366 | 2189 | LS2_WD03:BPM_D3366 | 3366
BPM | LS2_WD03:BPM_D3366 | BPM | D3366 | 34E000458 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 13 | 3 |  | 6 |  |  |  | 13 | 3 | LS2_WD03:BPM_D3366 | 2189 | LS2_WD03:BPM_D3366 | 3366
BPM | LS2_WD03:BPM_D3366 | BPM | D3366 | 34E000459 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 13 | 3 |  | 7 |  |  |  | 13 | 3 | LS2_WD03:BPM_D3366 | 2189 | LS2_WD03:BPM_D3366 | 3366
BPM | LS2_WD03:BPM_D3366 | BPM | D3366 | 34E000460 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 13 | 3 |  | 8 |  |  |  | 13 | 3 | LS2_WD03:BPM_D3366 | 2189 | LS2_WD03:BPM_D3366 | 3366
BPM | LS2_WD03:BPM_D3366 | BPM | D3366 | 34E000477 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 13 | 3 |  | Ref |  |  |  | 13 | 3 | LS2_WD03:BPM_D3366 | 2189 | LS2_WD03:BPM_D3366 | 3366
BPM | LS2_WD04:BPM_D3428 | BPM | D3428 | 34E000461 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 13 | 4 |  | 1 |  |  |  | 13 | 4 | LS2_WD04:BPM_D3428 | 2233 | LS2_WD04:BPM_D3428 | 3428
BPM | LS2_WD04:BPM_D3428 | BPM | D3428 | 34E000462 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 13 | 4 |  | 2 |  |  |  | 13 | 4 | LS2_WD04:BPM_D3428 | 2233 | LS2_WD04:BPM_D3428 | 3428
BPM | LS2_WD04:BPM_D3428 | BPM | D3428 | 34E000463 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 13 | 4 |  | 3 |  |  |  | 13 | 4 | LS2_WD04:BPM_D3428 | 2233 | LS2_WD04:BPM_D3428 | 3428
BPM | LS2_WD04:BPM_D3428 | BPM | D3428 | 34E000464 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 13 | 4 |  | 4 |  |  |  | 13 | 4 | LS2_WD04:BPM_D3428 | 2233 | LS2_WD04:BPM_D3428 | 3428
BPM | LS2_WD04:BPM_D3428 | BPM | D3428 | 34E000478 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | [17 empty] | LS2_WD04:BPM_D3428 | 2233 | LS2_WD04:BPM_D3428 | 3428
BPM | LS2_WD05:BPM_D3490 | BPM | D3490 | 34E000465 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 13 | 4 |  | 5 |  |  |  | 13 | 4 | LS2_WD05:BPM_D3490 | 2277 | LS2_WD05:BPM_D3490 | 3490
BPM | LS2_WD05:BPM_D3490 | BPM | D3490 | 34E000466 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 13 | 4 |  | 6 |  |  |  | 13 | 4 | LS2_WD05:BPM_D3490 | 2277 | LS2_WD05:BPM_D3490 | 3490
BPM | LS2_WD05:BPM_D3490 | BPM | D3490 | 34E000467 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 13 | 4 |  | 7 |  |  |  | 13 | 4 | LS2_WD05:BPM_D3490 | 2277 | LS2_WD05:BPM_D3490 | 3490
BPM | LS2_WD05:BPM_D3490 | BPM | D3490 | 34E000468 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 13 | 4 |  | 8 |  |  |  | 13 | 4 | LS2_WD05:BPM_D3490 | 2277 | LS2_WD05:BPM_D3490 | 3490
BPM | LS2_WD05:BPM_D3490 | BPM | D3490 | 34E000479 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 13 | 4 |  | Ref |  |  |  | 13 | 4 | LS2_WD05:BPM_D3490 | 2277 | LS2_WD05:BPM_D3490 | 3490
BPM | LS2_WD06:BPM_D3552 | BPM | D3552 | 34E000485 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 13 | 5 |  | 1 |  |  |  | 13 | 5 | LS2_WD06:BPM_D3552 | 2321 | LS2_WD06:BPM_D3552 | 3552
BPM | LS2_WD06:BPM_D3552 | BPM | D3552 | 34E000486 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 13 | 5 |  | 2 |  |  |  | 13 | 5 | LS2_WD06:BPM_D3552 | 2321 | LS2_WD06:BPM_D3552 | 3552
BPM | LS2_WD06:BPM_D3552 | BPM | D3552 | 34E000487 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 13 | 5 |  | 3 |  |  |  | 13 | 5 | LS2_WD06:BPM_D3552 | 2321 | LS2_WD06:BPM_D3552 | 3552
BPM | LS2_WD06:BPM_D3552 | BPM | D3552 | 34E000488 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 13 | 5 |  | 4 |  |  |  | 13 | 5 | LS2_WD06:BPM_D3552 | 2321 | LS2_WD06:BPM_D3552 | 3552
BPM | LS2_WD06:BPM_D3552 | BPM | D3552 | 34E000484 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | [17 empty] | LS2_WD06:BPM_D3552 | 2321 | LS2_WD06:BPM_D3552 | 3552
BPM | LS2_WD07:BPM_D3614 | BPM | D3614 | 34E000480 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 13 | 5 |  | 5 |  |  |  | 13 | 5 | LS2_WD07:BPM_D3614 | 2365 | LS2_WD07:BPM_D3614 | 3614
BPM | LS2_WD07:BPM_D3614 | BPM | D3614 | 34E000481 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 13 | 5 |  | 6 | [6 empty] | 2365 | LS2_WD07:BPM_D3614 | 3614
BPM | LS2_WD07:BPM_D3614 | BPM | D3614 | 34E000482 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 13 | 5 |  | 7 | [6 empty] | 2365 | LS2_WD07:BPM_D3614 | 3614
BPM | LS2_WD07:BPM_D3614 | BPM | D3614 | 34E000483 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 13 | 5 |  | 8 | [6 empty] | 2365 | LS2_WD07:BPM_D3614 | 3614
BPM | LS2_WD07:BPM_D3614 | BPM | D3614 | 34E000497 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 13 | 5 |  | Ref |  |  |  | 13 | 5 | LS2_WD07:BPM_D3614 | 2365 | LS2_WD07:BPM_D3614 | 3614
BPM | LS2_WD08:BPM_D3676 | BPM | D3676 | 34E000489 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 13 | 6 |  | 1 |  |  |  | 13 | 6 | LS2_WD08:BPM_D3676 | 2409 | LS2_WD08:BPM_D3676 | 3676
BPM | LS2_WD08:BPM_D3676 | BPM | D3676 | 34E000490 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 13 | 6 |  | 2 |  |  |  | 13 | 6 | LS2_WD08:BPM_D3676 | 2409 | LS2_WD08:BPM_D3676 | 3676
BPM | LS2_WD08:BPM_D3676 | BPM | D3676 | 34E000491 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 13 | 6 |  | 3 |  |  |  | 13 | 6 | LS2_WD08:BPM_D3676 | 2409 | LS2_WD08:BPM_D3676 | 3676
BPM | LS2_WD08:BPM_D3676 | BPM | D3676 | 34E000492 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 13 | 6 |  | 4 |  |  |  | 13 | 6 | LS2_WD08:BPM_D3676 | 2409 | LS2_WD08:BPM_D3676 | 3676
BPM | LS2_WD08:BPM_D3676 | BPM | D3676 | 34E000498 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | [17 empty] | LS2_WD08:BPM_D3676 | 2409 | LS2_WD08:BPM_D3676 | 3676
BPM | LS2_WD09:BPM_D3738 | BPM | D3738 | 34E000493 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 13 | 6 |  | 5 | [6 empty] | 2453 | LS2_WD09:BPM_D3738 | 3738
BPM | LS2_WD09:BPM_D3738 | BPM | D3738 | 34E000494 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 13 | 6 |  | 6 | [6 empty] | 2453 | LS2_WD09:BPM_D3738 | 3738
BPM | LS2_WD09:BPM_D3738 | BPM | D3738 | 34E000495 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 13 | 6 |  | 7 | [6 empty] | 2453 | LS2_WD09:BPM_D3738 | 3738
BPM | LS2_WD09:BPM_D3738 | BPM | D3738 | 34E000496 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 13 | 6 |  | 8 | [6 empty] | 2453 | LS2_WD09:BPM_D3738 | 3738
BPM | LS2_WD09:BPM_D3738 | BPM | D3738 | 34E000499 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 13 | 6 |  | Ref |  |  |  | 13 | 6 | LS2_WD09:BPM_D3738 | 2453 | LS2_WD09:BPM_D3738 | 3738

## LS2_N4202
LS2_N4202
Long Run Cable | [6 empty] | Inter-Rack Connections | [9 empty] | Terminating Device
Diagnostics Device | Lattice Beamline Location | Device Type | CableDB Location | Cable Number | Cable Type | Function | Termination Device | Port Number In | Pin Number | CPC pinout | CPC pinout | CPC pinout | Signal Pin | Shield Pin | Port Number Out | Patch Cable | MTCA Chassis | MTCA Slot | PICO Port | FGPDB Port | Bias Supply Chassis | Bias Supply Slot | Bias Supply Channel |  |  |  | search for row | device name | D number
BPMs | [23 empty] | PICO | CH[0-7] | Device
BPM | LS2_WD10:BPM_D3800 | BPM | D3800 | 34E000505 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 14 | 3 |  | 1 |  |  |  | 14 | 3 | LS2_WD10:BPM_D3800 | 2497 | LS2_WD10:BPM_D3800 | 3800
BPM | LS2_WD10:BPM_D3800 | BPM | D3800 | 34E000506 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 14 | 3 |  | 2 |  |  |  | 14 | 3 | LS2_WD10:BPM_D3800 | 2497 | LS2_WD10:BPM_D3800 | 3800
BPM | LS2_WD10:BPM_D3800 | BPM | D3800 | 34E000507 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 14 | 3 |  | 3 |  |  |  | 14 | 3 | LS2_WD10:BPM_D3800 | 2497 | LS2_WD10:BPM_D3800 | 3800
BPM | LS2_WD10:BPM_D3800 | BPM | D3800 | 34E000508 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 14 | 3 |  | 4 |  |  |  | 14 | 3 | LS2_WD10:BPM_D3800 | 2497 | LS2_WD10:BPM_D3800 | 3800
BPM | LS2_WD10:BPM_D3800 | BPM | D3800 | 34E000514 | 2C_1/4in_Hardline_036 | Ref | No terminiation | Ref | [6 empty] | Ref | SMA to SMA | 14 | 3 |  | Ref |  |  |  | 14 | 3 | LS2_WD10:BPM_D3800 | 2497 | LS2_WD10:BPM_D3800 | 3800
BPM | LS2_WD11:BPM_D3862 | BPM | D3862 | 34E000501 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 14 | 3 |  | 5 |  |  |  | 14 | 3 | LS2_WD11:BPM_D3862 | 2541 | LS2_WD11:BPM_D3862 | 3862
BPM | LS2_WD11:BPM_D3862 | BPM | D3862 | 34E000502 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 14 | 3 |  | 6 |  |  |  | 14 | 3 | LS2_WD11:BPM_D3862 | 2541 | LS2_WD11:BPM_D3862 | 3862
BPM | LS2_WD11:BPM_D3862 | BPM | D3862 | 34E000503 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 14 | 3 |  | 7 |  |  |  | 14 | 3 | LS2_WD11:BPM_D3862 | 2541 | LS2_WD11:BPM_D3862 | 3862
BPM | LS2_WD11:BPM_D3862 | BPM | D3862 | 34E000504 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 14 | 3 |  | 8 |  |  |  | 14 | 3 | LS2_WD11:BPM_D3862 | 2541 | LS2_WD11:BPM_D3862 | 3862
BPM | LS2_WD11:BPM_D3862 | BPM | D3862 | 34E000500 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | No term | [17 empty] | LS2_WD11:BPM_D3862 | 2541 | LS2_WD11:BPM_D3862 | 3862
BPM | LS2_WD12:BPM_D3924 | BPM | D3924 | 34E000509 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 14 | 4 |  | 1 |  |  |  | 14 | 4 | LS2_WD12:BPM_D3924 | 2585 | LS2_WD12:BPM_D3924 | 3924
BPM | LS2_WD12:BPM_D3924 | BPM | D3924 | 34E000510 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 14 | 4 |  | 2 |  |  |  | 14 | 4 | LS2_WD12:BPM_D3924 | 2585 | LS2_WD12:BPM_D3924 | 3924
BPM | LS2_WD12:BPM_D3924 | BPM | D3924 | 34E000511 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 14 | 4 |  | 3 |  |  |  | 14 | 4 | LS2_WD12:BPM_D3924 | 2585 | LS2_WD12:BPM_D3924 | 3924
BPM | LS2_WD12:BPM_D3924 | BPM | D3924 | 34E000512 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 14 | 4 |  | 4 |  |  |  | 14 | 4 | LS2_WD12:BPM_D3924 | 2585 | LS2_WD12:BPM_D3924 | 3924
BPM | LS2_WD12:BPM_D3924 | BPM | D3924 | 34E000513 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | [17 empty] | LS2_WD12:BPM_D3924 | 2585 | LS2_WD12:BPM_D3924 | 3924
BPM | FS2_BTS:BPM_D3943 | BPM | D3947 | 44E000192 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 14 | 4 |  | 5 |  |  |  | 14 | 4 | FS2_BTS:BPM_D3943 | 2605 | FS2_BTS:BPM_D3943 | 3943
BPM | FS2_BTS:BPM_D3943 | BPM | D3947 | 44E000193 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 14 | 4 |  | 6 |  |  |  | 14 | 4 | FS2_BTS:BPM_D3943 | 2605 | FS2_BTS:BPM_D3943 | 3943
BPM | FS2_BTS:BPM_D3943 | BPM | D3947 | 44E000194 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 14 | 4 |  | 7 |  |  |  | 14 | 4 | FS2_BTS:BPM_D3943 | 2605 | FS2_BTS:BPM_D3943 | 3943
BPM | FS2_BTS:BPM_D3943 | BPM | D3947 | 44E000195 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 14 | 4 |  | 8 |  |  |  | 14 | 4 | FS2_BTS:BPM_D3943 | 2605 | FS2_BTS:BPM_D3943 | 3943
BPM | FS2_BTS:BPM_D3943 | BPM | D3947 | 44E000184 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 14 | 4 |  | Ref |  |  |  | 14 | 4 | FS2_BTS:BPM_D3943 | 2605 | FS2_BTS:BPM_D3943 | 3943
BPM | FS2_BTS:BPM_D3958 | BPM | D3958 | 44E000179 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 14 | 5 |  | 1 | [6 empty] | 2618 | FS2_BTS:BPM_D3958 | 3958
BPM | FS2_BTS:BPM_D3958 | BPM | D3958 | 44E000185 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 14 | 5 |  | 2 | [6 empty] | 2618 | FS2_BTS:BPM_D3958 | 3958
BPM | FS2_BTS:BPM_D3958 | BPM | D3958 | 44E000186 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 14 | 5 |  | 3 | [6 empty] | 2618 | FS2_BTS:BPM_D3958 | 3958
BPM | FS2_BTS:BPM_D3958 | BPM | D3958 | 44E000187 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 14 | 5 |  | 4 | [6 empty] | 2618 | FS2_BTS:BPM_D3958 | 3958
BPM | FS2_BTS:BPM_D3958 | BPM | D3958 | 44E000200 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | [18 empty] | 2618 | FS2_BTS:BPM_D3958 | 3958
BPM | FS2_BTS:BPM_D4006 | BPM | D4012 | 44E000204 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 14 | 5 |  | 5 | [6 empty] | 3577 | FS2_BTS:BPM_D4006 | 4006
BPM | FS2_BTS:BPM_D4006 | BPM | D4012 | 44E000205 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 14 | 5 |  | 6 | [6 empty] | 3577 | FS2_BTS:BPM_D4006 | 4006
BPM | FS2_BTS:BPM_D4006 | BPM | D4012 | 44E000206 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 14 | 5 |  | 7 | [6 empty] | 3577 | FS2_BTS:BPM_D4006 | 4006
BPM | FS2_BTS:BPM_D4006 | BPM | D4012 | 44E000207 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 14 | 5 |  | 8 | [6 empty] | 3577 | FS2_BTS:BPM_D4006 | 4006
BPM | FS2_BTS:BPM_D4006 | BPM | D4012 | 44E000208 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 14 | 5 |  | Ref | [6 empty] | 3577 | FS2_BTS:BPM_D4006 | 4006
BPM | FS2_BBS:BPM_D4019 | BPM | D4017 | 44E000196 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 14 | 6 |  | 1 | [6 empty] | 2647 | FS2_BBS:BPM_D4019 | 4019
BPM | FS2_BBS:BPM_D4019 | BPM | D4017 | 44E000197 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 14 | 6 |  | 2 | [6 empty] | 2647 | FS2_BBS:BPM_D4019 | 4019
BPM | FS2_BBS:BPM_D4019 | BPM | D4017 | 44E000198 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 14 | 6 |  | 3 | [6 empty] | 2647 | FS2_BBS:BPM_D4019 | 4019
BPM | FS2_BBS:BPM_D4019 | BPM | D4017 | 44E000199 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 14 | 6 |  | 4 | [6 empty] | 2647 | FS2_BBS:BPM_D4019 | 4019
BPM | FS2_BBS:BPM_D4019 | BPM | D4017 | 44E000201 | 2C_1/4in_Hardline_036 | Ref | No terminiation | Ref | [6 empty] | Ref | SMA to SMA | 14 | 6 |  | Ref | [6 empty] | 2647 | FS2_BBS:BPM_D4019 | 4019
BPM | FS2_BBS:BPM_D4054 | BPM | D4054 | 44E000188 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 14 | 6 |  | 5 | [6 empty] | 2655 | FS2_BBS:BPM_D4054 | 4054
BPM | FS2_BBS:BPM_D4054 | BPM | D4054 | 44E000189 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 14 | 6 |  | 6 | [6 empty] | 2655 | FS2_BBS:BPM_D4054 | 4054
BPM | FS2_BBS:BPM_D4054 | BPM | D4054 | 44E000190 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 14 | 6 |  | 7 | [6 empty] | 2655 | FS2_BBS:BPM_D4054 | 4054
BPM | FS2_BBS:BPM_D4054 | BPM | D4054 | 44E000191 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 14 | 6 |  | 8 | [6 empty] | 2655 | FS2_BBS:BPM_D4054 | 4054
BPM | FS2_BBS:BPM_D4054 | BPM | D4054 | 44E000202 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | No term | [18 empty] | 2655 | FS2_BBS:BPM_D4054 | 4054
Neutron Detector Signal
Neutron Detector | LS2_CD10:ND_D3748 | ND | D3748 | 34E000555 | 2C_75LC_Coax_089 | Signal | N-Type to Lemo Box | 1 | [6 empty] | 1 | SMA to Lemo650 | 14 | 8 | CH0 - fmc2-1 |  |  |  |  | 14 | 8 | LS2_CD10:ND_D3748 | #N/A | #N/A | #N/A
Neutron Detector | LS2_CD12:ND_D3872 | ND | D3872 | 34E000557 | 2C_75LC_Coax_089 | Signal | N-Type to Lemo Box | 2 | [6 empty] | 2 | SMA to Lemo650 | 14 | 8 | CH1 - fmc2-2 |  |  |  |  | 14 | 8 | LS2_CD12:ND_D3872 | #N/A | #N/A | #N/A
Neutron Detector Test signal
Neutron Detector | LS2_CD10:BLM_D3748 |  | D3748 | 34E000554 | 2C_75LC_Coax_089 | Test Signal | N-Type to SMA Panel | 1 | [7 empty] | SMA to Ferrule | [9 empty] | LS2_CD10:BLM_D3748 | 2462 | LS2_CD10:BLM_D3748 | 3748
Neutron Detector | LS2_CD12:BLM_D3872 |  | D3872 | 34E000556 | 2C_75LC_Coax_089 | Test Signal | N-Type to SMA Panel | 2 | [7 empty] | SMA to Ferrule | [9 empty] | LS2_CD12:BLM_D3872 | 2550 | LS2_CD12:BLM_D3872 | 3872
Neutron Detector Bias
Neutron Detector | LS2_CD10:ND_D3748 |  | D3748 | 34H000074 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 1 | [6 empty] | 1 | SHV to SHV |  |  |  |  | 10 | 2 : EHS8020n | Ch0 |  |  | LS2_CD10:ND_D3748
Neutron Detector | LS2_CD12:ND_D3872 |  | D3872 | 34H000075 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 2 | [6 empty] | 2 | SHV to SHV |  |  |  |  | 10 | 2 : EHS8020n | Ch1 |  |  | LS2_CD12:ND_D3872
Halo Monitor Ring
Halo Monitor Ring,Spare | FS2_BTS:HMR_D3959 | HMR - Unused | D3959 | 44H000012 | 2C_1kV75LC_Coax_090 | Signal
| FS2_BTS:HMR_D3959 |  | D3959 | 44H000066 | 2C_5kVDC_Coax_035 | HV Bias
Halo Monitor Ring,Spare | LS2_WD08:HMR_D3677 | HMR - Unused | D3677 | 34H000050 | 2C_1kV75LC_Coax_090 | Signal
| LS2_WD08:HMR_D3677 |  | D3677 | 34H000137 | 2C_5kVDC_Coax_035 | HV Bias
Halo Monitor Ring | LS2_WD09:HMR_D3739 | HMR | D3739 | 34H000052 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J1 | [6 empty] | J1 | SHV to Lemo | 14 | 7 | CH0 - fmc2-1 |  | 10 | 1:EHS8003 | Ch0
| LS2_WD09:HMR_D3739 |  | D3739 | 34H000136 | 2C_5kVDC_Coax_035 | HV Bias
Halo Monitor Ring,Spare | LS2_WD10:HMR_D3801 | HMR - Unused | D3801 | 34H000051 | 2C_1kV75LC_Coax_090 | Signal
| LS2_WD10:HMR_D3801 |  | D3801 | 34H000138 | 2C_5kVDC_Coax_035 | HV Bias
Halo Monitor Ring,Spare | LS2_WD11:HMR_D3863 | HMR - Unused | D3863 | 34H000048 | 2C_1kV75LC_Coax_090 | Signal
| LS2_WD11:HMR_D3863 |  | D3863 | 34H000139 | 2C_5kVDC_Coax_035 | HV Bias
Halo Monitor Ring | LS2_WD12:HMR_D3925 | HMR | D3925 | 34H000053 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV Panel | J2 | [6 empty] | J2 | SHV to Lemo | 14 | 7 | CH1 - fmc2-2 |  | 10 | 1:EHS8003 | Ch0
| LS2_WD12:HMR_D3925 |  | D3925 | 34H000135 | 2C_5kVDC_Coax_035 | HV Bias
Lemo to Lemo Panel
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 1_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 10 | 2: EBS8003 | Ch0
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 1_OUT Port | [7 empty] | Lemo650/Lemo650 | 14 | 7 | CH3 - fmc2-4
Note on B5: Rodriguez, Pedro:
Note on E5: Rodriguez, Pedro:

## FS2_N0108
FS2_N0108
Long Run Cable | [6 empty] | Inter-Rack Connections | [9 empty] | Terminating Device
Diagnostics Device | Lattice Beamline Location | Device Type | CableDB Location | Cable Number | Cable Type | Function | Termination Device | Port Number In | Pin Number | CPC pinout | CPC pinout | CPC pinout | Signal Pin | Shield Pin | Port Number Out | Patch Cable | MTCA Chassis | MTCA Slot | PICO Port | FGPDB Port | Bias Supply Chassis | Bias Supply Slot | Bias Supply Channel | search for row | device name | D number
Panel 1
BPM | FS2_BBS:BPM_D4087 | BPM | D4088 | 44E000180 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 15 | 9 |  | 1 |  |  |  | 2661 | FS2_BBS:BPM_D4087 | 4087
BPM | FS2_BBS:BPM_D4087 | BPM | D4088 | 44E000181 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 15 | 9 |  | 2 |  |  |  | 2661 | FS2_BBS:BPM_D4087 | 4087
BPM | FS2_BBS:BPM_D4087 | BPM | D4088 | 44E000182 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 15 | 9 |  | 3 |  |  |  | 2661 | FS2_BBS:BPM_D4087 | 4087
BPM | FS2_BBS:BPM_D4087 | BPM | D4088 | 44E000183 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 15 | 9 |  | 4 |  |  |  | 2661 | FS2_BBS:BPM_D4087 | 4087
BPM | FS2_BBS:RFS_D4087 | BPM | D4088 | 44E000203 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term
BPM | FS2_BMS:BPM_D4142 | BPM | D4142 | 44E000083 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 15 | 9 |  | 5 |  |  |  | 2682 | FS2_BMS:BPM_D4142 | 4142
BPM | FS2_BMS:BPM_D4142 | BPM | D4142 | 44E000084 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 15 | 9 |  | 6 |  |  |  | 2682 | FS2_BMS:BPM_D4142 | 4142
BPM | FS2_BMS:BPM_D4142 | BPM | D4142 | 44E000085 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 15 | 9 |  | 7 |  |  |  | 2682 | FS2_BMS:BPM_D4142 | 4142
BPM | FS2_BMS:BPM_D4142 | BPM | D4142 | 44E000086 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 15 | 9 |  | 8 |  |  |  | 2682 | FS2_BMS:BPM_D4142 | 4142
BPM | FS2_BMS:RFS_D4142 | BPM | D4142 | 44E000078 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 15 | 9 |  | Ref
BPM | FS2_BMS:BPM_D4164 | BPM | D4164 | 44E000079 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 15 | 10 |  | 1 |  |  |  | 2698 | FS2_BMS:BPM_D4164 | 4164
BPM | FS2_BMS:BPM_D4164 | BPM | D4164 | 44E000080 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 15 | 10 |  | 2 |  |  |  | 2698 | FS2_BMS:BPM_D4164 | 4164
BPM | FS2_BMS:BPM_D4164 | BPM | D4164 | 44E000081 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 15 | 10 |  | 3 |  |  |  | 2698 | FS2_BMS:BPM_D4164 | 4164
BPM | FS2_BMS:BPM_D4164 | BPM | D4164 | 44E000082 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 15 | 10 |  | 4 |  |  |  | 2698 | FS2_BMS:BPM_D4164 | 4164
BPM | FS2_BMS:RFS_D4164 | BPM | D4164 | 44E000087 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term
BPM | FS2_BMS:BPM_D4177 | BPM | D4178 | 44E000074 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 15 | 10 |  | 5 |  |  |  | 2711 | FS2_BMS:BPM_D4177 | 4177
BPM | FS2_BMS:BPM_D4177 | BPM | D4178 | 44E000075 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 15 | 10 |  | 6 |  |  |  | 2711 | FS2_BMS:BPM_D4177 | 4177
BPM | FS2_BMS:BPM_D4177 | BPM | D4178 | 44E000076 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 15 | 10 |  | 7 |  |  |  | 2711 | FS2_BMS:BPM_D4177 | 4177
BPM | FS2_BMS:BPM_D4177 | BPM | D4178 | 44E000077 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 15 | 10 |  | 8 |  |  |  | 2711 | FS2_BMS:BPM_D4177 | 4177
BPM | FS2_BMS:RFS_D4177 | BPM | D4178 | 44E000088 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 15 | 10 |  | Ref
Ion Chamber | FS2_BBS:IC_D3979 | IC | D3979 | 44E000231 | 2C_1kV75LC_Coax_090 | Signal |  | J1 | [6 empty] | J1 | Lemo to Lemo | 15 | 7 | CH0 - fmc2-1 |  |  |  |  | #N/A | #N/A | #N/A
Ion Chamber | FS2_BBS:IC_D3979 |  | D3979 | 44E000015 | 2C_5kVDC_Coax_035 | HV Bias | SHV | 1 | [6 empty] | 1 | SHV to SHV |  |  |  |  | 11 | 2: EHS8020n | Ch0 | #N/A | #N/A | #N/A
Ion Chamber | FS2_BBS:IC_D4018 | IC | D4018 | 44E000232 | 2C_1kV75LC_Coax_090 | Signal |  | J2 | [6 empty] | J2 | Lemo to Lemo | 15 | 7 | CH1 - fmc2-2 |  |  |  |  | #N/A | #N/A | #N/A
Ion Chamber | FS2_BBS:IC_D4018 |  | D4018 | 44E000017 | 2C_5kVDC_Coax_035 | HV Bias | SHV | 2 | [6 empty] | 2 | SHV to SHV |  |  |  |  | 11 | 2: EHS8020n | Ch1 | #N/A | #N/A | #N/A
Ion Chamber | FS2_BBS:IC_D4055 | IC | D4055 | 44E000235 | 2C_1kV75LC_Coax_090 | Signal |  | J3 | [6 empty] | J3 | Lemo to Lemo | 15 | 7 | CH2 - fmc2-3 |  |  |  |  | #N/A | #N/A | #N/A
Ion Chamber | FS2_BBS:IC_D4055 |  | D4055 | 44E000019 | 2C_5kVDC_Coax_035 | HV Bias | SHV | 3 | [6 empty] | 3 | SHV to SHV |  |  |  |  | 11 | 2: EHS8020n | Ch2 | #N/A | #N/A | #N/A
Ion Chamber | FS2_BBS:IC_D4102 | IC | D4102 | 44E000236 | 2C_1kV75LC_Coax_090 | Signal |  | J4 | [6 empty] | J4 | Lemo to Lemo | 15 | 7 | CH3 - fmc2-4 |  |  |  |  | #N/A | #N/A | #N/A
Ion Chamber | FS2_BBS:IC_D4102 |  | D4102 | 44E000020 | 2C_5kVDC_Coax_035 | HV Bias | SHV | 4 | [6 empty] | 4 | SHV to SHV |  |  |  |  | 11 | 2: EHS8020n | Ch3 | #N/A | #N/A | #N/A
Ion Chamber | FS2_BTS:IC_D3940 | IC | D3940 | 44E000230 | 2C_1kV75LC_Coax_090 | Signal |  | J5 | [6 empty] | J5 | Lemo to Lemo | 15 | 7 | CH4 - fmc1-1 |  |  |  |  | #N/A | #N/A | #N/A
Ion Chamber | FS2_BTS:IC_D3940 |  | D3940 | 44E000014 | 2C_5kVDC_Coax_035 | HV Bias | SHV | 5 | [6 empty] | 5 | SHV to SHV |  |  |  |  | 11 | 2: EHS8020n | Ch4 | #N/A | #N/A | #N/A
Ion Chamber | FS2_BTS:IC_D4018 | IC | D4018 | 44E000233 | 2C_1kV75LC_Coax_090 | Signal |  | J6 | [6 empty] | J6 | Lemo to Lemo | 15 | 7 | CH5 - fmc1-2 |  |  |  |  | #N/A | #N/A | #N/A
Ion Chamber | FS2_BTS:IC_D4018 |  | D4018 | 44E000016 | 2C_5kVDC_Coax_035 | HV Bias | SHV | 6 | [6 empty] | 6 | SHV to SHV |  |  |  |  | 11 | 2: EHS8020n | Ch5 | #N/A | #N/A | #N/A
Ion Chamber | FS2_BTS:IC_D4019 | IC | D4019 | 44E000234 | 2C_1kV75LC_Coax_090 | Signal |  | J7 | [6 empty] | J7 | Lemo to Lemo | 15 | 7 | CH6 - fmc1-3 |  |  |  |  | #N/A | #N/A | #N/A
Ion Chamber | FS2_BTS:IC_D4019 |  | D4019 | 44E000018 | 2C_5kVDC_Coax_035 | HV Bias | SHV | 7 | [6 empty] | 7 | SHV to SHV |  |  |  |  | 11 | 2: EHS8020n | Ch6 | #N/A | #N/A | #N/A
PM 6" | FE_LEBT:PM_D0945 | PM | D0945 | 24F000107 | 4C_22AWG_Multi_006 | C Signal | CPC to LEMO650 Box | CPC-1 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | Lemo-1A | LEMO650/LEMO650 | 15 | 3 | CH2 - fmc2-3 |  | 11 | 1: EBS8005 | Ch0 | 190 | FE_LEBT:PM_D0945 | 945
PM 12" | FE_LEBT:PM_D0945 | PM | D0945 | 24F000110 | 6C_22AWG_Multi_010 | A Signal | CPC to LEMO650 Box | CPC-1 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | Lemo-1B | LEMO650/LEMO650 | 15 | 3 | CH0 - fmc2-1 |  | 11 | 1: EBS8005 | Ch0 | 190 | FE_LEBT:PM_D0945 | 945
PM 12" | FE_LEBT:PM_D0945 | PM | D0945 | 24F000110 | 6C_22AWG_Multi_010 | B Signal | CPC to LEMO650 Box | CPC-1 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | Lemo-1C | LEMO650/LEMO650 | 15 | 3 | CH1 - fmc2-2 |  | 11 | 1: EBS8005 | Ch0 | 190 | FE_LEBT:PM_D0945 | 945
PM 6" | FE_LEBT:PM_D0961 | PM | D0962 | 24F000102 | 4C_22AWG_Multi_006 | C Signal | CPC to LEMO650 Box | CPC-2 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | Lemo-2A | LEMO650/LEMO650 | 15 | 3 | CH6 - fmc1-3 |  | 11 | 1: EBS8005 | Ch1 | 202 | FE_LEBT:PM_D0961 | 961
PM 12" | FE_LEBT:PM_D0961 | PM | D0962 | 24F000053 | 6C_22AWG_Multi_010 | A Signal | CPC to LEMO650 Box | CPC-2 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | Lemo-2B | LEMO650/LEMO650 | 15 | 3 | CH4 - fmc1-1 |  | 11 | 1: EBS8005 | Ch1 | 202 | FE_LEBT:PM_D0961 | 961
PM 12" | FE_LEBT:PM_D0961 | PM | D0962 | 24F000053 | 6C_22AWG_Multi_010 | B Signal | CPC to LEMO650 Box | CPC-2 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | Lemo-2C | LEMO650/LEMO650 | 15 | 3 | CH5 - fmc1-2 |  | 11 | 1: EBS8005 | Ch1 | 202 | FE_LEBT:PM_D0961 | 961
PM 6" | FE_LEBT:PM_D0972 | PM | D0976 | 24F000094 | 4C_22AWG_Multi_006 | C Signal | CPC to LEMO650 Box | CPC-3 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | Lemo-3A | LEMO650/LEMO650 | 15 | 4 | CH2 - fmc2-3 |  | 11 | 1: EBS8005 | Ch2 | 209 | FE_LEBT:PM_D0972 | 972
PM 12" | FE_LEBT:PM_D0972 | PM | D0976 | 24F000098 | 6C_22AWG_Multi_010 | A Signal | CPC to LEMO650 Box | CPC-3 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | Lemo-3B | LEMO650/LEMO650 | 15 | 4 | CH0 - fmc2-1 |  | 11 | 1: EBS8005 | Ch2 | 209 | FE_LEBT:PM_D0972 | 972
PM 12" | FE_LEBT:PM_D0972 | PM | D0976 | 24F000098 | 6C_22AWG_Multi_010 | B Signal | CPC to LEMO650 Box | CPC-3 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | Lemo-3C | LEMO650/LEMO650 | 15 | 4 | CH1 - fmc2-2 |  | 11 | 1: EBS8005 | Ch2 | 209 | FE_LEBT:PM_D0972 | 972
PM 8" | FE_LEBT:PM_D0986 | PM | D0986 | 24F000091 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-4 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | Lemo-4A | LEMO650/LEMO650 | 15 | 4 | CH6 - fmc1-3 |  | 11 | 1: EBS8005 | Ch3 | 226 | FE_LEBT:PM_D0986 | 986
PM 8" | FE_LEBT:PM_D0986 | PM | D0986 | 24F000091 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-4 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | Lemo-4B | LEMO650/LEMO650 | 15 | 4 | CH5 - fmc1-2 |  | 11 | 1: EBS8005 | Ch3 | 226 | FE_LEBT:PM_D0986 | 986
PM 8" | FE_LEBT:PM_D0986 | PM | D0986 | 24F000091 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-4 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | Lemo-4C | LEMO650/LEMO650 | 15 | 4 | CH4 - fmc1-1 |  | 11 | 1: EBS8005 | Ch3 | 226 | FE_LEBT:PM_D0986 | 986
PM 6" | FS2_BBS:PM_D4051 | PM | D4051 | 44F000013 | 4C_22AWG_Multi_006 | C Signal | CPC to LEMO650 Box | CPC-5 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | Lemo-5A | LEMO650/LEMO650 | 15 | 5 | CH2 - fmc2-3 |  | 11 | 1: EBS8005 | Ch4 | 2651 | FS2_BBS:PM_D4051 | 4051
PM 12" | FS2_BBS:PM_D4051 | PM | D4051 | 44F000012 | 6C_22AWG_Multi_010 | A Signal | CPC to LEMO650 Box | CPC-5 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | Lemo-5B | LEMO650/LEMO650 | 15 | 5 | CH0 - fmc2-1 |  | 11 | 1: EBS8005 | Ch4 | 2651 | FS2_BBS:PM_D4051 | 4051
PM 12" | FS2_BBS:PM_D4051 | PM | D4051 | 44F000012 | 6C_22AWG_Multi_010 | B Signal | CPC to LEMO650 Box | CPC-5 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | Lemo-5C | LEMO650/LEMO650 | 15 | 5 | CH1 - fmc2-2 |  | 11 | 1: EBS8005 | Ch4 | 2651 | FS2_BBS:PM_D4051 | 4051
PM 8" | FS2_BMS:PM_D4144 | PM | D4143 | 44F000096 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-6 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | Lemo-6A | LEMO650/LEMO650 | 15 | 5 | CH6 - fmc1-3 |  | 11 | 1: EBS8005 | Ch5 | 2684 | FS2_BMS:PM_D4144 | 4144
PM 8" | FS2_BMS:PM_D4144 | PM | D4143 | 44F000096 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-6 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | Lemo-6B | LEMO650/LEMO650 | 15 | 5 | CH5 - fmc1-2 |  | 11 | 1: EBS8005 | Ch5 | 2684 | FS2_BMS:PM_D4144 | 4144
PM 8" | FS2_BMS:PM_D4144 | PM | D4143 | 44F000096 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-6 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | Lemo-6C | LEMO650/LEMO650 | 15 | 5 | CH4 - fmc1-1 |  | 11 | 1: EBS8005 | Ch5 | 2684 | FS2_BMS:PM_D4144 | 4144
PM 8" | FS2_BTS:PM_D3959 | PM | D3958 | 44F000003 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-7 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | Lemo-7A | LEMO650/LEMO650 | 15 | 6 | CH2 - fmc2-3 |  | 11 | 1: EBS8005 | Ch6 | 2620 | FS2_BTS:PM_D3959 | 3959
PM 8" | FS2_BTS:PM_D3959 | PM | D3958 | 44F000003 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-7 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | Lemo-7B | LEMO650/LEMO650 | 15 | 6 | CH1 - fmc2-2 |  | 11 | 1: EBS8005 | Ch6 | 2620 | FS2_BTS:PM_D3959 | 3959
PM 8" | FS2_BTS:PM_D3959 | PM | D3958 | 44F000003 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-7 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | Lemo-7C | LEMO650/LEMO650 | 15 | 6 | CH0 - fmc2-1 |  | 11 | 1: EBS8005 | Ch6 | 2620 | FS2_BTS:PM_D3959 | 3959
PM 6" | FS2_BTS:PM_D4009 | PM | D4013 | 44F000006 | 4C_22AWG_Multi_006 | C Signal | CPC to LEMO650 Box | CPC-8 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | Lemo-8A | LEMO650/LEMO650 | 15 | 6 | CH6 - fmc1-3 |  | 11 | 1: EBS8005 | Ch7 | 3579 | FS2_BTS:PM_D4009 | 4009
PM 12" | FS2_BTS:PM_D4009 | PM | D4013 | 44F000004 | 6C_22AWG_Multi_010 | A Signal | CPC to LEMO650 Box | CPC-8 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | Lemo-8B | LEMO650/LEMO650 | 15 | 6 | CH4 - fmc1-1 |  | 11 | 1: EBS8005 | Ch7 | 3579 | FS2_BTS:PM_D4009 | 4009
PM 12" | FS2_BTS:PM_D4009 | PM | D4013 | 44F000004 | 6C_22AWG_Multi_010 | B Signal | CPC to LEMO650 Box | CPC-8 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | Lemo-8C | LEMO650/LEMO650 | 15 | 6 | CH5 - fmc1-2 |  | 11 | 1: EBS8005 | Ch7 | 3579 | FS2_BTS:PM_D4009 | 4009
Lemo to Lemo Panel
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 1_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 11 | 2: EBS8003 | Ch0
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 1_OUT Port | [7 empty] | Lemo650/Lemo650 | 15 | 3 | CH3 - fmc2-4
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 2_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 11 | 2: EBS8003 | Ch1
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 2_OUT Port | [7 empty] | Lemo650/Lemo650 | 15 | 3 | CH7 - fmc1-4
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 3_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 11 | 2: EBS8003 | Ch2
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 3_OUT Port | [7 empty] | Lemo650/Lemo650 | 15 | 4 | CH3 - fmc2-4
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 4_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 11 | 2: EBS8003 | Ch3
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 4_OUT Port | [7 empty] | Lemo650/Lemo650 | 15 | 4 | CH7 - fmc1-4
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 5_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 11 | 2: EBS8003 | Ch4
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 5_OUT Port | [7 empty] | Lemo650/Lemo650 | 15 | 5 | CH3 - fmc2-4
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 6_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 11 | 2: EBS8003 | Ch5
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 6_OUT Port | [7 empty] | Lemo650/Lemo650 | 15 | 5 | CH7 - fmc1-4
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 7_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 11 | 2: EBS8003 | Ch6
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 7_OUT Port | [7 empty] | Lemo650/Lemo650 | 15 | 6 | CH3 - fmc2-4
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 8_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 11 | 2: EBS8003 | Ch7
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 8_OUT Port | [7 empty] | Lemo650/Lemo650 | 15 | 6 | CH7 - fmc1-4

## LS3_N1108
LS3_N1108
Long Run Cable | [6 empty] | Inter-Rack Connections | [9 empty] | Terminating Device
Diagnostics Device | Lattice Beamline Location | Device Type | CableDB Location | Cable Number | Cable Type | Function | Termination Device | Port Number In | Pin Number | CPC pinout | CPC pinout | CPC pinout | Signal Pin | Shield Pin | Port Number Out | Patch Cable | MTCA Chassis | MTCA Slot | PICO Port | FGPDB Port | Bias Supply Chassis | Bias Supply Slot | Bias Supply Channel |  |  |  | search for row | device name | D number
BPMs | [23 empty] | PICO | CH[0-7] | Device
BPM | LS3_WD02:BPM_D4451 | BPM | D4451 | 34E000088 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 16 | 3 |  | 1 |  |  |  | 16 | 3 | LS3_WD02:BPM_D4451 | 2858 | LS3_WD02:BPM_D4451 | 4451
BPM | LS3_WD02:BPM_D4451 | BPM | D4451 | 34E000089 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 16 | 3 |  | 2 |  |  |  | 16 | 3 | LS3_WD02:BPM_D4451 | 2858 | LS3_WD02:BPM_D4451 | 4451
BPM | LS3_WD02:BPM_D4451 | BPM | D4451 | 34E000090 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 16 | 3 |  | 3 |  |  |  | 16 | 3 | LS3_WD02:BPM_D4451 | 2858 | LS3_WD02:BPM_D4451 | 4451
BPM | LS3_WD02:BPM_D4451 | BPM | D4451 | 34E000091 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 16 | 3 |  | 4 |  |  |  | 16 | 3 | LS3_WD02:BPM_D4451 | 2858 | LS3_WD02:BPM_D4451 | 4451
BPM | LS3_WD02:BPM_D4451 | BPM | D4451 | 34E000098 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | [17 empty] | LS3_WD02:BPM_D4451 | 2858 | LS3_WD02:BPM_D4451 | 4451
BPM | LS3_WD03:BPM_D4513 | BPM | D4513 | 34E000084 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 16 | 3 |  | 5 |  |  |  | 16 | 3 | LS3_WD03:BPM_D4513 | 2902 | LS3_WD03:BPM_D4513 | 4513
BPM | LS3_WD03:BPM_D4513 | BPM | D4513 | 34E000085 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 16 | 3 |  | 6 |  |  |  | 16 | 3 | LS3_WD03:BPM_D4513 | 2902 | LS3_WD03:BPM_D4513 | 4513
BPM | LS3_WD03:BPM_D4513 | BPM | D4513 | 34E000086 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 16 | 3 |  | 7 |  |  |  | 16 | 3 | LS3_WD03:BPM_D4513 | 2902 | LS3_WD03:BPM_D4513 | 4513
BPM | LS3_WD03:BPM_D4513 | BPM | D4513 | 34E000087 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 16 | 3 |  | 8 |  |  |  | 16 | 3 | LS3_WD03:BPM_D4513 | 2902 | LS3_WD03:BPM_D4513 | 4513
BPM | LS3_WD03:BPM_D4513 | BPM | D4513 | 34E000097 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 16 | 3 |  | Ref |  |  |  | 16 | 3 | LS3_WD03:BPM_D4513 | 2902 | LS3_WD03:BPM_D4513 | 4513
BPM | LS3_WD04:BPM_D4575 | BPM | D4575 | 34E000080 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 16 | 4 |  | 1 |  |  |  | 16 | 4 | LS3_WD04:BPM_D4575 | 2946 | LS3_WD04:BPM_D4575 | 4575
BPM | LS3_WD04:BPM_D4575 | BPM | D4575 | 34E000081 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 16 | 4 |  | 2 |  |  |  | 16 | 4 | LS3_WD04:BPM_D4575 | 2946 | LS3_WD04:BPM_D4575 | 4575
BPM | LS3_WD04:BPM_D4575 | BPM | D4575 | 34E000082 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 16 | 4 |  | 3 |  |  |  | 16 | 4 | LS3_WD04:BPM_D4575 | 2946 | LS3_WD04:BPM_D4575 | 4575
BPM | LS3_WD04:BPM_D4575 | BPM | D4575 | 34E000083 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 16 | 4 |  | 4 |  |  |  | 16 | 4 | LS3_WD04:BPM_D4575 | 2946 | LS3_WD04:BPM_D4575 | 4575
BPM | LS3_WD04:BPM_D4575 | BPM | D4575 | 34E000096 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | [17 empty] | LS3_WD04:BPM_D4575 | 2946 | LS3_WD04:BPM_D4575 | 4575
BPM | LS3_WD05:BPM_D4637 | BPM | D4637 | 34E000075 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 16 | 4 |  | 5 |  |  |  | 16 | 4 | LS3_WD05:BPM_D4637 | 2990 | LS3_WD05:BPM_D4637 | 4637
BPM | LS3_WD05:BPM_D4637 | BPM | D4637 | 34E000076 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 16 | 4 |  | 6 |  |  |  | 16 | 4 | LS3_WD05:BPM_D4637 | 2990 | LS3_WD05:BPM_D4637 | 4637
BPM | LS3_WD05:BPM_D4637 | BPM | D4637 | 34E000077 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 16 | 4 |  | 7 |  |  |  | 16 | 4 | LS3_WD05:BPM_D4637 | 2990 | LS3_WD05:BPM_D4637 | 4637
BPM | LS3_WD05:BPM_D4637 | BPM | D4637 | 34E000078 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 16 | 4 |  | 8 |  |  |  | 16 | 4 | LS3_WD05:BPM_D4637 | 2990 | LS3_WD05:BPM_D4637 | 4637
BPM | LS3_WD05:BPM_D4637 | BPM | D4637 | 34E000079 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 16 | 4 |  | Ref |  |  |  | 16 | 4 | LS3_WD05:BPM_D4637 | 2990 | LS3_WD05:BPM_D4637 | 4637
BPM | LS3_WD06:BPM_D4699 | BPM | D4699 | 34E000025 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 16 | 5 |  | 1 |  |  |  | 16 | 5 | LS3_WD06:BPM_D4699 | 3034 | LS3_WD06:BPM_D4699 | 4699
BPM | LS3_WD06:BPM_D4699 | BPM | D4699 | 34E000026 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 16 | 5 |  | 2 |  |  |  | 16 | 5 | LS3_WD06:BPM_D4699 | 3034 | LS3_WD06:BPM_D4699 | 4699
BPM | LS3_WD06:BPM_D4699 | BPM | D4699 | 34E000027 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 16 | 5 |  | 3 |  |  |  | 16 | 5 | LS3_WD06:BPM_D4699 | 3034 | LS3_WD06:BPM_D4699 | 4699
BPM | LS3_WD06:BPM_D4699 | BPM | D4699 | 34E000028 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 16 | 5 |  | 4 |  |  |  | 16 | 5 | LS3_WD06:BPM_D4699 | 3034 | LS3_WD06:BPM_D4699 | 4699
BPM | LS3_WD06:BPM_D4699 | BPM | D4699 | 34E000029 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | [17 empty] | LS3_WD06:BPM_D4699 | 3034 | LS3_WD06:BPM_D4699 | 4699
BPM | LS3_BTS:BPM_D4753 | BPM | D4753 | 34E000069 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 16 | 5 |  | 5 |  |  |  | 16 | 5 | LS3_BTS:BPM_D4753 | 3054 | LS3_BTS:BPM_D4753 | 4753
BPM | LS3_BTS:BPM_D4753 | BPM | D4753 | 34E000070 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 16 | 5 |  | 6 |  |  |  | 16 | 5 | LS3_BTS:BPM_D4753 | 3054 | LS3_BTS:BPM_D4753 | 4753
BPM | LS3_BTS:BPM_D4753 | BPM | D4753 | 34E000071 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 16 | 5 |  | 7 |  |  |  | 16 | 5 | LS3_BTS:BPM_D4753 | 3054 | LS3_BTS:BPM_D4753 | 4753
BPM | LS3_BTS:BPM_D4753 | BPM | D4753 | 34E000072 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 16 | 5 |  | 8 |  |  |  | 16 | 5 | LS3_BTS:BPM_D4753 | 3054 | LS3_BTS:BPM_D4753 | 4753
BPM | LS3_BTS:BPM_D4753 | BPM | D4753 | 34E000074 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 16 | 5 |  | Ref |  |  |  | 16 | 5 | LS3_BTS:BPM_D4753 | 3054 | LS3_BTS:BPM_D4753 | 4753
BPM | LS3_BTS:BPM_D4769 | BPM | D4769 | 34E000064 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 16 | 6 |  | 1 |  |  |  | 16 | 6 | LS3_BTS:BPM_D4769 | 3061 | LS3_BTS:BPM_D4769 | 4769
BPM | LS3_BTS:BPM_D4769 | BPM | D4769 | 34E000065 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 16 | 6 |  | 2 |  |  |  | 16 | 6 | LS3_BTS:BPM_D4769 | 3061 | LS3_BTS:BPM_D4769 | 4769
BPM | LS3_BTS:BPM_D4769 | BPM | D4769 | 34E000066 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 16 | 6 |  | 3 |  |  |  | 16 | 6 | LS3_BTS:BPM_D4769 | 3061 | LS3_BTS:BPM_D4769 | 4769
BPM | LS3_BTS:BPM_D4769 | BPM | D4769 | 34E000067 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 16 | 6 |  | 4 |  |  |  | 16 | 6 | LS3_BTS:BPM_D4769 | 3061 | LS3_BTS:BPM_D4769 | 4769
BPM | LS3_BTS:BPM_D4769 | BPM | D4769 | 34E000073 | 2C_1/4in_Hardline_036 | Ref | No terminiation | Ref | [6 empty] | Ref | SMA to SMA | 16 | 6 |  | Ref |  |  |  | 16 | 6 | LS3_BTS:BPM_D4769 | 3061 | LS3_BTS:BPM_D4769 | 4769
Ion Chamber
Ion Chamber | LS3_CD05:IC_D4602 | IC | D4602 | 34E000560 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J1 | [6 empty] | J1 | Lemo to Lemo | 16 | 7 | CH0 - fmc2-1 |  |  |  |  | 16 | 7 | LS3_CD05:IC_D4602 | #N/A | #N/A | #N/A
Ion Chamber | LS3_BTS:IC_D4701 | IC | D4701 | 34E000563 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J2 | [6 empty] | J2 | Lemo to Lemo | 16 | 7 | CH1 - fmc2-2 |  |  |  |  | 16 | 7 | LS3_BTS:IC_D4701 | #N/A | #N/A | #N/A
Ion Chamber Bias
Ion Chamber | LS3_CD05:IC_D4602 |  | D4602 | 34H000078 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 1 | [6 empty] | 1 | SHV to SHV |  |  |  |  | 12 | 1 : EHS8020n | Ch0 |  |  | LS3_CD05:IC_D4602
Ion Chamber | LS3_BTS:IC_D4701 |  | D4701 | 34H000081 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 2 | [6 empty] | 2 | SHV to SHV |  |  |  |  | 12 | 1 : EHS8020n | Ch1 |  |  | LS3_BTS:IC_D4701
Note on B5: Rodriguez, Pedro:

## LS3_N2101
LS3_N2101
Long Run Cable | [6 empty] | Inter-Rack Connections | [9 empty] | Terminating Device
Diagnostics Device | Lattice Beamline Location | Device Type | CableDB Location | Cable Number | Cable Type | Function | Termination Device | Port Number In | Pin Number | CPC pinout | CPC pinout | CPC pinout | Signal Pin | Shield Pin | Port Number Out | Patch Cable | MTCA Chassis | MTCA Slot | PICO Port | FGPDB Port | Bias Supply Chassis | Bias Supply Slot | Bias Supply Channel |  |  |  | search for row | device name | D number
BPMs | [23 empty] | PICO | CH[0-7] | Device
BPM | LS3_BTS:BPM_D4843 | BPM | D4843 | 34E000060 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 17 | 3 |  | 1 |  |  |  | 17 | 3 | LS3_BTS:BPM_D4843 | 3083 | LS3_BTS:BPM_D4843 | 4843
BPM | LS3_BTS:BPM_D4843 | BPM | D4843 | 34E000061 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 17 | 3 |  | 2 |  |  |  | 17 | 3 | LS3_BTS:BPM_D4843 | 3083 | LS3_BTS:BPM_D4843 | 4843
BPM | LS3_BTS:BPM_D4843 | BPM | D4843 | 34E000062 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 17 | 3 |  | 3 |  |  |  | 17 | 3 | LS3_BTS:BPM_D4843 | 3083 | LS3_BTS:BPM_D4843 | 4843
BPM | LS3_BTS:BPM_D4843 | BPM | D4843 | 34E000063 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 17 | 3 |  | 4 |  |  |  | 17 | 3 | LS3_BTS:BPM_D4843 | 3083 | LS3_BTS:BPM_D4843 | 4843
BPM | LS3_BTS:BPM_D4843 | BPM | D4843 | 34E000059 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | [17 empty] | LS3_BTS:BPM_D4843 | 3083 | LS3_BTS:BPM_D4843 | 4843
BPM | LS3_BTS:BPM_D4886 | BPM | D4886 | 34E000055 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 17 | 3 |  | 5 |  |  |  | 17 | 3 | LS3_BTS:BPM_D4886 | 3095 | LS3_BTS:BPM_D4886 | 4886
BPM | LS3_BTS:BPM_D4886 | BPM | D4886 | 34E000056 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 17 | 3 |  | 6 |  |  |  | 17 | 3 | LS3_BTS:BPM_D4886 | 3095 | LS3_BTS:BPM_D4886 | 4886
BPM | LS3_BTS:BPM_D4886 | BPM | D4886 | 34E000057 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 17 | 3 |  | 7 |  |  |  | 17 | 3 | LS3_BTS:BPM_D4886 | 3095 | LS3_BTS:BPM_D4886 | 4886
BPM | LS3_BTS:BPM_D4886 | BPM | D4886 | 34E000058 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 17 | 3 |  | 8 |  |  |  | 17 | 3 | LS3_BTS:BPM_D4886 | 3095 | LS3_BTS:BPM_D4886 | 4886
BPM | LS3_BTS:BPM_D4886 | BPM | D4886 | 34E000068 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 17 | 3 |  | Ref |  |  |  | 17 | 3 | LS3_BTS:BPM_D4886 | 3095 | LS3_BTS:BPM_D4886 | 4886
BPM | LS3_BTS:BPM_D4968 | BPM | D4967 | 34E000032 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 17 | 4 |  | 1 |  |  |  | 17 | 4 | LS3_BTS:BPM_D4968 | 3118 | LS3_BTS:BPM_D4968 | 4968
BPM | LS3_BTS:BPM_D4968 | BPM | D4967 | 34E000033 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 17 | 4 |  | 2 |  |  |  | 17 | 4 | LS3_BTS:BPM_D4968 | 3118 | LS3_BTS:BPM_D4968 | 4968
BPM | LS3_BTS:BPM_D4968 | BPM | D4967 | 34E000034 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 17 | 4 |  | 3 |  |  |  | 17 | 4 | LS3_BTS:BPM_D4968 | 3118 | LS3_BTS:BPM_D4968 | 4968
BPM | LS3_BTS:BPM_D4968 | BPM | D4967 | 34E000035 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 17 | 4 |  | 4 |  |  |  | 17 | 4 | LS3_BTS:BPM_D4968 | 3118 | LS3_BTS:BPM_D4968 | 4968
BPM | LS3_BTS:BPM_D4968 | BPM | D4967 | 34E000054 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | [17 empty] | LS3_BTS:BPM_D4968 | 3118 | LS3_BTS:BPM_D4968 | 4968
BPM | LS3_BTS:BPM_D5010 | BPM | D5010 | 34E000047 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 17 | 4 |  | 5 |  |  |  | 17 | 4 | LS3_BTS:BPM_D5010 | 3125 | LS3_BTS:BPM_D5010 | 5010
BPM | LS3_BTS:BPM_D5010 | BPM | D5010 | 34E000048 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 17 | 4 |  | 6 |  |  |  | 17 | 4 | LS3_BTS:BPM_D5010 | 3125 | LS3_BTS:BPM_D5010 | 5010
BPM | LS3_BTS:BPM_D5010 | BPM | D5010 | 34E000049 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 17 | 4 |  | 7 |  |  |  | 17 | 4 | LS3_BTS:BPM_D5010 | 3125 | LS3_BTS:BPM_D5010 | 5010
BPM | LS3_BTS:BPM_D5010 | BPM | D5010 | 34E000050 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 17 | 4 |  | 8 |  |  |  | 17 | 4 | LS3_BTS:BPM_D5010 | 3125 | LS3_BTS:BPM_D5010 | 5010
BPM | LS3_BTS:BPM_D5010 | BPM | D5010 | 34E000053 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 17 | 4 |  | Ref |  |  |  | 17 | 4 | LS3_BTS:BPM_D5010 | 3125 | LS3_BTS:BPM_D5010 | 5010
BPM | LS3_BTS:BPM_D5092 | BPM | D5091 | 34E000043 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 17 | 5 |  | 1 |  |  |  | 17 | 5 | LS3_BTS:BPM_D5092 | 3149 | LS3_BTS:BPM_D5092 | 5092
BPM | LS3_BTS:BPM_D5092 | BPM | D5091 | 34E000044 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 17 | 5 |  | 2 |  |  |  | 17 | 5 | LS3_BTS:BPM_D5092 | 3149 | LS3_BTS:BPM_D5092 | 5092
BPM | LS3_BTS:BPM_D5092 | BPM | D5091 | 34E000045 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 17 | 5 |  | 3 |  |  |  | 17 | 5 | LS3_BTS:BPM_D5092 | 3149 | LS3_BTS:BPM_D5092 | 5092
BPM | LS3_BTS:BPM_D5092 | BPM | D5091 | 34E000046 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 17 | 5 |  | 4 |  |  |  | 17 | 5 | LS3_BTS:BPM_D5092 | 3149 | LS3_BTS:BPM_D5092 | 5092
BPM | LS3_BTS:BPM_D5092 | BPM | D5091 | 34E000052 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | [17 empty] | LS3_BTS:BPM_D5092 | 3149 | LS3_BTS:BPM_D5092 | 5092
BPM | LS3_BTS:BPM_D5134 | BPM | D5134 | 34E000039 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 17 | 5 |  | 5 |  |  |  | 17 | 5 | LS3_BTS:BPM_D5134 | 3156 | LS3_BTS:BPM_D5134 | 5134
BPM | LS3_BTS:BPM_D5134 | BPM | D5134 | 34E000040 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 17 | 5 |  | 6 |  |  |  | 17 | 5 | LS3_BTS:BPM_D5134 | 3156 | LS3_BTS:BPM_D5134 | 5134
BPM | LS3_BTS:BPM_D5134 | BPM | D5134 | 34E000041 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 17 | 5 |  | 7 |  |  |  | 17 | 5 | LS3_BTS:BPM_D5134 | 3156 | LS3_BTS:BPM_D5134 | 5134
BPM | LS3_BTS:BPM_D5134 | BPM | D5134 | 34E000042 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 17 | 5 |  | 8 |  |  |  | 17 | 5 | LS3_BTS:BPM_D5134 | 3156 | LS3_BTS:BPM_D5134 | 5134
BPM | LS3_BTS:BPM_D5134 | BPM | D5134 | 34E000051 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 17 | 5 |  | Ref |  |  |  | 17 | 5 | LS3_BTS:BPM_D5134 | 3156 | LS3_BTS:BPM_D5134 | 5134
BPM | LS3_BTS:BPM_D5216 | BPM | D5215 | 34E000030 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 17 | 6 |  | 1 |  |  |  | 17 | 6 | LS3_BTS:BPM_D5216 | 3180 | LS3_BTS:BPM_D5216 | 5216
BPM | LS3_BTS:BPM_D5216 | BPM | D5215 | 34E000031 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 17 | 6 |  | 2 |  |  |  | 17 | 6 | LS3_BTS:BPM_D5216 | 3180 | LS3_BTS:BPM_D5216 | 5216
BPM | LS3_BTS:BPM_D5216 | BPM | D5215 | 34E000037 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 17 | 6 |  | 3 |  |  |  | 17 | 6 | LS3_BTS:BPM_D5216 | 3180 | LS3_BTS:BPM_D5216 | 5216
BPM | LS3_BTS:BPM_D5216 | BPM | D5215 | 34E000038 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 17 | 6 |  | 4 |  |  |  | 17 | 6 | LS3_BTS:BPM_D5216 | 3180 | LS3_BTS:BPM_D5216 | 5216
BPM | LS3_BTS:BPM_D5216 | BPM | D5215 | 34E000036 | 2C_1/4in_Hardline_036 | Ref | No terminiation | Ref | [17 empty] | LS3_BTS:BPM_D5216 | 3180 | LS3_BTS:BPM_D5216 | 5216
BPM | LS3_BTS:BPM_D5259 | BPM | D5258 | 34E000009 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 17 | 6 |  | 5 |  |  |  | 17 | 6 | LS3_BTS:BPM_D5259 | 3187 | LS3_BTS:BPM_D5259 | 5259
BPM | LS3_BTS:BPM_D5259 | BPM | D5258 | 34E000010 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 17 | 6 |  | 6 |  |  |  | 17 | 6 | LS3_BTS:BPM_D5259 | 3187 | LS3_BTS:BPM_D5259 | 5259
BPM | LS3_BTS:BPM_D5259 | BPM | D5258 | 34E000011 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 17 | 6 |  | 7 |  |  |  | 17 | 6 | LS3_BTS:BPM_D5259 | 3187 | LS3_BTS:BPM_D5259 | 5259
BPM | LS3_BTS:BPM_D5259 | BPM | D5258 | 34E000012 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 17 | 6 |  | 8 |  |  |  | 17 | 6 | LS3_BTS:BPM_D5259 | 3187 | LS3_BTS:BPM_D5259 | 5259
BPM | LS3_BTS:BPM_D5259 | BPM | D5258 | 34E000024 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 17 | 6 |  | Ref |  |  |  | 17 | 6 | LS3_BTS:BPM_D5259 | 3187 | LS3_BTS:BPM_D5259 | 5259
Ion Chamber
Ion Chamber | LS3_BTS:IC_D4799 | IC | D4799 | 34E000562 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J1 | [6 empty] | J1 | Lemo to Lemo | 17 | 9 | CH0 - fmc2-1 |  |  |  |  | 17 | 9 | LS3_BTS:IC_D4799 | #N/A | #N/A | #N/A
Ion Chamber | LS3_BTS:IC_D4895 | IC | D4895 | 34E000561 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J2 | [6 empty] | J2 | Lemo to Lemo | 17 | 9 | CH1 - fmc2-2 |  |  |  |  | 17 | 9 | LS3_BTS:IC_D4895 | #N/A | #N/A | #N/A
Ion Chamber | LS3_BTS:IC_D5001 | IC | D5001 | 34E000567 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J3 | [6 empty] | J3 | Lemo to Lemo | 17 | 9 | CH2 - fmc2-3 |  |  |  |  | 17 | 9 | LS3_BTS:IC_D5001 | #N/A | #N/A | #N/A
Ion Chamber | LS3_BTS:IC_D5105 | IC | D5105 | 34E000566 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J4 | [6 empty] | J4 | Lemo to Lemo | 17 | 9 | CH3 - fmc2-4 |  |  |  |  | 17 | 9 | LS3_BTS:IC_D5105 | #N/A | #N/A | #N/A
Ion Chamber Bias
Ion Chamber | LS3_BTS:IC_D4799 |  | D4799 | 34H000080 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 1 | [6 empty] | 1 | SHV to SHV |  |  |  |  | 12 | 2 : EHS8020n | Ch0 |  |  | LS3_BTS:IC_D4799 | #N/A | #N/A | #N/A
Ion Chamber | LS3_BTS:IC_D4895 |  | D4895 | 34H000079 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 2 | [6 empty] | 2 | SHV to SHV |  |  |  |  | 12 | 2 : EHS8020n | Ch1 |  |  | LS3_BTS:IC_D4895 | #N/A | #N/A | #N/A
Ion Chamber | LS3_BTS:IC_D5001 |  | D5001 | 34H000085 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 3 | [6 empty] | 3 | SHV to SHV |  |  |  |  | 12 | 2 : EHS8020n | Ch2 |  |  | LS3_BTS:IC_D5001 | #N/A | #N/A | #N/A
Ion Chamber | LS3_BTS:IC_D5105 |  | D5105 | 34H000084 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 4 | [6 empty] | 4 | SHV to SHV |  |  |  |  | 12 | 2 : EHS8020n | Ch3 |  |  | LS3_BTS:IC_D5105 | #N/A | #N/A | #N/A
Profile Monitors
PM 8" | LS3_BTS:PM_D4771 | PM | D4769 | 34F000023 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-1 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | Lemo-1A | LEMO650/LEMO650 | 17 | 7 | CH0 - fmc2-1 |  | 12 | 1: EBS8005 | Ch0 | 17 | 7 | LS3_BTS:PM_D4771 | 3063 | LS3_BTS:PM_D4771 | 4771
PM 8" | LS3_BTS:PM_D4771 | PM | D4769 | 34F000023 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-1 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | Lemo-1B | LEMO650/LEMO650 | 17 | 7 | CH1 - fmc2-2 |  | 12 | 1: EBS8005 | Ch0 | 17 | 7 | LS3_BTS:PM_D4771 | 3063 | LS3_BTS:PM_D4771 | 4771
PM 8" | LS3_BTS:PM_D4771 | PM | D4769 | 34F000023 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-1 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | Lemo-1C | LEMO650/LEMO650 | 17 | 7 | CH2 - fmc2-3 |  | 12 | 1: EBS8005 | Ch0 | 17 | 7 | LS3_BTS:PM_D4771 | 3063 | LS3_BTS:PM_D4771 | 4771
PM 8" | LS3_BTS:PM_D4797 | PM | D4794 | 34F000001 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-2 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | Lemo-2A | LEMO650/LEMO650 | 17 | 7 | CH4 - fmc1-1 |  | 12 | 1: EBS8005 | Ch1 | 17 | 7 | LS3_BTS:PM_D4797 | 3073 | LS3_BTS:PM_D4797 | 4797
PM 8" | LS3_BTS:PM_D4797 | PM | D4794 | 34F000001 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-2 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield-6 | 4 | 6 | Lemo-2B | LEMO650/LEMO650 | 17 | 7 | CH5 - fmc1-2 |  | 12 | 1: EBS8005 | Ch1 | 17 | 7 | LS3_BTS:PM_D4797 | 3073 | LS3_BTS:PM_D4797 | 4797
PM 8" | LS3_BTS:PM_D4797 | PM | D4794 | 34F000001 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-2 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | Lemo-2C | LEMO650/LEMO650 | 17 | 7 | CH6 - fmc1-3 |  | 12 | 1: EBS8005 | Ch1 | 17 | 7 | LS3_BTS:PM_D4797 | 3073 | LS3_BTS:PM_D4797 | 4797
PM 8" | LS3_BTS:PM_D4827 | PM | D4824 | 34F000004 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-3 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | Lemo-3A | LEMO650/LEMO650 | 17 | 8 | CH0 - fmc2-1 |  | 12 | 1: EBS8005 | Ch2 | 17 | 8 | LS3_BTS:PM_D4827 | 3078 | LS3_BTS:PM_D4827 | 4827
PM 8" | LS3_BTS:PM_D4827 | PM | D4824 | 34F000004 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-3 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | Lemo-3B | LEMO650/LEMO650 | 17 | 8 | CH1 - fmc2-2 |  | 12 | 1: EBS8005 | Ch2 | 17 | 8 | LS3_BTS:PM_D4827 | 3078 | LS3_BTS:PM_D4827 | 4827
PM 8" | LS3_BTS:PM_D4827 | PM | D4824 | 34F000004 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-3 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | Lemo-3C | LEMO650/LEMO650 | 17 | 8 | CH2 - fmc2-3 |  | 12 | 1: EBS8005 | Ch2 | 17 | 8 | LS3_BTS:PM_D4827 | 3078 | LS3_BTS:PM_D4827 | 4827
PM 8" | LS3_BTS:PM_D4862 | PM | D4862 | 34F000008 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-4 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | Lemo-4A | LEMO650/LEMO650 | 17 | 8 | CH4 - fmc1-1 |  | 12 | 1: EBS8005 | Ch3 | 17 | 8 | LS3_BTS:PM_D4862 | 3087 | LS3_BTS:PM_D4862 | 4862
PM 8" | LS3_BTS:PM_D4862 | PM | D4862 | 34F000008 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-4 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | Lemo-4B | LEMO650/LEMO650 | 17 | 8 | CH5 - fmc1-2 |  | 12 | 1: EBS8005 | Ch3 | 17 | 8 |  | 3087 | LS3_BTS:PM_D4862 | 4862
PM 8" | LS3_BTS:PM_D4862 | PM | D4862 | 34F000008 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-4 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | Lemo-4C | LEMO650/LEMO650 | 17 | 8 | CH6 - fmc1-3 |  | 12 | 1: EBS8005 | Ch3 | 17 | 8 |  | 3087 | LS3_BTS:PM_D4862 | 4862
Halo Monitor Ring | LS3_WD06:HMR_D4700 | HMR | D4700 | 34H000012 | 2C_1kV75LC_Coax_090 | Signal | SHV to SHV | J1 | [6 empty] | J1 | SHV to Lemo | 17 | 10 | CH0 - fmc2-1 |  | 12 | 1: EBS8003 | Ch4
| LS3_WD06:HMR_D4700 |  | D4700 | 34H000097 | 2C_5kVDC_Coax_035 | HV Bias
Halo Monitor Ring,Spare | LS3_BTS:HMR_D4770 | HMR - Unused | D4770 | 34H000000 | 2C_1kV75LC_Coax_090 | Signal
| LS3_BTS:HMR_D4770 |  | D4770 | 34H000093 | 2C_5kVDC_Coax_035 | HV Bias
Halo Monitor Ring,Spare | LS3_BTS:HMR_D4887 | HMR - Unused | D4887 | 34H000002 | 2C_1kV75LC_Coax_090 | Signal
| LS3_BTS:HMR_D4887 |  | D4887 | 34H000091 | 2C_5kVDC_Coax_035 | HV Bias
Halo Monitor Ring,Spare | LS3_BTS:HMR_D5011 | HMR - Unused | D5011 | 34H000003 | 2C_1kV75LC_Coax_090 | Signal
| LS3_BTS:HMR_D5011 |  | D5011 | 34H000090 | 2C_5kVDC_Coax_035 | HV Bias
Halo Monitor Ring,Spare | LS3_BTS:HMR_D5135 | HMR - Unused | D5135 | 34H000005 | 2C_1kV75LC_Coax_090 | Signal
| LS3_BTS:HMR_D5135 |  | D5135 | 34H000087 | 2C_5kVDC_Coax_035 | HV Bias
Halo Monitor Ring,Spare | LS3_BTS:HMR_D5259 | HMR - Unused | D5259 | 34H000004 | 2C_1kV75LC_Coax_090 | Signal
| LS3_BTS:HMR_D5259 |  | D5259 | 34H000088 | 2C_5kVDC_Coax_035 | HV Bias
Lemo to Lemo Panel
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 1_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 12 | 2: EBS8003 | Ch0
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 1_OUT Port | [7 empty] | Lemo650/Lemo650 | 17 | 7 | CH3 - fmc2-4
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 2_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 12 | 2: EBS8003 | Ch1
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 2_OUT Port | [7 empty] | Lemo650/Lemo650 | 17 | 7 | CH7 - fmc1-4
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 3_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 12 | 2: EBS8003 | Ch2
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 3_OUT Port | [7 empty] | Lemo650/Lemo650 | 17 | 8 | CH3 - fmc2-4
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 4_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 12 | 2: EBS8003 | Ch3
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 4_OUT Port | [7 empty] | Lemo650/Lemo650 | 17 | 8 | CH7 - fmc1-4
Note on B5: Rodriguez, Pedro:

## BDS_N0104
BDS_N0104
Long Run Cable | [6 empty] | Inter-Rack Connections | [9 empty] | Terminating Device
Diagnostics Device | Lattice Beamline Location | Device Type | CableDB Location | Cable Number | Cable Type | Function | Termination Device | Port Number In | Pin Number | CPC pinout | CPC pinout | CPC pinout | Signal Pin | Shield Pin | Port Number Out | Patch Cable | MTCA Chassis | MTCA Slot | PICO Port | FGPDB Port | Bias Supply Chassis | Bias Supply Slot | Bias Supply Channel |  |  |  | search for row | device name | D number
BPMs | [23 empty] | PICO | CH[0-7] | Device
BPM | LS3_BTS:BPM_D5340 | BPM | D5339 | 34E000013 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 18 | 3 |  | 1 |  |  |  | 18 | 3 | LS3_BTS:BPM_D5340 | 3211 | LS3_BTS:BPM_D5340 | 5340
BPM | LS3_BTS:BPM_D5340 | BPM | D5339 | 34E000014 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 18 | 3 |  | 2 |  |  |  | 18 | 3 | LS3_BTS:BPM_D5340 | 3211 | LS3_BTS:BPM_D5340 | 5340
BPM | LS3_BTS:BPM_D5340 | BPM | D5339 | 34E000015 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 18 | 3 |  | 3 |  |  |  | 18 | 3 | LS3_BTS:BPM_D5340 | 3211 | LS3_BTS:BPM_D5340 | 5340
BPM | LS3_BTS:BPM_D5340 | BPM | D5339 | 34E000016 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 18 | 3 |  | 4 |  |  |  | 18 | 3 | LS3_BTS:BPM_D5340 | 3211 | LS3_BTS:BPM_D5340 | 5340
BPM | LS3_BTS:BPM_D5340 | BPM | D5339 | 34E000023 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | [17 empty] | LS3_BTS:BPM_D5340 | 3211 | LS3_BTS:BPM_D5340 | 5340
BPM | LS3_BTS:BPM_D5381 | BPM | D5381 | 34E000002 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 18 | 3 |  | 5 |  |  |  | 18 | 3 | LS3_BTS:BPM_D5381 | 3218 | LS3_BTS:BPM_D5381 | 5381
BPM | LS3_BTS:BPM_D5381 | BPM | D5381 | 34E000003 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 18 | 3 |  | 6 |  |  |  | 18 | 3 | LS3_BTS:BPM_D5381 | 3218 | LS3_BTS:BPM_D5381 | 5381
BPM | LS3_BTS:BPM_D5381 | BPM | D5381 | 34E000004 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 18 | 3 |  | 7 |  |  |  | 18 | 3 | LS3_BTS:BPM_D5381 | 3218 | LS3_BTS:BPM_D5381 | 5381
BPM | LS3_BTS:BPM_D5381 | BPM | D5381 | 34E000005 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 18 | 3 |  | 8 |  |  |  | 18 | 3 | LS3_BTS:BPM_D5381 | 3218 | LS3_BTS:BPM_D5381 | 5381
BPM | LS3_BTS:BPM_D5381 | BPM | D5381 | 34E000022 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 18 | 3 |  | Ref |  |  |  | 18 | 3 | LS3_BTS:BPM_D5381 | 3218 | LS3_BTS:BPM_D5381 | 5381
BPM | LS3_BTS:BPM_D5430 | BPM | D5430 | 34E000017 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 18 | 4 |  | 1 |  |  |  | 18 | 4 | LS3_BTS:BPM_D5430 | 3233 | LS3_BTS:BPM_D5430 | 5430
BPM | LS3_BTS:BPM_D5430 | BPM | D5430 | 34E000018 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 18 | 4 |  | 2 |  |  |  | 18 | 4 | LS3_BTS:BPM_D5430 | 3233 | LS3_BTS:BPM_D5430 | 5430
BPM | LS3_BTS:BPM_D5430 | BPM | D5430 | 34E000019 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 18 | 4 |  | 3 |  |  |  | 18 | 4 | LS3_BTS:BPM_D5430 | 3233 | LS3_BTS:BPM_D5430 | 5430
BPM | LS3_BTS:BPM_D5430 | BPM | D5430 | 34E000020 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 18 | 4 |  | 4 |  |  |  | 18 | 4 | LS3_BTS:BPM_D5430 | 3233 | LS3_BTS:BPM_D5430 | 5430
BPM | LS3_BTS:BPM_D5430 | BPM | D5430 | 34E000021 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | [17 empty] | LS3_BTS:BPM_D5430 | 3233 | LS3_BTS:BPM_D5430 | 5430
BPM | LS3_BTS:BPM_D5445 | BPM | D5444 | 34E000000 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 18 | 4 |  | 5 |  |  |  | 18 | 4 | LS3_BTS:BPM_D5445 | 3240 | LS3_BTS:BPM_D5445 | 5445
BPM | LS3_BTS:BPM_D5445 | BPM | D5444 | 34E000006 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 18 | 4 |  | 6 |  |  |  | 18 | 4 | LS3_BTS:BPM_D5445 | 3240 | LS3_BTS:BPM_D5445 | 5445
BPM | LS3_BTS:BPM_D5445 | BPM | D5444 | 34E000007 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 18 | 4 |  | 7 |  |  |  | 18 | 4 | LS3_BTS:BPM_D5445 | 3240 | LS3_BTS:BPM_D5445 | 5445
BPM | LS3_BTS:BPM_D5445 | BPM | D5444 | 34E000008 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 18 | 4 |  | 8 |  |  |  | 18 | 4 | LS3_BTS:BPM_D5445 | 3240 | LS3_BTS:BPM_D5445 | 5445
BPM | LS3_BTS:BPM_D5445 | BPM | D5444 | 34E000001 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 18 | 4 |  | Ref |  |  |  | 18 | 4 | LS3_BTS:BPM_D5445 | 3240 | LS3_BTS:BPM_D5445 | 5445
BPM | BDS_BTS:BPM_D5499 | BPM | D5499 | 44E000025 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 18 | 5 |  | 1 |  |  |  | 18 | 5 | BDS_BTS:BPM_D5499 | 3256 | BDS_BTS:BPM_D5499 | 5499
BPM | BDS_BTS:BPM_D5499 | BPM | D5499 | 44E000026 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 18 | 5 |  | 2 |  |  |  | 18 | 5 | BDS_BTS:BPM_D5499 | 3256 | BDS_BTS:BPM_D5499 | 5499
BPM | BDS_BTS:BPM_D5499 | BPM | D5499 | 44E000027 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 18 | 5 |  | 3 |  |  |  | 18 | 5 | BDS_BTS:BPM_D5499 | 3256 | BDS_BTS:BPM_D5499 | 5499
BPM | BDS_BTS:BPM_D5499 | BPM | D5499 | 44E000028 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 18 | 5 |  | 4 |  |  |  | 18 | 5 | BDS_BTS:BPM_D5499 | 3256 | BDS_BTS:BPM_D5499 | 5499
BPM | BDS_BTS:BPM_D5499 | BPM | D5499 | 44E000048 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | [17 empty] | BDS_BTS:BPM_D5499 | 3256 | BDS_BTS:BPM_D5499 | 5499
BPM | BDS_BTS:BPM_D5513 | BPM | D5512 | 44E000033 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 18 | 5 |  | 5 |  |  |  | 18 | 5 | BDS_BTS:BPM_D5513 | 3263 | BDS_BTS:BPM_D5513 | 5513
BPM | BDS_BTS:BPM_D5513 | BPM | D5512 | 44E000034 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 18 | 5 |  | 6 |  |  |  | 18 | 5 | BDS_BTS:BPM_D5513 | 3263 | BDS_BTS:BPM_D5513 | 5513
BPM | BDS_BTS:BPM_D5513 | BPM | D5512 | 44E000035 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 18 | 5 |  | 7 |  |  |  | 18 | 5 | BDS_BTS:BPM_D5513 | 3263 | BDS_BTS:BPM_D5513 | 5513
BPM | BDS_BTS:BPM_D5513 | BPM | D5512 | 44E000036 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 18 | 5 |  | 8 |  |  |  | 18 | 5 | BDS_BTS:BPM_D5513 | 3263 | BDS_BTS:BPM_D5513 | 5513
BPM | BDS_BTS:BPM_D5513 | BPM | D5512 | 44E000047 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 18 | 5 |  | Ref |  |  |  | 18 | 5 | BDS_BTS:BPM_D5513 | 3263 | BDS_BTS:BPM_D5513 | 5513
BPM | BDS_BTS:BPM_D5565 | BPM | D5566 | 44E000041 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 18 | 6 |  | 1 |  |  |  | 18 | 6 | BDS_BTS:BPM_D5565 | 3282 | BDS_BTS:BPM_D5565 | 5565
BPM | BDS_BTS:BPM_D5565 | BPM | D5566 | 44E000042 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 18 | 6 |  | 2 |  |  |  | 18 | 6 | BDS_BTS:BPM_D5565 | 3282 | BDS_BTS:BPM_D5565 | 5565
BPM | BDS_BTS:BPM_D5565 | BPM | D5566 | 44E000043 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 18 | 6 |  | 3 |  |  |  | 18 | 6 | BDS_BTS:BPM_D5565 | 3282 | BDS_BTS:BPM_D5565 | 5565
BPM | BDS_BTS:BPM_D5565 | BPM | D5566 | 44E000044 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 18 | 6 |  | 4 |  |  |  | 18 | 6 | BDS_BTS:BPM_D5565 | 3282 | BDS_BTS:BPM_D5565 | 5565
BPM | BDS_BTS:BPM_D5565 | BPM | D5566 | 44E000046 | 2C_1/4in_Hardline_036 | Ref | No terminiation | Ref | [17 empty] | BDS_BTS:BPM_D5565 | 3282 | BDS_BTS:BPM_D5565 | 5565
BPM | BDS_BBS:BPM_D5625 | BPM | D5626 | 44E000053 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 18 | 6 |  | 5 |  |  |  | 18 | 6 | BDS_BBS:BPM_D5625 | 3303 | BDS_BBS:BPM_D5625 | 5625
BPM | BDS_BBS:BPM_D5625 | BPM | D5626 | 44E000054 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 18 | 6 |  | 6 |  |  |  | 18 | 6 | BDS_BBS:BPM_D5625 | 3303 | BDS_BBS:BPM_D5625 | 5625
BPM | BDS_BBS:BPM_D5625 | BPM | D5626 | 44E000055 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 18 | 6 |  | 7 |  |  |  | 18 | 6 | BDS_BBS:BPM_D5625 | 3303 | BDS_BBS:BPM_D5625 | 5625
BPM | BDS_BBS:BPM_D5625 | BPM | D5626 | 44E000056 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 18 | 6 |  | 8 |  |  |  | 18 | 6 | BDS_BBS:BPM_D5625 | 3303 | BDS_BBS:BPM_D5625 | 5625
BPM | BDS_BBS:BPM_D5625 | BPM | D5626 | 44E000057 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 18 | 6 |  | Ref |  |  |  | 18 | 6 | BDS_BBS:BPM_D5625 | 3303 | BDS_BBS:BPM_D5625 | 5625
BPM | BDS_BBS:BPM_D5653 | BPM | D5653 | 44E000049 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 18 | 7 |  | 1 |  |  |  | 18 | 7 | BDS_BBS:BPM_D5653 | 3310 | BDS_BBS:BPM_D5653 | 5653
BPM | BDS_BBS:BPM_D5653 | BPM | D5653 | 44E000050 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 18 | 7 |  | 2 |  |  |  | 18 | 7 | BDS_BBS:BPM_D5653 | 3310 | BDS_BBS:BPM_D5653 | 5653
BPM | BDS_BBS:BPM_D5653 | BPM | D5653 | 44E000051 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 18 | 7 |  | 3 |  |  |  | 18 | 7 | BDS_BBS:BPM_D5653 | 3310 | BDS_BBS:BPM_D5653 | 5653
BPM | BDS_BBS:BPM_D5653 | BPM | D5653 | 44E000052 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 18 | 7 |  | 4 |  |  |  | 18 | 7 | BDS_BBS:BPM_D5653 | 3310 | BDS_BBS:BPM_D5653 | 5653
BPM | BDS_BBS:BPM_D5653 | BPM | D5653 | 44E000058 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | [17 empty] | BDS_BBS:BPM_D5653 | 3310 | BDS_BBS:BPM_D5653 | 5653
BPM | BDS_BBS:BPM_D5653 | BPM | D5655 | 44E000037 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 18 | 7 |  | 5 |  |  |  | 18 | 7 | BDS_BBS:BPM_D5653 | 3310 | BDS_BBS:BPM_D5653 | 5653
BPM | BDS_BBS:BPM_D5653 | BPM | D5655 | 44E000038 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 18 | 7 |  | 6 |  |  |  | 18 | 7 | BDS_BBS:BPM_D5653 | 3310 | BDS_BBS:BPM_D5653 | 5653
BPM | BDS_BBS:BPM_D5653 | BPM | D5655 | 44E000039 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 18 | 7 |  | 7 |  |  |  | 18 | 7 | BDS_BBS:BPM_D5653 | 3310 | BDS_BBS:BPM_D5653 | 5653
BPM | BDS_BBS:BPM_D5653 | BPM | D5655 | 44E000040 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 18 | 7 |  | 8 |  |  |  | 18 | 7 | BDS_BBS:BPM_D5653 | 3310 | BDS_BBS:BPM_D5653 | 5653
BPM | BDS_BBS:BPM_D5653 | BPM | D5655 | 44E000045 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 18 | 7 |  | Ref |  |  |  | 18 | 7 | BDS_BBS:BPM_D5653 | 3310 | BDS_BBS:BPM_D5653 | 5653
BPM | BDS_BBS:BPM_D5680 | BPM | D5681 | 44E000024 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 18 | 8 |  | 1 |  |  |  | 18 | 8 | BDS_BBS:BPM_D5680 | 3319 | BDS_BBS:BPM_D5680 | 5680
BPM | BDS_BBS:BPM_D5680 | BPM | D5681 | 44E000030 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 18 | 8 |  | 2 |  |  |  | 18 | 8 | BDS_BBS:BPM_D5680 | 3319 | BDS_BBS:BPM_D5680 | 5680
BPM | BDS_BBS:BPM_D5680 | BPM | D5681 | 44E000031 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 18 | 8 |  | 3 |  |  |  | 18 | 8 | BDS_BBS:BPM_D5680 | 3319 | BDS_BBS:BPM_D5680 | 5680
BPM | BDS_BBS:BPM_D5680 | BPM | D5681 | 44E000032 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 18 | 8 |  | 4 |  |  |  | 18 | 8 | BDS_BBS:BPM_D5680 | 3319 | BDS_BBS:BPM_D5680 | 5680
BPM | BDS_BBS:BPM_D5680 | BPM | D5681 | 44E000029 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | [17 empty] | BDS_BBS:BPM_D5680 | 3319 | BDS_BBS:BPM_D5680 | 5680
BPM | BDS_FFS:BPM_D5742 | BPM | D5742 | 44E000009 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 18 | 8 |  | 5 |  |  |  | 18 | 8 | BDS_FFS:BPM_D5742 | 3335 | BDS_FFS:BPM_D5742 | 5742
BPM | BDS_FFS:BPM_D5742 | BPM | D5742 | 44E000010 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 18 | 8 |  | 6 |  |  |  | 18 | 8 | BDS_FFS:BPM_D5742 | 3335 | BDS_FFS:BPM_D5742 | 5742
BPM | BDS_FFS:BPM_D5742 | BPM | D5742 | 44E000011 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 18 | 8 |  | 7 |  |  |  | 18 | 8 | BDS_FFS:BPM_D5742 | 3335 | BDS_FFS:BPM_D5742 | 5742
BPM | BDS_FFS:BPM_D5742 | BPM | D5742 | 44E000012 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 18 | 8 |  | 8 |  |  |  | 18 | 8 | BDS_FFS:BPM_D5742 | 3335 | BDS_FFS:BPM_D5742 | 5742
BPM | BDS_FFS:BPM_D5742 | BPM | D5742 | 44E000013 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 18 | 8 |  | Ref |  |  |  | 18 | 8 | BDS_FFS:BPM_D5742 | 3335 | BDS_FFS:BPM_D5742 | 5742
BPM | BDS_FFS:BPM_D5772 | BPM | D5772 | 44E000020 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 18 | 9 |  | 1 |  |  |  | 18 | 9 | BDS_FFS:BPM_D5772 | 3349 | BDS_FFS:BPM_D5772 | 5772
BPM | BDS_FFS:BPM_D5772 | BPM | D5772 | 44E000021 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 18 | 9 |  | 2 |  |  |  | 18 | 9 | BDS_FFS:BPM_D5772 | 3349 | BDS_FFS:BPM_D5772 | 5772
BPM | BDS_FFS:BPM_D5772 | BPM | D5772 | 44E000022 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 18 | 9 |  | 3 |  |  |  | 18 | 9 | BDS_FFS:BPM_D5772 | 3349 | BDS_FFS:BPM_D5772 | 5772
BPM | BDS_FFS:BPM_D5772 | BPM | D5772 | 44E000023 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 18 | 9 |  | 4 |  |  |  | 18 | 9 | BDS_FFS:BPM_D5772 | 3349 | BDS_FFS:BPM_D5772 | 5772
BPM | BDS_FFS:BPM_D5772 | BPM | D5772 | 44E000015 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | [17 empty] | BDS_FFS:BPM_D5772 | 3349 | BDS_FFS:BPM_D5772 | 5772
BPM | BDS_FFS:BPM_D5790 | BPM | D5799 | 44E000016 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 18 | 9 |  | 5 |  |  |  | 18 | 9 | BDS_FFS:BPM_D5790 | 3367 | BDS_FFS:BPM_D5790 | 5790
BPM | BDS_FFS:BPM_D5790 | BPM | D5799 | 44E000017 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 18 | 9 |  | 6 |  |  |  | 18 | 9 | BDS_FFS:BPM_D5790 | 3367 | BDS_FFS:BPM_D5790 | 5790
BPM | BDS_FFS:BPM_D5790 | BPM | D5799 | 44E000018 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 18 | 9 |  | 7 |  |  |  | 18 | 9 | BDS_FFS:BPM_D5790 | 3367 | BDS_FFS:BPM_D5790 | 5790
BPM | BDS_FFS:BPM_D5790 | BPM | D5799 | 44E000019 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 18 | 9 |  | 8 |  |  |  | 18 | 9 | BDS_FFS:BPM_D5790 | 3367 | BDS_FFS:BPM_D5790 | 5790
BPM | BDS_FFS:BPM_D5790 | BPM | D5799 | 44E000014 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 18 | 9 |  | Ref |  |  |  | 18 | 9 | BDS_FFS:BPM_D5790 | 3367 | BDS_FFS:BPM_D5790 | 5790
BPM | BDS_FFS:BPM_D5803 | BPM | D5803 | 44E000271 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 18 | 10 |  | 1 |  |  |  | 18 | 10 | BDS_FFS:BPM_D5803 | 3374 | BDS_FFS:BPM_D5803 | 5803
BPM | BDS_FFS:BPM_D5803 | BPM | D5803 | 44E000270 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 18 | 10 |  | 2 |  |  |  | 18 | 10 | BDS_FFS:BPM_D5803 | 3374 | BDS_FFS:BPM_D5803 | 5803
BPM | BDS_FFS:BPM_D5803 | BPM | D5803 | 44E000269 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 18 | 10 |  | 3 |  |  |  | 18 | 10 | BDS_FFS:BPM_D5803 | 3374 | BDS_FFS:BPM_D5803 | 5803
BPM | BDS_FFS:BPM_D5803 | BPM | D5803 | 44E000268 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 18 | 10 |  | 4 |  |  |  | 18 | 10 | BDS_FFS:BPM_D5803 | 3374 | BDS_FFS:BPM_D5803 | 5803
BPM | BDS_FFS:BPM_D5803 | BPM | D5803 | 44E000266 | 2C_1/4in_Hardline_036 | Ref | No terminiation | No term | [17 empty] | BDS_FFS:BPM_D5803 | 3374 | BDS_FFS:BPM_D5803 | 5803
BPM | BDS_FFS:BPM_D5818 | BPM | D5818 | 44E000265 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 18 | 10 |  | 5 |  |  |  | 18 | 10 | BDS_FFS:BPM_D5818 | 3382 | BDS_FFS:BPM_D5818 | 5818
BPM | BDS_FFS:BPM_D5818 | BPM | D5818 | 44E000264 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 18 | 10 |  | 6 |  |  |  | 18 | 10 | BDS_FFS:BPM_D5818 | 3382 | BDS_FFS:BPM_D5818 | 5818
BPM | BDS_FFS:BPM_D5818 | BPM | D5818 | 44E000263 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 18 | 10 |  | 7 |  |  |  | 18 | 10 | BDS_FFS:BPM_D5818 | 3382 | BDS_FFS:BPM_D5818 | 5818
BPM | BDS_FFS:BPM_D5818 | BPM | D5818 | 44E000262 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 18 | 10 |  | 8 |  |  |  | 18 | 10 | BDS_FFS:BPM_D5818 | 3382 | BDS_FFS:BPM_D5818 | 5818
BPM | BDS_FFS:BPM_D5818 | BPM | D5818 | 44E000267 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 18 | 10 |  | Ref |  |  |  | 18 | 10 | BDS_FFS:BPM_D5818 | 3382 | BDS_FFS:BPM_D5818 | 5818
Note on B5: Rodriguez, Pedro:

## BDS_N0105
BDS_N0105
Long Run Cable | [6 empty] | Inter-Rack Connections | [9 empty] | Terminating Device
Diagnostics Device | Lattice Beamline Location | Device Type | CableDB Location | Cable Number | Cable Type | Function | Termination Device | Port Number In | Pin Number | CPC pinout | CPC pinout | CPC pinout | Signal Pin | Shield Pin | Port Number Out | Patch Cable | MTCA Chassis | MTCA Slot | PICO Port | FGPDB Port | Bias Supply Chassis | Bias Supply Slot | Bias Supply Channel |  |  |  | search for row | device name | D number
Profile Monitors | [23 empty] | PICO | CH[0-7] | Device
PM 8" | BDS_BTS:PM_D5514 | PM | D5513 | 44F000020 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-1 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | Lemo-1A | LEMO650/LEMO650 | 19 | 3 | CH0 - fmc2-1 |  | 14 | 1: EBS8005 | Ch2 | 19 | 3 | BDS_BTS:PM_D5514 | 3265 | BDS_BTS:PM_D5514 | 5514
PM 8" | BDS_BTS:PM_D5514 | PM | D5513 | 44F000020 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-1 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | Lemo-1B | LEMO650/LEMO650 | 19 | 3 | CH1 - fmc2-2 |  | 14 | 1: EBS8005 | Ch2 | 19 | 3 | BDS_BTS:PM_D5514 | 3265 | BDS_BTS:PM_D5514 | 5514
PM 8" | BDS_BTS:PM_D5514 | PM | D5513 | 44F000020 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-1 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | Lemo-1C | LEMO650/LEMO650 | 19 | 3 | CH2 - fmc2-3 |  | 14 | 1: EBS8005 | Ch2 | 19 | 3 | BDS_BTS:PM_D5514 | 3265 | BDS_BTS:PM_D5514 | 5514
PM 8" | BDS_BTS:PM_D5567 | PM | D5567 | 44F000024 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-2 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | Lemo-2A | LEMO650/LEMO650 | 19 | 3 | CH4 - fmc1-1 |  | 14 | 1: EBS8005 | Ch3 | 19 | 3 | BDS_BTS:PM_D5567 | 3284 | BDS_BTS:PM_D5567 | 5567
PM 8" | BDS_BTS:PM_D5567 | PM | D5567 | 44F000024 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-2 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield-6 | 4 | 6 | Lemo-2B | LEMO650/LEMO650 | 19 | 3 | CH5 - fmc1-2 |  | 14 | 1: EBS8005 | Ch3 | 19 | 3 | BDS_BTS:PM_D5567 | 3284 | BDS_BTS:PM_D5567 | 5567
PM 8" | BDS_BTS:PM_D5567 | PM | D5567 | 44F000024 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-2 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | Lemo-2C | LEMO650/LEMO650 | 19 | 3 | CH6 - fmc1-3 |  | 14 | 1: EBS8005 | Ch3 | 19 | 3 | BDS_BTS:PM_D5567 | 3284 | BDS_BTS:PM_D5567 | 5567
PM 12" | BDS_BBS:PM_D5628 | PM | D5628 | 44F000108 | 6C_22AWG_Multi_010 | A Signal | CPC to LEMO650 Box | CPC-3 | [6 empty] | Lemo-3A | LEMO650/LEMO650 | 19 | 4 | CH0 - fmc2-1 |  | 14 | 1: EBS8005 | Ch0 | 19 | 4 | BDS_BBS:PM_D5628 | 3305 | BDS_BBS:PM_D5628 | 5628
PM 12" | BDS_BBS:PM_D5628 | PM | D5628 | 44F000108 | 6C_22AWG_Multi_010 | B Signal | CPC to LEMO650 Box | CPC-3 | [6 empty] | Lemo-3B | LEMO650/LEMO650 | 19 | 4 | CH1 - fmc2-2 |  | 14 | 1: EBS8005 | Ch0 | 19 | 4 | BDS_BBS:PM_D5628 | 3305 | BDS_BBS:PM_D5628 | 5628
PM 6" | BDS_BBS:PM_D5628 | PM | D5628 | 44F000112 | 4C_22AWG_Multi_006 | C Signal | CPC to LEMO650 Box | CPC-3 | [6 empty] | Lemo-3C | LEMO650/LEMO650 | 19 | 4 | CH2 - fmc2-3 |  | 14 | 1: EBS8005 | Ch0 | 19 | 4 | BDS_BBS:PM_D5628 | 3305 | BDS_BBS:PM_D5628 | 5628
PM 12" | BDS_BBS:PM_D5653 | PM | D5655 | 44F000101 | 6C_22AWG_Multi_010 | A Signal | CPC to LEMO650 Box | CPC-4 | [6 empty] | Lemo-4A | LEMO650/LEMO650 | 19 | 4 | CH4 - fmc1-1 |  | 14 | 1: EBS8005 | Ch1 | 19 | 4 | BDS_BBS:PM_D5653 | 3312 | BDS_BBS:PM_D5653 | 5653
PM 12" | BDS_BBS:PM_D5653 | PM | D5655 | 44F000101 | 6C_22AWG_Multi_010 | B Signal | CPC to LEMO650 Box | CPC-4 | [6 empty] | Lemo-4B | LEMO650/LEMO650 | 19 | 4 | CH5 - fmc1-2 |  | 14 | 1: EBS8005 | Ch1 | 19 | 4 | BDS_BBS:PM_D5653 | 3312 | BDS_BBS:PM_D5653 | 5653
PM 6" | BDS_BBS:PM_D5653 | PM | D5655 | 44F000100 | 4C_22AWG_Multi_006 | C Signal | CPC to LEMO650 Box | CPC-4 | [6 empty] | Lemo-4C | LEMO650/LEMO650 | 19 | 4 | CH6 - fmc1-3 |  | 14 | 1: EBS8005 | Ch1 | 19 | 4 | BDS_BBS:PM_D5653 | 3312 | BDS_BBS:PM_D5653 | 5653
PM 8" | BDS_FFS:PM_D5743 | PM | D5742 | 44F000028 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-5 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | Lemo-5A | LEMO650/LEMO650 | 19 | 5 | CH0 - fmc2-1 |  | 14 | 1: EBS8005 | Ch4 | 19 | 5 | BDS_FFS:PM_D5743 | 3337 | BDS_FFS:PM_D5743 | 5743
PM 8" | BDS_FFS:PM_D5743 | PM | D5742 | 44F000028 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-5 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | Lemo-5B | LEMO650/LEMO650 | 19 | 5 | CH1 - fmc2-2 |  | 14 | 1: EBS8005 | Ch4 | 19 | 5 | BDS_FFS:PM_D5743 | 3337 | BDS_FFS:PM_D5743 | 5743
PM 8" | BDS_FFS:PM_D5743 | PM | D5742 | 44F000028 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-5 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | Lemo-5C | LEMO650/LEMO650 | 19 | 5 | CH2 - fmc2-3 |  | 14 | 1: EBS8005 | Ch4 | 19 | 5 | BDS_FFS:PM_D5743 | 3337 | BDS_FFS:PM_D5743 | 5743
PM 8" | BDS_FFS:PM_D5774 | PM | D5773 | 44F000034 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-6 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | Lemo-6A | LEMO650/LEMO650 | 19 | 5 | CH4 - fmc1-1 |  | 14 | 1: EBS8005 | Ch5 | 19 | 5 | BDS_FFS:PM_D5774 | 3351 | BDS_FFS:PM_D5774 | 5774
PM 8" | BDS_FFS:PM_D5774 | PM | D5773 | 44F000034 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-6 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | Lemo-6B | LEMO650/LEMO650 | 19 | 5 | CH5 - fmc1-2 |  | 14 | 1: EBS8005 | Ch5 | 19 | 5 | BDS_FFS:PM_D5774 | 3351 | BDS_FFS:PM_D5774 | 5774
PM 8" | BDS_FFS:PM_D5774 | PM | D5773 | 44F000034 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-6 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | Lemo-6C | LEMO650/LEMO650 | 19 | 5 | CH6 - fmc1-3 |  | 14 | 1: EBS8005 | Ch5 | 19 | 5 | BDS_FFS:PM_D5774 | 3351 | BDS_FFS:PM_D5774 | 5774
PM 8" | BDS_FFS:PM_D5792 | PM | D5800 | 44F000116 | 8C_22AWG_Multi_016 | A Signal | CPC to LEMO650 Box | CPC-7 | 1,2,3 | Orange/Wh=1 | White/Or=2 | Shield=3 | 1 | 3 | Lemo-7A | LEMO650/LEMO650 | 19 | 6 | CH0 - fmc2-1 |  | 14 | 1: EBS8005 | Ch6 | 19 | 6 | BDS_FFS:PM_D5792 | 3369 | BDS_FFS:PM_D5792 | 5792
PM 8" | BDS_FFS:PM_D5792 | PM | D5800 | 44F000116 | 8C_22AWG_Multi_016 | B Signal | CPC to LEMO650 Box | CPC-7 | 4,5,6 | Blue/Wh=4 | White/Blue=5 | Shield=6 | 4 | 6 | Lemo-7B | LEMO650/LEMO650 | 19 | 6 | CH1 - fmc2-2 |  | 14 | 1: EBS8005 | Ch6 | 19 | 6 | BDS_FFS:PM_D5792 | 3369 | BDS_FFS:PM_D5792 | 5792
PM 8" | BDS_FFS:PM_D5792 | PM | D5800 | 44F000116 | 8C_22AWG_Multi_016 | C Signal | CPC to LEMO650 Box | CPC-7 | 7,8,9 | Green/Wh=7 | White/Gr=8 | Shield=9 | 7 | 9 | Lemo-7C | LEMO650/LEMO650 | 19 | 6 | CH2 - fmc2-3 |  | 14 | 1: EBS8005 | Ch6 | 19 | 6 | BDS_FFS:PM_D5792 | 3369 | BDS_FFS:PM_D5792 | 5792
Ion Chamber
Ion Chamber | LS3_BTS:IC_D5205 | IC | D5205 | 34E000565 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J1 | [6 empty] | J1 | Lemo to Lemo | 19 | 7 | CH0 - fmc2-1 |  |  |  |  | 19 | 7 | LS3_BTS:IC_D5205 | #N/A | #N/A | #N/A
Ion Chamber | LS3_BTS:IC_D5291 | IC | D5291 | 34E000564 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J2 | [6 empty] | J2 | Lemo to Lemo | 19 | 7 | CH1 - fmc2-2 |  |  |  |  | 19 | 7 | LS3_BTS:IC_D5291 | #N/A | #N/A | #N/A
Ion Chamber | BDS_BTS:IC_D5496 | IC | D5496 | 44E000253 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J3 | [6 empty] | J3 | Lemo to Lemo | 19 | 7 | CH2 - fmc2-3 |  |  |  |  | 19 | 7 | BDS_BTS:IC_D5496 | #N/A | #N/A | #N/A
Ion Chamber | BDS_BTS:IC_D5515 | IC | D5515 | 44E000250 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J4 | [6 empty] | J4 | Lemo to Lemo | 19 | 7 | CH3 - fmc2-4 |  |  |  |  | 19 | 7 | BDS_BTS:IC_D5515 | #N/A | #N/A | #N/A
Ion Chamber | BDS_BTS:IC_D5550 | IC | D5550 | 44E000249 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J5 | [6 empty] | J5 | Lemo to Lemo | 19 | 7 | CH4 - fmc1-1 |  |  |  |  | 19 | 7 | BDS_BTS:IC_D5550 | #N/A | #N/A | #N/A
Ion Chamber | BDS_BBS:IC_D5578 | IC | D5578 | 44E000248 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J6 | [6 empty] | J6 | Lemo to Lemo | 19 | 7 | CH5 - fmc1-2 |  |  |  |  | 19 | 7 | BDS_BBS:IC_D5578 | #N/A | #N/A | #N/A
Ion Chamber | BDS_BBS:IC_D5611 | IC | D5611 | 44E000252 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J7 | [6 empty] | J7 | Lemo to Lemo | 19 | 7 | CH6 - fmc1-3 |  |  |  |  | 19 | 7 | BDS_BBS:IC_D5611 | #N/A | #N/A | #N/A
Ion Chamber | BDS_BBS:IC_D5641 | IC | D5641 | 44E000251 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J8 | [6 empty] | J8 | Lemo to Lemo | 19 | 7 | CH7 - fmc1-4 |  |  |  |  | 19 | 7 | BDS_BBS:IC_D5641 | #N/A | #N/A | #N/A
Ion Chamber | BDS_BBS:IC_D5679 | IC | D5679 | 44E000247 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J1 | [6 empty] | J1 | Lemo to Lemo | 19 | 8 | CH0 - fmc2-1 |  |  |  |  | 19 | 8 | BDS_BBS:IC_D5679 | #N/A | #N/A | #N/A
Ion Chamber | BDS_BBS:IC_D5711 | IC | D5711 | 44E000246 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J2 | [6 empty] | J2 | Lemo to Lemo | 19 | 8 | CH1 - fmc2-2 |  |  |  |  | 19 | 8 | BDS_BBS:IC_D5711 | #N/A | #N/A | #N/A
Ion Chamber | BDS_FFS:IC_D5740 | IC | D5740 | 44E000245 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J3 | [6 empty] | J3 | Lemo to Lemo | 19 | 8 | CH2 - fmc2-3 |  |  |  |  | 19 | 8 | BDS_FFS:IC_D5740 | #N/A | #N/A | #N/A
Ion Chamber | BDS_FFS:IC_D5765 | IC | D5765 | 44E000244 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J4 | [6 empty] | J4 | Lemo to Lemo | 19 | 8 | CH3 - fmc2-4 |  |  |  |  | 19 | 8 | BDS_FFS:IC_D5765 | #N/A | #N/A | #N/A
Ion Chamber | BDS_FFS:IC_D5800 | IC | D5800 | 44E000243 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J5 | [6 empty] | J5 | Lemo to Lemo | 19 | 8 | CH4 - fmc1-1 |  |  |  |  | 19 | 8 | BDS_FFS:IC_D5800 | #N/A | #N/A | #N/A
Ion Chamber | BDS_FFS:IC_D5822 | IC | D5822 | 44E000242 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J6 | [6 empty] | J6 | Lemo to Lemo | 19 | 8 | CH5 - fmc1-2 |  |  |  |  | 19 | 8 | BDS_FFS:IC_D5822 | #N/A | #N/A | #N/A
Ion Chamber | BDS_BTS:IC_D5661 | IC | D5661 | 44E000277 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J7 | [6 empty] | J7 | Lemo to Lemo | 19 | 8 | CH6 - fmc1-3 |  |  |  |  | 19 | 8 | BDS_BTS:IC_D5661
Ion Chamber | BDS_BTS:IC_D5662 | IC | D5662 | 44E000278 | 2C_1kV75LC_Coax_090 | Signal | N-Type to LEMO Box | J8 | [6 empty] | J8 | Lemo to Lemo | 19 | 8 | CH7 - fmc1-4 |  |  |  |  | 19 | 8 | BDS_BTS:IC_D5662
Ion Chamber Bias
Ion Chamber | LS3_BTS:IC_D5205 |  | D5205 | 34H000083 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 1 | [6 empty] | 1 | SHV to SHV |  |  |  |  | 14 | 2 : EHS8020n | Ch0 |  |  | LS3_BTS:IC_D5205 | #N/A | #N/A | #N/A
Ion Chamber | LS3_BTS:IC_D5291 |  | D5291 | 34H000082 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 2 | [6 empty] | 2 | SHV to SHV |  |  |  |  | 14 | 2 : EHS8020n | Ch0 |  |  | LS3_BTS:IC_D5291 | #N/A | #N/A | #N/A
Ion Chamber | BDS_BTS:IC_D5496 |  | D5496 | 44H000037 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 3 | [6 empty] | 3 | SHV to SHV |  |  |  |  | 14 | 2 : EHS8020n | Ch1 |  |  | BDS_BTS:IC_D5496 | #N/A | #N/A | #N/A
Ion Chamber | BDS_BTS:IC_D5515 |  | D5515 | 44H000034 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 4 | [6 empty] | 4 | SHV to SHV |  |  |  |  | 14 | 2 : EHS8020n | Ch1 |  |  | BDS_BTS:IC_D5515 | #N/A | #N/A | #N/A
Ion Chamber | BDS_BTS:IC_D5550 |  | D5550 | 44H000033 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 5 | [6 empty] | 5 | SHV to SHV |  |  |  |  | 14 | 2 : EHS8020n | Ch2 |  |  | BDS_BTS:IC_D5550 | #N/A | #N/A | #N/A
Ion Chamber | BDS_BBS:IC_D5578 |  | D5578 | 44H000032 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 6 | [6 empty] | 6 | SHV to SHV |  |  |  |  | 14 | 2 : EHS8020n | Ch2 |  |  | BDS_BBS:IC_D5578 | #N/A | #N/A | #N/A
Ion Chamber | BDS_BBS:IC_D5611 |  | D5611 | 44H000035 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 7 | [6 empty] | 7 | SHV to SHV |  |  |  |  | 14 | 2 : EHS8020n | Ch3 |  |  | BDS_BBS:IC_D5611 | #N/A | #N/A | #N/A
Ion Chamber | BDS_BBS:IC_D5641 |  | D5641 | 44H000036 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 8 | [6 empty] | 8 | SHV to SHV |  |  |  |  | 14 | 2 : EHS8020n | Ch3 |  |  | BDS_BBS:IC_D5641 | #N/A | #N/A | #N/A
Ion Chamber | BDS_BBS:IC_D5679 |  | D5679 | 44H000031 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 1 | [6 empty] | 1 | SHV to SHV |  |  |  |  | 14 | 2 : EHS8020n | Ch4 |  |  | BDS_BBS:IC_D5679 | #N/A | #N/A | #N/A
Ion Chamber | BDS_BBS:IC_D5711 |  | D5711 | 44H000030 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 2 | [6 empty] | 2 | SHV to SHV |  |  |  |  | 14 | 2 : EHS8020n | Ch4 |  |  | BDS_BBS:IC_D5711 | #N/A | #N/A | #N/A
Ion Chamber | BDS_FFS:IC_D5740 |  | D5740 | 44H000029 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 3 | [6 empty] | 3 | SHV to SHV |  |  |  |  | 14 | 2 : EHS8020n | Ch5 |  |  | BDS_FFS:IC_D5740 | #N/A | #N/A | #N/A
Ion Chamber | BDS_FFS:IC_D5765 |  | D5765 | 44H000028 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 4 | [6 empty] | 4 | SHV to SHV |  |  |  |  | 14 | 2 : EHS8020n | Ch5 |  |  | BDS_FFS:IC_D5765 | #N/A | #N/A | #N/A
Ion Chamber | BDS_FFS:IC_D5800 |  | D5800 | 44H000027 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 5 | [6 empty] | 5 | SHV to SHV |  |  |  |  | 14 | 2 : EHS8020n | Ch6 |  |  | BDS_FFS:IC_D5800 | #N/A | #N/A | #N/A
Ion Chamber | BDS_FFS:IC_D5822 |  | D5822 | 44H000026 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 6 | [6 empty] | 6 | SHV to SHV |  |  |  |  | 14 | 2 : EHS8020n | Ch6 |  |  | BDS_FFS:IC_D5822 | #N/A | #N/A | #N/A
Ion Chamber | BDS_BTS:IC_D5661 |  | D5661 | 44H000068 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 7 | [6 empty] | 7 | SHV to SHV |  |  |  |  | 14 | 2 : EHS8020n | Ch7
Ion Chamber | BDS_BTS:IC_D5662 |  | D5662 | 44H000069 | 2C_5kVDC_Coax_035 | HV Bias | SHV to SHV | 8 | [6 empty] | 8 | SHV to SHV |  |  |  |  | 14 | 2 : EHS8020n | Ch7
Lemo to Lemo Panel
Halo Monitor Ring,Spare | LS3_BTS:HMR_D5382 | HMR - Unused | D5382 | 34H000001 | 2C_1kV75LC_Coax_090 | Signal | [19 empty] | LS3_BTS:HMR_D5382
| LS3_BTS:HMR_D5382 |  | D5382 | 34H000092 | 2C_5kVDC_Coax_035 | HV Bias | [19 empty] | LS3_BTS:HMR_D5382
Halo Monitor Ring,Spare | LS3_BTS:HMR_D5445 | HMR - Unused | D5445 | 34H000006 | 2C_1kV75LC_Coax_090 | Signal | [19 empty] | LS3_BTS:HMR_D5445
| LS3_BTS:HMR_D5445 |  | D5445 | 34H000089 | 2C_5kVDC_Coax_035 | HV Bias | [19 empty] | LS3_BTS:HMR_D5445
Lemo to Lemo Panel
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 1_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 14 | 2: EBS8003 | Ch0 |  |  | PM bias for PICO
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 1_OUT Port | [7 empty] | Lemo650/Lemo650 | 19 | 4 | CH3 - fmc2-4 |  |  |  |  | 19 | 4 | PM bias for PICO
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 2_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 14 | 2: EBS8003 | Ch1 |  |  | PM bias for PICO
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 2_OUT Port | [7 empty] | Lemo650/Lemo650 | 19 | 4 | CH7 - fmc1-4 |  |  |  |  | 19 | 4 | PM bias for PICO
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 3_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 14 | 2: EBS8003 | Ch2 |  |  | PM bias for PICO
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 3_OUT Port | [7 empty] | Lemo650/Lemo650 | 19 | 3 | CH3 - fmc2-4 |  |  |  |  | 19 | 3 | PM bias for PICO
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 4_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 14 | 2: EBS8003 | Ch3 |  |  | PM bias for PICO
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 4_OUT Port | [7 empty] | Lemo650/Lemo650 | 19 | 3 | CH7 - fmc1-4 |  |  |  |  | 19 | 3 | PM bias for PICO
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 5_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 14 | 2: EBS8003 | Ch4 |  |  | PM bias for PICO
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 5_OUT Port | [7 empty] | Lemo650/Lemo650 | 19 | 5 | CH3 - fmc2-4 |  |  |  |  | 19 | 5 | PM bias for PICO
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 6_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 14 | 2: EBS8003 | Ch5 |  |  | PM bias for PICO
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 6_OUT Port | [7 empty] | Lemo650/Lemo650 | 19 | 5 | CH7 - fmc1-4 |  |  |  |  | 19 | 5 | PM bias for PICO
Interconnect | PM bias for PICO | bias |  | - | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 7_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 14 | 2: EBS8003 | Ch6 |  |  | PM bias for PICO
Interconnect | PM bias for PICO | bias |  | - | Lemo650 Patch Cable | HV Bias | Pico-8 | 7_OUT Port | [7 empty] | Lemo650/Lemo650 | 19 | 6 | CH3 - fmc2-4 |  |  |  |  | 19 | 6 | PM bias for PICO

## BDS_N0403
BDS_N0403
Long Run Cable | [6 empty] | Inter-Rack Connections | [9 empty] | Terminating Device
Diagnostics Device | Lattice Beamline Location | Device Type | CableDB Location | Cable Number | Cable Type | Function | Termination Device | Port Number In | Pin Number | CPC pinout | CPC pinout | CPC pinout | Signal Pin | Shield Pin | Port Number Out | Patch Cable | MTCA Chassis | MTCA Slot | PICO Port | FGPDB Port | Bias Supply Chassis | Bias Supply Slot | Bias Supply Channel |  |  |  | search for row | device name | D number
BPMs | [23 empty] | PICO | CH[0-7] | Device
BPM | FS1_CSS:BPM_D2212 | BPM | D2212 | 44E000102 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 20 | 3 |  | 1 |  |  |  | 20 | 3 | FS1_CSS:BPM_D2212 | 1299 | FS1_CSS:BPM_D2212 | 2212
BPM | FS1_CSS:BPM_D2212 | BPM | D2212 | 44E000103 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 20 | 3 |  | 2 |  |  |  | 20 | 3 | FS1_CSS:BPM_D2212 | 1299 | FS1_CSS:BPM_D2212 | 2212
BPM | FS1_CSS:BPM_D2212 | BPM | D2212 | 44E000104 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 20 | 3 |  | 3 |  |  |  | 20 | 3 | FS1_CSS:BPM_D2212 | 1299 | FS1_CSS:BPM_D2212 | 2212
BPM | FS1_CSS:BPM_D2212 | BPM | D2212 | 44E000105 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 20 | 3 |  | 4 |  |  |  | 20 | 3 | FS1_CSS:BPM_D2212 | 1299 | FS1_CSS:BPM_D2212 | 2212
BPM | FS1_CSS:BPM_D2212 | BPM | D2212 | 44E000090 | 2C_1/4in_Hardline_036 | Ref | No terminiation | Ref | [6 empty] | Ref | SMA to SMA | 20 | 3 |  | Ref |  |  |  | 20 | 3 | FS1_CSS:BPM_D2212 | 1299 | FS1_CSS:BPM_D2212 | 2212
BPM | FS1_CSS:BPM_D2223 | BPM | D2223 | 44E000106 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 20 | 3 |  | 5 |  |  |  | 20 | 3 | FS1_CSS:BPM_D2223 | 1309 | FS1_CSS:BPM_D2223 | 2223
BPM | FS1_CSS:BPM_D2223 | BPM | D2223 | 44E000107 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 20 | 3 |  | 6 |  |  |  | 20 | 3 | FS1_CSS:BPM_D2223 | 1309 | FS1_CSS:BPM_D2223 | 2223
BPM | FS1_CSS:BPM_D2223 | BPM | D2223 | 44E000108 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 20 | 3 |  | 7 |  |  |  | 20 | 3 | FS1_CSS:BPM_D2223 | 1309 | FS1_CSS:BPM_D2223 | 2223
BPM | FS1_CSS:BPM_D2223 | BPM | D2223 | 44E000109 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 20 | 3 |  | 8 |  |  |  | 20 | 3 | FS1_CSS:BPM_D2223 | 1309 | FS1_CSS:BPM_D2223 | 2223
BPM | FS1_CSS:BPM_D2223 | BPM | D2223 | 44E000091 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | No term | [17 empty] | FS1_CSS:BPM_D2223 | 1309 | FS1_CSS:BPM_D2223 | 2223
BPM | FS1_CSS:BPM_D2248 | BPM | D2248 | 44E000098 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 20 | 4 |  | 1 |  |  |  | 20 | 4 | FS1_CSS:BPM_D2248 | 1328 | FS1_CSS:BPM_D2248 | 2248
BPM | FS1_CSS:BPM_D2248 | BPM | D2248 | 44E000099 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 20 | 4 |  | 2 |  |  |  | 20 | 4 | FS1_CSS:BPM_D2248 | 1328 | FS1_CSS:BPM_D2248 | 2248
BPM | FS1_CSS:BPM_D2248 | BPM | D2248 | 44E000100 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 20 | 4 |  | 3 |  |  |  | 20 | 4 | FS1_CSS:BPM_D2248 | 1328 | FS1_CSS:BPM_D2248 | 2248
BPM | FS1_CSS:BPM_D2248 | BPM | D2248 | 44E000101 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 20 | 4 |  | 4 |  |  |  | 20 | 4 | FS1_CSS:BPM_D2248 | 1328 | FS1_CSS:BPM_D2248 | 2248
BPM | FS1_CSS:BPM_D2248 | BPM | D2248 | 44E000089 | 2C_1/4in_Hardline_036 | Ref | No terminiation | Ref | [6 empty] | Ref | SMA to SMA | 20 | 4 |  | Ref |  |  |  | 20 | 4 | FS1_CSS:BPM_D2248 | 1328 | FS1_CSS:BPM_D2248 | 2248
BPM | FS1_CSS:BPM_D2278 | BPM | D2278 | 44E000110 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 20 | 4 |  | 5 |  |  |  | 20 | 4 | FS1_CSS:BPM_D2278 | 1359 | FS1_CSS:BPM_D2278 | 2278
BPM | FS1_CSS:BPM_D2278 | BPM | D2278 | 44E000111 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 20 | 4 |  | 6 |  |  |  | 20 | 4 | FS1_CSS:BPM_D2278 | 1359 | FS1_CSS:BPM_D2278 | 2278
BPM | FS1_CSS:BPM_D2278 | BPM | D2278 | 44E000112 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 20 | 4 |  | 7 |  |  |  | 20 | 4 | FS1_CSS:BPM_D2278 | 1359 | FS1_CSS:BPM_D2278 | 2278
BPM | FS1_CSS:BPM_D2278 | BPM | D2278 | 44E000117 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 20 | 4 |  | 8 |  |  |  | 20 | 4 | FS1_CSS:BPM_D2278 | 1359 | FS1_CSS:BPM_D2278 | 2278
BPM | FS1_CSS:BPM_D2278 | BPM | D2278 | 44E000092 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | No term | [17 empty] | FS1_CSS:BPM_D2278 | 1359 | FS1_CSS:BPM_D2278 | 2278
BPM | FS1_CSS:BPM_D2313 | BPM | D2313 | 44E000113 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 20 | 5 |  | 1 |  |  |  | 20 | 5 | FS1_CSS:BPM_D2313 | 1385 | FS1_CSS:BPM_D2313 | 2313
BPM | FS1_CSS:BPM_D2313 | BPM | D2313 | 44E000114 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 20 | 5 |  | 2 |  |  |  | 20 | 5 | FS1_CSS:BPM_D2313 | 1385 | FS1_CSS:BPM_D2313 | 2313
BPM | FS1_CSS:BPM_D2313 | BPM | D2313 | 44E000115 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 20 | 5 |  | 3 |  |  |  | 20 | 5 | FS1_CSS:BPM_D2313 | 1385 | FS1_CSS:BPM_D2313 | 2313
BPM | FS1_CSS:BPM_D2313 | BPM | D2313 | 44E000116 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 20 | 5 |  | 4 |  |  |  | 20 | 5 | FS1_CSS:BPM_D2313 | 1385 | FS1_CSS:BPM_D2313 | 2313
BPM | FS1_CSS:BPM_D2313 | BPM | D2313 | 44E000093 | 2C_1/4in_Hardline_036 | Ref | No terminiation | Ref | [6 empty] | Ref | SMA to SMA | 20 | 5 |  | Ref |  |  |  | 20 | 5 | FS1_CSS:BPM_D2313 | 1385 | FS1_CSS:BPM_D2313 | 2313
BPM | FS1_CSS:BPM_D2369 | BPM | D2369 | 44E000118 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 20 | 5 |  | 5 |  |  |  | 20 | 5 | FS1_CSS:BPM_D2369 | 1440 | FS1_CSS:BPM_D2369 | 2369
BPM | FS1_CSS:BPM_D2369 | BPM | D2369 | 44E000119 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 20 | 5 |  | 6 |  |  |  | 20 | 5 | FS1_CSS:BPM_D2369 | 1440 | FS1_CSS:BPM_D2369 | 2369
BPM | FS1_CSS:BPM_D2369 | BPM | D2369 | 44E000120 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 20 | 5 |  | 7 |  |  |  | 20 | 5 | FS1_CSS:BPM_D2369 | 1440 | FS1_CSS:BPM_D2369 | 2369
BPM | FS1_CSS:BPM_D2369 | BPM | D2369 | 44E000121 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 20 | 5 |  | 8 |  |  |  | 20 | 5 | FS1_CSS:BPM_D2369 | 1440 | FS1_CSS:BPM_D2369 | 2369
BPM | FS1_CSS:BPM_D2369 | BPM | D2369 | 44E000095 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | No term | [17 empty] | FS1_CSS:BPM_D2369 | 1440 | FS1_CSS:BPM_D2369 | 2369
BPM | FS1_CSS:BPM_D2383 | BPM | D2383 | 44E000122 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 20 | 6 |  | 1 |  |  |  | 20 | 6 | FS1_CSS:BPM_D2383 | 1449 | FS1_CSS:BPM_D2383 | 2383
BPM | FS1_CSS:BPM_D2383 | BPM | D2383 | 44E000123 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 20 | 6 |  | 2 |  |  |  | 20 | 6 | FS1_CSS:BPM_D2383 | 1449 | FS1_CSS:BPM_D2383 | 2383
BPM | FS1_CSS:BPM_D2383 | BPM | D2383 | 44E000124 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 20 | 6 |  | 3 |  |  |  | 20 | 6 | FS1_CSS:BPM_D2383 | 1449 | FS1_CSS:BPM_D2383 | 2383
BPM | FS1_CSS:BPM_D2383 | BPM | D2383 | 44E000125 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 20 | 6 |  | 4 |  |  |  | 20 | 6 | FS1_CSS:BPM_D2383 | 1449 | FS1_CSS:BPM_D2383 | 2383
BPM | FS1_CSS:BPM_D2383 | BPM | D2383 | 44E000094 | 2C_1/4in_Hardline_036 | Ref | No terminiation | Ref | [6 empty] | Ref | SMA to SMA | 20 | 6 |  | Ref |  |  |  | 20 | 6 | FS1_CSS:BPM_D2383 | 1449 | FS1_CSS:BPM_D2383 | 2383
BPM | FS1_BBS:BPM_D2421 | BPM | D2421 | 44E000130 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 20 | 6 |  | 5 |  |  |  | 20 | 6 | FS1_BBS:BPM_D2421 | 1472 | FS1_BBS:BPM_D2421 | 2421
BPM | FS1_BBS:BPM_D2421 | BPM | D2421 | 44E000131 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 20 | 6 |  | 6 |  |  |  | 20 | 6 | FS1_BBS:BPM_D2421 | 1472 | FS1_BBS:BPM_D2421 | 2421
BPM | FS1_BBS:BPM_D2421 | BPM | D2421 | 44E000132 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 20 | 6 |  | 7 |  |  |  | 20 | 6 | FS1_BBS:BPM_D2421 | 1472 | FS1_BBS:BPM_D2421 | 2421
BPM | FS1_BBS:BPM_D2421 | BPM | D2421 | 44E000133 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 20 | 6 |  | 8 |  |  |  | 20 | 6 | FS1_BBS:BPM_D2421 | 1472 | FS1_BBS:BPM_D2421 | 2421
BPM | FS1_BBS:BPM_D2421 | BPM | D2421 | 44E000097 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | No term | [17 empty] | FS1_BBS:BPM_D2421 | 1472 | FS1_BBS:BPM_D2421 | 2421
BPM | FS1_BTS:BPM_D2436 | BPM | D2436 | 44E000126 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 1T | [6 empty] | 1T | SMA to SMA | 20 | 7 |  | 1 |  |  |  | 20 | 7 | FS1_BTS:BPM_D2436 | 3534 | FS1_BTS:BPM_D2436 | 2436
BPM | FS1_BTS:BPM_D2436 | BPM | D2436 | 44E000127 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 2L | [6 empty] | 2L | SMA to SMA | 20 | 7 |  | 2 |  |  |  | 20 | 7 | FS1_BTS:BPM_D2436 | 3534 | FS1_BTS:BPM_D2436 | 2436
BPM | FS1_BTS:BPM_D2436 | BPM | D2436 | 44E000128 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 3R | [6 empty] | 3R | SMA to SMA | 20 | 7 |  | 3 |  |  |  | 20 | 7 | FS1_BTS:BPM_D2436 | 3534 | FS1_BTS:BPM_D2436 | 2436
BPM | FS1_BTS:BPM_D2436 | BPM | D2436 | 44E000129 | 2C_1/4in_Hardline_036 | Signal | N-Type to SMA Panel | 4B | [6 empty] | 4B | SMA to SMA | 20 | 7 |  | 4 |  |  |  | 20 | 7 | FS1_BTS:BPM_D2436 | 3534 | FS1_BTS:BPM_D2436 | 2436
BPM | FS1_BTS:BPM_D2436 | BPM | D2436 | 44E000096 | 2C_1/4in_Hardline_036 | Ref | N-Type to SMA Panel | Ref | [6 empty] | Ref | SMA to SMA | 20 | 7 |  | Ref |  |  |  | 20 | 7 | FS1_BTS:BPM_D2436 | 3534 | FS1_BTS:BPM_D2436 | 2436
Note on B5: Rodriguez, Pedro:

## dStation Rack 3
LS1_RSB608:CAB_N0101
Long Run Cable | [6 empty] | Inter-Rack Connections | [9 empty] | Terminating Device
Diagnostics Device | Lattice Beamline Location | Device Type | CableDB Location | Cable Number | Cable Type | Function | Termination Device | Port Number In | Pin Number | CPC pinout | CPC pinout | CPC pinout | Signal Pin | Shield Pin | Port Number Out | Patch Cable | MTCA Chassis | MTCA Slot | PICO Port | FGPDB Port | Bias Supply Chassis | Bias Supply Slot | Bias Supply Channel
Profile Monitor Signal Wiring
PM 8" | LS1_D:PM_D1234 | PM |  | D014_PM | 8C_22AWG_Multi_016 | 8" Signal | CPC to LEMO650 | 1 | [6 empty] | 1A | LEMO650/LEMO650 |  | 3 | CH0 - fmc2-1 |  |  | 1: EBS8003 | Ch0
PM 8" | LS1_D:PM_D1234 | PM |  | D014_PM | 8C_22AWG_Multi_016 | 8" Signal | CPC to LEMO650 | 1 | [6 empty] | 1B | LEMO650/LEMO650 |  | 3 | CH1 - fmc2-2 |  |  | 1: EBS8003 | Ch0
PM 8" | LS1_D:PM_D1234 | PM |  | D014_PM | 8C_22AWG_Multi_016 | 8" Signal | CPC to LEMO650 | 1 | [6 empty] | 1C | LEMO650/LEMO650 |  | 3 | CH2 - fmc2-3 |  |  | 1: EBS8003 | Ch0
PM 8" | LS1_D:PM_D1240 | PM |  | D031_PM | 8C_22AWG_Multi_016 | 8" Signal | CPC to LEMO650 | 2 | [6 empty] | 2A | LEMO650/LEMO650 |  | 3 | CH4 - fmc1-1 |  |  | 1: EBS8003 | Ch1
PM 8" | LS1_D:PM_D1240 | PM |  | D031_PM | 8C_22AWG_Multi_016 | 8" Signal | CPC to LEMO650 | 2 | [6 empty] | 2B | LEMO650/LEMO650 |  | 3 | CH5 - fmc1-2 |  |  | 1: EBS8003 | Ch1
PM 8" | LS1_D:PM_D1240 | PM |  | D031_PM | 8C_22AWG_Multi_016 | 8" Signal | CPC to LEMO650 | 2 | [6 empty] | 2C | LEMO650/LEMO650 |  | 3 | CH6 - fmc1-3 |  |  | 1: EBS8003 | Ch1
Profile Monitor Bias Wiring
| LS1_D:PM_D1234 |  |  | D018_PM | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 1_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 1 | 2: EBS8003 | Ch0
| LS1_D:PM_D1234 |  |  | D019_PM | Lemo650 Patch Cable | HV Bias | Pico-8 | 1_OUT Port | [7 empty] | Lemo650/Lemo650 | 1 | 3 | CH3 - fmc2-4
| LS1_D:PM_D1240 |  |  | D035_PM | Lemo250 Patch Cable | HV Bias | ISEG 8003 | 2_IN Port | [7 empty] | Lemo250/Lemo250 |  |  |  |  | 1 | 2: EBS8003 | Ch1
| LS1_D:PM_D1240 |  |  | D036_PM | Lemo650 Patch Cable | HV Bias | Pico-8 | 2_OUT Port | [7 empty] | Lemo650/Lemo650 | 1 | 3 | CH7 - fmc1-4
Faraday Cup
Faraday Cup | LS1_D:FC_D1244 | FC |  | D095_FC | 2C_75LC_Coax_089 | Signal | Pico-8 | [10 empty] | 4 | CH0 - fmc2-1

## Sheet1
Date | Initals | Desciption of Change (adding/removing/etc.) | WCP/Jira# | Updated Trackers (ü)